## Cover
Smart Cold Chain Monitor
Master Product Lifecycle Record & Traceability File
A reusable device that rides inside vaccine shipping boxes and safeguards the cold-chain record from factory to patient.
STATUS: Conditionally Ready — production not yet authorised, pending testing
Revision | A (planning complete)
Prepared by | Product Engineer LLC — full lifecycle review
What this is | The complete, connected record of the product's development so far
How to move around | Open the Contents tab; every section is one click away
Open Contents ▶

## Contents
Contents
Every section is one click away. Follow the flow top to bottom, or jump anywhere.
| Section | Part of | What you'll find
Start here
| Cover | Start here | The product, its status, and the headline.
| Executive Summary | Start here | The whole programme on one page.
Foundation
| 0 Concept & Decisions | Foundation | What the product is, who it is for, and the four foundational choices that shaped everything that follows.
The Case & the Customer
| 1 Market Requirements | The Case & the Customer | What the market is asking for — the opportunity, the customers, their problems, and what the product must deliver to win.
| 2 Voice of Customer | The Case & the Customer | What customers, partners and front-line staff actually told us, in their own words.
| 3 Stakeholder Needs | The Case & the Customer | What every group touched by the product needs to succeed — and where those needs pull in different directions.
| 4 User Needs | The Case & the Customer | What the people who use the product need it to do for them, described as outcomes rather than features.
What We Will Build
| 5 Product Requirements | What We Will Build | The single source of truth for what the product must do, with a plain description and how the parts fit together.
| 6 System Requirements | What We Will Build | How those product promises are split across the working parts of the system.
| 7 Detailed Requirements | What We Will Build | The detailed, part-by-part requirements for the hardware, software, and connections.
| 8 System Architecture | What We Will Build | The blueprint: how the parts, information, power and behaviour are organised.
| 9 Design Outputs | What We Will Build | The actual set of drawings, models, software and specifications from which the product can be reproduced.
Making It Safe & Reliable
| 10 Hazard Analysis | Making It Safe & Reliable | How the product could expose people or vaccines to harm, and the safeguards that prevent it.
| 11 Failure & Risk Analyses | Making It Safe & Reliable | How each part could fail, how likely and serious that is, and what we do about it.
| 12 Risk Controls | Making It Safe & Reliable | The firm, testable safeguards drawn from the safety and failure work — the promises the product must keep to be safe.
Proving It Works
| 13 Verification | Proving It Works | The plan to prove, by test and analysis, that the product meets every promise.
| 14 Validation | Proving It Works | The plan to prove the product works for real people, in real conditions, for its real purpose.
Making & Launching
| 15 Manufacturing & Transfer | Making & Launching | How the design becomes a product that can be built the same way, correctly, every time.
| 16 Production Release | Making & Launching | The formal go / no-go decision and the exact version of the product it applies to.
| 17 Field Monitoring | Making & Launching | How real-world performance is watched after launch, and how lessons feed back into the design.
The Connected Record
| Traceability Overview | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Traceability Index | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Object Register | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Relationship Register | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T1 End-to-End | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T4 Hazards & Controls | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T7 Verification Coverage | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T8 Validation Coverage | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T9 Manufacturing | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T10 Release Readiness | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| T11 Field Feedback | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Coverage Dashboard | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Gaps & Orphans | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Inconsistencies | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Closure Action Plan | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.
| Release Decision | The Connected Record | The connected record: start from any item and follow it both ways — why it exists and how it is proven.

## Executive Summary
Executive Summary
The whole programme on one page, in plain terms.
↩ Contents
The opportunity
Vaccines lose potency if they get too warm or too cold in transit, and a large share of doses are wasted or, worse, given after they have spoiled. This product is a small, reusable monitor that rides inside vaccine shipping boxes, keeps an honest record of the journey, and warns people the moment conditions go wrong — so good doses are trusted and spoiled doses are caught.
What it is
A battery-powered device that measures temperature, humidity and rough handling; stores a complete trip history even with no signal; sends data to a cloud dashboard on a schedule; raises an immediate alert on a problem; charges over USB-C; and updates its own software securely. It is designed to run up to about 90 days and work from minus 30 to plus 60 degrees Celsius.
The single most important idea
Because this is an information device, its biggest risk is not breaking down — it is telling people the wrong thing. The design is built around one promise: never show 'all clear' unless it is true, and make it easy for a trained person to act on a real warning. Two independent safeguards protect that promise.
The two make-or-break items
(1) Proving the battery truly lasts about 90 days in deep cold, where batteries are weakest and cannot be safely charged when frozen. (2) Proving alerts reliably escape from inside an insulated, sometimes metal-lined box. Both are carried openly and will be settled by testing.
Where we are
The planning and design work is complete and fully connected — every customer need traces to a requirement, a design, a safeguard, a test, and a field check. Building, independent review, and testing are the next phases. Nothing has been tested yet, so no claim is marked 'proven'.
The recommendation
Approve progression into review, build and test. The thinking is thorough, the risks are known and owned, and there is a clear, evidence-based path to a confident production decision. Production itself is deliberately not yet authorised — that gate comes after the evidence is in.

## 0 Concept & Decisions
Concept & Decisions
What the product is, who it is for, and the four foundational choices that shaped everything that follows.
↩ Contents
In this section:
• Smart Cold Chain Monitor (SCM) — Concept & Foundational Decisions
Smart Cold Chain Monitor (SCM) — Concept & Foundational Decisions
Project: smart-cold-chain-monitor-lifecycle · Owner: Chief Systems Engineer (EM-ENG-001) · Traceability: EM-ENG-012 / ORCH-018 · Status: Concept locked; full lifecycle in progress.
This is the single source of truth for the full Product Engineer LLC lifecycle run. Every downstream stage (Market Requirements → … → Traceability) is based on this document. Evidence posture per 00-governance/evidence-standard.md and hallucination-policy.md: every claim is Cited, Standard, or Judgment/Assumption; unknowns are recorded as gaps, never invented.
---
1. Concept (as provided)
A battery-operated IoT device that monitors vaccine shipments from inside insulated shipping containers: logs temperature/humidity/shock, alerts on temperature excursion, uploads data on a cadence, supports OTA firmware updates, and charges over USB-C. Target endurance 90 days; operating range −30 °C to +60 °C. Feature set: BLE, Wi-Fi, LTE-M, GPS, rechargeable battery, temperature + humidity sensors, accelerometer, OLED, USB-C, OTA, cloud dashboard.
2. Intended use (working statement)
The SCM is a reusable, cellular-connected environmental data logger placed inside insulated vaccine shipping containers/cartons to continuously monitor and record temperature, humidity, and shock, alert stakeholders on excursions that could compromise vaccine potency, and provide an auditable cold-chain record for release/acceptance decisions across the distribution chain. It is a monitoring and information device — it does not control temperature and is not the primary means of preserving the vaccine; it supports human/organizational decisions about product disposition.
3. Foundational decisions (D1–D4) — resolved with rationale
These were the four ambiguities in the concept. They are resolved here as program decisions to unblock the lifecycle; each residual tension is carried forward as a key risk/assumption.
D1 — Lifecycle model: REUSABLE / RECHARGEABLE (multi-trip) (Decision)
Rationale: the concept explicitly includes USB-C charging, a rechargeable battery, and OTA — all of which imply a multi-trip reusable device, not a single-trip disposable. Reusable also fits the 90-day endurance and cloud-dashboard fleet model. Residual: reverse-logistics/return, cleaning between trips, and configuration re-provisioning become lifecycle requirements (carried to User/Product/Service reqs).
D2 — Data-upload model: LOCAL LOGGING + SCHEDULED CADENCE UPLOAD + IMMEDIATE EXCURSION ALERTS (Decision)
Rationale: continuous live cellular every 10 min for 90 days is energy-prohibitive for the target endurance (Judgment — confirmed in the power budget). Adopt store-and-forward: log at a configurable interval (default 5–10 min), upload on a cadence (default hourly, configurable), and push an immediate alert on a detected excursion/shock event when connectivity is available; buffer when offline and upload opportunistically (LTE-M primary; Wi-Fi/BLE offload where available). Residual: alert latency depends on connectivity — carried as a hazard (delayed/missing alert) and a user need.
D3 — Primary energy source: RECHARGEABLE, LOW-TEMPERATURE-CAPABLE Li-ion, duty-cycled (Decision, with high-risk tension)
Rationale: reusability (D1) requires a rechargeable chemistry. The −30 °C + 90-day combination is in genuine tension (Judgment): standard Li-ion loses usable capacity and cannot be charged below 0 °C. Direction: select a low-temperature-tolerant chemistry (e.g., low-temp Li-ion or LTO-class), achieve 90 days via aggressive duty-cycling (sleep between samples, batched uploads), inhibit charging below 0 °C, and evaluate self-heating before charge. Residual (KEY RISK): endurance at −30 °C, charge safety at low temperature, and capacity fade over reuse cycles — carried to hazard analysis, reliability, and the power budget as top risks; validate by test.
D4 — Regulatory market(s): MULTI-MARKET — WHO PQS + EU GDP + US + AIR-TRANSPORT (Decision)
Rationale: global vaccine distribution. Target regime set (Standard/Judgment — confirm current revisions during the regulatory stage):
• WHO PQS temperature-monitoring device performance specs (e.g., cold-chain monitor/logger category).
• EU GDP (Good Distribution Practice) for medicinal products; data-integrity expectations.
• US — FDA cold-chain/data-integrity expectations for distribution monitoring (device is an accessory/ information tool, not a therapeutic; confirm classification).
• Air/ground transport — UN 38.3 (lithium battery transport), IATA dangerous-goods for Li batteries; radio — FCC Part 15 / CE RED (BLE/Wi-Fi/LTE-M/GPS); EMC/safety — IEC 62368-1, CISPR/IEC 61000; environmental — IEC 60529 (IP), IEC 60068 environmental tests. Residual: exact WHO PQS category and US classification to be confirmed and ingested during the Regulatory/Market stages.
4. Product summary for downstream stages
Class | reusable connected environmental monitor (information device; not life-supporting; supports cold-chain integrity decisions for vaccines — a public-health-significant but non-therapeutic role).
Subsystems (initial) | enclosure/mechanical (insulated-container-internal, sealed, IP-rated); power (low-temp rechargeable battery, BMS, USB-C charge path); compute/MCU + firmware; sensing (temperature × (int/ext probe), humidity, 3-axis accelerometer/shock, GPS); connectivity (LTE-M primary, Wi-Fi, BLE); HMI (OLED + button/indicator); storage (non-volatile log); OTA; cloud platform + dashboard; mobile/BLE provisioning.
Operating range | −30 °C to +60 °C operating; endurance target 90 days; storage/transport per IEC 60068.
Key performance drivers | temperature-measurement accuracy & calibration; excursion-detection logic; 90-day energy budget at −30 °C; alert latency; data integrity & security; cold-chain auditability.
5. ID scheme (applies to all stages — for end-to-end traceability, ORCH-018/EM-ENG-012)
• Market Requirements → MR-<CODE>-NNN (EXEC/OPP/SEG/PROB/JTBD/COMP/…)
• Voice of Customer → VOC-### · Stakeholder Needs → STN-###
• User Needs → UN-<CODE>-NNN (VIS/STK/PER/JTBD/…)
• Product Requirements → PRD-<CODE>-NNN (FUN/PERF/ELEC/MECH/…)
• System Requirements → SYS-###; Subsystem/HW/SW/Interface → SSR-/HWR-/SWR-/IFR-###; Software lib → SRS-<CODE>-NNN
• Architecture elements → FUN-###/ARC-###/CMP-###/INT-###
• Design Outputs → DO-<CODE>-NNN (SYS/MECH/EE/PCB/FW/SW/CLD/APP/…)
• Hazard Analysis → HA-<CODE>-NNN; hazard taxonomy → HAZ-<CLASS>
• Failure/risk (ERFKL) → FA-<CODE>-NNN (DFMEA/PFMEA/SWFMEA/HWFMEA/FTA/…)
• Risk-Control Requirements → RCR-### (trace to HA + requirement + design output + verification)
• Verification → PV-<CODE>-NNN; Validation → VAL-<CODE>-NNN
• Manufacturing/transfer → PROC-/CP-/STEP-###; Release → REL-/DR-###; Field → FF-/CMP-/CAPA-###
6. Governing standards & tools (built this session)
Market (market-requirements-standard / ORCH-009), User Needs (user-needs-standard / ORCH-010), Product Req (product-requirements-standard / ORCH-011), Software Req (software-requirements-standard / ORCH-012), Design Outputs (design-output-standard / ORCH-013), Verification (product-verification-standard / ORCH-014), Validation (validation-standard / ORCH-015), ERFKL (erfkl-standard / ORCH-016), Hazard (hazard-analysis-standard / ORCH-017), Traceability (traceability-standard / ORCH-018, EM-ENG-012). Severity-first Action Priority; ISO 14971 risk chain; evidence & honesty throughout.
7. Open items to confirm during the run (gaps, not invented)
WHO PQS exact device category & spec revision; US regulatory classification; final battery chemistry & low-temp datasheet; antenna performance inside insulated/metallized containers; certification test houses.

## 1 Market Requirements
Market Requirements
What the market is asking for — the opportunity, the customers, their problems, and what the product must deliver to win.
↩ Contents
In this section:
• Master Market Requirements Library — Smart Cold Chain Monitor (SCM)
• MR-000 — Executive Summary (EXEC)
• MR-100 — Market Opportunity (OPP)
• MR-1000 — Success Metrics (MET)
• MR-1100 — Regulatory Landscape (REG)
• MR-1200 — Personas (PER)
• MR-1300 — Use Environments (ENV)
• MR-1400 — Product Portfolio (PORT)
• MR-1500 — Ecosystem Strategy (ECO)
• MR-1600 — Innovation Roadmap (INNO)
• MR-1700 — Sustainability (SUS)
• MR-1800 — Strategic Risks (RISK)
• MR-1900 — Market Validation (VAL)
• MR-200 — Customer Segmentation (SEG)
• MR-300 — Customer Problems (PROB)
• MR-400 — Jobs To Be Done (JTBD)
• MR-500 — Competitive Landscape (COMP)
• MR-600 — State of the Art (SOTA)
• MR-700 — Product Positioning (POS)
• MR-800 — Commercial Requirements (COMM)
• MR-900 — Business Constraints (CON)
Master Market Requirements Library — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 01 — Market Requirements (ORCH-009) · Owner: Principal Product Strategist (with EM-ENG-001 Chief Systems Engineer as integrator; EM-ENG-006 Quality/Regulatory for MR-COMP/MR-REG evidence) · Status: Authored — instantiated, product-specific.
This library is the highest-level business/market baseline for the SCM lifecycle. It is based on 00-concept/concept-and-decisions.md (the single source of truth), including foundational decisions D1–D4, and follows 00-governance/market-requirements-standard.md and the 20-module map at 14-tools/market-copilot/market-requirements-library/references/module-map.md.
Every requirement uses the semantic ID scheme MR-<CODE>-NNN and the mandated 20-field block (ID · Title · Requirement Statement · Business Objective · Customer Segment · Geographic Applicability · Market Priority · Importance · Competitive Driver · Business Value · Revenue Impact · Risk if Unmet · Evidence Source · Supporting Research · VOC Reference · Strategic Theme · Product Family · Lifecycle Stage · Traceability · Notes). Requirement Statements describe what the market demands, never how the product implements it (the cardinal rule).
---
Product in one line
The SCM is a reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shipping containers to continuously log temperature/humidity/shock, alert stakeholders on excursions, and provide an auditable cold-chain record for product-disposition decisions. It is a monitoring/information device — it does not control temperature and is not therapeutic.
Anchor facts (D1–D4): reusable/rechargeable multi-trip (D1) · local logging + scheduled cadence upload + immediate excursion alerts (D2) · low-temperature-capable rechargeable Li-ion, duty-cycled, charge inhibited below 0 °C (D3) · multi-market WHO PQS + EU GDP + US FDA + air/ground transport (D4). Target endurance 90 days; operating range −30 °C to +60 °C.
---
Module table
Module | Code | File | Requirements | Applicability
MR-000 Executive Summary | EXEC | [MR-000-executive-summary.md](MR-000-executive-summary.md) | 9 | Applicable
MR-100 Market Opportunity | OPP | [MR-100-market-opportunity.md](MR-100-market-opportunity.md) | 11 | Applicable
MR-200 Customer Segmentation | SEG | [MR-200-customer-segmentation.md](MR-200-customer-segmentation.md) | 11 | Applicable
MR-300 Customer Problems | PROB | [MR-300-customer-problems.md](MR-300-customer-problems.md) | 11 | Applicable
MR-400 Jobs To Be Done | JTBD | [MR-400-jobs-to-be-done.md](MR-400-jobs-to-be-done.md) | 10 | Applicable
MR-500 Competitive Landscape | COMP | [MR-500-competitive-landscape.md](MR-500-competitive-landscape.md) | 9 (+ competitive table) | Applicable
MR-600 State of the Art | SOTA | [MR-600-state-of-the-art.md](MR-600-state-of-the-art.md) | 9 | Applicable
MR-700 Product Positioning | POS | [MR-700-product-positioning.md](MR-700-product-positioning.md) | 9 | Applicable
MR-800 Commercial Requirements | COMM | [MR-800-commercial.md](MR-800-commercial.md) | 11 | Applicable
MR-900 Business Constraints | CON | [MR-900-business-constraints.md](MR-900-business-constraints.md) | 10 | Applicable
MR-1000 Market Success Metrics | MET | [MR-1000-success-metrics.md](MR-1000-success-metrics.md) | 12 | Applicable
MR-1100 Regulatory Landscape | REG | [MR-1100-regulatory-landscape.md](MR-1100-regulatory-landscape.md) | 12 | Applicable
MR-1200 Personas | PER | [MR-1200-personas.md](MR-1200-personas.md) | 11 (+ personas table) | Applicable
MR-1300 Use Environments | ENV | [MR-1300-use-environments.md](MR-1300-use-environments.md) | 11 (+ environments table) | Applicable
MR-1400 Product Portfolio | PORT | [MR-1400-product-portfolio.md](MR-1400-product-portfolio.md) | 9 | Applicable
MR-1500 Ecosystem Strategy | ECO | [MR-1500-ecosystem-strategy.md](MR-1500-ecosystem-strategy.md) | 11 | Applicable
MR-1600 Innovation Roadmap | INNO | [MR-1600-innovation-roadmap.md](MR-1600-innovation-roadmap.md) | 10 | Applicable
MR-1700 Sustainability | SUS | [MR-1700-sustainability.md](MR-1700-sustainability.md) | 10 | Applicable
MR-1800 Strategic Risks | RISK | [MR-1800-strategic-risks.md](MR-1800-strategic-risks.md) | 11 | Applicable
MR-1900 Market Validation | VAL | [MR-1900-market-validation.md](MR-1900-market-validation.md) | 11 | Applicable
Total |  | 20 modules | 208 | 20 / 20 applicable
---
Applicability Register
Per the market-requirements standard, scope is never silently dropped: every module is emitted, and any exclusion is recorded with a rationale. For the SCM, all 20 default modules are Applicable — the product is a globally distributed, regulated, reusable, connected medical-adjacent IoT device, so every module carries real market demand. No modules are marked Not Applicable.
Module (code) | Status | Rationale
All 20 (EXEC … VAL) | Applicable | The SCM is a multi-market (WHO PQS/EU GDP/US/air-transport), reusable, connected device with recurring-revenue, regulatory, sustainability, portfolio, and ecosystem dimensions — every module maps to a real market demand for this product.
Borderline topics flagged inside modules (confirm, not excluded):
• MR-REG — exact WHO PQS device category/code and US FDA classification/pathway are Evidence Pending; to be confirmed during the Regulatory stage (ORCH / regulatory ingestion). Standards are cited by designation; record IDs are not fabricated.
• MR-COMP — specific competitor regulatory records (FDA listings, WHO PQS prequalification codes) are Evidence Pending; competitor facts are described as reported/industry-known pending source acquisition and parsing (market-analysis/competitors/).
• MR-800 / MR-1000 — price, cost, margin, share, and attach-rate figures are labelled Judgment/Assumption ("Evidence Pending"); no precise market financials are asserted as fact.
• MR-1400 (ultra-cold −70 °C variant) — treated as a boundary/roadmap topic, not a committed launch SKU.
---
Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, every claim is tagged Cited, Standard, or Judgment/Assumption. Widely reported public-health facts (e.g., WHO/UNICEF/GAVI statements that a material share of vaccines are exposed to out-of-range temperatures) are cited as reported/estimated, not as exact figures. Precise market sizes (TAM/SAM/SOM), spoilage costs, willingness-to-pay, competitor record IDs, and the exact WHO PQS category are Evidence Pending — never invented. The D3 low-temperature battery tension (90-day endurance + charge safety at −30 °C + capacity fade over reuse) is carried as the top strategic risk (MR-RISK-001) with a dedicated closure-by-test in MR-VAL-006.
---
Traceability note (downstream flow)
This library is the authoritative market baseline and the primary upstream input for the next lifecycle stages. Forward flow:
MR (this library) → Voice of Customer (VOC-###) / Stakeholder Needs (STN-###) → User Needs Library (UN-<CODE>-NNN, ORCH-010) → Product Requirements Library (PRD-<CODE>-NNN, ORCH-011) → System / Subsystem Requirements → Architecture → Design Outputs (ORCH-013) → Hazard Analysis (ORCH-017) / ERFKL (ORCH-016) → Risk-Control Requirements → Verification (ORCH-014) / Validation (ORCH-015).
Each requirement's Traceability field carries forward placeholders (UN-SCM-…, PRD-SCM-…) and the governance anchors ORCH-018 (Traceability orchestrator) and EM-ENG-012 (Traceability engineer), which will assemble the end-to-end trace matrix. VOC references are placeholders (VOC-### (pending)) until the Voice-of-Customer library is instantiated.
Convening: ORCH-009 (Market Requirements). Integrator: EM-ENG-001. Regulatory/competitive evidence owner: EM-ENG-006. Human-factors contributor owns PER/ENV/JTBD topics. This baseline feeds ORCH-018 / EM-ENG-012 for lifecycle traceability.
MR-000 — Executive Summary (EXEC)
Traceability: this module rolls up to the SCM Master Market Requirements Library and traces forward to PRD-SCM-001 (Product Requirements), UN-SCM-001 (User Needs), and back to ORCH-018 / EM-ENG-012 (orchestration + originating engineering mandate).
1. Purpose
This module states, at the business level, why the Smart Cold Chain Monitor (SCM) should exist, what market outcome it is chartered to produce, and the criteria by which the program will be judged a commercial and mission success. It is the apex of the market-requirements tree: every downstream opportunity, segmentation, and problem requirement (MR-100/200/300) must be reconcilable with the vision, mission, and success criteria asserted here. It is written for an executive sponsor, an investment committee, and the cross-functional program leadership who must commit capital and headcount.
2. Scope
In scope: the strategic framing of the SCM as a reusable, battery-operated IoT temperature/humidity/shock monitor that rides inside insulated vaccine shipping containers, logs locally, uploads on a cadence, and raises immediate excursion alerts. The module covers business opportunity, market vision, mission, product strategy, success criteria, strategic objectives, investment rationale, and expected outcomes.
Out of scope: engineering implementation, hardware architecture, firmware behaviour, regulatory test plans, and detailed financial models. Those live in the PRD, the discipline engineering modules, and the business-case workbook respectively. This module also does not select a specific WHO PQS category code (marked TBC) or commit exact market figures where those are not evidence-backed.
3. Applicable Standards & References
• WHO PQS (Performance, Quality and Safety) — prequalification framework for cold-chain equipment and temperature monitoring devices/loggers; exact monitor/logger category code TBC.
• EU GDP (Good Distribution Practice, 2013/C 343/01) — distribution-integrity expectations for medicinal products.
• US FDA data integrity / 21 CFR Part 11 — electronic records and signatures.
• WHO / GAVI / UNICEF Supply Division — cold-chain policy, immunization-supply, and procurement references (cited by name for market context, not exact figures).
• UN 38.3, IATA DG for lithium batteries — transport of the device's Li-ion cell.
• Cross-references: ROSTER.md, ORCHESTRATION.md, and the SCM engineering discipline modules.
4. Background
Vaccines lose potency when held outside their labelled temperature window. WHO and immunization-supply bodies have repeatedly reported that a substantial share of vaccines are exposed to temperatures outside recommended ranges somewhere along the distribution chain, and that both heat excursions and inadvertent freezing are common failure modes (reported/estimated, not exact). The economic and human cost is compounded in last-mile and low- and middle-income-country (LMIC) settings, where visibility is thinnest.
Legacy monitoring is fragmented: single-use chemical indicators and vaccine vial monitors (VVMs) give a coarse, non-quantitative signal; USB dataloggers must be physically recovered and manually offloaded; and connected devices are often too costly, too power-hungry in cold, or too market-specific to deploy across the WHO PQS, EU GDP, and US FDA regimes at once. The SCM is positioned to close this gap with one reusable, multi-market, connected device that delivers quantitative, tamper-evident data and real-time excursion alerts from inside the shipper.
5. Requirement Categories
• Business Opportunity — the addressable pain and value pool the program targets.
• Market Vision & Mission — the durable statement of intent and change the SCM effects.
• Product Strategy — how the SCM wins (reusable, multi-market, data-integrity-grade).
• Strategic Objectives — the measurable business goals for the horizon.
• Success Criteria — the thresholds that define program success.
• Investment Rationale & Expected Outcomes — why capital is justified and what it returns.
6. Detailed Market Requirements
6.1 Business Opportunity & Vision
MR-EXEC-001 — Reusable Connected Cold-Chain Monitoring Opportunity
Requirement Statement | The market demands a single reusable, connected monitoring device that delivers quantitative, tamper-evident temperature/humidity/shock data and real-time excursion alerts from inside vaccine shippers, displacing the fragmented mix of single-use indicators, offline dataloggers, and market-specific connected loggers.
Business Objective | Establish the SCM as the default cross-market cold-chain monitor and capture recurring value across the vaccine distribution network.
Customer Segment | Pharma manufacturers, distributors/wholesalers, 3PLs, national immunization programs, and global procurement bodies (see MR-200).
Geographic Applicability | Global, with distinct value propositions in regulated (EU/US) and last-mile/LMIC markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Incumbents force a trade between low-cost single-use indicators and expensive, siloed connected loggers; no incumbent spans WHO PQS + GDP + FDA on one reusable platform.
Business Value | Anchors the entire program; every downstream requirement inherits from this opportunity.
Revenue Impact | Primary revenue engine (device + connectivity/SaaS); magnitude Evidence Pending pending business-case workbook.
Risk if Unmet | No coherent product thesis; program fragments into point solutions with weak differentiation.
Evidence Source | Cited — WHO/GAVI cold-chain reporting that vaccines are frequently exposed to out-of-range temperatures (reported/estimated); Judgment/Assumption for market-capture magnitude (Evidence Pending).
Supporting Research | WHO immunization-supply-chain publications; GAVI cold-chain equipment optimization materials; UNICEF Supply Division cold-chain guidance.
VOC Reference | VOC-001 (pending) — "we cannot prove chain integrity across borders on one system."
Strategic Theme | Platform consolidation.
Product Family | SCM (Smart Cold Chain Monitor) — reusable monitoring platform.
Lifecycle Stage | Concept
Traceability | UN-SCM-001, PRD-SCM-001; ORCH-018 / EM-ENG-012.
Notes | Based on device facts D1–D4; do not overstate market size without the workbook.
MR-EXEC-002 — Market Vision: Total Chain-of-Custody Visibility
Requirement Statement | The market demands continuous, provable chain-of-custody for vaccine consignments from packout to point-of-use, such that any excursion is detected, attributed, and actionable in near-real-time rather than discovered after delivery.
Business Objective | Position the SCM as the visibility backbone of the vaccine cold chain.
Customer Segment | All commercial and public-health segments in MR-200.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Offline loggers reveal excursions only on recovery; the market increasingly expects live visibility comparable to premium logistics tracking.
Business Value | Frames the product's north star and disciplines scope decisions.
Revenue Impact | Indirect — enables premium positioning; magnitude Evidence Pending.
Risk if Unmet | Product reduced to a commodity logger; loses strategic pricing power.
Evidence Source | Judgment/Assumption — vision statement based on observed market direction toward real-time visibility (Evidence Pending for quantification).
Supporting Research | GDP guidance on continuous monitoring; 3PL market movement toward real-time shipment telemetry.
VOC Reference | VoC theme — "we find out too late."
Strategic Theme | Real-time visibility.
Product Family | SCM — vision anchor.
Lifecycle Stage | Concept
Traceability | UN-SCM-002, PRD-SCM-002; ORCH-018 / EM-ENG-012.
Notes | Aligns to D2 (scheduled upload + immediate excursion alerts).
6.2 Mission & Product Strategy
MR-EXEC-003 — Mission: Protect Vaccine Potency and Human Access
Requirement Statement | The market demands solutions that measurably reduce vaccine loss and wasted immunization opportunities caused by undetected temperature excursions, especially in last-mile and LMIC delivery.
Business Objective | Tie commercial success to a defensible public-health mission that unlocks procurement and donor channels.
Customer Segment | National immunization programs, GAVI/UNICEF/WHO procurement, MoH.
Geographic Applicability | Global, weighted to LMICs for mission framing.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Donor and public-health buyers increasingly require evidence of impact, not just features.
Business Value | Mission alignment is a gate to institutional procurement and grant-funded deployment.
Revenue Impact | Enables large-volume tender access; magnitude Evidence Pending.
Risk if Unmet | Excluded from donor-funded and PQS-gated procurement; loses the highest-volume channel.
Evidence Source | Cited — WHO/UNICEF reporting on vaccine wastage and cold-chain losses (reported/estimated).
Supporting Research | WHO vaccine-management assessments; UNICEF Supply Division cold-chain reports.
VOC Reference | VoC theme — "we throw away vaccine we cannot vouch for."
Strategic Theme | Mission-led growth.
Product Family | SCM — mission statement.
Lifecycle Stage | Concept
Traceability | UN-SCM-003, PRD-SCM-003; ORCH-018 / EM-ENG-012.
Notes | Do not claim exact wastage percentages as fact.
MR-EXEC-004 — Product Strategy: One Reusable Multi-Market Platform
Requirement Statement | The market demands a monitoring product that is reusable across many trips and compliant across WHO PQS, EU GDP, and US FDA regimes simultaneously, avoiding the cost and complexity of maintaining separate single-use or region-specific devices.
Business Objective | Win on total cost of ownership and single-platform compliance rather than unit price alone.
Customer Segment | Distributors, 3PLs, manufacturers operating across borders.
Geographic Applicability | Multi-market (EU, US, air/ground transport, LMIC).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Fragmented incumbent portfolios force customers to run parallel systems per market.
Business Value | Reusability + multi-market compliance is the core differentiator and margin lever.
Revenue Impact | Drives per-trip economics favourable to reusable model; magnitude Evidence Pending.
Risk if Unmet | Reduced to competing on single-use price, eroding margin and defensibility.
Evidence Source | Standard — based on WHO PQS / EU GDP / FDA multi-regime requirements (D4).
Supporting Research | WHO PQS device categories; EU GDP; 21 CFR Part 11.
VOC Reference | VoC theme — "we run three systems for three markets."
Strategic Theme | Platform consolidation / TCO.
Product Family | SCM — reusable multi-market platform.
Lifecycle Stage | Concept
Traceability | UN-SCM-004, PRD-SCM-010; ORCH-018 / EM-ENG-012.
Notes | Reusability based on D1; charging-below-0°C inhibition (D3) is an engineering constraint, not a market claim.
MR-EXEC-005 — Product Strategy: Data-Integrity-Grade Evidence
Requirement Statement | The market demands monitoring data that stands as audit-grade, tamper-evident legal evidence for product-disposition and regulatory purposes, not merely indicative readings.
Business Objective | Make SCM data the accepted system of record for disposition decisions and disputes.
Customer Segment | Manufacturers, distributors, regulators, QA/QP release functions.
Geographic Applicability | Global, weighted to FDA/GDP-regulated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Buyers in regulated markets reject data that cannot survive an audit or a liability dispute.
Business Value | Converts telemetry into defensible evidence, raising switching costs and price.
Revenue Impact | Enables premium regulated-market pricing; magnitude Evidence Pending.
Risk if Unmet | Data dismissed in audits/disputes; product fails its core promise.
Evidence Source | Standard — 21 CFR Part 11 / FDA data integrity; EU GDP.
Supporting Research | FDA data-integrity guidance; GDP monitoring expectations.
VOC Reference | VoC theme — "our logger data doesn't hold up in a dispute."
Strategic Theme | Trust and compliance.
Product Family | SCM — data-integrity backbone.
Lifecycle Stage | Concept
Traceability | UN-SCM-005, PRD-SCM-020; ORCH-018 / EM-ENG-012.
Notes | Device is informational (supports disposition); it does not itself release product.
6.3 Strategic Objectives, Success Criteria & Investment
MR-EXEC-006 — Strategic Objectives for the Program Horizon
Requirement Statement | The market demands that the SCM program commit to explicit, measurable objectives — reduction in undetected excursions, reduction in manual logging effort, and cross-market compliance coverage — against which buyers can hold the product accountable.
Business Objective | Convert vision into measurable, buyer-verifiable objectives.
Customer Segment | Program sponsors and enterprise buyers across MR-200.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Enterprise procurement increasingly demands outcome-based commitments.
Business Value | Aligns the organization and gives sales a defensible value narrative.
Revenue Impact | Enables outcome-linked contracting; magnitude Evidence Pending.
Risk if Unmet | Program drifts; no shared definition of "good."
Evidence Source | Judgment/Assumption — objectives to be quantified in the business case (Evidence Pending).
Supporting Research | GDP continuous-monitoring expectations; enterprise 3PL SLAs.
VOC Reference | VoC theme — "prove it will actually cut our losses."
Strategic Theme | Accountable outcomes.
Product Family | SCM — objective framework.
Lifecycle Stage | Concept
Traceability | UN-SCM-006, PRD-SCM-030; ORCH-018 / EM-ENG-012.
Notes | Targets to be set with data owners; endurance 90d and range -30/+60°C are product targets, not market objectives.
MR-EXEC-007 — Success Criteria for Launch and Adoption
Requirement Statement | The market demands clear success thresholds — verified multi-market compliance, demonstrated excursion-detection reliability, and evidence of reduced spoilage/dispute cost — before it will standardize on a new monitoring platform.
Business Objective | Define the bar for declaring the SCM launch and scale-up successful.
Customer Segment | Enterprise and institutional buyers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Buyers de-risk platform switches by demanding proof before standardization.
Business Value | Focuses launch investment on the few metrics buyers actually gate on.
Revenue Impact | Gates access to enterprise standardization; magnitude Evidence Pending.
Risk if Unmet | Ambiguous launch; stalled adoption despite technical readiness.
Evidence Source | Judgment/Assumption — thresholds pending pilot data (Evidence Pending).
Supporting Research | Pilot design references; PQS field-evaluation practice.
VOC Reference | VoC theme — "show us a clean pilot first."
Strategic Theme | Evidence-led launch.
Product Family | SCM — success criteria.
Lifecycle Stage | Launch
Traceability | UN-SCM-007, PRD-SCM-031; ORCH-018 / EM-ENG-012.
Notes | Tie to pilot and PQS field-evaluation milestones.
MR-EXEC-008 — Investment Rationale
Requirement Statement | The market demands that reusable, connected monitoring deliver a lower per-trip cost and lower loss/dispute exposure than incumbent single-use and offline approaches, justifying capital investment in a reusable fleet plus connectivity/SaaS.
Business Objective | Justify program capital via TCO advantage and recurring revenue.
Customer Segment | Distributors, 3PLs, manufacturers (fleet buyers); internal investment committee.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Reusable economics beat single-use over multi-trip lifecycles when recovery/reverse-logistics is managed.
Business Value | Frames the device + recurring-revenue model that underpins valuation.
Revenue Impact | Hardware plus recurring connectivity/SaaS; magnitude Evidence Pending pending workbook.
Risk if Unmet | Capital request unsupported; program not funded.
Evidence Source | Judgment/Assumption — per-trip and recurring economics to be modelled (Evidence Pending).
Supporting Research | Reusable-vs-single-use logistics TCO literature; GAVI cold-chain investment framing.
VOC Reference | VoC theme — "make the reusable math work including recovery."
Strategic Theme | Recurring-revenue platform.
Product Family | SCM — investment case.
Lifecycle Stage | Concept
Traceability | UN-SCM-008, PRD-SCM-040; ORCH-018 / EM-ENG-012.
Notes | Reverse-logistics loss risk is material to the reusable model (see MR-300).
MR-EXEC-009 — Expected Outcomes and Portfolio Value
Requirement Statement | The market demands that a successful monitoring platform yield reduced vaccine wastage, faster and more confident disposition decisions, cleaner audits, and fewer liability disputes — outcomes that compound into durable customer retention.
Business Objective | Articulate the outcome set that sustains growth and retention beyond launch.
Customer Segment | All MR-200 segments.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Sticky outcomes (audit confidence, dispute reduction) create high switching costs.
Business Value | Defines the retention flywheel and expansion narrative.
Revenue Impact | Drives renewal/expansion revenue; magnitude Evidence Pending.
Risk if Unmet | Product wins deals but fails to retain; churn erodes recurring model.
Evidence Source | Judgment/Assumption — outcomes to be validated in-field (Evidence Pending); wastage-reduction premise Cited to WHO/UNICEF reporting.
Supporting Research | WHO/UNICEF cold-chain outcome studies; GDP audit practice.
VOC Reference | VoC theme — "we keep what actually reduces our losses."
Strategic Theme | Retention flywheel.
Product Family | SCM — outcome model.
Lifecycle Stage | Growth
Traceability | UN-SCM-009, PRD-SCM-050; ORCH-018 / EM-ENG-012.
Notes | Outcomes trace directly to the problems in MR-300.
7. Evidence & Validation Notes
• Anchor facts D1–D4 and the stated device specification are the single source of truth and are treated as given, not re-litigated here.
• Cited market context (vaccine exposure to out-of-range temperatures; wastage) is drawn from WHO/GAVI/UNICEF reporting and is labelled reported/estimated — never as an exact figure.
• All magnitude and financial claims are marked Evidence Pending until the business-case workbook and pilot data exist. No precise market number is asserted as fact in this module.
Validation path | pilot/field-evaluation, PQS field evaluation, and the business-case workbook will convert the Judgment/Assumption items into cited evidence.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-EXEC-001 | UN-SCM-001 | PRD-SCM-001 | ORCH-018 / EM-ENG-012
MR-EXEC-002 | UN-SCM-002 | PRD-SCM-002 | ORCH-018 / EM-ENG-012
MR-EXEC-003 | UN-SCM-003 | PRD-SCM-003 | ORCH-018 / EM-ENG-012
MR-EXEC-004 | UN-SCM-004 | PRD-SCM-010 | ORCH-018 / EM-ENG-012
MR-EXEC-005 | UN-SCM-005 | PRD-SCM-020 | ORCH-018 / EM-ENG-012
MR-EXEC-006 | UN-SCM-006 | PRD-SCM-030 | ORCH-018 / EM-ENG-012
MR-EXEC-007 | UN-SCM-007 | PRD-SCM-031 | ORCH-018 / EM-ENG-012
MR-EXEC-008 | UN-SCM-008 | PRD-SCM-040 | ORCH-018 / EM-ENG-012
MR-EXEC-009 | UN-SCM-009 | PRD-SCM-050 | ORCH-018 / EM-ENG-012
Downstream modules: MR-100 (OPP), MR-200 (SEG), MR-300 (PROB) inherit from and must remain consistent with this Executive Summary.
MR-100 — Market Opportunity (OPP)
Traceability: this module elaborates the opportunity asserted in MR-EXEC-001 and traces forward to PRD-SCM-1xx and UN-SCM-1xx, and back to ORCH-018 / EM-ENG-012.
1. Purpose
This module sizes and characterizes the market opportunity for the Smart Cold Chain Monitor (SCM): the addressable value pools, growth vectors, regional opportunities (including LMICs), emerging markets, and the industry, technology, and economic drivers that make the timing favourable. It gives investment and go-to-market leadership a evidence-backed, honestly-caveated view of where demand concentrates and why.
2. Scope
In scope: TAM/SAM/SOM framing, growth dynamics, regional and LMIC opportunity, emerging markets, industry trends, technology trends, and economic drivers relevant to reusable connected vaccine cold-chain monitoring. Out of scope: precise financial forecasts (business-case workbook), competitor teardown (competitive-analysis module), and segment definitions (MR-200). Where a precise market figure is not evidence-backed, it is given as a labelled Judgment/Assumption range or marked "Evidence Pending" — no invented numbers are stated as fact.
3. Applicable Standards & References
• WHO immunization and cold-chain publications — global immunization coverage, cold-chain equipment, and temperature-exposure reporting (reported/estimated).
• GAVI, the Vaccine Alliance — cold-chain equipment optimization platform (CCEOP) and immunization-market context.
• UNICEF Supply Division — vaccine procurement volumes and cold-chain supply reporting.
• IQVIA / industry pharma-logistics reports — pharma cold-chain and biologics market context (cited by name).
• EU GDP; US FDA 21 CFR Part 11 — regulatory drivers shaping demand in regulated markets.
• WHO PQS — prequalification as a market-access gate (category code TBC).
4. Background
Two structural shifts widen the opportunity. First, the biologics and temperature-sensitive share of the pharmaceutical pipeline has grown, and industry cold-chain logistics reporting (e.g., IQVIA and pharma-logistics trade sources) describes cold-chain pharma as a large and faster-growing segment of pharma logistics (reported, not exact). Second, immunization programs and their funders — WHO, GAVI, UNICEF — have prioritized cold-chain visibility and equipment modernization, particularly for last-mile and LMIC delivery, where temperature-exposure losses concentrate (reported/estimated).
Against that demand, connectivity and power technology have matured: LTE-M/NB-IoT low-power wide-area networks, low-temperature-capable Li-ion chemistries with duty-cycling, and cloud data platforms make a 90-day-endurance, -30 to +60°C reusable connected monitor feasible where it previously was not. The SCM opportunity sits at the intersection of rising demand and newly-enabling technology.
5. Requirement Categories
• Market Sizing — TAM / SAM / SOM framing and boundaries.
• Growth Dynamics — demand drivers and growth vectors.
• Regional & LMIC Opportunity — geography-specific demand concentration.
• Emerging Markets — adjacency and expansion beyond vaccines.
• Industry Trends — regulatory and supply-chain structural shifts.
• Technology Trends — enabling technology maturation.
• Economic Drivers — cost, loss, and reimbursement dynamics.
6. Detailed Market Requirements
6.1 Market Sizing (TAM / SAM / SOM)
MR-OPP-001 — Total Addressable Market (TAM) Framing
Requirement Statement | The market demands monitoring for the full population of temperature-sensitive vaccine and biologic shipments moving through global distribution, establishing a large total addressable pool of monitored shipments and reusable-device fleets.
Business Objective | Establish the outer boundary of the opportunity to justify platform-scale investment.
Customer Segment | All temperature-sensitive pharma shippers (see MR-200).
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | No incumbent addresses the full multi-market population on one reusable platform.
Business Value | Frames the ceiling of the opportunity.
Revenue Impact | Defines TAM ceiling; precise value Evidence Pending pending business-case workbook.
Risk if Unmet | Investment scoped too small; program under-resourced relative to opportunity.
Evidence Source | Cited — IQVIA/pharma-logistics reporting that cold-chain pharma is a large, growing logistics segment (reported); exact TAM figure Evidence Pending.
Supporting Research | IQVIA pharma-logistics reports; WHO/UNICEF vaccine-volume reporting.
VOC Reference | VoC theme — "we monitor only a fraction of what we ship."
Strategic Theme | Platform-scale opportunity.
Product Family | SCM — market sizing.
Lifecycle Stage | Concept
Traceability | UN-SCM-101, PRD-SCM-101; ORCH-018 / EM-ENG-012.
Notes | TAM to be quantified in workbook; do not assert a precise dollar figure here.
MR-OPP-002 — Serviceable Addressable Market (SAM) Framing
Requirement Statement | The market demands reusable connected monitoring specifically for vaccine and high-value biologic cold-chain shipments across WHO PQS, EU GDP, and US FDA regimes — the subset the SCM is designed and certified to serve.
Business Objective | Bound the realistically-serviceable market given the SCM's design and certification scope.
Customer Segment | Vaccine manufacturers, distributors, 3PLs, immunization programs.
Geographic Applicability | Multi-market (EU, US, PQS/LMIC, air/ground).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Multi-regime certification narrows the field of credible competitors.
Business Value | Focuses go-to-market on the reachable population.
Revenue Impact | Defines the near-term revenue frontier; precise value Evidence Pending.
Risk if Unmet | Effort diffused across non-serviceable segments.
Evidence Source | Judgment/Assumption — SAM to be derived from workbook (Evidence Pending); certification scope Standard (WHO PQS/GDP/FDA).
Supporting Research | WHO PQS device categories; GDP; FDA data-integrity guidance.
VOC Reference | VoC theme — "certify once, sell everywhere."
Strategic Theme | Multi-market SAM.
Product Family | SCM — market sizing.
Lifecycle Stage | Concept
Traceability | UN-SCM-102, PRD-SCM-102; ORCH-018 / EM-ENG-012.
Notes | SAM excludes shipment types outside the -30/+60°C and 90-day design envelope.
MR-OPP-003 — Serviceable Obtainable Market (SOM) and Beachhead
Requirement Statement | The market demands an initial, winnable beachhead — regulated-market distributors and 3PLs plus donor-funded immunization deployments — where reusable connected monitoring can displace incumbents fastest.
Business Objective | Identify the obtainable near-term share and entry wedge.
Customer Segment | GDP/FDA-regulated distributors, 3PLs; GAVI/UNICEF-funded programs.
Geographic Applicability | EU, US, and selected PQS/LMIC corridors.
Market Priority | High
Importance | High
Competitive Driver | Beachhead segments feel the pain acutely and have budget to switch.
Business Value | Concentrates launch investment where conversion is highest.
Revenue Impact | Defines Year 1–3 obtainable revenue; precise value Evidence Pending.
Risk if Unmet | Diffuse launch; slow traction and capital burn.
Evidence Source | Judgment/Assumption — SOM/beachhead selection pending pilot and pipeline data (Evidence Pending).
Supporting Research | GAVI CCEOP deployment patterns; regulated-market 3PL adoption of telemetry.
VOC Reference | VoC theme — "start where the audit pain is worst."
Strategic Theme | Beachhead entry.
Product Family | SCM — market sizing.
Lifecycle Stage | Launch
Traceability | UN-SCM-103, PRD-SCM-103; ORCH-018 / EM-ENG-012.
Notes | Beachhead aligns with MR-EXEC-007 success criteria.
6.2 Growth Dynamics & Industry Trends
MR-OPP-004 — Growth of Temperature-Sensitive Biologics
Requirement Statement | The market demands scalable cold-chain monitoring as the temperature-sensitive share of the pharma pipeline (biologics, mRNA vaccines, cell/gene therapies) grows, expanding the volume of shipments requiring rigorous monitoring.
Business Objective | Ride secular growth in temperature-sensitive product volumes.
Customer Segment | Manufacturers, CDMOs, distributors.
Geographic Applicability | Global, concentrated in regulated markets.
Market Priority | High
Importance | High
Competitive Driver | Growth outpaces incumbent single-use capacity and manual processes.
Business Value | Tailwind that compounds SCM adoption over time.
Revenue Impact | Expands SAM year over year; magnitude Evidence Pending.
Risk if Unmet | Miss the growth wave; incumbents entrench.
Evidence Source | Cited — IQVIA/pharma-logistics reporting of growing cold-chain biologics share (reported); exact rates Evidence Pending.
Supporting Research | IQVIA pharma-logistics reports; industry cold-chain trade sources.
VOC Reference | VoC theme — "more of our pipeline needs cold chain every year."
Strategic Theme | Secular growth.
Product Family | SCM — growth vector.
Lifecycle Stage | Growth
Traceability | UN-SCM-104, PRD-SCM-104; ORCH-018 / EM-ENG-012.
Notes | Do not cite exact CAGR as fact.
MR-OPP-005 — Regulatory Tightening as Demand Driver
Requirement Statement | The market demands audit-grade continuous monitoring as regulators (EU GDP, US FDA data integrity/Part 11) and PQS raise expectations for documented, tamper-evident temperature control across distribution.
Business Objective | Convert regulatory pressure into demand for SCM's data-integrity advantage.
Customer Segment | Manufacturers, distributors, QA/QP release, 3PLs.
Geographic Applicability | EU, US, and PQS-gated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Compliance is non-negotiable; devices that fail audit are disqualified.
Business Value | Makes compliance the buying trigger SCM is built to satisfy.
Revenue Impact | Accelerates replacement of non-compliant methods; magnitude Evidence Pending.
Risk if Unmet | Demand captured by compliant competitors.
Evidence Source | Standard — EU GDP, FDA 21 CFR Part 11, WHO PQS.
Supporting Research | GDP guidelines; FDA data-integrity guidance; PQS framework.
VOC Reference | VoC theme — "auditors keep raising the bar."
Strategic Theme | Compliance-driven demand.
Product Family | SCM — growth vector.
Lifecycle Stage | Growth
Traceability | UN-SCM-105, PRD-SCM-105; ORCH-018 / EM-ENG-012.
Notes | Links to MR-EXEC-005 data-integrity strategy.
MR-OPP-006 — Shift Toward Reusable / Circular Logistics
Requirement Statement | The market demands reusable monitoring assets as shippers pursue lower per-trip cost and sustainability goals, moving away from disposable indicators and single-use loggers where reverse logistics can be managed.
Business Objective | Capture the shift to reusable, circular cold-chain assets.
Customer Segment | Distributors, 3PLs, packaging/asset-pooling providers.
Geographic Applicability | Global, strongest in mature logistics networks.
Market Priority | High
Importance | High
Competitive Driver | Reusable economics + ESG mandates disadvantage single-use incumbents.
Business Value | Directly favours the SCM's reusable design (D1).
Revenue Impact | Improves per-trip economics and recurring revenue; magnitude Evidence Pending.
Risk if Unmet | Loss of the sustainability/TCO narrative to reusable competitors.
Evidence Source | Judgment/Assumption — reusable-logistics trend observed; quantification Evidence Pending.
Supporting Research | Reusable-container / asset-pooling logistics literature; ESG procurement trends.
VOC Reference | VoC theme — "cut waste and cost per shipment."
Strategic Theme | Circular logistics.
Product Family | SCM — growth vector.
Lifecycle Stage | Growth
Traceability | UN-SCM-106, PRD-SCM-106; ORCH-018 / EM-ENG-012.
Notes | Reverse-logistics recovery is a dependency (see MR-300).
6.3 Regional, LMIC & Emerging-Market Opportunity
MR-OPP-007 — LMIC Last-Mile Opportunity
Requirement Statement | The market demands affordable, durable connected monitoring for last-mile immunization in low- and middle-income countries, where temperature-exposure losses concentrate and visibility is thinnest.
Business Objective | Win the high-impact, donor-funded LMIC last-mile segment.
Customer Segment | National immunization programs/MoH, GAVI/UNICEF/WHO procurement.
Geographic Applicability | LMICs across Sub-Saharan Africa, South/Southeast Asia, and comparable regions.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Few devices combine ruggedness, low-temperature battery capability, PQS eligibility, and affordability.
Business Value | Highest mission impact and largest unit volumes via donor channels.
Revenue Impact | Large-volume tender potential; magnitude Evidence Pending.
Risk if Unmet | Cede the highest-volume, highest-impact channel to competitors.
Evidence Source | Cited — WHO/UNICEF/GAVI reporting on last-mile cold-chain losses and equipment gaps (reported/estimated).
Supporting Research | WHO EPI/cold-chain assessments; UNICEF Supply Division; GAVI CCEOP.
VOC Reference | VoC theme — "the last mile is a black box."
Strategic Theme | Mission-led LMIC growth.
Product Family | SCM — regional opportunity.
Lifecycle Stage | Growth
Traceability | UN-SCM-107, PRD-SCM-107; ORCH-018 / EM-ENG-012.
Notes | LTE-M primary with BLE/Wi-Fi fallback matters where coverage is patchy.
MR-OPP-008 — Regulated-Market (EU/US) Opportunity
Requirement Statement | The market demands premium, audit-grade monitoring in EU and US distribution, where GDP and FDA data-integrity requirements and high product values justify higher device and service spend.
Business Objective | Capture the high-margin regulated-market opportunity.
Customer Segment | Manufacturers, distributors/wholesalers, 3PLs, hospitals/pharmacies.
Geographic Applicability | EU, US.
Market Priority | Must-have
Importance | High
Competitive Driver | High product value and strict audit rules reward data-integrity-grade devices.
Business Value | Margin anchor that funds LMIC affordability.
Revenue Impact | High per-unit and recurring revenue; magnitude Evidence Pending.
Risk if Unmet | Lose the margin pool that cross-subsidizes mission volume.
Evidence Source | Standard — EU GDP, FDA 21 CFR Part 11; market value Judgment/Assumption (Evidence Pending).
Supporting Research | GDP; FDA guidance; IQVIA regulated-market cold-chain reporting.
VOC Reference | VoC theme — "we'll pay for data that survives an audit."
Strategic Theme | Premium regulated markets.
Product Family | SCM — regional opportunity.
Lifecycle Stage | Launch
Traceability | UN-SCM-108, PRD-SCM-108; ORCH-018 / EM-ENG-012.
Notes | Regulated-market margin underwrites the dual-market model.
MR-OPP-009 — Emerging Adjacent Markets (Biologics, Clinical Trials, Specialty Rx)
Requirement Statement | The market demands the same reusable connected monitoring for adjacent high-value cold-chain categories — cell/gene therapy, clinical-trial material, and specialty biologics — expanding the opportunity beyond vaccines.
Business Objective | Extend the platform into adjacent, higher-value cold-chain segments.
Customer Segment | CDMOs, specialty distributors, clinical-trial logistics, manufacturers.
Geographic Applicability | Global, concentrated in regulated markets.
Market Priority | Medium
Importance | Medium
Competitive Driver | High-value therapies demand even tighter monitoring and tolerate premium pricing.
Business Value | Diversifies revenue and raises average selling price.
Revenue Impact | Incremental high-margin expansion; magnitude Evidence Pending.
Risk if Unmet | Platform confined to vaccines; misses adjacent growth.
Evidence Source | Judgment/Assumption — adjacency plausible given design fit; quantification Evidence Pending.
Supporting Research | IQVIA specialty/biologics cold-chain reporting; clinical-trial logistics sources.
VOC Reference | VoC theme — "our therapies are even more temperature-fragile than vaccines."
Strategic Theme | Adjacent expansion.
Product Family | SCM — emerging markets.
Lifecycle Stage | Growth
Traceability | UN-SCM-109, PRD-SCM-109; ORCH-018 / EM-ENG-012.
Notes | Some therapies need ultra-low temperatures beyond the -30°C design floor; scope carefully.
6.4 Technology & Economic Drivers
MR-OPP-010 — Enabling Technology Maturation
Requirement Statement | The market demands connected monitoring that is only now practical because LPWAN (LTE-M/NB-IoT), low-temperature Li-ion with duty-cycling, and cloud data platforms have matured enough to deliver 90-day endurance across -30 to +60°C at acceptable cost.
Business Objective | Exploit the technology inflection that makes the SCM feasible and timely.
Customer Segment | All segments benefiting from real-time, long-endurance telemetry.
Geographic Applicability | Global, subject to LPWAN coverage.
Market Priority | High
Importance | High
Competitive Driver | First movers on the technology inflection set the reference standard.
Business Value | Justifies the timing of the investment.
Revenue Impact | Enables the product category; magnitude Evidence Pending.
Risk if Unmet | Mistimed entry; either too early (immature tech) or too late (commoditized).
Evidence Source | Judgment/Assumption — technology-readiness assessment; based on device spec (LTE-M, Li-ion duty-cycled, 90d/-30-+60°C).
Supporting Research | LPWAN deployment reporting; low-temperature Li-ion literature.
VOC Reference | VoC theme — "past connected loggers gave out in the cold or ran out of battery."
Strategic Theme | Technology timing.
Product Family | SCM — technology driver.
Lifecycle Stage | Concept
Traceability | UN-SCM-110, PRD-SCM-110; ORCH-018 / EM-ENG-012.
Notes | Charging inhibited below 0°C (D3) is a safety constraint shaping duty-cycle strategy.
MR-OPP-011 — Economic Driver: Cost of Wasted Vaccine and Disputes
Requirement Statement | The market demands monitoring that pays for itself by reducing the cost of wasted vaccine, avoidable re-shipments, and liability disputes — the economic losses that connected, provable monitoring can prevent.
Business Objective | Anchor SCM value in avoided-loss economics, not feature spend.
Customer Segment | Manufacturers, distributors, 3PLs, immunization programs, insurers.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | ROI framed on avoided loss beats feature-based competition.
Business Value | Turns loss avoidance into the primary purchase justification.
Revenue Impact | Underpins value-based pricing; magnitude Evidence Pending.
Risk if Unmet | Product judged on price alone; ROI story absent.
Evidence Source | Cited — WHO/UNICEF reporting on vaccine wastage and cold-chain losses (reported/estimated); per-customer ROI Evidence Pending.
Supporting Research | WHO vaccine-management assessments; UNICEF cold-chain cost reporting.
VOC Reference | VoC theme — "one saved shipment pays for the fleet."
Strategic Theme | Avoided-loss ROI.
Product Family | SCM — economic driver.
Lifecycle Stage | Growth
Traceability | UN-SCM-111, PRD-SCM-111; ORCH-018 / EM-ENG-012.
Notes | Loss figures reported/estimated; do not assert exact spoilage cost per shipment.
7. Evidence & Validation Notes
• Cited market context relies on WHO, GAVI, UNICEF Supply Division, and IQVIA/industry pharma-logistics reporting, and is labelled reported/estimated throughout. No exact TAM/SAM/SOM value is stated as fact.
• All sizing and financial magnitudes are Evidence Pending until the business-case workbook is built; ranges, where offered, are explicitly Judgment/Assumption.
• Technology feasibility claims are based on the stated device specification (LTE-M, low-temperature Li-ion, 90-day endurance, -30 to +60°C) and in public LPWAN/battery literature.
Validation path | business-case workbook (sizing), pilot ROI measurement (economic drivers), and PQS field evaluation (LMIC opportunity).
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-OPP-001 | UN-SCM-101 | PRD-SCM-101 | ORCH-018 / EM-ENG-012
MR-OPP-002 | UN-SCM-102 | PRD-SCM-102 | ORCH-018 / EM-ENG-012
MR-OPP-003 | UN-SCM-103 | PRD-SCM-103 | ORCH-018 / EM-ENG-012
MR-OPP-004 | UN-SCM-104 | PRD-SCM-104 | ORCH-018 / EM-ENG-012
MR-OPP-005 | UN-SCM-105 | PRD-SCM-105 | ORCH-018 / EM-ENG-012
MR-OPP-006 | UN-SCM-106 | PRD-SCM-106 | ORCH-018 / EM-ENG-012
MR-OPP-007 | UN-SCM-107 | PRD-SCM-107 | ORCH-018 / EM-ENG-012
MR-OPP-008 | UN-SCM-108 | PRD-SCM-108 | ORCH-018 / EM-ENG-012
MR-OPP-009 | UN-SCM-109 | PRD-SCM-109 | ORCH-018 / EM-ENG-012
MR-OPP-010 | UN-SCM-110 | PRD-SCM-110 | ORCH-018 / EM-ENG-012
MR-OPP-011 | UN-SCM-111 | PRD-SCM-111 | ORCH-018 / EM-ENG-012
Upstream: MR-000 (EXEC). Lateral: MR-200 (SEG), MR-300 (PROB).
MR-1000 — Success Metrics (MET)
Traceability note: This module (code MET) forms part of the Smart Cold Chain Monitor (SCM) Master Market Requirements Library. Metrics trace forward to User Needs (UN-…) and Product Requirements (PRD-…) placeholders; governed under ORCH-018 (Traceability), authored under ORCH-009 / EM-ENG-012.
1. Purpose
This module defines the market-level success metrics for the Smart Cold Chain Monitor (SCM): the measurable outcomes by which the market and business will judge whether the product wins — revenue, margin, market share, adoption/attach, monitoring efficacy (excursion detection and false-alarm rate), customer satisfaction, SaaS retention, device utilization and reliability, support cost, and achievement of the WHO PQS listing that gates institutional demand.
Metrics here state what success looks like as a market demand — the market demands the product achieve X outcome — rather than prescribing engineering targets. Where a specific numeric target is not yet evidence-backed, the requirement names the metric and marks the target as Evidence Pending, to be set from Voice-of-Customer, competitive benchmarks, and the financial model.
2. Scope
In scope: revenue, gross margin, market share, adoption rate, accessory/SaaS attach rate, excursion-detection efficacy, false-alarm rate, customer satisfaction/NPS, SaaS retention/renewal, device utilization (trips per year), device failure/return rate, support cost per device, and WHO PQS listing achievement.
Out of scope: the commercial model that generates these metrics (MR-800 COMM), the constraints that bound achievability (MR-900 CON), and the regulatory approvals themselves (MR-1100 REG) — though PQS listing appears here as a success metric because the market treats it as an outcome gate.
3. Applicable Standards & References
• WHO PQS prequalification — listing is both a regulatory gate (MR-1100 REG) and a market-success metric (institutional demand is conditioned on it).
• EU GDP and FDA data-integrity / 21 CFR Part 11 — frame the efficacy and data-trust metrics buyers care about (excursion detection, defensible records).
• NPS (Net Promoter Score) and standard SaaS metrics (gross/net revenue retention, churn) — industry-standard measurement frameworks referenced for customer-satisfaction and retention metrics.
• Cold-chain excursion-management practice (WHO/PQS, VVM context) — grounds why detection efficacy and false-alarm rate are decisive metrics.
• All numeric targets are Judgment/Assumption / Evidence Pending unless explicitly cited; they are to be set by financial modelling, competitive benchmarking, and VOC — never asserted as fact here.
4. Background
For a reusable, subscription-based cold-chain monitor, success is not a single number but a system of interacting metrics. Device revenue is one-time and volume-driven; SaaS/data-plan revenue is recurring and retention-driven; and the reusable model means device utilization (trips per year) and return/failure rates directly determine whether the unit economics work. Meanwhile the product's core job — catching temperature excursions on vaccine shipments — is judged by excursion-detection efficacy and, just as importantly, a low false-alarm rate: over-alerting causes alert fatigue and erodes trust, while missed excursions cause spoiled vaccines and lost credibility.
Institutional demand (GAVI/UNICEF/EPI) is effectively gated by WHO PQS listing, so achieving that listing is itself a headline success metric. Enterprise/distributor demand is gated by data integrity and reliability, measured through satisfaction/NPS, retention, and support cost per device. This module names each metric, ties it to a business objective, and holds a placeholder for the evidence-backed target.
5. Requirement Categories
1. Financial — revenue, margin, market share.
2. Adoption — adoption rate, attach rate.
3. Product Efficacy — excursion detection, false-alarm rate.
4. Customer — satisfaction/NPS, SaaS retention/renewal.
5. Fleet Economics — utilization, failure/return rate, support cost.
6. Market Gate — WHO PQS listing achieved.
6. Detailed Market Requirements
6.1 Financial
MR-MET-001 — Revenue Growth (Device + Recurring)
Requirement Statement | The market demands that the SCM achieve a combined device and recurring-revenue trajectory that establishes it as a viable, growing product line.
Business Objective | Reach and sustain revenue scale sufficient to fund operations and growth.
Customer Segment | All (aggregate outcome).
Geographic Applicability | Global blended.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Revenue scale signals viability to buyers standardising on a platform.
Business Value | Primary top-line success measure.
Revenue Impact | Direct — this is the revenue metric.
Risk if Unmet | Product judged a failure; funding/continuation at risk.
Evidence Source | Judgment/Assumption — revenue targets Evidence Pending (financial model).
Supporting Research | Hardware+SaaS revenue-mix logic; targets pending.
VOC Reference | n/a (business outcome).
Strategic Theme | Growth.
Product Family | SCM — device + SaaS.
Lifecycle Stage | Growth
Traceability | forward to MR-800 COMM (pricing/subscription), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Split reporting: one-time device vs recurring SaaS/data-plan.
MR-MET-002 — Gross-Margin Achievement
Requirement Statement | The market/business demands achievement of the blended gross-margin target across device and subscription, net of reverse-logistics and calibration/support costs.
Business Objective | Prove the reusable+SaaS model is profitable at scale.
Customer Segment | Internal/investor.
Geographic Applicability | Global blended.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Margin durability underpins valuation and pricing flexibility.
Business Value | Confirms unit economics.
Revenue Impact | Direct on profitability.
Risk if Unmet | Model economically unviable despite revenue.
Evidence Source | Judgment/Assumption — margin target Evidence Pending (ties to MR-COMM-005).
Supporting Research | Reverse-logistics + calibration cost model pending.
VOC Reference | n/a.
Strategic Theme | Profitable growth.
Product Family | SCM — blended.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-005, MR-900 CON (cost); ORCH-018 / EM-ENG-012.
Notes | Must include amortised reverse-logistics cost per trip.
MR-MET-003 — Market Share in Target Segments
Requirement Statement | The market demands that the SCM capture a meaningful share of the reusable cold-chain-monitoring segment in its priority geographies.
Business Objective | Establish category position and reference credibility.
Customer Segment | All target segments.
Geographic Applicability | Priority markets (EU, US, key LMIC/donor programs).
Market Priority | High
Importance | High
Competitive Driver | Share signals category leadership and drives network/reference effects.
Business Value | Positioning and pricing power.
Revenue Impact | High — share correlates with sustained revenue.
Risk if Unmet | Marginal player; squeezed by incumbents.
Evidence Source | Judgment/Assumption — share target and market sizing Evidence Pending.
Supporting Research | Cold-chain-monitoring market sizing pending.
VOC Reference | n/a.
Strategic Theme | Category leadership.
Product Family | SCM.
Lifecycle Stage | Growth
Traceability | forward to MR-800 COMM, PRD-…; ORCH-018 / EM-ENG-012.
Notes | Requires validated TAM/SAM/SOM.
6.2 Adoption
MR-MET-004 — Adoption Rate / Deployment Scale
Requirement Statement | The market demands that deployed active-device count and program adoptions grow to a scale that proves operational fit across real cold-chain routes.
Business Objective | Convert pilots into scaled deployments.
Customer Segment | Distributors, 3PLs, programs.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | High
Competitive Driver | Scaled live deployments are the strongest proof point for new buyers.
Business Value | Drives revenue and reference base.
Revenue Impact | High — active devices drive SaaS revenue.
Risk if Unmet | Stuck in pilot purgatory; no scale economics.
Evidence Source | Judgment/Assumption — adoption target Evidence Pending.
Supporting Research | Pilot-to-scale conversion benchmarks pending.
VOC Reference | VOC-### (pending) — theme: pilot-to-rollout.
Strategic Theme | Adoption.
Product Family | SCM.
Lifecycle Stage | Growth
Traceability | forward to MR-MET-001, PRD-…; ORCH-018 / EM-ENG-012.
Notes | Track active vs shipped devices.
MR-MET-005 — Attach Rate (SaaS & Accessories)
Requirement Statement | The market demands a high attach rate of SaaS subscriptions and accessories (probes/mounts) to deployed devices, confirming the service+ecosystem model.
Business Objective | Maximise recurring and accessory revenue per device.
Customer Segment | All device buyers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Attach rate proves the platform (not just hardware) is valued.
Business Value | Increases revenue per device and stickiness.
Revenue Impact | High — recurring + accessory revenue.
Risk if Unmet | Product sold as bare hardware; recurring thesis fails.
Evidence Source | Judgment/Assumption — attach target Evidence Pending.
Supporting Research | Ties to MR-COMM-006/009.
VOC Reference | VOC-### (pending).
Strategic Theme | Recurring attach.
Product Family | SCM — SaaS + accessories.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-006, MR-COMM-009; ORCH-018 / EM-ENG-012.
Notes | Separate SaaS attach from accessory attach.
6.3 Product Efficacy
MR-MET-006 — Excursion-Detection Efficacy
Requirement Statement | The market demands that the SCM reliably detect and correctly flag genuine temperature (and where relevant humidity/shock) excursions on monitored shipments, with high sensitivity to true excursions.
Business Objective | Deliver the product's core job — catch real excursions that threaten vaccine integrity.
Customer Segment | All (this is the core value).
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Detection efficacy is the fundamental reason to buy the product.
Business Value | Prevents vaccine loss; the basis of trust and TCO case.
Revenue Impact | High — efficacy failures destroy credibility and renewals.
Risk if Unmet | Missed excursions → spoiled vaccines administered or waste → safety and reputational failure.
Evidence Source | Standard/Judgment — detection is based on device function (D2); target sensitivity Evidence Pending (validation testing).
Supporting Research | Evidence: immediate excursion alerts + logging (D2); WHO/PQS excursion-management context.
VOC Reference | VOC-### (pending) — theme: "never miss a real excursion."
Strategic Theme | Trustworthy monitoring.
Product Family | SCM — sensing + alerting.
Lifecycle Stage | Launch
Traceability | forward to PRD-… (sensor accuracy/alert logic), MR-1100 REG (GDP); ORCH-018 / EM-ENG-012.
Notes | Pair with MR-MET-007 — sensitivity vs false alarms must be balanced.
MR-MET-007 — Low False-Alarm Rate
Requirement Statement | The market demands a low false-alarm/nuisance-alert rate, because over-alerting causes alert fatigue and erodes trust in the monitoring system.
Business Objective | Keep alerts actionable so operators act on them.
Customer Segment | Operations teams receiving alerts.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Alert quality differentiates trusted platforms from noisy ones.
Business Value | Sustains operator trust and correct excursion response.
Revenue Impact | Medium/High — nuisance alerts drive churn and ignored real alerts.
Risk if Unmet | Alert fatigue → real excursions ignored → the efficacy metric is undermined in practice.
Evidence Source | Judgment/Assumption — false-alarm target Evidence Pending (field validation).
Supporting Research | Alert-fatigue literature (general); cold-chain alerting practice.
VOC Reference | VOC-### (pending) — theme: "too many false alarms and we stop looking."
Strategic Theme | Trustworthy monitoring.
Product Family | SCM — alerting logic.
Lifecycle Stage | Launch
Traceability | forward to PRD-… (threshold/algorithm), MR-MET-006; ORCH-018 / EM-ENG-012.
Notes | Report as paired metric with detection efficacy (sensitivity/specificity trade-off).
6.4 Customer
MR-MET-008 — Customer Satisfaction / NPS
Requirement Statement | The market demands high customer satisfaction (e.g., NPS) across device, platform, and support, reflecting a product buyers would recommend and re-buy.
Business Objective | Sustain reputation, referrals, and renewals.
Customer Segment | All customers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Satisfaction drives references in a trust-sensitive market.
Business Value | Referral-led growth and retention.
Revenue Impact | Medium/High — correlates with retention/expansion.
Risk if Unmet | Poor reputation stalls adoption and renewals.
Evidence Source | Standard/Judgment — NPS is a standard framework; target Evidence Pending.
Supporting Research | NPS methodology; benchmark targets pending.
VOC Reference | VOC-### (pending).
Strategic Theme | Customer trust.
Product Family | SCM.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-010 (support), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Segment NPS by buyer type.
MR-MET-009 — SaaS Retention / Renewal
Requirement Statement | The market demands high SaaS/data-plan retention and renewal (low churn, healthy net revenue retention), confirming durable recurring value.
Business Objective | Secure predictable recurring revenue and lifetime value.
Customer Segment | SaaS subscribers.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Retention is the core health metric of a subscription business.
Business Value | Compounds recurring revenue; underpins valuation.
Revenue Impact | High — churn directly reduces recurring revenue.
Risk if Unmet | Leaky bucket; growth cannot outrun churn.
Evidence Source | Standard/Judgment — retention/NRR are standard SaaS metrics; targets Evidence Pending.
Supporting Research | SaaS retention benchmarks; ties to MR-COMM-006.
VOC Reference | VOC-### (pending).
Strategic Theme | Recurring durability.
Product Family | SCM — SaaS.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-006, MR-MET-001; ORCH-018 / EM-ENG-012.
Notes | Track gross and net revenue retention.
6.5 Fleet Economics
MR-MET-010 — Device Utilization (Trips per Year)
Requirement Statement | The market/business demands that reusable devices achieve high utilization (trips per device per year) so per-trip economics and margin targets are met.
Business Objective | Maximise revenue-generating trips per deployed asset.
Customer Segment | Internal (drives buyer TCO too).
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Utilization is what makes reusable beat disposable on cost.
Business Value | Improves amortisation, margin, and TCO story.
Revenue Impact | High — utilization drives unit economics.
Risk if Unmet | Low reuse → economics resemble a costly disposable.
Evidence Source | Judgment/Assumption — utilization target Evidence Pending (reverse-logistics pilot).
Supporting Research | Depends on return loop (MR-CON-007, MR-COMM-007).
VOC Reference | n/a.
Strategic Theme | Reuse economics.
Product Family | SCM — fleet.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-001/007, MR-CON-007; ORCH-018 / EM-ENG-012.
Notes | Depends on 90-day endurance and successful returns.
MR-MET-011 — Device Failure / Return Rate & Support Cost per Device
Requirement Statement | The market/business demands low field-failure and non-return rates and a low support cost per device, since reliability and serviceability directly govern reusable margin and customer trust.
Business Objective | Keep devices reliable, recoverable, and cheap to support across their service life.
Customer Segment | All (reliability affects trust); internal (cost).
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Reliability and low support burden distinguish enterprise-grade offerings.
Business Value | Protects margin and reputation; enables scale support.
Revenue Impact | High — failures/returns and support cost erode margin.
Risk if Unmet | High RMA and support cost destroy unit economics and trust.
Evidence Source | Judgment/Assumption — failure/return/support targets Evidence Pending (reliability data + field pilot).
Supporting Research | Ties to MR-COMM-010 (warranty/support), MR-CON-007 (returns).
VOC Reference | VOC-### (pending) — theme: "how often do they fail?"
Strategic Theme | Reliable, serviceable fleet.
Product Family | SCM — fleet/support.
Lifecycle Stage | Sustaining
Traceability | forward to MR-COMM-010, MR-CON-007, PRD-… (reliability); ORCH-018 / EM-ENG-012.
Notes | Distinguish device failure from operational non-return.
6.6 Market Gate
MR-MET-012 — WHO PQS Listing Achieved
Requirement Statement | The market demands that the SCM achieve WHO PQS prequalification listing, which functions as the outcome gate for institutional/donor-funded demand.
Business Objective | Unlock eligibility for the largest cold-chain buyers.
Customer Segment | Donor-funded programs; UNICEF/GAVI; national EPI.
Geographic Applicability | Global donor/LMIC markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | PQS listing is effectively mandatory for donor-funded procurement.
Business Value | Opens the institutional volume channel; strong credibility signal.
Revenue Impact | High — gates access to the largest volume pool.
Risk if Unmet | Excluded from institutional tenders; capped addressable market.
Evidence Source | Standard — PQS listing is the recognised gate; exact device category code and record ID are Evidence Pending (see MR-1100 REG / MR-REG-001).
Supporting Research | WHO PQS scheme for cold-chain monitors/loggers; category code TBC.
VOC Reference | VOC-### (pending) — theme: "must be PQS-listed to bid."
Strategic Theme | Institutional access.
Product Family | SCM.
Lifecycle Stage | Launch
Traceability | forward to MR-1100 REG (MR-REG-001), MR-COMM-003; ORCH-018 / EM-ENG-012.
Notes | Achievement is binary (listed/not); do not assert a record ID as fact.
7. Evidence & Validation Notes
• Every numeric target (revenue, margin, share, adoption, attach, detection sensitivity, false-alarm rate, NPS, retention, utilization, failure/return/support cost) is Judgment/Assumption / Evidence Pending and must be set from: the financial model, validated market sizing, competitive benchmarks, VOC/willingness-to-pay, and — for efficacy metrics — design verification and field validation testing (feeding the V&V modules).
• Efficacy metrics (MR-MET-006/007) must be measured as a paired sensitivity/specificity trade-off, not in isolation, and validated against reference-standard temperature data.
• WHO PQS listing (MR-MET-012) is Standard-backed as a gate; its category code and record ID remain Evidence Pending and are owned by MR-1100 REG.
8. Traceability Placeholders
MR ID | Metric | Cross-module | Governance
MR-MET-001 | Revenue | MR-800 COMM | ORCH-018 / EM-ENG-012
MR-MET-002 | Gross margin | MR-COMM-005, MR-900 CON | ORCH-018 / EM-ENG-012
MR-MET-003 | Market share | MR-800 COMM | ORCH-018 / EM-ENG-012
MR-MET-004 | Adoption | PRD-… | ORCH-018 / EM-ENG-012
MR-MET-005 | Attach rate | MR-COMM-006/009 | ORCH-018 / EM-ENG-012
MR-MET-006 | Detection efficacy | PRD-…, MR-1100 REG | ORCH-018 / EM-ENG-012
MR-MET-007 | False-alarm rate | PRD-… | ORCH-018 / EM-ENG-012
MR-MET-008 | NPS/satisfaction | MR-COMM-010 | ORCH-018 / EM-ENG-012
MR-MET-009 | SaaS retention | MR-COMM-006 | ORCH-018 / EM-ENG-012
MR-MET-010 | Utilization | MR-COMM-007, MR-CON-007 | ORCH-018 / EM-ENG-012
MR-MET-011 | Failure/return/support | MR-COMM-010, MR-CON-007 | ORCH-018 / EM-ENG-012
MR-MET-012 | PQS listing | MR-1100 REG (MR-REG-001) | ORCH-018 / EM-ENG-012
Cross-module links: MR-1000 MET measures the success of MR-800 COMM's model, is bounded by MR-900 CON's constraints, and depends on MR-1100 REG approvals (notably PQS).
MR-1100 — Regulatory Landscape (REG)
Traceability note: This module (code REG) forms part of the Smart Cold Chain Monitor (SCM) Master Market Requirements Library. Regulatory requirements trace forward to User Needs (UN-…) and Product Requirements (PRD-…) placeholders; governed under ORCH-018 (Traceability), authored under ORCH-009 / EM-ENG-012. This module is research-driven: each requirement identifies the obligation, cites the governing standard/regime, and marks exact category codes, classifications, and record IDs as Evidence Pending rather than asserting them.
1. Purpose
This module maps the regulatory landscape the Smart Cold Chain Monitor (SCM) must satisfy to be lawfully sold, transported, connected, and operated across its target markets: WHO PQS prequalification, EU GDP, US FDA data-integrity / 21 CFR Part 11 and device-classification posture, lithium-battery transport (UN 38.3 + IATA DG), radio/EMC/safety (FCC, CE RED, IEC 62368-1, CISPR/IEC 61000), ingress and environmental robustness (IEC 60529, IEC 60068), data privacy (GDPR/regional), cybersecurity (IEC 62443 / ETSI EN 303 645), and country-specific import/type-approval.
It expresses regulatory obligations as market requirements — the market demands the product meet these regimes to be purchasable — and deliberately does not assert specific record numbers, PQS category codes, or a final FDA classification as fact. Those are flagged Evidence Pending for the regulatory team to confirm.
2. Scope
In scope: all regulatory, safety, transport, radio/EMC, environmental, privacy, cybersecurity, and market-access approval regimes that gate sale/operation of the SCM as a monitoring/information device across WHO PQS, EU, US, and air/ground transport contexts (anchor fact D4).
Out of scope: the commercial packaging of compliance (MR-800 COMM, e.g., calibration-as-a-service), the business constraints these regimes create (MR-900 CON, e.g., approval sequencing), and the success metric of achieving PQS listing (MR-1000 MET) — each cross-referenced here.
3. Applicable Standards & References
• WHO PQS — Performance, Quality and Safety prequalification for cold-chain equipment (temperature monitoring devices / loggers). Exact device category code TBC / Evidence Pending.
• EU GDP — Guidelines on Good Distribution Practice of medicinal products for human use (2013/C 343/01).
• US FDA — 21 CFR Part 11 (electronic records and signatures) and FDA data-integrity guidance; device-classification posture for a monitoring/information (non-therapeutic) product Evidence Pending.
• UN Manual of Tests and Criteria, Section 38.3 (UN 38.3) — lithium battery transport testing.
• IATA Dangerous Goods Regulations — lithium-battery air-transport provisions (and applicable ground/sea modal rules: ADR, IMDG).
• FCC Part 15 (US unintentional/intentional radiators) and CE RED — Directive 2014/53/EU (EU radio equipment).
• IEC 62368-1 — audio/video, information and communication technology equipment safety.
• CISPR / IEC 61000 series — EMC emissions and immunity.
• IEC 60529 — degrees of protection provided by enclosures (IP code).
• IEC 60068 — environmental testing.
• GDPR — Regulation (EU) 2016/679 and regional privacy/data-residency regimes.
• IEC 62443 (industrial/IoT cybersecurity) and ETSI EN 303 645 (consumer IoT cybersecurity baseline).
• Country-specific radio type-approval and import regimes (e.g., referenced examples require verification per market) — Evidence Pending per country.
4. Background
The SCM is explicitly a monitoring/information device — it does not control temperature and is not therapeutic (anchor fact). That framing matters regulatorily: it typically keeps the device out of high-risk therapeutic-device classification, but it does not exempt it. The data it produces is used in GDP- and FDA-regulated distribution and may be relied upon for product-disposition decisions, which pulls it firmly into data-integrity and record-controls obligations (EU GDP; FDA 21 CFR Part 11). Its precise FDA device-classification posture as a monitoring/information product must be confirmed, not assumed.
As a globally-shipped, rechargeable, cellular-connected product (anchor fact D4: WHO PQS + EU GDP + US FDA + air/ground transport), the SCM simultaneously triggers lithium-battery transport rules (UN 38.3, IATA DG), radio and EMC regimes (FCC Part 15, CE RED, CISPR/IEC 61000), product-safety (IEC 62368-1), ingress/environmental robustness (IEC 60529 for the −30 °C to +60 °C, humidity/shock in-container environment; IEC 60068), data-privacy (GDPR/regional), and IoT-cybersecurity expectations (IEC 62443 / ETSI EN 303 645). Donor-funded institutional demand is additionally gated by WHO PQS listing, and each country imposes its own radio type-approval and import controls. This module inventories those obligations so downstream requirements and the certification plan (MR-900 CON) are complete and correctly sequenced.
5. Requirement Categories
1. Prequalification & Market Access — WHO PQS; country type-approval/import.
2. Distribution & Data Integrity — EU GDP; FDA Part 11 / data integrity; FDA classification posture.
3. Battery Transport — UN 38.3; IATA/modal DG.
4. Radio, EMC & Safety — FCC Part 15; CE RED; IEC 62368-1; CISPR/IEC 61000.
5. Environmental & Ingress — IEC 60529; IEC 60068.
6. Data Protection & Cybersecurity — GDPR/regional; IEC 62443 / ETSI EN 303 645.
6. Detailed Market Requirements
6.1 Prequalification & Market Access
MR-REG-001 — WHO PQS Prequalification Listing
Requirement Statement | The market demands that the SCM obtain WHO PQS prequalification under the applicable cold-chain temperature-monitoring/logger category, as this listing gates donor-funded procurement.
Business Objective | Achieve eligibility for WHO/UNICEF/GAVI-funded cold-chain procurement.
Customer Segment | Donor-funded programs; UNICEF Supply Division; national EPI.
Geographic Applicability | Global donor/LMIC markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | PQS listing is effectively mandatory to bid in donor procurement.
Business Value | Unlocks the largest institutional volume channel and credibility.
Revenue Impact | High — gates institutional demand.
Risk if Unmet | Excluded from donor tenders; addressable market capped.
Evidence Source | Standard — WHO PQS scheme applies to cold-chain monitoring devices; exact PQS device category code and any record ID are Evidence Pending (must be confirmed with WHO PQS).
Supporting Research | WHO PQS prequalification scheme for cold-chain equipment; category code TBC per anchor fact.
VOC Reference | VOC-### (pending) — theme: "must be PQS-listed to bid."
Strategic Theme | Institutional market access.
Product Family | SCM.
Lifecycle Stage | Launch
Traceability | forward to MR-1000 MET (MR-MET-012), MR-COMM-003, PRD-…; ORCH-018 / EM-ENG-012.
Notes | Do NOT assert a PQS code/record number as fact until confirmed.
MR-REG-002 — Country-Specific Radio Type-Approval & Import
Requirement Statement | The market demands per-country radio type-approval and import/customs compliance for the cellular/Wi-Fi/BLE/GPS device before it can be legally sold and operated in each target country.
Business Objective | Legally place the connected device on market in each priority country.
Customer Segment | All buyers in each national market.
Geographic Applicability | Per-country (varies widely).
Market Priority | Must-have
Importance | High
Competitive Driver | Type-approval gates lawful operation of the radios in-country.
Business Value | Enables lawful multi-country rollout.
Revenue Impact | High — each unapproved country is closed.
Risk if Unmet | Illegal operation; seizure; blocked sales.
Evidence Source | Standard — national radio type-approval and import regimes exist; specific country requirements/labels are Evidence Pending per market.
Supporting Research | National telecom regulators impose radio type-approval; requires per-country legal review.
VOC Reference | n/a (regulatory).
Strategic Theme | Multi-market access.
Product Family | SCM — radio/connectivity.
Lifecycle Stage | Growth
Traceability | forward to MR-900 CON (sequencing), PRD-… (radio); ORCH-018 / EM-ENG-012.
Notes | Sequence approvals by priority-market plan.
6.2 Distribution & Data Integrity
MR-REG-003 — EU GDP Compliance
Requirement Statement | The market demands that the SCM support compliance with EU Good Distribution Practice for temperature monitoring and record-keeping in the distribution of medicinal products.
Business Objective | Make the device usable within GDP-regulated EU distribution.
Customer Segment | EU pharma distributors, wholesalers, 3PLs.
Geographic Applicability | EU/EEA (and GDP-aligned markets).
Market Priority | Must-have
Importance | Critical
Competitive Driver | GDP defensibility is a purchase precondition for EU distributors.
Business Value | Enables sale into the regulated EU distribution market.
Revenue Impact | High — gates EU distributor demand.
Risk if Unmet | Product unusable for GDP-regulated distribution; EU market lost.
Evidence Source | Standard — EU GDP Guidelines 2013/C 343/01 govern temperature control and monitoring in distribution.
Supporting Research | EU GDP (2013/C 343/01) temperature-monitoring and traceability provisions.
VOC Reference | VOC-### (pending) — theme: GDP audit defensibility.
Strategic Theme | Regulated distribution.
Product Family | SCM — device + records.
Lifecycle Stage | Launch
Traceability | forward to MR-COMM-011 (calibration), PRD-… (records); ORCH-018 / EM-ENG-012.
Notes | GDP interlocks with calibration traceability (MR-CON-003).
MR-REG-004 — FDA Data Integrity & 21 CFR Part 11
Requirement Statement | The market demands that the SCM's electronic records and data-handling support US FDA data-integrity expectations and 21 CFR Part 11 (electronic records/signatures) where its data is used in regulated distribution/decisions.
Business Objective | Make the device's records defensible for US pharma/regulated use.
Customer Segment | US pharma manufacturers, distributors, 3PLs.
Geographic Applicability | United States (and Part-11-aligned quality systems globally).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Part-11-defensible records are a US purchase precondition.
Business Value | Enables US regulated-market sale.
Revenue Impact | High — gates US regulated demand.
Risk if Unmet | Records rejected in audits; US market blocked.
Evidence Source | Standard — 21 CFR Part 11 and FDA data-integrity guidance apply to relied-upon electronic records.
Supporting Research | 21 CFR Part 11; FDA data-integrity guidance (ALCOA+ principles).
VOC Reference | VOC-### (pending) — theme: audit-ready electronic records.
Strategic Theme | Data integrity.
Product Family | SCM — records/platform.
Lifecycle Stage | Launch
Traceability | forward to MR-CON-009 (data), PRD-… (audit trail/e-records); ORCH-018 / EM-ENG-012.
Notes | Audit trails, access control, time-stamping are downstream PRD items.
MR-REG-005 — FDA Device-Classification Posture (Monitoring/Information Device)
Requirement Statement | The market demands clarity and compliance on the SCM's US FDA regulatory posture as a monitoring/information device (non-therapeutic, does not control temperature), including confirmation of whether/how it is regulated as a device.
Business Objective | Confirm the correct FDA regulatory pathway/posture and avoid mis-classification.
Customer Segment | US regulated buyers; internal regulatory.
Geographic Applicability | United States.
Market Priority | High
Importance | High
Competitive Driver | Classification clarity de-risks US market entry and buyer confidence.
Business Value | Correct, defensible market-entry pathway.
Revenue Impact | Medium/High — mis-classification could delay or block US sales.
Risk if Unmet | Wrong pathway → enforcement risk or launch delay.
Evidence Source | Judgment/Assumption — device is based as monitoring/information (non-therapeutic); exact FDA classification/pathway is Evidence Pending (requires regulatory determination).
Supporting Research | Anchor fact: monitoring/information device, not therapeutic, does not control temperature. Classification determination pending.
VOC Reference | n/a (regulatory).
Strategic Theme | Regulatory clarity.
Product Family | SCM.
Lifecycle Stage | Concept
Traceability | forward to MR-900 CON (sequencing), MR-REG-004; ORCH-018 / EM-ENG-012.
Notes | Do NOT assert a class/regulation number until confirmed by regulatory counsel.
6.3 Battery Transport
MR-REG-006 — UN 38.3 Lithium-Battery Testing
Requirement Statement | The market demands that the SCM's rechargeable lithium battery meet UN 38.3 transport-testing requirements so the device can be legally transported worldwide.
Business Objective | Enable compliant global shipment (forward and reverse) of the battery-containing device.
Customer Segment | All (transport underpins delivery and returns).
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | UN 38.3 is mandatory to ship lithium batteries.
Business Value | Enables the reusable device's forward/reverse logistics.
Revenue Impact | High — non-compliance halts all shipping.
Risk if Unmet | Cannot legally ship; reverse-logistics model impossible.
Evidence Source | Standard — UN Manual of Tests and Criteria Section 38.3 governs Li-battery transport.
Supporting Research | UN 38.3; anchor facts D3 (low-temp Li-ion) and reverse-logistics need (MR-900 CON).
VOC Reference | n/a (regulatory).
Strategic Theme | Transportability.
Product Family | SCM — battery.
Lifecycle Stage | Launch
Traceability | forward to MR-CON-004 (battery supply), MR-CON-007 (reverse logistics); ORCH-018 / EM-ENG-012.
Notes | Cell/pack changes trigger UN 38.3 retest (schedule/cost constraint, MR-CON-004).
MR-REG-007 — IATA / Modal Dangerous-Goods Transport (Li Batteries)
Requirement Statement | The market demands compliance with IATA Dangerous Goods (air) and applicable modal rules (ground/sea) for packaging, marking, and documentation of the lithium-battery device across air and ground transport.
Business Objective | Ship and return the device compliantly across all transport modes.
Customer Segment | All; especially air-freight and cross-border logistics.
Geographic Applicability | Global; mode-dependent.
Market Priority | Must-have
Importance | Critical
Competitive Driver | DG compliance is mandatory for lawful shipment.
Business Value | Enables air/ground transport and returns at scale.
Revenue Impact | High — DG non-compliance blocks or fines shipments.
Risk if Unmet | Shipment refusal, penalties, broken reverse loop.
Evidence Source | Standard — IATA DGR lithium provisions; ADR (road) / IMDG (sea) as applicable.
Supporting Research | IATA DGR; evidence-backed multi-market/air+ground fact D4.
VOC Reference | n/a (regulatory).
Strategic Theme | Transportability.
Product Family | SCM — packaging/logistics.
Lifecycle Stage | Launch
Traceability | forward to MR-CON-007 (reverse logistics), PRD-… (packaging); ORCH-018 / EM-ENG-012.
Notes | Packaging/marking design must support cost-effective returns.
6.4 Radio, EMC & Safety
MR-REG-008 — FCC Part 15 (US) & CE RED (EU) Radio Compliance
Requirement Statement | The market demands that the SCM's radios (LTE-M, Wi-Fi, BLE, GPS) comply with FCC Part 15 in the US and CE RED (2014/53/EU) in the EU.
Business Objective | Legally sell/operate the connected device in the two anchor markets.
Customer Segment | All US and EU buyers.
Geographic Applicability | US (FCC) and EU/EEA (RED); models other national regimes.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Radio certification is mandatory for lawful sale.
Business Value | Enables US/EU market entry.
Revenue Impact | High — gates the two anchor markets.
Risk if Unmet | Illegal to sell/operate; blocked markets.
Evidence Source | Standard — FCC Part 15; CE RED Directive 2014/53/EU.
Supporting Research | FCC Part 15; RED 2014/53/EU; evidence-backed connectivity facts (LTE-M/Wi-Fi/BLE/GPS).
VOC Reference | n/a (regulatory).
Strategic Theme | Market access.
Product Family | SCM — radio.
Lifecycle Stage | Launch
Traceability | forward to MR-CON-005 (module recert), MR-REG-002 (type-approval); ORCH-018 / EM-ENG-012.
Notes | Cellular-module change likely triggers re-certification (MR-CON-005).
MR-REG-009 — Product Safety (IEC 62368-1) & EMC (CISPR / IEC 61000)
Requirement Statement | The market demands compliance with product-safety (IEC 62368-1) and electromagnetic-compatibility (CISPR/IEC 61000 emissions and immunity) standards for the rechargeable electronic device.
Business Objective | Demonstrate the device is electrically safe and electromagnetically compatible.
Customer Segment | All buyers; a precondition of CE/other conformity.
Geographic Applicability | EU and markets referencing IEC; models global norms.
Market Priority | Must-have
Importance | High
Competitive Driver | Safety/EMC conformity is mandatory for CE and equivalent marks.
Business Value | Enables conformity marking and safe operation near sensitive cargo.
Revenue Impact | High — required for market placement.
Risk if Unmet | No CE mark; safety/EMC non-conformity blocks sale.
Evidence Source | Standard — IEC 62368-1; CISPR/IEC 61000 series.
Supporting Research | IEC 62368-1 safety; CISPR/IEC 61000 EMC; USB-C recharge and radios raise both.
VOC Reference | n/a (regulatory).
Strategic Theme | Safety & compatibility.
Product Family | SCM — hardware.
Lifecycle Stage | Launch
Traceability | forward to PRD-… (safety/EMC design), MR-REG-008; ORCH-018 / EM-ENG-012.
Notes | EMC immunity matters in transport/aircraft-adjacent environments.
6.5 Environmental & Ingress
MR-REG-010 — Ingress Protection (IEC 60529) & Environmental Robustness (IEC 60068)
Requirement Statement | The market demands demonstrated ingress protection (IEC 60529 IP rating) and environmental robustness (IEC 60068 tests) appropriate to the −30 °C to +60 °C, humidity, condensation, and shock environment inside insulated shipping containers.
Business Objective | Prove the reusable device survives its real operating environment across many trips.
Customer Segment | All (reliability/robustness affects trust and reuse).
Geographic Applicability | Global (spans full climatic range).
Market Priority | Must-have
Importance | High
Competitive Driver | Robustness underpins the multi-trip reusability claim.
Business Value | Supports reliability, reuse economics, and trust.
Revenue Impact | Medium/High — environmental failures drive returns/losses.
Risk if Unmet | Moisture/condensation/shock failures in field; reuse economics undermined.
Evidence Source | Standard — IEC 60529 (IP code); IEC 60068 (environmental testing); evidence-backed range −30 °C to +60 °C.
Supporting Research | IEC 60529; IEC 60068; evidence-backed environmental facts (temp/humidity/shock in-container).
VOC Reference | VOC-### (pending) — theme: "survives cold boxes and rough handling."
Strategic Theme | Durable reuse.
Product Family | SCM — enclosure.
Lifecycle Stage | Launch
Traceability | forward to MR-MET-011 (failure/return), PRD-… (IP/enviro); ORCH-018 / EM-ENG-012.
Notes | Target IP rating Evidence Pending (set by use-environment analysis).
6.6 Data Protection & Cybersecurity
MR-REG-011 — Data Privacy (GDPR & Regional Regimes)
Requirement Statement | The market demands that the SCM cloud/mobile platform comply with GDPR and applicable regional data-protection and residency regimes for the telemetry, account, and location data it processes.
Business Objective | Operate the platform lawfully across all served jurisdictions.
Customer Segment | EU and privacy-regulated/government buyers.
Geographic Applicability | EU/EEA and data-localisation jurisdictions; models global.
Market Priority | High
Importance | High
Competitive Driver | Demonstrable data protection is an enterprise/government precondition.
Business Value | Enables lawful SaaS across markets; unlocks privacy-sensitive buyers.
Revenue Impact | Medium/High — non-compliance blocks EU/gov revenue and risks fines.
Risk if Unmet | Fines; loss of enterprise/government contracts.
Evidence Source | Standard — GDPR (EU) 2016/679; regional privacy/residency regimes; specific residency requirements Evidence Pending.
Supporting Research | GDPR; regional data-localisation trends; ties to MR-CON-009.
VOC Reference | VOC-### (pending) — theme: "where does our data live?"
Strategic Theme | Trust & compliance.
Product Family | SCM — platform.
Lifecycle Stage | Launch
Traceability | forward to MR-CON-009, PRD-… (privacy/residency); ORCH-018 / EM-ENG-012.
Notes | GPS/location data raises additional privacy considerations.
MR-REG-012 — IoT Cybersecurity (IEC 62443 / ETSI EN 303 645)
Requirement Statement | The market demands that the SCM meet recognised IoT cybersecurity expectations (e.g., ETSI EN 303 645 consumer-IoT baseline and/or IEC 62443 for connected industrial systems), covering secure OTA firmware update, provisioning, and data protection.
Business Objective | Secure the connected device and data against tampering and compromise across its service life.
Customer Segment | Enterprise, government, and regulated buyers.
Geographic Applicability | Global; emerging IoT-security mandates (EU, UK, US) sharpen this.
Market Priority | High
Importance | High
Competitive Driver | Cybersecurity assurance is increasingly mandated and a buyer precondition.
Business Value | Protects data integrity, OTA, and buyer trust; aligns with emerging mandates.
Revenue Impact | Medium/High — security gaps block enterprise/gov sales and risk recalls.
Risk if Unmet | Compromise of firmware/data; regulatory non-compliance; loss of trust.
Evidence Source | Standard — ETSI EN 303 645; IEC 62443 series; applicable IoT-security regulations Evidence Pending per market.
Supporting Research | ETSI EN 303 645; IEC 62443; evidence-backed OTA firmware + BLE provisioning facts.
VOC Reference | VOC-### (pending) — theme: secure firmware and data.
Strategic Theme | Secure-by-design.
Product Family | SCM — firmware/platform.
Lifecycle Stage | Launch
Traceability | forward to MR-REG-004 (data integrity), PRD-… (secure OTA/provisioning); ORCH-018 / EM-ENG-012.
Notes | Secure OTA is essential to the reusable multi-year service life.
7. Evidence & Validation Notes
• This module is deliberately research-driven: obligations are Standard-backed, but the following are explicitly Evidence Pending and must be confirmed by the regulatory team before being asserted as fact: the exact WHO PQS device category code and any listing record ID (MR-REG-001); the FDA device-classification/pathway (MR-REG-005); per-country radio type-approval and import requirements (MR-REG-002); target IP rating (MR-REG-010); and applicable regional data-residency and IoT-security mandates (MR-REG-011/012).
• No regulatory record number, certificate ID, or classification code is asserted as fact anywhere in this module.
• Validation approach: engage accredited test labs and notified/recognised bodies for FCC/CE/IEC testing; engage WHO PQS and FDA regulatory consultants for prequalification and classification; conduct per-country type-approval and data-residency legal reviews aligned to the priority-market sequence (MR-900 CON, MR-CON-008).
• Certification results feed the Verification (PVL) and Validation (PVaL) modules and the manufacturing/quality system (MR-CON-003).
8. Traceability Placeholders
MR ID | Regime | Cross-module | Governance
MR-REG-001 | WHO PQS | MR-MET-012, MR-COMM-003 | ORCH-018 / EM-ENG-012
MR-REG-002 | Country type-approval/import | MR-CON-008, MR-REG-008 | ORCH-018 / EM-ENG-012
MR-REG-003 | EU GDP | MR-COMM-011 | ORCH-018 / EM-ENG-012
MR-REG-004 | FDA Part 11 / data integrity | MR-CON-009, MR-REG-012 | ORCH-018 / EM-ENG-012
MR-REG-005 | FDA classification posture | MR-CON-008 | ORCH-018 / EM-ENG-012
MR-REG-006 | UN 38.3 | MR-CON-004, MR-CON-007 | ORCH-018 / EM-ENG-012
MR-REG-007 | IATA / modal DG | MR-CON-007 | ORCH-018 / EM-ENG-012
MR-REG-008 | FCC Part 15 / CE RED | MR-CON-005, MR-REG-002 | ORCH-018 / EM-ENG-012
MR-REG-009 | IEC 62368-1 / EMC | PRD-… | ORCH-018 / EM-ENG-012
MR-REG-010 | IEC 60529 / IEC 60068 | MR-MET-011 | ORCH-018 / EM-ENG-012
MR-REG-011 | GDPR / regional privacy | MR-CON-009 | ORCH-018 / EM-ENG-012
MR-REG-012 | IEC 62443 / ETSI EN 303 645 | MR-REG-004 | ORCH-018 / EM-ENG-012
Cross-module links: MR-1100 REG details the regulatory obligations that MR-900 CON sequences as constraints, that MR-800 COMM packages commercially (e.g., calibration-as-a-service), and whose achievement MR-1000 MET tracks (notably PQS listing).
MR-1200 — Personas (PER)
Master Market Requirements Library — Smart Cold Chain Monitor (SCM). Traceability: this module (code PER) feeds forward to UN-/PRD- placeholders and is convened under ORCH-018 / EM-ENG-012.
1. Purpose
This module defines the human actors — the personas — whose needs, workflows, qualifications, and failure modes the Smart Cold Chain Monitor (SCM) must serve across the global vaccine cold chain. It converts "who touches the product, and what do they demand of it" into discrete, testable market requirements. Personas are expressed as market demands (what the class of user needs the market to deliver), never as product features. The module exists so that downstream product requirements (PRD) and system requirements (UN) can be traced to a named user population rather than to an assumed generic "operator."
A cold-chain monitor is unusual in that a single device passes through the hands of a dozen distinct roles on one trip — from a QA manager who provisions it, through logistics coordinators and airline handlers who never open the box, to a rural nurse who reads its screen at the point of care, and finally to a service technician who cleans and recalibrates it for the next trip. Each role has a different literacy, a different regulatory obligation, a different tolerance for complexity, and a different definition of "the device worked." This module makes those differences explicit.
2. Scope
In scope: The identification and characterization of every persona class that interacts with, depends on, procures, administers, or is held accountable by the SCM across its reusable multi-trip lifecycle — provisioning, transit, receipt, excursion response, audit, reverse logistics, and fleet management. Included are the persona summary table and the persona-as-requirement blocks.
Out of scope: Specific use environments and their physical/connectivity conditions (see MR-1300 ENV); product/SKU strategy (MR-1400 PORT); the software ecosystem those personas use (MR-1500 ECO); and detailed regulatory-clause mapping (handled in the standards and compliance modules of the library). Patients/vaccine recipients are stakeholders in outcome but are not device users and are excluded as personas.
3. Applicable Standards & References
• WHO PQS (Performance, Quality and Safety) — prequalification framework and E006 temperature-monitoring device specifications; establishes the operator and cold-chain-manager expectations in EPI settings.
• EU GDP (Good Distribution Practice, 2013/C 343/01) — defines responsibilities of distributors, responsible persons, and personnel handling medicinal products.
• FDA 21 CFR Part 11 — electronic records/signatures; drives the IT/data-administrator and quality-auditor personas' access-control and audit-trail needs.
• WHO/GAVI/UNICEF supply-chain guidance — Effective Vaccine Management (EVM) assessment, and procurement roles for pooled buyers.
• IATA Perishable Cargo Regulations (PCR) / Temperature Control Regulations (TCR) — air-cargo handler expectations.
• ISPE GDP / GAMP 5 — data-governance and computerized-system roles.
• References cited above are real published frameworks; persona quantitative attributes not derivable from them are marked Judgment/Assumption or Evidence Pending in Section 6.
4. Background
The vaccine cold chain is a relay, not a single custodian. A reusable SCM is provisioned by a pharma or distributor QA function, shipped by 3PLs and airlines, received by national immunization programs, deployed to the last mile by health workers, and returned, cleaned, recharged, recalibrated, and re-pooled by service teams. Because the device is reusable, personas who never appear in a single-use logger's life — the field-service/calibration technician, the fleet/IT administrator managing device pooling — become first-class users.
Persona literacy spans an enormous range: from a supply-chain director reading fleet KPIs on a cloud dashboard, to a last-mile nurse in an LMIC health post with intermittent power and limited device familiarity, reading a two-line OLED at the point of vaccine administration. The device's OLED-plus-single-button interface must be legible to the least-trained user while its data model must satisfy the most exacting auditor. This tension — one device, many literacies, many accountabilities — is the central design constraint that this persona module surfaces for downstream engineering.
5. Requirement Categories
• 5.1 Custody & Provisioning personas — those who commission a device to a trip and own the compliance record (QA/cold-chain manager, distributor coordinator, IT/data admin).
• 5.2 Transit & Handling personas — those who move the shipment without opening it (3PL operations lead, warehouse cold-store operator, air-cargo handler represented via logistics coordinator).
• 5.3 Receipt & Point-of-Use personas — those who accept the shipment and act on its status (EPI/immunization program manager, last-mile health worker/nurse).
• 5.4 Procurement & Governance personas — those who buy, fund, or audit (procurement officer GAVI/UNICEF, regulatory/quality auditor, executive/supply-chain director).
• 5.5 Sustainment personas — those who keep the reusable fleet trustworthy (field-service/calibration technician).
6. Detailed Market Requirements
6.1 Custody & Provisioning Personas
MR-PER-001 — Pharma Cold-Chain / QA Manager
Requirement Statement | The market demands that a device serve a pharmaceutical cold-chain/QA manager who must provision, release, and defend the temperature-monitoring record for each shipment against GDP and Part 11 expectations without becoming a bottleneck to dispatch.
Business Objective | Win the trust of the accountable quality function that authorizes purchase and repeat use across a distributor's network.
Customer Segment | Pharma manufacturers and GDP-licensed distributors; regulated wholesalers.
Geographic Applicability | EU/US/regulated markets primarily; global for multinational pharma.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Incumbent single-use loggers and reusable competitors (e.g., established cold-chain data-logger vendors) compete primarily on QA defensibility and audit-readiness.
Business Value | The QA manager is the economic approver; satisfying this persona unlocks fleet-scale adoption.
Revenue Impact | High — gates enterprise and distributor contracts; anchor of recurring SaaS attach.
Risk if Unmet | Device rejected in supplier qualification; excursion records deemed inadmissible; deauthorization for GDP use.
Evidence Source | Standard — EU GDP §2 (personnel/Responsible Person), 21 CFR Part 11.
Supporting Research | GDP Responsible Person accountability model; VVM/temperature-record defensibility literature.
VOC Reference | VOC-### (pending) — theme: "I have to be able to defend every excursion to an inspector."
Strategic Theme | Compliance-grade trust as the primary purchase driver.
Product Family | SCM (Smart Cold Chain Monitor) — professional reusable line.
Lifecycle Stage | Launch
Traceability | forward to UN-QA-REC / PRD-PROVISION placeholders; ORCH-018 / EM-ENG-012.
Notes | This persona owns release-to-ship decisions; provisioning speed is a covert requirement — audit rigor must not slow dispatch.
MR-PER-002 — Distributor Logistics Coordinator
Requirement Statement | The market demands that a device serve a distributor logistics coordinator who books, tracks, and reconciles many concurrent shipments and needs shipment status and excursion alerts without interpreting raw sensor data.
Business Objective | Make the device the default monitoring choice embedded in routine dispatch workflows.
Customer Segment | Pharma distributors, wholesalers, consolidators.
Geographic Applicability | Global; density in EU/US/regional hubs.
Market Priority | Must-have
Importance | High
Competitive Driver | Ease-of-use and status-at-a-glance differentiate against loggers requiring manual download.
Business Value | Daily active user whose adoption drives per-trip consumable/SaaS pull-through.
Revenue Impact | High — volume driver via shipment cadence.
Risk if Unmet | Coordinator reverts to familiar single-use loggers; low utilization erodes reusable ROI.
Evidence Source | Judgment/Assumption — based on GDP dispatch workflow; quantitative volume Evidence Pending.
Supporting Research | GDP shipment reconciliation practice; 3PL dispatch workflow observation (pending VOC).
VOC Reference | VOC-### (pending) — theme: "I manage 200 boxes a day; tell me only what needs action."
Strategic Theme | Frictionless daily operations.
Product Family | SCM — professional reusable line.
Lifecycle Stage | Launch
Traceability | forward to UN-ALERT / PRD-STATUS placeholders; ORCH-018 / EM-ENG-012.
Notes | Alert fatigue is the failure mode; needs role-appropriate, action-only signaling.
MR-PER-003 — IT / Data Administrator
Requirement Statement | The market demands that a device and its platform serve an IT/data administrator who must manage user access, data integrity, retention, and system validation across a device fleet under Part 11 and data-governance obligations.
Business Objective | Clear the enterprise IT and validation gate that blocks platform deployment.
Customer Segment | Enterprise pharma/distributor IT and quality-systems functions.
Geographic Applicability | Regulated markets; global enterprise.
Market Priority | High
Importance | High
Competitive Driver | Validated, integrable platforms beat closed proprietary loggers in enterprise procurement.
Business Value | Enables multi-year platform contracts and integration lock-in.
Revenue Impact | Medium-High — gates SaaS platform expansion and integration revenue.
Risk if Unmet | Failed CSV/validation review; security veto; no enterprise rollout.
Evidence Source | Standard — 21 CFR Part 11, GAMP 5, ISPE data-integrity guidance.
Supporting Research | GAMP 5 computerized-system validation roles; data-integrity (ALCOA+) principles.
VOC Reference | VOC-### (pending) — theme: "It must pass our CSV and SSO/access policy."
Strategic Theme | Enterprise-grade data governance.
Product Family | SCM — platform/fleet tier.
Lifecycle Stage | Growth
Traceability | forward to UN-ACCESS / PRD-AUDITTRAIL placeholders; ORCH-018 / EM-ENG-012.
Notes | Coordinates with regulatory auditor (MR-PER-010) on audit-trail evidence.
6.2 Transit & Handling Personas
MR-PER-004 — 3PL Operations Lead
Requirement Statement | The market demands that a device serve a third-party-logistics operations lead who runs high-throughput handling operations and needs the device to be invisible-until-relevant, non-disruptive to scanning/handling, and safe for air/ground transport.
Business Objective | Secure endorsement from 3PLs who influence which monitors shippers standardize on.
Customer Segment | Pharma-focused 3PLs, freight forwarders, cold-chain logistics providers.
Geographic Applicability | Global logistics corridors.
Market Priority | High
Importance | High
Competitive Driver | Transport safety (UN 38.3) and handling neutrality are table stakes for 3PL acceptance.
Business Value | 3PL standardization multiplies device pull across many shippers.
Revenue Impact | High — channel-scale adoption lever.
Risk if Unmet | Device flagged as dangerous goods handling risk or handling nuisance; excluded from 3PL SOPs.
Evidence Source | Standard — UN 38.3, IATA DGR / PCR.
Supporting Research | IATA lithium-battery and perishable-cargo handling rules.
VOC Reference | VOC-### (pending) — theme: "It can't slow the dock or trip a battery-shipping rule."
Strategic Theme | Logistics-network neutrality and transport compliance.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-TRANSPORT / PRD-BATTERY placeholders; ORCH-018 / EM-ENG-012.
Notes | Li-ion transport classification (D3) is decisive for this persona; coordinate with battery/regulatory engineers.
MR-PER-005 — Warehouse Cold-Store Operator
Requirement Statement | The market demands that a device serve a warehouse/cold-store operator working in +2/+8°C rooms and -20°C freezers who must place, retrieve, and verify devices quickly while gloved and in low-dexterity conditions.
Business Objective | Ensure the device is usable at the physical staging points where trips begin and end.
Customer Segment | Distribution-center and cold-store operations teams.
Geographic Applicability | Global; concentrated at distribution hubs.
Market Priority | High
Importance | Medium
Competitive Driver | Glove-friendly, cold-legible interaction differentiates against fiddly loggers.
Business Value | Reduces handling errors that corrupt trip records and drive returns/support cost.
Revenue Impact | Medium — protects margin via lower support and re-work.
Risk if Unmet | Mis-provisioned or mis-read devices; operator workarounds defeat monitoring.
Evidence Source | Judgment/Assumption — cold-store ergonomics; quantitative dexterity thresholds Evidence Pending.
Supporting Research | Cold-environment human-factors literature (gloved operation, OLED legibility at low temp).
VOC Reference | VOC-### (pending) — theme: "I'm gloved and freezing; make it one action."
Strategic Theme | Human factors at temperature.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-HMI / PRD-OLED placeholders; ORCH-018 / EM-ENG-012.
Notes | OLED low-temp legibility (range to -30°C) intersects display-engineering; flag for HMI review.
6.3 Receipt & Point-of-Use Personas
MR-PER-006 — National Immunization Program / EPI Manager
Requirement Statement | The market demands that a device serve a national immunization program (EPI) manager who must accept or reject arriving vaccine consignments and needs an unambiguous, WHO-PQS-aligned trip verdict to make the accept/reject call.
Business Objective | Establish SCM as a trusted instrument within national and PQS-governed programs.
Customer Segment | Ministries of Health, national/subnational EPI, WHO/UNICEF-supported programs.
Geographic Applicability | Global; LMIC emphasis.
Market Priority | Must-have
Importance | Critical
Competitive Driver | WHO PQS alignment and clear excursion verdicts are decisive in program procurement.
Business Value | Program adoption drives large pooled volumes and reference credibility.
Revenue Impact | High — pooled-procurement scale.
Risk if Unmet | Consignments wrongly accepted/rejected; program distrust; PQS non-eligibility.
Evidence Source | Standard — WHO PQS E006; WHO vaccine arrival/excursion decision guidance.
Supporting Research | WHO PQS device specs; EVM assessment criteria.
VOC Reference | VOC-### (pending) — theme: "I need a clear go/no-go, not a data dump."
Strategic Theme | Decision-grade clarity for immunization programs.
Product Family | SCM — value and professional lines.
Lifecycle Stage | Launch
Traceability | forward to UN-VERDICT / PRD-ALARM placeholders; ORCH-018 / EM-ENG-012.
Notes | Accept/reject rule sets should map to configurable, PQS-consistent thresholds; coordinate with regulatory.
MR-PER-007 — Last-Mile Health Worker / Nurse
Requirement Statement | The market demands that a device serve a last-mile health worker or nurse at the point of vaccine administration who has minimal device training and needs to read cold-chain status instantly at the OLED without a phone, network, or manual.
Business Objective | Make the device trustworthy and self-evident to the least-trained, most-remote user.
Customer Segment | Rural health posts, mobile vaccination teams, community health workers.
Geographic Applicability | Global; strong LMIC last-mile emphasis.
Market Priority | Must-have
Importance | Critical
Competitive Driver | On-device, offline-legible status differentiates against app-dependent monitors.
Business Value | Protects patient safety at the point of use and builds field trust that drives program scale.
Revenue Impact | Medium-High — program credibility and reorder driver.
Risk if Unmet | Compromised vaccine administered, or safe vaccine discarded; loss of program confidence.
Evidence Source | Standard/Judgment — WHO last-mile monitoring guidance; literacy/training assumptions Evidence Pending.
Supporting Research | WHO last-mile cold-chain and VVM interpretation practice; low-literacy HMI design.
VOC Reference | VOC-### (pending) — theme: "No signal, no phone — the box itself must tell me."
Strategic Theme | Offline, glanceable point-of-care trust.
Product Family | SCM — value/entry and professional lines.
Lifecycle Stage | Launch
Traceability | forward to UN-OFFLINE-STATUS / PRD-OLED placeholders; ORCH-018 / EM-ENG-012.
Notes | Must function with intermittent/no connectivity (see MR-1300 ENV rural health post); OLED+button is the sole interface here.
6.4 Procurement & Governance Personas
MR-PER-008 — Procurement Officer (GAVI / UNICEF / Pooled Buyer)
Requirement Statement | The market demands that a device serve institutional procurement officers running pooled, tendered buys who require PQS prequalification, total-cost-of-ownership evidence for reuse, and supply assurance at scale.
Business Objective | Qualify SCM into pooled tenders that determine LMIC market access at volume.
Customer Segment | GAVI, UNICEF Supply Division, national procurement, donor-funded programs.
Geographic Applicability | Global; LMIC-directed procurement.
Market Priority | Must-have
Importance | High
Competitive Driver | PQS listing and reuse-based TCO are the gating criteria in institutional tenders.
Business Value | Tender qualification opens the largest single volume channel for the reusable model.
Revenue Impact | High — potential large-volume framework contracts.
Risk if Unmet | Excluded from tenders; ceded LMIC market to PQS-listed competitors.
Evidence Source | Standard/Cited — WHO PQS process; UNICEF/GAVI procurement frameworks.
Supporting Research | UNICEF Supply Division cold-chain tender criteria; GAVI cold-chain equipment platform.
VOC Reference | VOC-### (pending) — theme: "Show me PQS listing and cost-per-trip over the device life."
Strategic Theme | Institutional market access via prequalification and TCO.
Product Family | SCM — value/entry and professional lines.
Lifecycle Stage | Concept
Traceability | forward to UN-PQS / PRD-TCO placeholders; ORCH-018 / EM-ENG-012.
Notes | Reuse economics (D1) must be documented as cost-per-trip evidence; coordinate with cost/finance.
MR-PER-009 — Regulatory / Quality Auditor
Requirement Statement | The market demands that a device and its records serve internal and external auditors/inspectors who must verify data integrity, calibration currency, and excursion handling against GDP and Part 11 without ambiguity.
Business Objective | Make the SCM record the easiest artifact in the audit — a differentiator, not a liability.
Customer Segment | QA auditors, regulatory inspectors, notified/competent authorities, customer auditors.
Geographic Applicability | Regulated markets; global for multinationals.
Market Priority | High
Importance | Critical
Competitive Driver | Tamper-evident, ALCOA+-aligned records win audits and reduce findings.
Business Value | Reduces customers' audit risk, a decisive B2B trust factor.
Revenue Impact | Medium-High — de-risks enterprise commitment; supports premium positioning.
Risk if Unmet | Audit findings; records challenged; product blamed for compliance gaps.
Evidence Source | Standard — EU GDP, 21 CFR Part 11, ALCOA+ data-integrity principles.
Supporting Research | MHRA/FDA data-integrity guidance; GDP inspection expectations.
VOC Reference | VOC-### (pending) — theme: "I need traceable, tamper-evident, exportable records."
Strategic Theme | Audit-grade evidentiary integrity.
Product Family | SCM — professional/platform tiers.
Lifecycle Stage | Launch
Traceability | forward to UN-DATAINTEGRITY / PRD-EXPORT placeholders; ORCH-018 / EM-ENG-012.
Notes | Calibration currency links to field-service technician (MR-PER-011); export links to MR-1500 ECO GDP export.
MR-PER-010 — Executive / Supply-Chain Director
Requirement Statement | The market demands that a device program serve an executive/supply-chain director who is accountable for cold-chain loss, compliance exposure, and network performance and needs fleet-level visibility and ROI evidence, not device-level detail.
Business Objective | Convert the strategic buyer who authorizes fleet-scale, multi-year commitments.
Customer Segment | VP/Director Supply Chain, Quality, and Logistics at pharma/distributors/3PLs.
Geographic Applicability | Global enterprise.
Market Priority | High
Importance | High
Competitive Driver | Fleet analytics and demonstrable loss/ROI reporting differentiate platforms from point loggers.
Business Value | Executive sponsorship secures budget and standardization decisions.
Revenue Impact | High — anchors large, multi-year platform contracts.
Risk if Unmet | No executive sponsor; pilots stall without scaling budget.
Evidence Source | Judgment/Assumption — B2B buying-committee dynamics; ROI figures Evidence Pending.
Supporting Research | Cold-chain loss-cost literature (WHO/IQVIA estimates of wastage); enterprise buying-committee research.
VOC Reference | VOC-### (pending) — theme: "Show me network-wide loss reduction and cost-per-trip."
Strategic Theme | Fleet-level ROI and risk visibility.
Product Family | SCM — platform/fleet tier.
Lifecycle Stage | Growth
Traceability | forward to UN-FLEET-KPI / PRD-DASHBOARD placeholders; ORCH-018 / EM-ENG-012.
Notes | Consumes outputs of MR-1500 ECO fleet management; distinct needs from QA manager (MR-PER-001).
6.5 Sustainment Personas
MR-PER-011 — Field Service / Calibration Technician
Requirement Statement | The market demands that a device serve a field-service/calibration technician responsible for cleaning, recharging, re-provisioning, calibration verification, and pooling of reusable devices between trips at scale.
Business Objective | Make reuse operationally sustainable so the reusable business model actually delivers its promised TCO.
Customer Segment | Distributor/3PL service depots, pooling operators, third-party calibration labs.
Geographic Applicability | Global; concentrated at return/pooling hubs.
Market Priority | High
Importance | High
Competitive Driver | Serviceable, calibration-manageable fleets are what make reuse credible versus single-use.
Business Value | Turnaround efficiency directly determines device utilization and margin per unit.
Revenue Impact | Medium-High — drives fleet asset productivity and service-revenue attach.
Risk if Unmet | Slow/uncertain returns; out-of-cal devices in circulation; reuse economics collapse.
Evidence Source | Standard/Judgment — calibration traceability (ISO/IEC 17025 concepts); throughput figures Evidence Pending.
Supporting Research | Metrology/calibration-interval practice; reverse-logistics turnaround benchmarks (pending).
VOC Reference | VOC-### (pending) — theme: "I need to clean, recharge, recal, and redeploy fast and verifiably."
Strategic Theme | Sustainable reuse and fleet trustworthiness.
Product Family | SCM — all reusable SKUs; core to platform commonality.
Lifecycle Stage | Growth
Traceability | forward to UN-CALIB / PRD-SERVICE placeholders; ORCH-018 / EM-ENG-012.
Notes | Cleaning-between-trips and charge-inhibit-below-0°C (D3) intersect materials and battery engineering; flag both.
Personas Summary Table
Persona | Role/Context | Training/Qualification | Key needs
MR-PER-001 Pharma Cold-Chain / QA Manager | Provisions & releases shipments; owns compliance record | GDP-trained quality professional; Responsible Person-level | Defensible, audit-ready records; fast provisioning that doesn't block dispatch
MR-PER-002 Distributor Logistics Coordinator | Books/tracks many concurrent shipments | Logistics/ops trained; not a data analyst | Action-only alerts; status at a glance; no manual data handling
MR-PER-003 IT / Data Administrator | Manages fleet platform, access, data integrity | IT/CSV/GAMP-literate | Part 11 access control, validation, retention, integration
MR-PER-004 3PL Operations Lead | High-throughput handling; never opens box | Logistics/DG-handling trained | Transport compliance (UN 38.3); handling-neutral; no dock slowdown
MR-PER-005 Warehouse Cold-Store Operator | Places/retrieves devices in +2/+8°C & -20°C | Warehouse SOP-trained; gloved | Glove-friendly, cold-legible, single-action interaction
MR-PER-006 EPI / Immunization Program Manager | Accepts/rejects arriving consignments | PQS/EVM-trained program staff | Unambiguous PQS-aligned accept/reject verdict
MR-PER-007 Last-Mile Health Worker / Nurse | Reads status at point of administration | Minimal device training; clinical | Offline, glanceable OLED status; no phone/network needed
MR-PER-008 Procurement Officer (GAVI/UNICEF) | Runs pooled tenders | Procurement professional | PQS listing; reuse TCO/cost-per-trip; supply assurance
MR-PER-009 Regulatory / Quality Auditor | Verifies integrity, calibration, excursions | GDP/Part 11 inspector-level | Tamper-evident, ALCOA+ traceable, exportable records
MR-PER-010 Executive / Supply-Chain Director | Accountable for loss, compliance, network KPIs | Senior leadership | Fleet-level ROI, loss reduction, standardization evidence
MR-PER-011 Field Service / Calibration Technician | Cleans, recharges, recals, re-pools devices | Metrology/service trained | Fast verifiable turnaround; calibration management; serviceability
7. Evidence & Validation Notes
Evidence-backed anchors | Persona accountabilities are anchored to real frameworks (WHO PQS E006, EU GDP personnel clauses, 21 CFR Part 11, IATA DGR/PCR, UNICEF/GAVI procurement). These are cited as Standard.
Pending validation | Quantitative persona attributes (shipment volumes per coordinator, service throughput per technician, training-hour baselines, ROI figures) are marked Judgment/Assumption or Evidence Pending and must be confirmed via a VOC program (contextual interviews and field observation across at least EU, US, and one LMIC corridor).
Recommended validation methods | contextual inquiry at distributor dispatch and cold-store; ride-along with last-mile teams; audit-simulation with QA/regulatory personas; service-depot time-and-motion study for MR-PER-011.
Confidence | High that these persona classes exist and matter; Medium on relative weighting until VOC quantifies frequency and influence per segment/geography.
8. Traceability Placeholders
Upstream | ORCH-018 (orchestration convening), EM-ENG-012 (product-strategy mandate), anchor facts D1–D4.
Peer modules | MR-1300 ENV (environments these personas operate in), MR-1400 PORT (SKUs serving each persona), MR-1500 ECO (software each persona uses).
Downstream (placeholders) | UN-QA-REC, UN-ALERT, UN-ACCESS, UN-TRANSPORT, UN-HMI, UN-VERDICT, UN-OFFLINE-STATUS, UN-PQS, UN-DATAINTEGRITY, UN-FLEET-KPI, UN-CALIB → PRD-PROVISION, PRD-STATUS, PRD-AUDITTRAIL, PRD-BATTERY, PRD-OLED, PRD-ALARM, PRD-EXPORT, PRD-TCO, PRD-DASHBOARD, PRD-SERVICE.
VOC linkage | VOC-### identifiers pending; each MR-PER block to be back-filled once VOC study completes.
MR-1300 — Use Environments (ENV)
Master Market Requirements Library — Smart Cold Chain Monitor (SCM). Traceability: this module (code ENV) feeds forward to UN-/PRD- placeholders and is convened under ORCH-018 / EM-ENG-012.
1. Purpose
This module defines the physical, thermal, mechanical, and connectivity environments through which a reusable Smart Cold Chain Monitor (SCM) travels across a vaccine shipment's life, and expresses each environment as a market requirement — what the market demands the device withstand, operate within, and communicate through. It exists so that downstream system and product requirements can be traced to a named environment with characterized conditions, rather than to a generic "operating range." The SCM's stated envelope (endurance 90 days; -30°C to +60°C; LTE-M primary with Wi-Fi/BLE and GPS) is only meaningful when mapped to the concrete places it must survive.
2. Scope
In scope: Every distinct use environment across provisioning, multi-modal transit, staging, receipt, point-of-use, and reverse logistics — including thermal extremes, mechanical/shock exposure, connectivity conditions, and regulatory/handling contexts of each. Includes the use-environments table and environment-as-requirement blocks.
Out of scope: The people operating in these environments (MR-1200 PER); SKU strategy that addresses environment variants (MR-1400 PORT); connectivity software/partners (MR-1500 ECO). Detailed IEC 60068/60529 test-level specification belongs to the environmental-qualification module downstream; this module states the market-level environmental demand.
3. Applicable Standards & References
• IEC 60529 (IP ratings) — ingress protection for dust/water in cold-store, tarmac, and ocean-freight exposure.
• IEC 60068 — environmental test methods (cold, dry heat, damp heat, vibration, shock, temperature cycling).
• ISTA / ASTM D4169 — distribution/transport simulation profiles (shock, vibration, drop) for multi-modal shipping.
• IATA PCR/TCR & DGR — air-cargo temperature-control and dangerous-goods (lithium battery) handling conditions.
• WHO PQS E006 / WHO cold-chain guidance — passive cold box / shipper and last-mile conditions; temperature-monitoring device environmental expectations.
• CISPR / IEC 61000, FCC Part 15, CE RED — EMC/radio environments (reefer, aircraft, customs scanning).
• Environmental extremes cited to standards where possible; site-specific values (e.g., tarmac dwell durations, LMIC connectivity gaps) are marked Judgment/Assumption or Evidence Pending in Section 6.
4. Background
A vaccine shipment is a multi-modal journey. A device is sealed inside a passive cold box or active reefer, loaded into an aircraft hold, left on a tropical or arctic tarmac, cleared through customs (where radios may be restricted and packages scanned), stored in +2/+8°C rooms or -20°C freezers, and finally carried the last mile by motorbike or on foot to a rural health post with intermittent power and no reliable network. Some vaccines demand ultra-cold (-70°C dry-ice) contexts that sit at or beyond the device's specified boundary and must be treated explicitly as a limit, not a claim.
The device therefore experiences the widest environmental envelope of almost any consumer-class IoT product: -30°C to +60°C operating, potential -70°C boundary exposure, tropical humidity, condensation on cold-to-warm transitions, transport vibration and drop, and connectivity ranging from LTE-M cellular to complete radio silence. Critically, its most safety-relevant readings happen when it is least connected — inside a sealed box, off-grid, in the last mile. This module makes each environment's demands explicit so the design does not optimize for the easy (connected warehouse) case at the expense of the hard (sealed, cold, offline) one.
5. Requirement Categories
• 5.1 Enclosed thermal-payload environments — inside the shipper/cold box, ultra-cold boundary context.
• 5.2 Transit environments — reefer truck, air cargo hold, ocean freight.
• 5.3 Interchange & staging environments — tarmac/customs, cross-border customs, ambient staging.
• 5.4 Storage environments — warehouse cold room (+2/+8°C) and freezer (-20°C).
• 5.5 Last-mile & point-of-use environments — LMIC motorbike/on-foot transport, rural health post.
6. Detailed Market Requirements
6.1 Enclosed Thermal-Payload Environments
MR-ENV-001 — Inside Insulated Shipper / Passive Cold Box
Requirement Statement | The market demands that a device operate accurately and log continuously while fully enclosed inside an insulated passive shipper or cold box, where it is thermally coupled to the payload, RF-attenuated, and physically inaccessible for the duration of transit.
Business Objective | Guarantee trustworthy monitoring in the device's primary operating location — the sealed payload space.
Customer Segment | All cold-chain shippers; PQS programs; distributors.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Reliable in-box logging with store-and-forward beats monitors that assume constant connectivity.
Business Value | This is the core value moment; failure here voids the product's reason to exist.
Revenue Impact | High — foundational to all sales.
Risk if Unmet | Blind spots inside the box; excursions missed; safety and trust collapse.
Evidence Source | Standard/Judgment — WHO PQS cold-box practice; RF attenuation of insulated enclosures (physics), attenuation magnitude Evidence Pending.
Supporting Research | WHO PQS passive-container specifications; RF propagation through EPS/VIP insulation.
VOC Reference | VOC-### (pending) — theme: "Inside the box is exactly where I can't see, and where it matters most."
Strategic Theme | Trustworthy monitoring in the sealed payload space.
Product Family | SCM (Smart Cold Chain Monitor) — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-LOGGING / UN-STOREFWD / PRD-RF placeholders; ORCH-018 / EM-ENG-012.
Notes | Store-and-forward (D2 cadence upload) is mandatory here; external probe placement vs. payload coupling intersects sensor engineering.
MR-ENV-002 — Ultra-Cold / Dry-Ice (-70°C) Boundary Context
Requirement Statement | The market demands clarity on device behavior in ultra-cold (dry-ice, approx. -70°C) shipments — an environment at or beyond the device's -30°C rated floor — such that the device either safely serves the context via a defined variant or is unambiguously scoped out to prevent misuse.
Business Objective | Address ultra-cold demand honestly, capturing the segment where feasible and protecting the brand where not.
Customer Segment | mRNA/ultra-cold vaccine shippers; specialty pharma; research logistics.
Geographic Applicability | Global; concentrated in developed and specialty corridors.
Market Priority | Medium
Importance | High
Competitive Driver | Ultra-cold-capable monitoring is a distinct, higher-value niche where few reusable devices compete.
Business Value | Optional premium segment; also a liability boundary if left ambiguous.
Revenue Impact | Medium — premium variant upside; risk-avoidance value.
Risk if Unmet | Devices used out of spec at -70°C; false confidence; battery/sensor failure; brand damage.
Evidence Source | Judgment/Assumption — dry-ice sublimation ~ -78.5°C (physics, cited); device behavior below -30°C Evidence Pending.
Supporting Research | Ultra-cold vaccine handling guidance; Li-ion low-temperature performance limits.
VOC Reference | VOC-### (pending) — theme: "Does it work on dry-ice lanes, yes or no?"
Strategic Theme | Honest boundary definition; optional premium reach.
Product Family | SCM — ultra-cold variant note (see MR-1400 PORT).
Lifecycle Stage | Concept
Traceability | forward to UN-TEMPFLOOR / PRD-ULTRACOLD placeholders; ORCH-018 / EM-ENG-012.
Notes | Charge-inhibit-below-0°C (D3) and low-temp Li-ion behavior are decisive; coordinate battery/regulatory. This is a boundary requirement, not a capability claim.
6.2 Transit Environments
MR-ENV-003 — Refrigerated Truck / Reefer
Requirement Statement | The market demands that a device operate through active reefer transport, tolerating sustained cold, defrost cycles, engine/compressor vibration, and the electrically noisy reefer environment while maintaining logging and periodic upload.
Business Objective | Ensure continuous trust across the ground-transport leg that dominates regional distribution.
Customer Segment | Distributors, 3PLs, regional carriers.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | High
Competitive Driver | Robust operation through vibration/EMC and defrost transients is baseline reliability.
Business Value | Protects the highest-frequency transport leg.
Revenue Impact | High — volume leg.
Risk if Unmet | Vibration-induced faults, EMC-driven data loss, missed defrost-related excursions.
Evidence Source | Standard — IEC 60068 vibration; CISPR/IEC 61000 EMC; ISTA transport profiles.
Supporting Research | Reefer defrost-cycle temperature transients; road-transport vibration spectra.
VOC Reference | VOC-### (pending) — theme: "Reefers rattle and cycle; it can't drop data."
Strategic Theme | Ruggedness through active transport.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-VIB / UN-EMC / PRD-MOUNT placeholders; ORCH-018 / EM-ENG-012.
Notes | Defrost transients must be distinguishable from true excursions in alert logic; flag to firmware/algorithm.
MR-ENV-004 — Air Cargo Hold
Requirement Statement | The market demands that a device transit pressurized/unpressurized air-cargo holds in compliance with air-transport rules, tolerating altitude, temperature swings, and mandated radio-silence or transmit-restricted conditions without compromising the logged record.
Business Objective | Enable trusted air freight — the backbone of international vaccine distribution.
Customer Segment | International shippers, freight forwarders, airlines' cargo operations.
Geographic Applicability | Global air corridors.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Air-transport compliance (lithium battery + radio behavior) is a gating requirement.
Business Value | Unlocks the international lane essential to global vaccine supply.
Revenue Impact | High — international volume.
Risk if Unmet | Shipment refused; regulatory violation; data gaps over the longest, coldest leg.
Evidence Source | Standard — IATA DGR (Li battery), IATA PCR/TCR, UN 38.3; RTCA/altitude context.
Supporting Research | IATA lithium-battery transport provisions; cargo-hold thermal/altitude conditions.
VOC Reference | VOC-### (pending) — theme: "It must be flight-legal and still capture the whole leg."
Strategic Theme | Air-transport compliance and continuity.
Product Family | SCM — all reusable SKUs; connectivity variants relevant.
Lifecycle Stage | Launch
Traceability | forward to UN-FLIGHTMODE / UN-TRANSPORT / PRD-RADIO placeholders; ORCH-018 / EM-ENG-012.
Notes | Transmit behavior must be configurable/compliant in-flight; store-and-forward covers the radio-silent window.
MR-ENV-005 — Ocean Freight
Requirement Statement | The market demands that a device endure long-duration ocean-freight legs — extended endurance, salt/humidity exposure, container condensation, and prolonged connectivity loss — while preserving a complete logged record for later upload.
Business Objective | Serve long-haul sea lanes where trip durations approach the endurance envelope.
Customer Segment | International shippers using sea freight; consolidators.
Geographic Applicability | Global maritime corridors; cost-sensitive long-haul.
Market Priority | Medium
Importance | Medium
Competitive Driver | 90-day endurance and offline logging make long sea legs viable versus shorter-life loggers.
Business Value | Extends addressable trips to slow, high-volume, cost-sensitive lanes.
Revenue Impact | Medium — expands lane coverage.
Risk if Unmet | Battery exhaustion mid-voyage; corrosion; multi-week data gaps.
Evidence Source | Standard/Judgment — IEC 60068 damp-heat/salt; endurance 90 days (D-spec); voyage durations Evidence Pending.
Supporting Research | Marine container humidity/condensation profiles; salt-fog corrosion testing.
VOC Reference | VOC-### (pending) — theme: "Sea legs run weeks with no signal — it still has to remember everything."
Strategic Theme | Endurance and integrity over long, offline hauls.
Product Family | SCM — professional reusable line.
Lifecycle Stage | Growth
Traceability | forward to UN-ENDURANCE / UN-CORROSION / PRD-SEAL placeholders; ORCH-018 / EM-ENG-012.
Notes | Condensation on temperature transitions intersects IP rating and sensor accuracy; coordinate mechanical.
6.3 Interchange & Staging Environments
MR-ENV-006 — Airport Tarmac / Customs Dwell
Requirement Statement | The market demands that a device survive unconditioned tarmac and holding-area dwell — direct solar load or arctic cold, precipitation, and handling shocks — during the uncontrolled interchange windows that produce many real-world excursions.
Business Objective | Protect trust during the exposed interchange gaps where control is weakest.
Customer Segment | International shippers; airlines' ground handling; customs-bonded operations.
Geographic Applicability | Global; extremes in tropical and arctic hubs.
Market Priority | High
Importance | High
Competitive Driver | Capturing tarmac/dwell excursions is a key differentiator versus point-only sampling.
Business Value | Detects the excursions most likely to be disputed between parties.
Revenue Impact | Medium-High — evidentiary value drives adoption.
Risk if Unmet | Excursions during dwell missed or mis-attributed; liability disputes.
Evidence Source | Standard/Judgment — IEC 60068 dry heat/solar; IP rain ingress; dwell durations Evidence Pending.
Supporting Research | Tarmac solar-load and cold-soak data; airport handling-shock profiles.
VOC Reference | VOC-### (pending) — theme: "Boxes bake or freeze on the tarmac — that's where we lose product."
Strategic Theme | Coverage of the uncontrolled interchange gap.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-SOLAR / UN-IP / PRD-ENCLOSURE placeholders; ORCH-018 / EM-ENG-012.
Notes | Enclosure solar absorptivity and sensor self-heating intersect thermal/mechanical review.
MR-ENV-007 — Cross-Border Customs (Scanning / Radio Restrictions)
Requirement Statement | The market demands that a device pass through cross-border customs — X-ray/scanning exposure, potential radio restrictions, and extended unpowered holds — without data corruption, security compromise, or unexplained gaps in the chain of custody.
Business Objective | Preserve an unbroken, defensible record across jurisdictional boundaries.
Customer Segment | International shippers; customs brokers; regulated importers.
Geographic Applicability | All cross-border lanes; multi-market (D4).
Market Priority | High
Importance | High
Competitive Driver | Multi-market radio certification (FCC/CE/regional) and scan-tolerance enable border-crossing use.
Business Value | Ensures the record survives the highest-scrutiny, multi-authority segment.
Revenue Impact | Medium-High — international lane enablement.
Risk if Unmet | Radio non-compliance in a market; data gaps at borders undermine GDP chain of custody.
Evidence Source | Standard — FCC Part 15, CE RED, regional radio approvals; customs-hold practice.
Supporting Research | Multi-region radio certification requirements; customs X-ray effects on electronics.
VOC Reference | VOC-### (pending) — theme: "It has to be legal in every country the box passes through."
Strategic Theme | Multi-market compliance and continuity across borders.
Product Family | SCM — regional SKU/certification variants (see MR-1400 PORT).
Lifecycle Stage | Growth
Traceability | forward to UN-RADIOCERT / UN-CUSTODY / PRD-CERT placeholders; ORCH-018 / EM-ENG-012.
Notes | Regional radio approvals tie directly to SKU/regional-variant strategy; coordinate regulatory.
MR-ENV-008 — Ambient Staging / Cross-Dock
Requirement Statement | The market demands that a device operate through ambient (uncontrolled room-temperature) staging and cross-dock windows — where payloads await loading — capturing the short but frequent ambient exposures that erode temperature budget.
Business Objective | Close the everyday gaps between controlled legs where cumulative exposure accrues.
Customer Segment | Distributors, 3PLs, warehouses.
Geographic Applicability | Global.
Market Priority | High
Importance | Medium
Competitive Driver | Continuous cross-dock capture differentiates against interval-only monitors.
Business Value | Surfaces cumulative-exposure risk invisible to coarse sampling.
Revenue Impact | Medium — reinforces continuous-record value proposition.
Risk if Unmet | Cumulative ambient exposure undetected; late-stage excursions missed.
Evidence Source | Judgment/Assumption — cross-dock dwell frequency/duration Evidence Pending.
Supporting Research | Warehouse cross-dock cycle-time studies (pending); cumulative thermal-budget modeling.
VOC Reference | VOC-### (pending) — theme: "It's the ten-minute waits on the dock that add up."
Strategic Theme | Continuous coverage of the mundane gaps.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-CONTINUOUS / PRD-SAMPLING placeholders; ORCH-018 / EM-ENG-012.
Notes | Sampling cadence vs. battery life trade-off intersects power engineering.
6.4 Storage Environments
MR-ENV-009 — Warehouse Cold Room (+2/+8°C) and Freezer (-20°C)
Requirement Statement | The market demands that a device operate and be handled reliably in +2/+8°C cold rooms and -20°C freezers — including condensation on entry/exit, gloved handling, and dense storage — while remaining connectable for provisioning and status checks.
Business Objective | Serve the fixed storage nodes that bracket every transit leg and where fleets are staged.
Customer Segment | Distribution centers, cold-stores, national/regional stores, hospitals.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | High
Competitive Driver | Reliable low-temp operation plus connectivity in metal-dense cold rooms is baseline.
Business Value | Covers the staging and pooling nodes central to reuse operations.
Revenue Impact | High — ubiquitous node.
Risk if Unmet | Condensation faults; -20°C operation failure; connectivity dead zones defeat provisioning.
Evidence Source | Standard — IEC 60068 cold/damp-heat/temperature-cycling; IEC 60529 (condensation-adjacent).
Supporting Research | Cold-room RF dead-zone behavior; condensation cycling reliability data.
VOC Reference | VOC-### (pending) — theme: "In the freezer it fogs up and loses signal."
Strategic Theme | Reliable storage-node operation and connectivity.
Product Family | SCM — all reusable SKUs.
Lifecycle Stage | Launch
Traceability | forward to UN-CONDENSATION / UN-CONNECT / PRD-CONFORMAL placeholders; ORCH-018 / EM-ENG-012.
Notes | Freezer -20°C is well within the -30°C floor; condensation and RF in metal rooms are the real risks. Coordinate mechanical/RF.
6.5 Last-Mile & Point-of-Use Environments
MR-ENV-010 — LMIC Last Mile (Motorbike / On-Foot Transport)
Requirement Statement | The market demands that a device withstand rugged last-mile transport — motorbike vibration/shock, drops, dust, tropical heat and humidity — while continuing to log and to display trustworthy status without external power or reliable network.
Business Objective | Deliver dependable monitoring where the cold chain most often breaks and matters most.
Customer Segment | EPI programs, NGOs, community health teams, rural distributors.
Geographic Applicability | LMIC last-mile; rural corridors globally.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Ruggedized, offline-capable last-mile monitoring is a decisive PQS/program differentiator.
Business Value | Protects vaccine integrity at the point where wastage is highest.
Revenue Impact | High — core to institutional/pooled volume.
Risk if Unmet | Physical failure or blackout in the last mile; wasted vaccine or unsafe administration.
Evidence Source | Standard/Judgment — ISTA/IEC 60068 shock/vibration/drop; WHO last-mile guidance; road/dust profiles Evidence Pending.
Supporting Research | WHO/UNICEF last-mile cold-chain studies; motorbike vibration/drop spectra.
VOC Reference | VOC-### (pending) — theme: "It rides a motorbike on dirt roads and still has to work."
Strategic Theme | Rugged, offline last-mile dependability.
Product Family | SCM — value/entry and professional lines.
Lifecycle Stage | Launch
Traceability | forward to UN-DROP / UN-IPDUST / PRD-RUGGED placeholders; ORCH-018 / EM-ENG-012.
Notes | Shock sensing (accelerometer) doubles as data here; drop/IP rating intersects mechanical. Pairs with MR-PER-007.
MR-ENV-011 — Rural Health Post (Intermittent Power & Connectivity)
Requirement Statement | The market demands that a device function at a rural health post with unreliable grid power and intermittent or absent cellular coverage — providing offline on-device status, tolerant recharge behavior, and deferred upload when connectivity returns.
Business Objective | Make the device fully useful at the terminal node even when nothing else works.
Customer Segment | Rural/remote health posts; mobile clinics; off-grid facilities.
Geographic Applicability | LMIC and remote regions globally.
Market Priority | Must-have
Importance | Critical
Competitive Driver | True offline-first operation with store-and-forward differentiates from always-connected assumptions.
Business Value | Ensures point-of-care trust and data eventual-delivery at the hardest node.
Revenue Impact | Medium-High — program credibility and reorder driver.
Risk if Unmet | Device unusable off-grid; data never uploads; last-mile decisions made blind.
Evidence Source | Standard/Judgment — LTE-M coverage limits; USB-C recharge (D-spec); charge-inhibit-below-0°C (D3); local power quality Evidence Pending.
Supporting Research | LMIC connectivity/electrification statistics (ITU/World Bank); store-and-forward IoT design.
VOC Reference | VOC-### (pending) — theme: "No grid, no bars — it still has to tell me and remember."
Strategic Theme | Offline-first resilience at the terminal node.
Product Family | SCM — value/entry and professional lines.
Lifecycle Stage | Launch
Traceability | forward to UN-OFFLINEFIRST / UN-CHARGE / PRD-RECHARGE placeholders; ORCH-018 / EM-ENG-012.
Notes | Recharge from unstable/solar power and charge-inhibit-below-0°C are decisive; coordinate battery/power engineering. Pairs with MR-PER-007.
Use Environments Table
Environment | Description / Conditions | Implications for the Device
MR-ENV-001 Inside insulated shipper / cold box | Sealed passive container; thermally coupled to payload; RF-attenuated; inaccessible | Continuous offline logging + store-and-forward; RF strategy for enclosed operation
MR-ENV-002 Ultra-cold / dry-ice (-70°C) boundary | Approx. -78.5°C sublimation; below -30°C rated floor | Defined variant or explicit scope-out; low-temp Li-ion & charge-inhibit limits
MR-ENV-003 Refrigerated truck / reefer | Sustained cold; defrost cycles; compressor vibration; EMC-noisy | Vibration/EMC robustness; defrost vs. excursion discrimination
MR-ENV-004 Air cargo hold | Altitude; temp swings; mandated radio restrictions; Li-battery rules | Flight-compliant radio behavior; UN 38.3; store-and-forward over radio-silent leg
MR-ENV-005 Ocean freight | Weeks-long; salt/humidity; condensation; long connectivity loss | 90-day endurance; corrosion/IP; large offline buffer
MR-ENV-006 Airport tarmac / customs dwell | Solar load or cold-soak; rain; handling shock; uncontrolled | Solar/thermal survivability; IP rain; capture dwell excursions
MR-ENV-007 Cross-border customs | X-ray scanning; radio restrictions; unpowered holds; multi-jurisdiction | Multi-market radio certification; scan tolerance; unbroken custody record
MR-ENV-008 Ambient staging / cross-dock | Uncontrolled room temp; short frequent dwells | Continuous sampling to catch cumulative exposure
MR-ENV-009 Warehouse cold room / freezer | +2/+8°C and -20°C; condensation; gloved; metal-dense | Low-temp operation; conformal coating; connectivity in RF dead zones
MR-ENV-010 LMIC last mile (motorbike/on-foot) | Vibration, drops, dust, tropical heat/humidity; off-grid | Rugged/drop/IP; offline logging; on-device status
MR-ENV-011 Rural health post | Intermittent grid/solar power; no/patchy cellular | Offline-first; tolerant recharge; deferred upload; charge-inhibit-below-0°C
7. Evidence & Validation Notes
Evidence-backed anchors | Thermal, mechanical, ingress, EMC, and transport conditions are anchored to real standards (IEC 60068, IEC 60529, ISTA/ASTM D4169, IATA PCR/TCR/DGR, UN 38.3, FCC/CE RED, WHO PQS). Dry-ice sublimation (~ -78.5°C) is a cited physical constant.
Pending validation | Site-specific magnitudes (tarmac dwell durations and solar loads, LMIC connectivity/electrification gaps, cross-dock cycle frequency, ocean-voyage durations, RF attenuation inside shippers, cold-room dead-zone maps) are Judgment/Assumption or Evidence Pending and require field instrumentation studies.
Recommended validation methods | instrumented "shadow shipment" trials across at least one intercontinental air lane, one reefer ground lane, one ocean lane, and one LMIC last-mile corridor; environmental chamber testing to IEC 60068 profiles; RF mapping inside representative shippers and cold rooms.
Confidence | High that these environments and their qualitative demands are correct; Medium on quantitative thresholds pending shadow-shipment data.
8. Traceability Placeholders
Upstream | ORCH-018, EM-ENG-012; anchor facts D1–D4 and endurance/range specs.
Peer modules | MR-1200 PER (operators in each environment), MR-1400 PORT (variants addressing environments), MR-1500 ECO (connectivity/upload behavior).
Downstream (placeholders) | UN-LOGGING, UN-STOREFWD, UN-TEMPFLOOR, UN-VIB, UN-EMC, UN-FLIGHTMODE, UN-TRANSPORT, UN-ENDURANCE, UN-CORROSION, UN-SOLAR, UN-IP, UN-RADIOCERT, UN-CUSTODY, UN-CONTINUOUS, UN-CONDENSATION, UN-CONNECT, UN-DROP, UN-IPDUST, UN-OFFLINEFIRST, UN-CHARGE → PRD-RF, PRD-ULTRACOLD, PRD-MOUNT, PRD-RADIO, PRD-SEAL, PRD-ENCLOSURE, PRD-CERT, PRD-SAMPLING, PRD-CONFORMAL, PRD-RUGGED, PRD-RECHARGE.
VOC linkage | VOC-### identifiers pending; to be back-filled from shadow-shipment and field studies.
MR-1400 — Product Portfolio (PORT)
Master Market Requirements Library — Smart Cold Chain Monitor (SCM). Traceability: this module (code PORT) feeds forward to UN-/PRD- placeholders and is convened under ORCH-018 / EM-ENG-012.
1. Purpose
This module defines the market-level portfolio strategy for the Smart Cold Chain Monitor (SCM): the set of product variants, tiers, and configuration axes the market demands so that a single reusable platform can address segments ranging from cost-sensitive LMIC pooled procurement to premium regulated pharma and ultra-cold niches. It expresses portfolio structure as market demand — what configurations the market requires to be served — not as bill-of-materials or engineering architecture. It exists so downstream product-line requirements can trace each SKU or variant back to a segment-level market need and a deliberate commonality strategy.
2. Scope
In scope: Portfolio tiering (entry/value vs. professional reusable), the ultra-cold variant boundary, connectivity and sensor/probe configuration axes, the single-use-vs-reusable line strategy, platform/firmware commonality, regional certification variants, and roadmap alignment across lifecycle stages.
Out of scope: Personas (MR-1200 PER), environments (MR-1300 ENV), and ecosystem/software strategy (MR-1500 ECO). Detailed pricing, COGS targets, and engineering architecture belong to downstream PRD and business-case modules; this module states the market demand for portfolio breadth and commonality.
3. Applicable Standards & References
• WHO PQS E006 — device categories eligible for prequalification; drives value-tier and PQS-listed SKU definition.
• EU GDP / FDA 21 CFR Part 11 — professional-tier compliance features that separate regulated SKUs from value SKUs.
• UN 38.3 / IATA DGR — Li-ion transport classification applying across all reusable variants.
• FCC Part 15 / CE RED / regional radio approvals — basis for regional SKU/certification variants.
• IEC 62368-1, IEC 60529, IEC 60068 — safety, ingress, and environmental grading that differentiate rugged/ultra-cold variants.
• UNICEF/GAVI procurement frameworks — pooled-tender criteria shaping the value-line and reuse-TCO strategy.
• Portfolio quantitative targets (price points, volume mix, margin) are Judgment/Assumption or Evidence Pending; standards cited are real frameworks.
4. Background
The SCM is a reusable, rechargeable, multi-trip platform (D1) serving a multi-market world (WHO PQS + EU GDP + US FDA + air/ground transport, D4). No single fixed device can efficiently serve both a GAVI-funded LMIC pooled buyer optimizing cost-per-trip and a regulated EU pharma distributor optimizing audit defensibility, nor both an LTE-M-only ground lane and a multi-radio international lane, nor a +2/+8°C vaccine and an ultra-cold mRNA product. The market therefore demands a portfolio — but the reuse business model demands that the portfolio share a common platform, firmware base, and service/pooling model, or the operational cost of managing many variants destroys the reuse economics.
The central strategic tension in this module is breadth versus commonality: enough variants to win each segment, few enough shared platforms to keep the reusable fleet serviceable, calibratable, and OTA-maintainable at scale. This module frames each portfolio axis as a market demand and explicitly ties it to the commonality requirement (MR-PORT-007) that keeps the reusable model viable.
5. Requirement Categories
• 5.1 Tier structure — entry/value SKU, professional reusable SKU.
• 5.2 Capability variants — ultra-cold boundary variant, connectivity variants, probe/sensor variants.
• 5.3 Line strategy — single-use vs. reusable positioning.
• 5.4 Platform discipline — firmware/platform commonality, regional certification variants.
• 5.5 Roadmap — lifecycle-staged portfolio evolution.
6. Detailed Market Requirements
6.1 Tier Structure
MR-PORT-001 — Entry / Value Reusable SKU
Requirement Statement | The market demands a cost-optimized reusable SKU that meets essential WHO-PQS monitoring needs at the lowest defensible cost-per-trip for price-sensitive, pooled, and LMIC buyers.
Business Objective | Win pooled/institutional tenders and high-volume LMIC deployment with a credible low-TCO offer.
Customer Segment | GAVI/UNICEF pooled buyers, EPI programs, cost-sensitive distributors.
Geographic Applicability | Global; LMIC emphasis.
Market Priority | Must-have
Importance | High
Competitive Driver | Cost-per-trip and PQS eligibility are decisive against single-use loggers and value competitors.
Business Value | Anchors volume and market access; reference base for the platform.
Revenue Impact | High — largest unit volume channel.
Risk if Unmet | Priced out of pooled tenders; ceded LMIC scale.
Evidence Source | Standard/Judgment — WHO PQS E006 minimum specs; price/TCO targets Evidence Pending.
Supporting Research | UNICEF cold-chain tender criteria; single-use-vs-reusable TCO modeling.
VOC Reference | VOC-### (pending) — theme: "Cheapest trustworthy cost-per-trip wins the tender."
Strategic Theme | Volume and access via low-TCO reuse.
Product Family | SCM (Smart Cold Chain Monitor) — value/entry line.
Lifecycle Stage | Launch
Traceability | forward to UN-VALUESPEC / PRD-VALUESKU placeholders; ORCH-018 / EM-ENG-012.
Notes | Must share platform with professional tier (MR-PORT-007) to keep reuse-service costs low. Serves MR-PER-006/007/008.
MR-PORT-002 — Professional Reusable SKU
Requirement Statement | The market demands a full-featured professional reusable SKU delivering GDP/Part 11-grade records, richer sensing, and multi-radio connectivity for regulated pharma, distributors, and 3PLs.
Business Objective | Capture premium regulated-market revenue and anchor recurring platform/SaaS attach.
Customer Segment | Pharma manufacturers, GDP distributors, cold-chain 3PLs.
Geographic Applicability | EU/US/regulated markets; global multinationals.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Audit-grade compliance and multi-sensor coverage differentiate against basic loggers.
Business Value | Highest margin per unit; drives enterprise platform contracts.
Revenue Impact | High — premium ASP and SaaS anchor.
Risk if Unmet | Loss of the profitable regulated segment to established data-logger vendors.
Evidence Source | Standard — EU GDP, 21 CFR Part 11; feature scope evidence-backed, targets Evidence Pending.
Supporting Research | Regulated cold-chain monitoring competitive landscape; GDP feature expectations.
VOC Reference | VOC-### (pending) — theme: "Give me the audit-proof, fully-sensed device."
Strategic Theme | Premium compliance-grade positioning.
Product Family | SCM — professional reusable line.
Lifecycle Stage | Launch
Traceability | forward to UN-PROSPEC / PRD-PROSKU placeholders; ORCH-018 / EM-ENG-012.
Notes | Serves MR-PER-001/002/004/009. Differentiation from value tier should be feature/firmware-gated on a common platform where possible.
6.2 Capability Variants
MR-PORT-003 — Ultra-Cold / -70°C Variant (Boundary Note)
Requirement Statement | The market demands either a defined ultra-cold-capable variant or an explicit scope-out for -70°C dry-ice lanes, so that ultra-cold demand is served or refused deliberately rather than by accidental out-of-spec use.
Business Objective | Address the ultra-cold niche where viable and protect the brand from misuse where not.
Customer Segment | mRNA/ultra-cold vaccine shippers; specialty pharma; research logistics.
Geographic Applicability | Developed/specialty corridors globally.
Market Priority | Medium
Importance | High
Competitive Driver | Ultra-cold-capable reusable monitoring is a scarce, higher-value niche.
Business Value | Optional premium upside; risk mitigation on the -30°C boundary.
Revenue Impact | Medium — premium niche.
Risk if Unmet | Out-of-spec use at -70°C; battery/sensor failure; brand and safety exposure.
Evidence Source | Judgment/Assumption — dry-ice ~ -78.5°C (physics, cited); sub -30°C device behavior Evidence Pending.
Supporting Research | Ultra-cold vaccine handling guidance; low-temp Li-ion and sensor limits.
VOC Reference | VOC-### (pending) — theme: "Do you cover dry-ice lanes or not?"
Strategic Theme | Deliberate boundary management; optional premium reach.
Product Family | SCM — ultra-cold variant note.
Lifecycle Stage | Concept
Traceability | forward to UN-ULTRACOLDSPEC / PRD-ULTRACOLD placeholders; ORCH-018 / EM-ENG-012.
Notes | Directly linked to MR-ENV-002 and charge-inhibit-below-0°C (D3). Decision required before any ultra-cold marketing claim.
MR-PORT-004 — Connectivity Variants (LTE-M vs. Multi-Radio)
Requirement Statement | The market demands selectable connectivity configurations — from LTE-M-primary economy to full multi-radio (LTE-M + Wi-Fi + BLE + GPS) — so buyers match radio capability and cost to their lane and coverage profile.
Business Objective | Right-size connectivity cost per segment while maximizing lane coverage where it is valued.
Customer Segment | All segments; multinational and international-lane buyers prioritize multi-radio.
Geographic Applicability | Global; region-dependent LTE-M availability.
Market Priority | High
Importance | High
Competitive Driver | Configurable connectivity avoids overpaying for radios on simple lanes and underserving complex ones.
Business Value | Broadens addressable lanes and optimizes per-unit cost/margin.
Revenue Impact | Medium-High — expands lane coverage and configurability upsell.
Risk if Unmet | Either overpriced simple-lane SKUs or coverage gaps on complex lanes.
Evidence Source | Standard/Judgment — FCC/CE RED radio approvals; regional LTE-M coverage Evidence Pending.
Supporting Research | Regional LTE-M/eSIM coverage maps; carrier roaming availability.
VOC Reference | VOC-### (pending) — theme: "I need GPS+cellular internationally but just cellular domestically."
Strategic Theme | Cost-matched connectivity breadth.
Product Family | SCM — connectivity variants across value and professional lines.
Lifecycle Stage | Growth
Traceability | forward to UN-RADIOCONFIG / PRD-CONNECTIVITY placeholders; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-1500 ECO carrier/eSIM strategy and MR-ENV-004/007 radio compliance. Prefer common RF platform with populate-on-demand modules.
MR-PORT-005 — Probe / External-Sensor Variants
Requirement Statement | The market demands external-probe and sensor configuration options (external temperature probe, humidity, shock/accelerometer emphasis) so the device can be placed on the payload while sensing the load's true condition.
Business Objective | Match sensing configuration to product and packaging so the record reflects the vaccine, not the device shell.
Customer Segment | Pharma/distributors with diverse packaging; specialty and high-value loads.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | External-probe accuracy on the payload differentiates against single-point internal sensors.
Business Value | Improves record fidelity and defensibility for high-value shipments.
Revenue Impact | Medium-High — configuration upsell and fidelity-driven adoption.
Risk if Unmet | Device-shell readings misrepresent payload; disputed or invalid records.
Evidence Source | Standard/Judgment — probe placement affects accuracy (metrology); accuracy targets Evidence Pending.
Supporting Research | Temperature-mapping/probe-placement practice in GDP validation.
VOC Reference | VOC-### (pending) — theme: "I need to sense the vial, not the case."
Strategic Theme | Payload-true sensing fidelity.
Product Family | SCM — probe/sensor variants, primarily professional line.
Lifecycle Stage | Launch
Traceability | forward to UN-PROBE / PRD-SENSOR placeholders; ORCH-018 / EM-ENG-012.
Notes | External probe (int+ext, D-spec) intersects sensor engineering and calibration (MR-PER-011). Serves MR-ENV-001.
6.3 Line Strategy
MR-PORT-006 — Single-Use vs. Reusable Line Strategy
Requirement Statement | The market demands a clear positioning of the reusable SCM against entrenched single-use loggers, including a migration path and TCO evidence that justifies switching to reuse without a single-use fallback where reverse logistics are infeasible.
Business Objective | Convert single-use incumbency by proving reuse economics while acknowledging where single-use still fits.
Customer Segment | Distributors and programs currently standardized on single-use loggers.
Geographic Applicability | Global; reuse strongest where reverse logistics exist.
Market Priority | High
Importance | High
Competitive Driver | Sustainability and cost-per-trip advantages of reuse versus single-use disposables.
Business Value | Defines the core disruptive value proposition and adoption narrative.
Revenue Impact | High — determines the shift from consumable to durable-plus-SaaS revenue.
Risk if Unmet | Reuse value unproven; buyers default to familiar single-use; stranded fleet risk where returns fail.
Evidence Source | Judgment/Assumption — reuse TCO and reverse-logistics feasibility Evidence Pending.
Supporting Research | Reusable-vs-single-use cold-chain TCO and e-waste/sustainability studies.
VOC Reference | VOC-### (pending) — theme: "Reuse only pays if I can actually get devices back."
Strategic Theme | Reuse-led disruption with honest applicability limits.
Product Family | SCM — reusable line strategy (with single-use-context acknowledgment).
Lifecycle Stage | Concept
Traceability | forward to UN-REUSEMODEL / PRD-REVLOGISTICS placeholders; ORCH-018 / EM-ENG-012.
Notes | Reverse logistics, cleaning, re-provisioning, pooling (D1) are prerequisites; ties to MR-PER-011 and MR-1500 ECO fleet management.
6.4 Platform Discipline
MR-PORT-007 — Platform / Firmware Commonality
Requirement Statement | The market (via the reuse business model) demands that all SCM variants share a common hardware platform, firmware base, and OTA/service model, so the reusable fleet remains serviceable, calibratable, and updatable at scale across SKUs.
Business Objective | Preserve reuse economics and OTA maintainability by preventing variant proliferation from fragmenting the fleet.
Customer Segment | Internal (product/ops) and fleet buyers who depend on consistent behavior across a mixed fleet.
Geographic Applicability | Global; all SKUs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | A common platform enables faster OTA, uniform compliance, and lower service cost than fragmented product lines.
Business Value | Multiplies engineering/service investment across all SKUs; foundation of margin.
Revenue Impact | High — protects margin and speed across the entire portfolio.
Risk if Unmet | Variant sprawl inflates service/calibration/OTA cost and breaks reuse economics.
Evidence Source | Judgment/Assumption — platform-strategy best practice; cost-of-fragmentation Evidence Pending.
Supporting Research | Product-platform/modularity literature; OTA fleet-management cost models.
VOC Reference | VOC-### (pending) — theme (internal/strategic): "one platform, many configurations."
Strategic Theme | Platform commonality as the enabler of reuse at scale.
Product Family | SCM — platform/firmware core across all lines.
Lifecycle Stage | Concept
Traceability | forward to UN-PLATFORM / UN-OTA / PRD-COMMONALITY placeholders; ORCH-018 / EM-ENG-012.
Notes | The gating constraint for MR-PORT-001..005; feature-gating preferred over separate hardware where feasible. Ties to MR-1500 ECO OTA/fleet.
MR-PORT-008 — Regional SKU / Certification Variants
Requirement Statement | The market demands regional configuration variants that satisfy each market's radio, safety, and regulatory approvals (FCC/CE/regional RED equivalents, LTE-M bands, PQS/GDP/FDA expectations) while preserving platform commonality.
Business Objective | Enable lawful, certified sale and cross-border operation in each target market with minimal variant overhead.
Customer Segment | Multinational and cross-border shippers; regional distributors; importers.
Geographic Applicability | All target markets (D4).
Market Priority | Must-have
Importance | High
Competitive Driver | Broad, current certifications enable border-crossing lanes competitors can't legally serve.
Business Value | Unlocks market access and cross-border continuity.
Revenue Impact | High — gates entry to each regulated market.
Risk if Unmet | Illegal to sell/operate in a market; cross-border lanes blocked; recalls.
Evidence Source | Standard — FCC Part 15, CE RED, regional radio/safety approvals; PQS/GDP/FDA frameworks.
Supporting Research | Regional radio band plans; multi-market certification roadmaps.
VOC Reference | VOC-### (pending) — theme: "It must be certified everywhere the box travels."
Strategic Theme | Certified multi-market reach on a common platform.
Product Family | SCM — regional certification variants across lines.
Lifecycle Stage | Growth
Traceability | forward to UN-REGIONCERT / PRD-CERTVARIANT placeholders; ORCH-018 / EM-ENG-012.
Notes | Ties directly to MR-ENV-007 cross-border customs and MR-PORT-004 connectivity variants.
6.5 Roadmap
MR-PORT-009 — Roadmap Alignment Across Lifecycle
Requirement Statement | The market demands a coherent portfolio roadmap that sequences value and professional launch, connectivity/regional expansion, and any ultra-cold decision, so buyers can commit to a fleet knowing the platform will evolve without stranding their investment.
Business Objective | Give fleet buyers confidence to standardize now by making the evolution path explicit and non-disruptive.
Customer Segment | Fleet buyers, executives (MR-PER-010), procurement (MR-PER-008).
Geographic Applicability | Global.
Market Priority | High
Importance | Medium
Competitive Driver | A credible, commonality-preserving roadmap reduces buyer switching risk versus one-off products.
Business Value | De-risks large fleet commitments and supports multi-year contracts.
Revenue Impact | Medium-High — enables long-term standardization deals.
Risk if Unmet | Buyers delay standardization fearing obsolescence; fragmented ad-hoc purchasing.
Evidence Source | Judgment/Assumption — roadmap strategy; sequencing timing Evidence Pending.
Supporting Research | Platform-roadmap and technology-adoption lifecycle practice.
VOC Reference | VOC-### (pending) — theme: "Will today's fleet still be supported and extensible in three years?"
Strategic Theme | Investment-protecting roadmap coherence.
Product Family | SCM — portfolio-wide roadmap.
Lifecycle Stage | Sustaining
Traceability | forward to UN-ROADMAP / PRD-LIFECYCLE placeholders; ORCH-018 / EM-ENG-012.
Notes | Roadmap must honor platform commonality (MR-PORT-007) and OTA (MR-1500 ECO) so upgrades reach installed fleets.
7. Evidence & Validation Notes
Evidence-backed anchors | Tier and variant compliance features are anchored to real frameworks (WHO PQS E006, EU GDP, 21 CFR Part 11, UN 38.3, FCC/CE RED, IEC 62368-1/60529/60068, UNICEF/GAVI procurement). Dry-ice temperature is a cited constant.
Pending validation | All price points, cost-per-trip and TCO figures, volume mix by tier, reverse-logistics feasibility rates, cost-of-variant-fragmentation, and regional coverage specifics are Judgment/Assumption or Evidence Pending. No precise financial numbers are asserted as fact.
Recommended validation methods | segment-level willingness-to-pay and TCO studies; pooled-tender requirement analysis (UNICEF/GAVI); reverse-logistics feasibility assessment by corridor; platform-cost modeling for feature-gating vs. hardware variants; regional certification gap analysis.
Confidence | High that these portfolio axes are the right ones; Medium on tier boundaries and variant scope pending WTP/TCO and tender-requirement data; ultra-cold inclusion explicitly undecided (Concept).
8. Traceability Placeholders
Upstream | ORCH-018, EM-ENG-012; anchor facts D1–D4.
Peer modules | MR-1200 PER (segments each SKU serves), MR-1300 ENV (environments variants address), MR-1500 ECO (platform/OTA/fleet enabling commonality).
Downstream (placeholders) | UN-VALUESPEC, UN-PROSPEC, UN-ULTRACOLDSPEC, UN-RADIOCONFIG, UN-PROBE, UN-REUSEMODEL, UN-PLATFORM, UN-OTA, UN-REGIONCERT, UN-ROADMAP → PRD-VALUESKU, PRD-PROSKU, PRD-ULTRACOLD, PRD-CONNECTIVITY, PRD-SENSOR, PRD-REVLOGISTICS, PRD-COMMONALITY, PRD-CERTVARIANT, PRD-LIFECYCLE.
VOC linkage | VOC-### identifiers pending; to be back-filled from segment WTP/TCO and tender-requirement studies.
MR-1500 — Ecosystem Strategy (ECO)
Master Market Requirements Library — Smart Cold Chain Monitor (SCM). Traceability: this module (code ECO) feeds forward to UN-/PRD- placeholders and is convened under ORCH-018 / EM-ENG-012.
1. Purpose
This module defines the market-level ecosystem strategy for the Smart Cold Chain Monitor (SCM): the software, connectivity, integration, data-sharing, and partner relationships that surround the device and turn it from a data logger into a networked cold-chain platform. It expresses each ecosystem element as a market demand — what the market requires the surrounding system to deliver — not as software architecture. A reusable device (D1) with cadence upload and OTA (D2) is only as valuable as the cloud, apps, integrations, and partnerships that provision it, deliver its data to the systems of record, keep its fleet managed, and connect it across carriers and borders. This module exists so downstream software and integration requirements can trace to a named ecosystem need.
2. Scope
In scope: Cloud dashboard, mobile/BLE provisioning app, open APIs and integration to TMS/WMS/ERP/pharma QMS, GDP audit-record export, third-party 3PL platform integration, developer ecosystem, fleet/device management, accessory ecosystem, data-sharing with consignees/regulators, carrier/eSIM connectivity partners, and interoperability standards.
Out of scope: Personas (MR-1200 PER), environments (MR-1300 ENV), and device/SKU portfolio (MR-1400 PORT). Detailed API specifications, data schemas, and security architecture belong to downstream software requirements (SRS/PRD); this module states the market demand.
3. Applicable Standards & References
• FDA 21 CFR Part 11 — electronic records/signatures, audit trails, access controls for cloud and export functions.
• EU GDP — record-keeping, traceability, and data-availability obligations that shape export and data-sharing needs.
• WHO PQS / WHO DDL guidance — remote temperature-monitoring and data-reporting expectations.
• GS1 standards — identification/serialization interoperability for device and shipment linkage.
• ISO/IEC 27001, SOC 2 — information-security expectations enterprise buyers require of the cloud platform.
• GSMA eSIM / 3GPP LTE-M / CBRS-equivalent — cellular connectivity and eSIM provisioning frameworks.
• HL7/EPCIS (context) — supply-chain event and health-data interoperability references where applicable.
• Ecosystem quantitative claims (integration counts, uptime, partner coverage) are Judgment/Assumption or Evidence Pending; standards cited are real frameworks.
4. Background
Modern cold-chain buyers do not buy loggers; they buy visibility and defensible records inside the systems they already run. A distributor's data must land in its TMS/WMS/ERP and its pharma QMS; a 3PL wants monitoring inside its own control tower; an auditor wants a one-click GDP export; a consignee wants shared visibility of the shipment approaching them; a regulator may need controlled access. Meanwhile the reusable model (D1) makes fleet/device management — provisioning, pooling, calibration tracking, OTA, battery-health monitoring across thousands of returning devices — a first-class ecosystem requirement rather than an afterthought.
Connectivity is itself an ecosystem: LTE-M primary across borders implies carrier and eSIM partnerships and roaming coverage; Wi-Fi and BLE imply provisioning apps and offline sync. The strategic choice is between a closed appliance and an open platform. The market — enterprise pharma, 3PLs, and institutional programs — strongly favors open, integrable, secure platforms over proprietary silos. This module frames the ecosystem so that integration, data-sharing, fleet management, and connectivity partnerships are treated as core value, not accessories.
5. Requirement Categories
• 5.1 First-party software — cloud dashboard, mobile/BLE provisioning app.
• 5.2 Integration & data exchange — open APIs to TMS/WMS/ERP/QMS, GDP audit export, 3PL platform integration, data-sharing with consignees/regulators.
• 5.3 Platform extensibility — developer ecosystem, interoperability standards.
• 5.4 Fleet & connectivity operations — fleet/device management, carrier/eSIM partners.
• 5.5 Physical ecosystem — accessory ecosystem.
6. Detailed Market Requirements
6.1 First-Party Software
MR-ECO-001 — Cloud Dashboard
Requirement Statement | The market demands a cloud dashboard that presents shipment status, excursion history, and fleet health with role-appropriate views and secure, compliant access for operational, quality, and executive users.
Business Objective | Deliver the primary visibility surface that anchors recurring SaaS revenue and daily engagement.
Customer Segment | All B2B segments; distributors, pharma, 3PLs, programs.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | A clear, compliant, multi-role dashboard differentiates a platform from bare loggers.
Business Value | The recurring-revenue engine and central engagement point.
Revenue Impact | High — core SaaS ARR.
Risk if Unmet | Device seen as a commodity logger; no recurring revenue; low stickiness.
Evidence Source | Standard/Judgment — 21 CFR Part 11 access/audit; ISO 27001/SOC 2 expectations; UX specifics Evidence Pending.
Supporting Research | Cold-chain visibility platform competitive analysis; Part 11 dashboard patterns.
VOC Reference | VOC-### (pending) — theme: "One screen: what's in trouble, what's fine, what needs me."
Strategic Theme | Visibility as the recurring-revenue core.
Product Family | SCM (Smart Cold Chain Monitor) — platform/SaaS layer.
Lifecycle Stage | Launch
Traceability | forward to UN-DASHBOARD / PRD-CLOUD placeholders; ORCH-018 / EM-ENG-012.
Notes | Role views map to MR-PER-001/002/010; security ties to MR-PER-003.
MR-ECO-002 — Mobile / BLE Provisioning App
Requirement Statement | The market demands a mobile app that provisions, configures, and reads devices over BLE — including offline provisioning and status readout at the warehouse, dock, or last mile without network dependence.
Business Objective | Make device setup and readout fast and field-usable, driving utilization of the reusable fleet.
Customer Segment | Warehouse operators, logistics coordinators, field/service technicians, health workers.
Geographic Applicability | Global; essential where connectivity is intermittent.
Market Priority | Must-have
Importance | High
Competitive Driver | BLE offline provisioning differentiates against cloud-only setup that fails off-grid.
Business Value | Reduces setup friction and support cost; enables reuse turnaround.
Revenue Impact | Medium-High — utilization and reuse-efficiency driver.
Risk if Unmet | Slow/failed provisioning off-grid; low fleet utilization; high support burden.
Evidence Source | Standard/Judgment — BLE provisioning practice; offline sync patterns; specifics Evidence Pending.
Supporting Research | BLE commissioning UX; offline-first mobile design.
VOC Reference | VOC-### (pending) — theme: "Set it up on the dock with no signal in seconds."
Strategic Theme | Frictionless, offline-capable provisioning.
Product Family | SCM — mobile/provisioning layer.
Lifecycle Stage | Launch
Traceability | forward to UN-PROVISIONAPP / PRD-BLE placeholders; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-PER-005/007/011 and MR-ENV-009/011; BLE + Wi-Fi provisioning per D-spec.
6.2 Integration & Data Exchange
MR-ECO-003 — Open APIs / Integration to TMS/WMS/ERP/Pharma QMS
Requirement Statement | The market demands open, documented APIs that push shipment and excursion data into customers' TMS, WMS, ERP, and pharma QMS so monitoring data lives in the systems of record, not a separate silo.
Business Objective | Embed the platform into enterprise workflows, creating integration stickiness and standardization.
Customer Segment | Enterprise pharma, distributors, 3PLs with existing supply-chain/quality systems.
Geographic Applicability | Global enterprise.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Open integration beats closed platforms in enterprise procurement and is often a hard requirement.
Business Value | Integration lock-in and workflow embedding drive retention and expansion.
Revenue Impact | High — enables enterprise deals and integration revenue.
Risk if Unmet | Data silo rejected by enterprise IT; blocked from large accounts.
Evidence Source | Standard/Judgment — GS1/EPCIS interoperability; Part 11 for QMS data; API scope Evidence Pending.
Supporting Research | Supply-chain integration patterns; pharma QMS data-exchange practice.
VOC Reference | VOC-### (pending) — theme: "It has to feed our ERP/QMS, not another login."
Strategic Theme | Open, embedded integration as enterprise moat.
Product Family | SCM — integration/API layer.
Lifecycle Stage | Growth
Traceability | forward to UN-API / PRD-INTEGRATION placeholders; ORCH-018 / EM-ENG-012.
Notes | Gated by IT/data admin (MR-PER-003); foundation for developer ecosystem (MR-ECO-006).
MR-ECO-004 — GDP Audit-Record Export
Requirement Statement | The market demands one-click, tamper-evident, GDP/Part 11-conformant export of the complete temperature/excursion record per shipment, in formats auditors and regulators accept without further processing.
Business Objective | Make audit evidence effortless — a decisive trust and differentiation factor for regulated buyers.
Customer Segment | QA/regulatory functions, auditors, GDP distributors, pharma.
Geographic Applicability | Regulated markets; global multinationals.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Audit-ready export with data integrity is a top decision factor versus manual logger downloads.
Business Value | Reduces customers' audit risk; premium trust differentiator.
Revenue Impact | Medium-High — de-risks and accelerates regulated adoption.
Risk if Unmet | Records rejected in audit; product blamed for compliance gaps.
Evidence Source | Standard — EU GDP record-keeping; 21 CFR Part 11; ALCOA+ data integrity.
Supporting Research | GDP inspection evidence expectations; audit-trail export practice.
VOC Reference | VOC-### (pending) — theme: "Give the inspector one file, tamper-evident, done."
Strategic Theme | Effortless audit-grade evidence.
Product Family | SCM — compliance/export layer.
Lifecycle Stage | Launch
Traceability | forward to UN-EXPORT / PRD-AUDITEXPORT placeholders; ORCH-018 / EM-ENG-012.
Notes | Serves MR-PER-009; depends on device-side data integrity (MR-1200/1300 logging).
MR-ECO-005 — Third-Party 3PL Platform Integration
Requirement Statement | The market demands that SCM data flow into 3PLs' own control-tower/visibility platforms so 3PLs can offer monitoring as part of their service without operating a separate system.
Business Objective | Turn 3PLs into a distribution channel by making SCM native to their platforms.
Customer Segment | Cold-chain 3PLs, freight forwarders, logistics control towers.
Geographic Applicability | Global logistics networks.
Market Priority | High
Importance | High
Competitive Driver | 3PL-native integration wins channel adoption competitors' standalone tools can't match.
Business Value | 3PL channel multiplies reach across many shippers at once.
Revenue Impact | High — channel-scale volume lever.
Risk if Unmet | 3PLs choose integrated competitors; SCM excluded from control towers.
Evidence Source | Judgment/Assumption — 3PL platform integration practice; partner specifics Evidence Pending.
Supporting Research | Logistics control-tower integration patterns; 3PL visibility ecosystems.
VOC Reference | VOC-### (pending) — theme: "Show it inside our control tower, not another portal."
Strategic Theme | Channel enablement via native integration.
Product Family | SCM — integration/partner layer.
Lifecycle Stage | Growth
Traceability | forward to UN-3PLINT / PRD-CONTROLTOWER placeholders; ORCH-018 / EM-ENG-012.
Notes | Serves MR-PER-004; builds on open APIs (MR-ECO-003).
MR-ECO-009 — Data-Sharing with Consignees & Regulators
Requirement Statement | The market demands controlled, permissioned data-sharing that lets shippers grant consignees, receiving programs, and regulators scoped visibility into a shipment's condition and record without exposing the full fleet.
Business Objective | Extend trust across the custody chain and to authorities, increasing platform value and stickiness.
Customer Segment | Shippers, consignees, receiving programs (EPI), regulators/authorities.
Geographic Applicability | Global; cross-border emphasis.
Market Priority | High
Importance | High
Competitive Driver | Permissioned multi-party sharing differentiates from single-tenant loggers.
Business Value | Multi-party visibility increases network value and switching cost.
Revenue Impact | Medium-High — network-effect and retention driver.
Risk if Unmet | Fragmented handoffs; disputes over shipment condition; regulator friction.
Evidence Source | Standard/Judgment — GDP traceability, Part 11 access control; sharing model Evidence Pending.
Supporting Research | Multi-party supply-chain visibility and consent models.
VOC Reference | VOC-### (pending) — theme: "Let my consignee and the ministry see just this shipment."
Strategic Theme | Permissioned multi-party trust network.
Product Family | SCM — data-sharing/platform layer.
Lifecycle Stage | Growth
Traceability | forward to UN-DATASHARE / PRD-PERMISSIONS placeholders; ORCH-018 / EM-ENG-012.
Notes | Serves MR-PER-006/009; scoped access ties to MR-PER-003 governance.
6.3 Platform Extensibility
MR-ECO-006 — Developer Ecosystem
Requirement Statement | The market demands developer-facing resources — documented APIs, SDKs, webhooks, and sandbox access — so customers and partners can build custom integrations and workflows on the SCM platform.
Business Objective | Extend platform reach and stickiness beyond first-party software through partner-built value.
Customer Segment | Enterprise IT, system integrators, ISV/tech partners.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | A developer ecosystem compounds platform value the way closed products cannot.
Business Value | Partner-built integrations expand use cases at low marginal cost.
Revenue Impact | Medium — long-tail expansion and retention.
Risk if Unmet | Every integration is bespoke and vendor-delivered; scaling integration is slow and costly.
Evidence Source | Judgment/Assumption — platform/ecosystem strategy; demand sizing Evidence Pending.
Supporting Research | API-platform and developer-ecosystem growth patterns.
VOC Reference | VOC-### (pending) — theme: "Give my team the API and we'll build the rest."
Strategic Theme | Extensibility as compounding platform value.
Product Family | SCM — developer/platform layer.
Lifecycle Stage | Sustaining
Traceability | forward to UN-DEVPORTAL / PRD-SDK placeholders; ORCH-018 / EM-ENG-012.
Notes | Builds directly on MR-ECO-003 open APIs; governed with MR-PER-003.
MR-ECO-010 — Interoperability Standards
Requirement Statement | The market demands adherence to recognized supply-chain and data interoperability standards (e.g., GS1 identification/EPCIS events, standard data formats) so SCM data integrates predictably across heterogeneous systems and partners.
Business Objective | Reduce integration cost and friction by speaking the ecosystem's common languages.
Customer Segment | Enterprise supply-chain, 3PLs, programs, integrators.
Geographic Applicability | Global.
Market Priority | High
Importance | Medium
Competitive Driver | Standards conformance lowers buyer integration risk versus proprietary-only formats.
Business Value | Faster, cheaper integrations and broader partner compatibility.
Revenue Impact | Medium — accelerates deals by cutting integration cost.
Risk if Unmet | Costly bespoke mappings; slower integration; buyer resistance.
Evidence Source | Standard — GS1/EPCIS; common supply-chain data formats.
Supporting Research | GS1 EPCIS adoption in pharma serialization/cold chain.
VOC Reference | VOC-### (pending) — theme: "Use GS1 so it just plugs into our stack."
Strategic Theme | Standards-based interoperability.
Product Family | SCM — integration/standards layer.
Lifecycle Stage | Growth
Traceability | forward to UN-INTEROP / PRD-STANDARDS placeholders; ORCH-018 / EM-ENG-012.
Notes | Underpins MR-ECO-003/005/009; aligns with pharma serialization mandates.
6.4 Fleet & Connectivity Operations
MR-ECO-007 — Fleet / Device Management
Requirement Statement | The market demands fleet/device-management capability — provisioning, pooling, location, battery/calibration-health tracking, OTA update management, and lifecycle status across thousands of reusable devices.
Business Objective | Make the reusable fleet operable at scale, delivering the reuse economics the business model promises.
Customer Segment | Fleet owners: distributors, 3PLs, programs, pooling operators.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Fleet management is what makes reuse viable; single-use competitors have no equivalent need or capability.
Business Value | Directly determines device utilization, uptime, and margin per unit.
Revenue Impact | High — enabler of the entire reuse model and asset productivity.
Risk if Unmet | Lost/uncalibrated/dead devices; collapsed reuse ROI; stranded fleet.
Evidence Source | Standard/Judgment — OTA/device-management practice; calibration traceability; scale figures Evidence Pending.
Supporting Research | IoT fleet-management and OTA cost models; calibration-interval management.
VOC Reference | VOC-### (pending) — theme: "I manage thousands of returning devices — show me every one's health."
Strategic Theme | Fleet operability as the backbone of reuse.
Product Family | SCM — fleet-management/platform layer.
Lifecycle Stage | Growth
Traceability | forward to UN-FLEETMGMT / UN-OTA / PRD-DEVICEMGMT placeholders; ORCH-018 / EM-ENG-012.
Notes | Serves MR-PER-010/011; realizes MR-PORT-007 commonality/OTA; OTA per D2.
MR-ECO-008 — Carrier / eSIM Connectivity Partners
Requirement Statement | The market demands managed, multi-market cellular connectivity — carrier and eSIM partnerships with roaming coverage — so devices connect out-of-the-box across borders without customers sourcing SIMs per country.
Business Objective | Deliver turnkey global connectivity as part of the offer, removing a major adoption barrier.
Customer Segment | International shippers, multinationals, cross-border lanes.
Geographic Applicability | Global; strongest value on cross-border lanes.
Market Priority | High
Importance | High
Competitive Driver | Bundled global eSIM connectivity beats per-country SIM management competitors impose.
Business Value | Turnkey connectivity increases adoption and supports recurring connectivity revenue.
Revenue Impact | Medium-High — connectivity recurring revenue and adoption enabler.
Risk if Unmet | Coverage gaps and SIM-logistics burden deter international buyers.
Evidence Source | Standard/Judgment — GSMA eSIM, 3GPP LTE-M; per-market coverage Evidence Pending.
Supporting Research | IoT eSIM/roaming coverage; LTE-M global availability maps.
VOC Reference | VOC-### (pending) — theme: "It should connect the moment it lands, in any country."
Strategic Theme | Turnkey borderless connectivity.
Product Family | SCM — connectivity/partner layer.
Lifecycle Stage | Growth
Traceability | forward to UN-ESIM / PRD-CONNECTIVITYPARTNER placeholders; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-PORT-004 connectivity variants and MR-ENV-004/007 radio compliance; LTE-M primary per D-spec.
6.5 Physical Ecosystem
MR-ECO-011 — Accessory Ecosystem
Requirement Statement | The market demands a supporting accessory ecosystem — external probes, mounts/holders, charging/docking stations for fleet recharge, and cases — that makes deployment, charging, and reuse turnaround efficient.
Business Objective | Complete the reuse operation and open an accessory revenue line while improving usability.
Customer Segment | Fleet operators, warehouses, service depots, all device users.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Multi-bay charging/docking and probe accessories streamline reuse in ways bare loggers don't.
Business Value | Improves turnaround efficiency and adds attach revenue.
Revenue Impact | Medium — accessory attach and reuse-efficiency support.
Risk if Unmet | Inefficient recharge/turnaround; ad-hoc mounting; friction in reuse operations.
Evidence Source | Judgment/Assumption — accessory demand; USB-C recharge (D-spec); specifics Evidence Pending.
Supporting Research | Fleet charging/docking practice; probe-accessory usage.
VOC Reference | VOC-### (pending) — theme: "I need to recharge fifty devices overnight, easily."
Strategic Theme | Physical ecosystem for efficient reuse.
Product Family | SCM — accessory line.
Lifecycle Stage | Growth
Traceability | forward to UN-ACCESSORY / PRD-DOCK placeholders; ORCH-018 / EM-ENG-012.
Notes | Multi-bay charging supports MR-PER-011 turnaround; probes tie to MR-PORT-005; charge-inhibit-below-0°C (D3) applies to docking behavior.
7. Evidence & Validation Notes
Evidence-backed anchors | Compliance, security, and interoperability elements are anchored to real frameworks (21 CFR Part 11, EU GDP, WHO PQS/DDL, GS1/EPCIS, ISO/IEC 27001, SOC 2, GSMA eSIM, 3GPP LTE-M).
Pending validation | Integration breadth (which TMS/WMS/ERP/QMS matter most), 3PL platform priorities, developer-ecosystem demand, carrier/eSIM coverage per corridor, fleet-management scale requirements, and accessory demand are Judgment/Assumption or Evidence Pending. No uptime, coverage, or integration-count figures are asserted as fact.
Recommended validation methods | enterprise IT/integration requirement interviews; 3PL control-tower partnership discovery; connectivity coverage testing across target corridors; fleet-operator workflow studies; accessory concept testing at service depots.
Confidence | High that these ecosystem elements are the right strategic set and that open/integrable/secure is the winning posture; Medium on prioritization and partner specifics pending discovery and VOC.
8. Traceability Placeholders
Upstream | ORCH-018, EM-ENG-012; anchor facts D1–D4 (reuse, cadence upload/OTA, LTE-M/Wi-Fi/BLE/GPS, multi-market).
Peer modules | MR-1200 PER (users of each ecosystem element), MR-1300 ENV (connectivity/upload conditions), MR-1400 PORT (platform commonality/OTA/connectivity variants that the ecosystem realizes).
Downstream (placeholders) | UN-DASHBOARD, UN-PROVISIONAPP, UN-API, UN-EXPORT, UN-3PLINT, UN-DEVPORTAL, UN-FLEETMGMT, UN-OTA, UN-ESIM, UN-DATASHARE, UN-INTEROP, UN-ACCESSORY → PRD-CLOUD, PRD-BLE, PRD-INTEGRATION, PRD-AUDITEXPORT, PRD-CONTROLTOWER, PRD-SDK, PRD-DEVICEMGMT, PRD-CONNECTIVITYPARTNER, PRD-PERMISSIONS, PRD-STANDARDS, PRD-DOCK.
VOC linkage | VOC-### identifiers pending; to be back-filled from enterprise-integration and partner-discovery studies.
MR-1600 — Innovation Roadmap (INNO)
Traceability: this module rolls up to the SCM Master Market Requirements Library and traces forward to PRD-SCM-16xx (Product Requirements) and UN-SCM-16xx (User Needs), and back to ORCH-018 / EM-ENG-012 (orchestration + originating engineering mandate). It defines the market's forward-looking demand for capabilities that will differentiate the SCM beyond its launch feature set.
1. Purpose
This module captures the innovation the market will reward over the SCM's multi-year lifecycle — capabilities that are not required to ship the first product, but that the cold-chain market is signalling it will pay for, standardize on, or demand from any credible platform vendor within the growth and sustaining horizons. It exists so the program funds a deliberate roadmap rather than reacting to competitors, and so that today's architecture decisions (compute headroom, sensor abstraction, connectivity modularity, data model) leave room for tomorrow's differentiation. It is written for the product-strategy owner, the CTO/architecture council, and the investment committee that must sequence R&D spend.
2. Scope
In scope: market demand for predictive and AI-driven analytics (excursion prediction, anomaly detection, alarm-reduction), remaining-stability / cumulative-thermal-load modelling, shipment digital twins, energy harvesting and endurance extension, next-generation low-temperature cell chemistries, tamper and chain-of-custody assurance (including distributed-ledger approaches), edge machine learning for false-alarm suppression, ultra-cold (-70 °C) sensing, and non-terrestrial / satellite connectivity for remote last-mile. Each is expressed as a market demand the SCM roadmap should answer, not as a committed feature.
Out of scope: the concrete engineering implementation, algorithm selection, silicon choices, and validation protocols — those belong to the PRD and discipline engineering modules. This module also does not commit delivery dates or R&D budgets; sequencing lives in the roadmap governance process. The SCM remains an informational/monitoring device throughout: none of these innovations make it a therapeutic or temperature-controlling product.
3. Applicable Standards & References
• WHO PQS (Performance, Quality and Safety) — prequalification framework for cold-chain monitoring devices/loggers; roadmap features must not jeopardize PQS conformance (category code TBC).
• EU GDP (2013/C 343/01) and US FDA 21 CFR Part 11 — any new analytics or data feature must preserve data-integrity and audit-grade evidence.
• WHO / GAVI / UNICEF Supply Division — cold-chain policy and immunization-supply direction, cited by name for market/technology context, not exact figures.
• CTLC / cumulative thermal exposure and vaccine heat/freeze stability concepts — WHO vaccine management and stability literature (referenced generically; specific stability curves are product/vaccine-specific and Evidence Pending).
• 3GPP NTN (Release 17+) / satellite-IoT direction and LTE-M/NB-IoT — connectivity evolution context (cited as industry direction).
• UN 38.3, IATA DG — any new cell chemistry remains bound by lithium-battery transport qualification.
• Cross-references: MR-000 (EXEC), MR-1800 (RISK), MR-1900 (VAL); ROSTER.md, ORCHESTRATION.md, and the SCM engineering discipline modules.
4. Background
The launch SCM competes on being reusable, multi-market, and data-integrity-grade. That advantage is durable only if the platform keeps moving up the value curve from recording what happened toward predicting and preventing loss. Buyers across pharma manufacturers, distributors, 3PLs, and immunization programs are increasingly asking not merely "was there an excursion?" but "will this consignment still be potent?", "which alerts actually matter?", and "can I trust the chain of custody end to end?". At the same time, the physics of the device — a low-temperature Li-ion pack that must be charge-inhibited below 0 °C (D3) — makes endurance and energy the binding constraints on how much on-device intelligence and connectivity the SCM can afford. Innovation in energy harvesting, cell chemistry, and edge efficiency is therefore not a luxury; it is the enabler for the analytics the market wants.
This roadmap deliberately separates near-term, low-regret bets (edge false-alarm suppression, remaining-stability modelling that reuses existing sensor data) from capital-intensive or ecosystem-dependent bets (satellite/NTN connectivity, next-generation chemistries, ultra-cold sensing). Every item is based on an observed market signal or an explicit assumption, and quantitative claims are withheld until validated (see MR-1900).
5. Requirement Categories
• Predictive & Stability Analytics — turning telemetry into potency and disposition foresight.
• AI / Edge Intelligence — anomaly detection, alarm reduction, and on-device ML within the energy budget.
• Chain-of-Custody & Trust — tamper assurance and provable custody.
• Energy & Endurance Innovation — harvesting and next-generation low-temperature chemistries.
• Sensing & Connectivity Frontier — ultra-cold sensing and non-terrestrial connectivity for the true last mile.
6. Detailed Market Requirements
6.1 Predictive & Stability Analytics
MR-INNO-001 — Predictive Excursion Warning
Requirement Statement | The market demands early, predictive warning that a consignment is trending toward a temperature excursion — before the limit is breached — so operators can intervene while the shipment can still be saved, rather than only being told after the fact.
Business Objective | Move the SCM up the value curve from detection to prevention, creating a defensible analytics differentiator and a recurring-revenue tier.
Customer Segment | Distributors, 3PLs, manufacturers, and immunization-program logisticians managing time-critical consignments.
Geographic Applicability | Global, with highest value on long or interrupted last-mile routes.
Market Priority | High
Importance | High
Competitive Driver | Offline loggers and basic connected loggers report excursions only after they occur; premium logistics telemetry is normalizing "predict and intervene" expectations.
Business Value | Enables an intervention-grade service tier and reduces salvageable-loss events, deepening switching costs.
Revenue Impact | Supports premium analytics/SaaS pricing; magnitude Evidence Pending.
Risk if Unmet | SCM stays a commodity recorder; competitors with predictive analytics define the category.
Evidence Source | Judgment/Assumption — market direction toward predictive logistics telemetry (Evidence Pending for adoption/uplift quantification).
Supporting Research | GDP continuous-monitoring expectations; premium 3PL predictive-ETA/telemetry practice.
VOC Reference | VOC-160 (pending) — "tell us before it goes out of range, not after."
Strategic Theme | Prevention over detection.
Product Family | SCM (Smart Cold Chain Monitor) — analytics tier.
Lifecycle Stage | Growth
Traceability | UN-SCM-1601, PRD-SCM-1601; ORCH-018 / EM-ENG-012.
Notes | Prediction depends on adequate upload cadence/connectivity (see MR-1800 RISK-002); on-device prediction must respect the energy budget (D3).
MR-INNO-002 — Remaining-Stability / Cumulative Thermal Load Modelling
Requirement Statement | The market demands a cumulative-thermal-load (CTLC) style assessment that translates the full temperature history into a remaining-potency / disposition indication for the specific product carried, so release decisions rest on cumulative exposure rather than a single threshold breach.
Business Objective | Make SCM data the basis for smarter, less wasteful disposition — releasing product that is still good and quarantining product that is not.
Customer Segment | Manufacturers, QA/QP release functions, distributors, immunization program managers.
Geographic Applicability | Global, weighted to regulated markets and high-value biologics.
Market Priority | High
Importance | High
Competitive Driver | Threshold-only alarms cause both over-discard (still-good product) and under-detection (cumulative harm); buyers want exposure-aware disposition.
Business Value | Reduces avoidable wastage and strengthens the "system of record for disposition" position from MR-EXEC-005.
Revenue Impact | Premium analytics feature; wastage-reduction ROI Evidence Pending.
Risk if Unmet | Disposition stays crude; buyers discount the platform's data value.
Evidence Source | Cited (concept) — WHO cumulative-heat/VVM and vaccine-stability principles; Judgment/Assumption for per-product model accuracy (Evidence Pending).
Supporting Research | WHO vaccine management / VVM and heat-stability literature; product-specific stability data (owner-provided).
VOC Reference | VoC theme — "one spike shouldn't scrap a whole shipment that's still potent."
Strategic Theme | Exposure-aware disposition.
Product Family | SCM — stability analytics.
Lifecycle Stage | Growth
Traceability | UN-SCM-1602, PRD-SCM-1602; ORCH-018 / EM-ENG-012.
Notes | Model outputs are informational decision support; the SCM does not release product. Stability curves are product-specific — do not assert generic potency numbers.
MR-INNO-003 — Shipment Digital Twin
Requirement Statement | The market demands a live digital representation of each consignment — combining location, thermal/humidity/shock history, remaining-stability state, and custody events — that logistics and quality teams can query as a single shipment "twin" across its journey.
Business Objective | Consolidate fragmented shipment data into one authoritative model, raising the platform's strategic stickiness and integration value.
Customer Segment | 3PLs, distributors, manufacturers, and control-tower/logistics-visibility teams.
Geographic Applicability | Global.
Market Priority | Medium
Importance | High
Competitive Driver | Enterprise logistics is consolidating around control-tower/"digital twin" visibility; a monitor that feeds a twin is more defensible than a bare datalogger.
Business Value | Positions SCM as a data backbone with integration/API revenue and high switching cost.
Revenue Impact | Enables platform/API and enterprise-tier revenue; magnitude Evidence Pending.
Risk if Unmet | SCM data is siloed and commoditized; value accrues to third-party visibility platforms instead.
Evidence Source | Judgment/Assumption — enterprise control-tower/digital-twin market direction (Evidence Pending).
Supporting Research | Supply-chain control-tower and digital-twin industry direction; GDP end-to-end monitoring expectations.
VOC Reference | VoC theme — "give us one live view of the whole shipment."
Strategic Theme | Data backbone / platform.
Product Family | SCM — shipment twin.
Lifecycle Stage | Growth
Traceability | UN-SCM-1603, PRD-SCM-1603; ORCH-018 / EM-ENG-012.
Notes | Depends on connectivity cadence and cloud data model; must preserve Part 11 data integrity.
6.2 AI / Edge Intelligence
MR-INNO-004 — AI Anomaly Detection & Alarm Reduction
Requirement Statement | The market demands intelligent anomaly detection that distinguishes true, action-worthy events from benign fluctuations, so operators receive fewer, higher-quality alerts and are not desensitized by noise.
Business Objective | Protect the trust and usability of the alerting system — the core of the SCM value promise — as fleet scale grows.
Customer Segment | Distributors, 3PLs, immunization-program and warehouse operators managing many devices.
Geographic Applicability | Global.
Market Priority | High
Importance | Critical
Competitive Driver | Alert fatigue is a top complaint against connected loggers; buyers reward platforms that cut false alarms without missing real ones.
Business Value | Sustains adoption and renewal by keeping alerts credible; differentiates on signal quality, not just sensing.
Revenue Impact | Retention/expansion lever; churn-avoidance value Evidence Pending.
Risk if Unmet | Alarm fatigue erodes trust (see MR-1800 RISK-011) and drives churn.
Evidence Source | Judgment/Assumption — widely reported alarm-fatigue problem in monitoring (Evidence Pending for SCM-specific rates).
Supporting Research | Alarm-management literature (analogous domains); GDP alert-handling expectations.
VOC Reference | VOC-161 (pending) — "we get so many alerts we stopped reading them."
Strategic Theme | Signal quality / trust.
Product Family | SCM — intelligent alerting.
Lifecycle Stage | Growth
Traceability | UN-SCM-1604, PRD-SCM-1604; ORCH-018 / EM-ENG-012.
Notes | Must never suppress a genuine excursion alert; safety-of-signal governance required. Coordinates tightly with MR-INNO-005.
MR-INNO-005 — Edge ML for False-Alarm Suppression
Requirement Statement | The market demands that basic intelligent filtering run on the device at the edge, so that transient sensor artefacts and non-events are resolved locally without depending on a live connection and without draining upload budget on noise.
Business Objective | Deliver alarm quality even in disconnected last-mile conditions while conserving energy and data cost.
Customer Segment | LMIC last-mile programs, remote-route 3PLs, and any low-connectivity deployment.
Geographic Applicability | Global, highest value in low-connectivity / LMIC last-mile.
Market Priority | Medium
Importance | High
Competitive Driver | Cloud-only analytics fail where connectivity is intermittent — exactly where cold-chain risk is highest.
Business Value | Extends analytics value into disconnected environments; reduces connectivity/data spend.
Revenue Impact | Lowers per-trip data cost and supports LMIC deployment economics; magnitude Evidence Pending.
Risk if Unmet | Intelligence is unavailable precisely where it matters most; energy wasted on noisy uploads.
Evidence Source | Judgment/Assumption — edge-ML/TinyML direction and LMIC connectivity gaps (Evidence Pending).
Supporting Research | Edge/TinyML industry direction; WHO/UNICEF last-mile connectivity constraints.
VOC Reference | VoC theme — "half our routes have no signal — it still has to be smart."
Strategic Theme | Intelligence at the edge.
Product Family | SCM — edge intelligence.
Lifecycle Stage | Growth
Traceability | UN-SCM-1605, PRD-SCM-1605; ORCH-018 / EM-ENG-012.
Notes | Bounded by the D3 energy budget; edge inference must fit duty-cycled power. Coordinates with hardware/firmware engineering on compute headroom.
6.3 Chain-of-Custody & Trust
MR-INNO-006 — Tamper & Chain-of-Custody Assurance (incl. Distributed Ledger)
Requirement Statement | The market demands provable, tamper-evident chain-of-custody records — including cryptographically verifiable or distributed-ledger-anchored custody events — so that data integrity and handling accountability can be demonstrated across organizational and border handoffs.
Business Objective | Elevate the SCM from a data source to a trust anchor for multi-party cold-chain custody.
Customer Segment | Manufacturers, distributors, 3PLs, customs/regulatory and donor-audit stakeholders.
Geographic Applicability | Global, weighted to cross-border and donor-funded consignments.
Market Priority | Medium
Importance | High
Competitive Driver | Multi-party chains lack a neutral, tamper-proof custody record; buyers and donors increasingly demand verifiable accountability.
Business Value | Strengthens the data-integrity moat (MR-EXEC-005) and unlocks donor/audit-driven procurement.
Revenue Impact | Supports premium regulated/donor-tier pricing; magnitude Evidence Pending.
Risk if Unmet | Custody disputes remain unresolvable; competitors or third-party ledgers capture the trust layer.
Evidence Source | Standard + Judgment/Assumption — 21 CFR Part 11 / GDP integrity requirements (Standard); distributed-ledger adoption in supply chain (Judgment/Assumption, Evidence Pending).
Supporting Research | FDA data-integrity guidance; GDP custody expectations; supply-chain provenance/ledger pilots (industry).
VOC Reference | VoC theme — "nobody can prove who mishandled it."
Strategic Theme | Trust layer.
Product Family | SCM — chain-of-custody assurance.
Lifecycle Stage | Growth
Traceability | UN-SCM-1606, PRD-SCM-1606; ORCH-018 / EM-ENG-012.
Notes | Ledger complexity must be justified against a simpler signed-record approach; avoid technology for its own sake. Coordinates with cybersecurity engineer.
6.4 Energy & Endurance Innovation
MR-INNO-007 — Energy Harvesting & Endurance Extension
Requirement Statement | The market demands longer unattended endurance and reduced recharge dependence — potentially via energy harvesting — so devices can complete longer or delayed journeys without a mid-trip charge, especially where recharge below 0 °C is inhibited.
Business Objective | Relax the binding endurance constraint that limits sensing, analytics, and connectivity, and reduce operational recharge burden.
Customer Segment | Long-haul 3PLs, LMIC last-mile programs, and reverse-logistics operators.
Geographic Applicability | Global, highest value on long/remote routes and cold environments.
Market Priority | Medium
Importance | High
Competitive Driver | Endurance and cold-charge limits are a shared industry pain; extending them is a tangible, defensible advantage.
Business Value | Enables richer on-device features and reduces field-charging labor; supports the reusable TCO story.
Revenue Impact | Reduces operational cost and enables premium endurance SKUs; magnitude Evidence Pending.
Risk if Unmet | Endurance ceiling caps analytics/connectivity ambitions and constrains long-route use.
Evidence Source | Judgment/Assumption — energy-harvesting feasibility for this form factor is unproven here (Evidence Pending). Cold-charge inhibition below 0 °C is a evidence-backed constraint (D3).
Supporting Research | Energy-harvesting IoT literature; low-temperature battery behaviour references.
VOC Reference | VoC theme — "it can't die halfway through a delayed shipment."
Strategic Theme | Endurance as enabler.
Product Family | SCM — endurance platform.
Lifecycle Stage | Growth
Traceability | UN-SCM-1607, PRD-SCM-1607; ORCH-018 / EM-ENG-012.
Notes | Directly linked to the D3 KEY RISK (MR-1800 RISK-001); harvesting yield in cold/dark shippers is highly uncertain and must be validated by test.
MR-INNO-008 — Next-Generation Low-Temperature Cell Chemistries
Requirement Statement | The market demands battery chemistries that retain capacity and can be safely charged (or safely operated) at lower temperatures than today's Li-ion, so the device performs reliably at -30 °C and resists capacity fade across many reuse cycles.
Business Objective | Attack the SCM's top physical constraint (D3) at the source, converting a key risk into a differentiator.
Customer Segment | All segments operating in ultra-cold/frozen chains and high-reuse fleets.
Geographic Applicability | Global, weighted to frozen/ultra-cold logistics.
Market Priority | Medium
Importance | High
Competitive Driver | Whoever solves cold-temperature endurance and cycle-life first sets the reliability benchmark for reusable monitors.
Business Value | Reduces cold-related failures and attrition; strengthens reusable economics and reliability claims.
Revenue Impact | Lowers warranty/attrition cost and enables premium cold-rated SKUs; magnitude Evidence Pending.
Risk if Unmet | SCM stays exposed to cold-endurance failure and capacity fade — the program's defining technical vulnerability.
Evidence Source | Judgment/Assumption — emerging low-temperature chemistry maturity is uncertain (Evidence Pending). Low-temperature Li-ion limitations and cold-charge safety are evidence-backed (D3).
Supporting Research | Low-temperature Li-ion / alternative-chemistry literature; cell-vendor roadmaps.
VOC Reference | VoC theme — "it has to survive -30 and keep surviving after 200 trips."
Strategic Theme | Solve the cold constraint.
Product Family | SCM — cell technology.
Lifecycle Stage | Growth
Traceability | UN-SCM-1608, PRD-SCM-1608; ORCH-018 / EM-ENG-012.
Notes | Any new chemistry must re-qualify under UN 38.3 / IATA DG. Coordinates with battery/power engineer; ties to MR-1800 RISK-001 and RISK-006 (single-source cells).
6.5 Sensing & Connectivity Frontier
MR-INNO-009 — Ultra-Cold (-70 °C) Sensing
Requirement Statement | The market demands accurate temperature sensing into the ultra-cold (-70 °C class) regime, so the SCM platform can address frozen/ultra-cold biologics and mRNA-class vaccine logistics, not only the -30 °C to +60 °C mainstream range.
Business Objective | Extend the addressable market into the fastest-growing ultra-cold segment without a separate product line.
Customer Segment | Manufacturers and distributors of ultra-cold biologics; specialty cold-chain 3PLs.
Geographic Applicability | Global, concentrated in developed-market ultra-cold logistics.
Market Priority | Medium
Importance | Medium
Competitive Driver | Ultra-cold demand surged with mRNA vaccines; platforms that span standard and ultra-cold ranges win consolidation.
Business Value | Expands TAM and reinforces the one-platform-many-markets thesis (MR-EXEC-004).
Revenue Impact | Opens a new segment; magnitude Evidence Pending.
Risk if Unmet | Ceded ultra-cold segment; buyers run a separate ultra-cold monitoring system.
Evidence Source | Judgment/Assumption — ultra-cold demand growth is observed but SCM sensing feasibility at -70 °C is unvalidated (Evidence Pending). Launch range -30 °C to +60 °C is based.
Supporting Research | mRNA/ultra-cold logistics reporting; external-probe sensing at cryogenic-adjacent ranges.
VOC Reference | VoC theme — "we need one monitor for both our -20 and our -70 shipments."
Strategic Theme | Range expansion.
Product Family | SCM — ultra-cold sensing.
Lifecycle Stage | Growth
Traceability | UN-SCM-1609, PRD-SCM-1609; ORCH-018 / EM-ENG-012.
Notes | Likely requires external-probe and battery-behaviour rethink at ultra-cold; interacts strongly with D3. Sensor accuracy at ultra-cold must be validated (MR-1900).
MR-INNO-010 — Satellite / Non-Terrestrial (NTN) Connectivity for Remote Last-Mile
Requirement Statement | The market demands connectivity that reaches consignments beyond terrestrial cellular coverage — via satellite / non-terrestrial networks — so excursion alerts and telemetry are not lost on remote routes, ocean legs, or coverage-poor LMIC last-mile.
Business Objective | Close the coverage gap that most undermines the real-time-alert promise where cold-chain risk is highest.
Customer Segment | LMIC last-mile programs, remote-region distributors, and long-haul/ocean 3PLs.
Geographic Applicability | Global, concentrated in coverage-poor and remote corridors.
Market Priority | Medium
Importance | High
Competitive Driver | Cellular-only monitors go dark exactly where visibility is most valuable; satellite-IoT is maturing toward affordable low-rate messaging.
Business Value | Extends the SCM's core alert promise to the true last mile; strong mission/donor alignment.
Revenue Impact | Enables premium remote-coverage tier and new corridors; magnitude Evidence Pending.
Risk if Unmet | Alerts are lost on remote routes; competitors with NTN own the hardest, most mission-critical segment.
Evidence Source | Judgment/Assumption — 3GPP NTN / satellite-IoT direction is real but cost/power/latency for this device are unproven (Evidence Pending).
Supporting Research | 3GPP NTN (Rel-17+) direction; commercial satellite-IoT messaging services; WHO/UNICEF remote last-mile connectivity gaps.
VOC Reference | VoC theme — "on the boat and in the bush, it goes silent."
Strategic Theme | Coverage everywhere.
Product Family | SCM — connectivity roadmap.
Lifecycle Stage | Growth
Traceability | UN-SCM-1610, PRD-SCM-1610; ORCH-018 / EM-ENG-012.
Notes | Power draw of satellite links is a real tension with D3 endurance; low-rate store-and-forward messaging is the likely fit. Ties to MR-1800 RISK-002 and RISK-005.
7. Evidence & Validation Notes
• Anchor facts D1–D4 and the stated device specification are treated as given. The D3 low-temperature battery constraint is the central gating physics for most items here (energy harvesting, next-gen chemistry, edge ML, satellite connectivity) and must be validated by test — see MR-1800 RISK-001 and MR-1900.
• Cited concepts (cumulative thermal/heat-stability, VVM principles, NTN direction) are drawn from WHO and industry references and labelled as concept/direction — never as exact SCM performance figures.
• All adoption, uplift, and financial magnitudes are marked Evidence Pending until VOC, pilot, and business-case data exist. No precise number is asserted as fact in this module.
Validation path | VOC studies (MR-1900) confirm demand priority; technical feasibility studies and pilots convert Judgment/Assumption items into cited evidence before any roadmap commitment. Roadmap governance sequences items by (market pull × feasibility × energy-budget fit).
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-INNO-001 | UN-SCM-1601 | PRD-SCM-1601 | ORCH-018 / EM-ENG-012
MR-INNO-002 | UN-SCM-1602 | PRD-SCM-1602 | ORCH-018 / EM-ENG-012
MR-INNO-003 | UN-SCM-1603 | PRD-SCM-1603 | ORCH-018 / EM-ENG-012
MR-INNO-004 | UN-SCM-1604 | PRD-SCM-1604 | ORCH-018 / EM-ENG-012
MR-INNO-005 | UN-SCM-1605 | PRD-SCM-1605 | ORCH-018 / EM-ENG-012
MR-INNO-006 | UN-SCM-1606 | PRD-SCM-1606 | ORCH-018 / EM-ENG-012
MR-INNO-007 | UN-SCM-1607 | PRD-SCM-1607 | ORCH-018 / EM-ENG-012
MR-INNO-008 | UN-SCM-1608 | PRD-SCM-1608 | ORCH-018 / EM-ENG-012
MR-INNO-009 | UN-SCM-1609 | PRD-SCM-1609 | ORCH-018 / EM-ENG-012
MR-INNO-010 | UN-SCM-1610 | PRD-SCM-1610 | ORCH-018 / EM-ENG-012
Related modules: MR-000 (EXEC) sets the vision this roadmap extends; MR-1800 (RISK) governs the risks these bets carry; MR-1900 (VAL) validates the demand before commitment.
MR-1700 — Sustainability (SUS)
Traceability: this module rolls up to the SCM Master Market Requirements Library and traces forward to PRD-SCM-17xx (Product Requirements) and UN-SCM-17xx (User Needs), and back to ORCH-018 / EM-ENG-012 (orchestration + originating engineering mandate). It expresses the market's environmental, circular-economy, and materials-compliance demands on a reusable, battery-operated cold-chain device.
1. Purpose
This module states what the market — buyers, regulators, and donor/procurement bodies — will demand of the SCM on environmental and circular-economy grounds across its whole life: manufacture, multi-trip reuse, reverse logistics, service, and end-of-life. The SCM's reusable, rechargeable design (D1) is itself a sustainability thesis: displacing single-use loggers and chemical indicators. That thesis only holds if battery lifecycle, repairability, reverse-logistics carbon, materials compliance, and end-of-life take-back are managed deliberately. This module is written for the product-strategy owner, sustainability/ESG and regulatory-affairs stakeholders, and procurement leads who increasingly gate on environmental criteria.
2. Scope
In scope: market demand for reusable-vs-single-use waste reduction, battery lifecycle and recycling (UN 38.3 transport, WEEE/battery-directive handling), repairability and replaceable battery, sustainable packaging, the carbon footprint of reverse logistics, restricted-materials compliance (RoHS/REACH), e-waste take-back programs, energy efficiency in use, and circular/device-pooling business models. Each is a market demand the SCM must answer, not an engineering solution.
Out of scope: the specific bill-of-materials, cell selection, enclosure design, and recycling process engineering — those live in the PRD and discipline engineering modules. This module does not compute the device's actual carbon or waste numbers (those are Evidence Pending pending LCA); it states what the market requires the program to measure and improve. The SCM remains a monitoring device; sustainability requirements never compromise the D2 excursion-alert or D3 low-temperature safety functions.
3. Applicable Standards & References
• EU WEEE Directive (2012/19/EU) — waste electrical and electronic equipment collection/recycling obligations.
• EU Batteries Regulation (2023/1542) / battery directive lineage — battery collection, recycling, carbon-footprint, and removability/replaceability expectations.
• EU RoHS (2011/65/EU) and REACH (EC 1907/2006) — restricted substances and substances-of-very-high-concern declarations.
• UN 38.3 / IATA DG — lithium-battery transport qualification, including for end-of-life return shipping.
• WHO PQS / EU GDP — must remain conformant; sustainability changes cannot degrade monitoring performance or data integrity.
• ISO 14040/14044 (LCA) and GHG Protocol — life-cycle assessment and carbon-accounting frameworks (referenced as method; SCM figures Evidence Pending).
• WHO / GAVI / UNICEF Supply Division — environmental-sustainability and circular-procurement direction in immunization supply (cited by name for context).
• Cross-references: MR-000 (EXEC), MR-1800 (RISK), MR-1900 (VAL); ROSTER.md, ORCHESTRATION.md, and the SCM engineering discipline modules.
4. Background
Single-use temperature indicators, chemical VVMs, and disposable USB loggers generate recurring waste and, in the case of connected disposables, recurring e-waste and embedded batteries. A reusable SCM promises to amortize its manufacturing and battery footprint over many trips — but only if it is actually recovered, reused many times, repaired rather than scrapped, and responsibly retired. The same reverse-logistics chain that enables reuse also carries a carbon and loss cost (return shipping, cleaning, re-provisioning), and the same Li-ion pack that enables endurance is the device's principal end-of-life hazard and the focus of tightening battery regulation.
Procurement is moving in step: EU public and GDP-regulated buyers face WEEE/battery/RoHS/REACH obligations, and donor/immunization procurement (GAVI, UNICEF, WHO) increasingly weighs environmental impact and total lifecycle cost, not just unit price. The market will therefore reward an SCM that can prove its reusable advantage with a credible LCA, offer a replaceable battery and take-back, and minimize the reverse-logistics footprint — and will penalize one that quietly recreates the e-waste problem it claims to solve.
5. Requirement Categories
• Circular Advantage — reusable-vs-single-use waste reduction and pooling/circular models.
• Battery Lifecycle — cell longevity, safe end-of-life, recycling, and replaceability.
• Materials & Packaging — RoHS/REACH compliance and sustainable packaging.
• Reverse-Logistics Footprint — carbon of recovery, cleaning, and re-provisioning; take-back.
• In-Use Efficiency — energy efficiency that compounds the sustainability case.
6. Detailed Market Requirements
6.1 Circular Advantage
MR-SUS-001 — Reusable-vs-Single-Use Waste Reduction
Requirement Statement | The market demands a monitoring solution that measurably reduces monitoring-related waste versus single-use indicators, chemical VVMs, and disposable loggers, with a credible life-cycle basis for the claim rather than an unsubstantiated "green" assertion.
Business Objective | Convert the reusable design (D1) into a defensible, auditable sustainability differentiator that survives procurement scrutiny.
Customer Segment | Distributors, 3PLs, manufacturers, and immunization/donor procurement with ESG mandates.
Geographic Applicability | Global, weighted to EU and ESG-driven institutional buyers.
Market Priority | High
Importance | High
Competitive Driver | Buyers increasingly reject unquantified green claims; a proven waste-reduction story differentiates against single-use incumbents.
Business Value | Anchors the environmental value proposition and supports premium/tender positioning.
Revenue Impact | Enables ESG-gated tender access; magnitude Evidence Pending.
Risk if Unmet | "Greenwashing" exposure; reusable advantage dismissed as marketing.
Evidence Source | Judgment/Assumption — reuse-reduces-waste premise is sound but requires SCM LCA to quantify (Evidence Pending). Reusable design is based (D1).
Supporting Research | ISO 14040/14044 LCA method; reusable-vs-single-use packaging/logistics literature.
VOC Reference | VOC-170 (pending) — "prove it actually creates less waste than our disposables."
Strategic Theme | Circular advantage.
Product Family | SCM (Smart Cold Chain Monitor) — reusable platform.
Lifecycle Stage | Growth
Traceability | UN-SCM-1701, PRD-SCM-1701; ORCH-018 / EM-ENG-012.
Notes | Advantage depends on real recovery/reuse rates (see MR-1800 RISK-009 reverse-logistics attrition). Do not claim a specific waste-reduction percentage without LCA.
MR-SUS-002 — Circular / Device-Pooling Economy
Requirement Statement | The market demands flexible ownership and circular models — device pooling, sharing, and managed-fleet return — so buyers can maximize utilization and reuse rather than each party owning under-utilized stock.
Business Objective | Maximize per-device trips (the numerator of the sustainability and TCO case) via pooling/managed-fleet models.
Customer Segment | 3PLs, distributor networks, and multi-party immunization supply chains.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Reusable economics and sustainability both improve with utilization; pooling is emerging in returnable-asset logistics.
Business Value | Higher utilization improves both carbon-per-trip and unit economics; supports as-a-service revenue.
Revenue Impact | Enables pooling/subscription revenue; magnitude Evidence Pending.
Risk if Unmet | Devices sit idle; reuse and sustainability advantage under-realized.
Evidence Source | Judgment/Assumption — pooling model demand and viability unproven for this device (Evidence Pending).
Supporting Research | Returnable transport asset (RTA) pooling models; as-a-service logistics.
VOC Reference | VoC theme — "we don't want to own idle monitors between shipments."
Strategic Theme | Circular / as-a-service.
Product Family | SCM — pooling model.
Lifecycle Stage | Growth
Traceability | UN-SCM-1702, PRD-SCM-1702; ORCH-018 / EM-ENG-012.
Notes | Pooling raises cleaning, re-provisioning, and custody demands (MR-SUS-007, MR-1800 RISK-009).
6.2 Battery Lifecycle
MR-SUS-003 — Battery Lifecycle & Cycle Longevity
Requirement Statement | The market demands that the device's battery retain useful capacity across many reuse cycles, so devices are not retired prematurely by capacity fade — the reuse count is the foundation of both the sustainability and TCO cases.
Business Objective | Protect the reuse-count denominator that every sustainability and cost claim depends on.
Customer Segment | Fleet owners, 3PLs, distributors, and procurement evaluating lifecycle cost.
Geographic Applicability | Global, weighted to cold/frozen operations where fade accelerates.
Market Priority | High
Importance | High
Competitive Driver | Premature battery-driven retirement destroys the reusable value proposition; long cycle life is a tangible differentiator.
Business Value | More trips per device improves carbon-per-trip and unit economics simultaneously.
Revenue Impact | Reduces replacement cost and warranty exposure; magnitude Evidence Pending.
Risk if Unmet | Early retirement erodes both the green and the economic case (see MR-1800 RISK-001).
Evidence Source | Judgment/Assumption — cycle life at low temperature is uncertain and must be tested (Evidence Pending). Capacity fade over reuse cycles is a evidence-backed KEY RISK (D3).
Supporting Research | Low-temperature Li-ion cycle-life literature; battery-management best practice.
VOC Reference | VoC theme — "how many trips before the battery is done?"
Strategic Theme | Longevity = sustainability.
Product Family | SCM — battery lifecycle.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1703, PRD-SCM-1703; ORCH-018 / EM-ENG-012.
Notes | Directly tied to D3 KEY RISK and MR-INNO-008 (next-gen chemistry). Validate cycle life at -30 °C (MR-1900).
MR-SUS-004 — Safe Battery End-of-Life & Recycling
Requirement Statement | The market demands a compliant, safe pathway for end-of-life battery handling — recycling and disposal that meet transport (UN 38.3), WEEE, and battery-regulation obligations — so retired devices do not become a hazard or a compliance liability for the buyer.
Business Objective | Remove end-of-life risk and compliance burden from the buyer, and meet tightening battery regulation.
Customer Segment | All fleet owners; especially EU-regulated and institutional/donor buyers.
Geographic Applicability | Global, with binding obligations in the EU.
Market Priority | Must-have
Importance | High
Competitive Driver | EU battery/WEEE regulation and buyer ESG policy make credible end-of-life handling a gate, not a nicety.
Business Value | Compliance and liability protection; reinforces the responsible-reuse narrative.
Revenue Impact | Avoids fines/exclusion and enables EU tender access; magnitude Evidence Pending.
Risk if Unmet | Regulatory non-compliance, safety incidents in disposal, and reputational damage (see MR-1800 RISK-008 regulatory).
Evidence Source | Standard — UN 38.3 / IATA DG, EU WEEE (2012/19/EU), EU Batteries Regulation (2023/1542).
Supporting Research | WEEE and battery-regulation guidance; lithium-battery end-of-life handling references.
VOC Reference | VoC theme — "what do we do with 500 dead devices, legally and safely?"
Strategic Theme | Responsible end-of-life.
Product Family | SCM — end-of-life handling.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1704, PRD-SCM-1704; ORCH-018 / EM-ENG-012.
Notes | Return shipping of end-of-life Li-ion is itself DG-regulated; coordinate with logistics/regulatory engineers.
MR-SUS-005 — Repairability & Replaceable Battery
Requirement Statement | The market demands that serviceable elements — especially the battery — be replaceable, so a single worn component does not force the whole device to be scrapped, extending service life and cutting waste.
Business Objective | Extend usable device life and reduce whole-unit scrappage, improving both sustainability and TCO.
Customer Segment | Fleet owners, service operators, and repair-conscious/EU buyers.
Geographic Applicability | Global, weighted to EU (right-to-repair / battery-removability direction).
Market Priority | High
Importance | High
Competitive Driver | EU battery-removability rules and right-to-repair sentiment favor serviceable designs over sealed disposables.
Business Value | Lower lifetime cost per device and a stronger circular narrative; supports refurbishment revenue.
Revenue Impact | Enables battery-replacement/refurbishment service revenue and reduces attrition; magnitude Evidence Pending.
Risk if Unmet | Sealed design forces early scrappage; conflicts with EU battery removability direction.
Evidence Source | Standard + Judgment/Assumption — EU Batteries Regulation removability direction (Standard); repair-model demand for this device (Judgment/Assumption, Evidence Pending).
Supporting Research | EU Batteries Regulation removability provisions; right-to-repair policy direction.
VOC Reference | VoC theme — "don't make us bin a good device for a dead cell."
Strategic Theme | Repairable / circular.
Product Family | SCM — serviceable design.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1705, PRD-SCM-1705; ORCH-018 / EM-ENG-012.
Notes | Replaceable battery must not compromise IP sealing (IEC 60529), low-temperature safety (D3), or data integrity. Coordinate with mechanical and battery engineers.
6.3 Materials & Packaging
MR-SUS-006 — RoHS / REACH Restricted-Materials Compliance
Requirement Statement | The market demands that the device be free of restricted and high-concern substances per RoHS/REACH, with declarations available, so buyers can place it on regulated markets without materials-compliance risk.
Business Objective | Guarantee market access in regulated regions and pre-empt materials-compliance objections in procurement.
Customer Segment | All EU/regulated-market buyers and their compliance/procurement functions.
Geographic Applicability | Global, binding in the EU and adopting jurisdictions.
Market Priority | Must-have
Importance | High
Competitive Driver | Materials compliance is a hard gate for EU market placement; failure blocks sale outright.
Business Value | Ensures unimpeded market access and de-risks enterprise procurement.
Revenue Impact | Protects addressable market; failure forfeits EU revenue; magnitude Evidence Pending.
Risk if Unmet | Market exclusion, recalls, and compliance liability.
Evidence Source | Standard — EU RoHS (2011/65/EU), REACH (EC 1907/2006).
Supporting Research | RoHS/REACH substance lists and declaration practice.
VOC Reference | VoC theme — "send us the RoHS/REACH declarations before we buy."
Strategic Theme | Compliant materials.
Product Family | SCM — materials compliance.
Lifecycle Stage | Launch
Traceability | UN-SCM-1706, PRD-SCM-1706; ORCH-018 / EM-ENG-012.
Notes | Declarations flow from BOM control; coordinate with supply-chain and regulatory engineers.
MR-SUS-007 — Sustainable Packaging & Reusable Transit Protection
Requirement Statement | The market demands minimal, recyclable, and where possible reusable packaging for device shipment, return, and pooling cycles, so packaging waste does not offset the device's reuse benefit.
Business Objective | Prevent packaging from undermining the net environmental story, especially across repeated return/pooling cycles.
Customer Segment | Fleet owners, 3PLs, and ESG-conscious procurement.
Geographic Applicability | Global, weighted to EU packaging-waste regulation.
Market Priority | Medium
Importance | Medium
Competitive Driver | Packaging waste is increasingly regulated and scrutinized in ESG procurement.
Business Value | Protects the net-waste-reduction claim and reduces packaging cost across many cycles.
Revenue Impact | Marginal cost reduction plus ESG-tender support; magnitude Evidence Pending.
Risk if Unmet | Repeated-cycle packaging waste erodes the sustainability narrative.
Evidence Source | Judgment/Assumption — packaging optimization benefit is sound but unquantified for this program (Evidence Pending).
Supporting Research | EU packaging-waste direction; reusable transit-packaging practice.
VOC Reference | VoC theme — "the box shouldn't create the waste the device saves."
Strategic Theme | Whole-system waste.
Product Family | SCM — packaging.
Lifecycle Stage | Growth
Traceability | UN-SCM-1707, PRD-SCM-1707; ORCH-018 / EM-ENG-012.
Notes | Reusable transit packaging ties to the pooling model (MR-SUS-002) and reverse logistics (MR-SUS-008).
6.4 Reverse-Logistics Footprint
MR-SUS-008 — Reverse-Logistics Carbon & Efficiency
Requirement Statement | The market demands that the carbon and operational cost of recovering, cleaning, and re-provisioning reusable devices be minimized and accounted for, so the reverse-logistics burden does not cancel the benefit of reuse.
Business Objective | Ensure the net life-cycle footprint (including recovery) is genuinely lower than single-use, and account for it honestly.
Customer Segment | Fleet owners, 3PLs, and ESG procurement evaluating full life-cycle impact.
Geographic Applicability | Global, weighted to long-distance return corridors.
Market Priority | Medium
Importance | High
Competitive Driver | Sophisticated buyers know reuse only wins net-of-recovery; an honest, optimized reverse chain is a differentiator.
Business Value | Validates the net-sustainability claim and lowers reverse-logistics operating cost.
Revenue Impact | Reduces operating cost and defends the ESG value story; magnitude Evidence Pending.
Risk if Unmet | Reverse-logistics carbon/cost undermines the reuse case and invites greenwashing critique.
Evidence Source | Judgment/Assumption — reverse-logistics footprint must be measured via LCA (Evidence Pending). Reverse logistics is a evidence-backed consequence of reusability.
Supporting Research | GHG Protocol; reusable-asset reverse-logistics LCA literature.
VOC Reference | VoC theme — "count the return trips before you call it green."
Strategic Theme | Honest net footprint.
Product Family | SCM — reverse logistics.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1708, PRD-SCM-1708; ORCH-018 / EM-ENG-012.
Notes | Reverse-logistics loss/attrition (MR-1800 RISK-009) also affects the footprint per successful reuse.
MR-SUS-009 — E-Waste Take-Back Program
Requirement Statement | The market demands a manufacturer/vendor take-back pathway for retired devices, so end-of-life units are collected and responsibly recycled rather than becoming uncontrolled e-waste in the field, particularly in LMIC settings with limited recycling infrastructure.
Business Objective | Close the loop on device end-of-life and meet extended-producer-responsibility expectations.
Customer Segment | All fleet owners; especially LMIC/donor programs and EU EPR-obligated buyers.
Geographic Applicability | Global, critical where local e-waste recycling is weak (LMIC).
Market Priority | Medium
Importance | High
Competitive Driver | Extended producer responsibility and donor ESG requirements increasingly expect vendor take-back.
Business Value | Compliance, reputation, and reclaimed-material value; strengthens donor/institutional trust.
Revenue Impact | Enables EPR-compliant and donor-funded deployment; magnitude Evidence Pending.
Risk if Unmet | Uncontrolled e-waste, EPR non-compliance, and reputational/mission harm.
Evidence Source | Standard + Judgment/Assumption — WEEE/EPR obligations (Standard); take-back logistics feasibility in LMIC (Judgment/Assumption, Evidence Pending).
Supporting Research | WEEE/EPR frameworks; LMIC e-waste-handling constraints (WHO/UNICEF context).
VOC Reference | VoC theme — "who takes them back when they die in the field?"
Strategic Theme | Closed loop.
Product Family | SCM — take-back program.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1709, PRD-SCM-1709; ORCH-018 / EM-ENG-012.
Notes | Take-back shipping of Li-ion is DG-regulated (UN 38.3). Ties to MR-SUS-004 and reverse-logistics chain.
6.5 In-Use Efficiency
MR-SUS-010 — In-Use Energy Efficiency
Requirement Statement | The market demands that the device operate energy-efficiently in use — extracting maximum monitoring and connectivity value per unit of stored energy — so recharge frequency, energy consumption, and cell stress over life are minimized.
Business Objective | Reduce recharge energy and cell stress, compounding both the sustainability and cycle-life advantages.
Customer Segment | Fleet owners and operations teams managing charging at scale.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Efficient duty-cycling extends endurance and battery life — a compound advantage over power-hungry connected loggers.
Business Value | Lower operating energy, fewer recharges, and reduced cell degradation.
Revenue Impact | Reduces operating cost and extends device life; magnitude Evidence Pending.
Risk if Unmet | Frequent recharge and cell stress shorten life and increase footprint and labor.
Evidence Source | Judgment/Assumption — efficiency benefit is sound; SCM figures Evidence Pending. Duty-cycled operation is based (D3).
Supporting Research | Low-power IoT design practice; duty-cycling and energy-budgeting references.
VOC Reference | VoC theme — "we can't recharge hundreds of these every week."
Strategic Theme | Efficiency compounds.
Product Family | SCM — energy efficiency.
Lifecycle Stage | Sustaining
Traceability | UN-SCM-1710, PRD-SCM-1710; ORCH-018 / EM-ENG-012.
Notes | Efficiency gains must not suppress the D2 immediate-excursion-alert function. Ties to MR-INNO-005/007 (edge ML, energy harvesting).
7. Evidence & Validation Notes
• Anchor facts D1–D4 and the device specification are given. Reusability (D1) is the root of the sustainability thesis; the D3 low-temperature battery constraints (cycle fade, cold-charge inhibition) are the principal limiters of that thesis and must be validated by test (see MR-1800 RISK-001, MR-1900).
• Standards (WEEE, EU Batteries Regulation, RoHS, REACH, UN 38.3) are cited as binding obligations, not as achieved compliance — conformance is Evidence Pending until demonstrated.
• All environmental magnitudes (carbon, waste, cycle counts) are marked Evidence Pending until a program LCA (ISO 14040/14044) and field cycle-life data exist. No specific green number is asserted as fact.
Validation path | LCA quantifies the net reuse advantage and reverse-logistics footprint; field data confirms reuse counts and battery cycle life; compliance testing/declarations close the RoHS/REACH/WEEE/battery items. VOC (MR-1900) confirms which sustainability demands are purchase-gating.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-SUS-001 | UN-SCM-1701 | PRD-SCM-1701 | ORCH-018 / EM-ENG-012
MR-SUS-002 | UN-SCM-1702 | PRD-SCM-1702 | ORCH-018 / EM-ENG-012
MR-SUS-003 | UN-SCM-1703 | PRD-SCM-1703 | ORCH-018 / EM-ENG-012
MR-SUS-004 | UN-SCM-1704 | PRD-SCM-1704 | ORCH-018 / EM-ENG-012
MR-SUS-005 | UN-SCM-1705 | PRD-SCM-1705 | ORCH-018 / EM-ENG-012
MR-SUS-006 | UN-SCM-1706 | PRD-SCM-1706 | ORCH-018 / EM-ENG-012
MR-SUS-007 | UN-SCM-1707 | PRD-SCM-1707 | ORCH-018 / EM-ENG-012
MR-SUS-008 | UN-SCM-1708 | PRD-SCM-1708 | ORCH-018 / EM-ENG-012
MR-SUS-009 | UN-SCM-1709 | PRD-SCM-1709 | ORCH-018 / EM-ENG-012
MR-SUS-010 | UN-SCM-1710 | PRD-SCM-1710 | ORCH-018 / EM-ENG-012
Related modules: MR-000 (EXEC) frames reusable-platform value; MR-1800 (RISK) governs battery, reverse-logistics, and regulatory risks; MR-1900 (VAL) validates which sustainability demands gate procurement.
MR-1800 — Strategic Risks (RISK)
Traceability: this module rolls up to the SCM Master Market Requirements Library and traces forward to PRD-SCM-18xx (Product Requirements) and UN-SCM-18xx (User Needs), and back to ORCH-018 / EM-ENG-012 (orchestration + originating engineering mandate). It records the market-level strategic risks that could prevent the SCM from achieving its commercial and mission goals, and the market demands that mitigate them.
1. Purpose
This module identifies the strategic, market-level risks that could cause the SCM program to fail commercially or in mission even if the engineering is sound — and states, as market requirements, what the market demands the program do to de-risk. It exists so leadership funds mitigations deliberately and so the single most dangerous technical-commercial tension — the D3 low-temperature battery endurance and cold-charge-safety problem — is treated as a top strategic risk to be closed by test, not assumed away. It is written for the program leadership, investment committee, and risk owners who must decide what to prove, insure, or design around before scaling. Risks here are severity-first (per org rule 3): safety and loss-of-core-function dominate.
2. Scope
In scope: the D3 low-temperature battery endurance/charge-safety risk; delayed or missing excursion alerts from connectivity gaps; WHO PQS non-qualification; competitive disruption by incumbents; cellular carrier/roaming/eSIM dependency; single-source low-temperature cell supply; data-integrity and cyber breach; regulatory divergence across markets; reverse-logistics loss/device attrition; economic dependence on GAVI/donor tender funding; and false-alarm erosion of user trust. Each risk is paired with the market demand that mitigates it.
Out of scope: the detailed engineering FMEA/hazard analysis (that lives in the HAL/ERFKL and discipline modules), the quantitative risk register with likelihood/severity scoring, and program-level contingency budgets. This module states market-level risk demands; it does not itself perform the safety analysis. Risk scoring uses Action Priority framing (severity-dominant), consistent with org standards, but detailed S/O/D scoring lives downstream.
3. Applicable Standards & References
• WHO PQS — prequalification is a market gate; non-qualification is a top commercial risk (category code TBC).
• EU GDP, US FDA / 21 CFR Part 11 — data-integrity and monitoring obligations whose breach is both a compliance and trust risk.
• UN 38.3, IATA DG — lithium-battery transport qualification; cold-charge safety and cell behaviour bound the D3 risk.
• IEC 62368-1 (product safety), IEC 60068 (environmental testing), IEC 60529 (IP) — safety/environment validation relevant to the battery and endurance risks.
• FCC Part 15 / CE RED, CISPR / IEC 61000 — radio and EMC conformity affecting connectivity reliability and market access.
• WHO / GAVI / UNICEF Supply Division — procurement and funding context; donor-funding dependence is a market risk cited by name.
• Cross-references: MR-000 (EXEC), MR-1600 (INNO), MR-1700 (SUS), MR-1900 (VAL); HAL / ERFKL, ROSTER.md, ORCHESTRATION.md, and the SCM engineering discipline modules.
4. Background
The SCM's promise — reliable, provable, real-time cold-chain visibility from inside the shipper — collides with hard realities the market already understands. The most acute is physics: a low-temperature Li-ion pack that must be charge-inhibited below 0 °C (D3) and that fades over reuse cycles, operating in a device that claims 90-day endurance across -30 °C to +60 °C. If endurance or cold-charge safety fails, the device goes dark or becomes unsafe precisely in the frozen chains it targets. This is the program's defining risk and the one the market will most insist be proven by test.
Around that core sit market and ecosystem risks: connectivity gaps that swallow the real-time-alert promise on remote routes; PQS non-qualification that locks the SCM out of the largest procurement channel; well-capitalized incumbents; dependence on carrier roaming/eSIM ecosystems and on single-source specialty cells; the ever-present cyber and data-integrity threat to an audit-grade evidence product; regulatory divergence across WHO/EU/US/transport regimes; attrition of expensive reusable devices in reverse logistics; concentration of demand in donor/GAVI-funded tenders; and the insidious erosion of trust from false alarms. Each is stated below as a risk plus the market demand that mitigates it, severity-first.
5. Requirement Categories
• Physical / Battery Risk — the D3 low-temperature endurance and charge-safety KEY RISK.
• Function-Delivery Risk — connectivity gaps, carrier dependence, false-alarm trust erosion.
• Market-Access Risk — PQS non-qualification, regulatory divergence.
• Competitive & Supply Risk — incumbent disruption, single-source cells.
• Data & Security Risk — data-integrity/cyber breach.
• Commercial Model Risk — reverse-logistics attrition, donor-funding dependence.
6. Detailed Market Requirements
6.1 Physical / Battery Risk (KEY)
MR-RISK-001 — Low-Temperature Battery Endurance & Charge-Safety (KEY RISK)
Requirement Statement | The market demands proven, safe device operation across the full -30 °C to +60 °C range for the stated endurance — including safe handling of the below-0 °C charge inhibition and resilience to capacity fade over reuse cycles — so the device does not go dark, mislog, or become unsafe in the frozen chains it is sold to protect.
Business Objective | Close the program's defining technical-commercial risk by proving cold endurance and charge safety before scaling.
Customer Segment | All segments, critically frozen/ultra-cold operators, LMIC last-mile, and long-haul 3PLs.
Geographic Applicability | Global, acute in cold-climate and frozen-chain operations.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Cold-temperature reliability is the make-or-break credibility test buyers apply to any battery-powered in-shipper monitor.
Business Value | Proving it converts the program's biggest liability into its strongest reliability claim.
Revenue Impact | Existential — failure invalidates the core use case and reusable model; magnitude Evidence Pending.
Risk if Unmet | Devices die or misbehave at low temperature; loss of the D2 alert function, unsafe charging, safety incident, and total loss of market trust.
Evidence Source | Judgment/Assumption — endurance at -30 °C, cold-charge safety, and cycle fade are the stated KEY RISKS and are UNPROVEN until tested (Evidence Pending). Cold-charge inhibition below 0 °C and Li-ion low-temperature limits are evidence-backed (D3).
Supporting Research | Low-temperature Li-ion behaviour and charge-safety literature; IEC 60068 environmental test methods; UN 38.3.
VOC Reference | VOC-180 (pending) — "will it survive and stay safe at -30 for the whole trip, trip after trip?"
Strategic Theme | Prove the physics.
Product Family | SCM (Smart Cold Chain Monitor) — battery/endurance.
Lifecycle Stage | Concept
Traceability | UN-SCM-1801, PRD-SCM-1801; ORCH-018 / EM-ENG-012.
Notes | TOP STRATEGIC RISK. Must be validated by test (MR-1900 endurance/charge-safety validation). Ties to MR-INNO-007/008 (harvesting, next-gen chemistry) and MR-SUS-003 (cycle life). Coordinate with battery/power, thermal, and safety engineers; feeds HAL/ERFKL.
6.2 Function-Delivery Risk
MR-RISK-002 — Delayed or Missing Excursion Alert (Connectivity Gap)
Requirement Statement | The market demands that excursion alerts still be captured and ultimately delivered even when connectivity is intermittent or absent, so a coverage gap never silently swallows the single most important signal the product exists to deliver.
Business Objective | Guarantee the integrity of the core alert promise regardless of network conditions.
Customer Segment | LMIC last-mile, remote/ocean routes, and any low-coverage deployment.
Geographic Applicability | Global, acute in coverage-poor corridors.
Market Priority | Must-have
Importance | Critical
Competitive Driver | A monitor that goes silent where risk is highest fails its purpose; store-and-forward reliability is a differentiator.
Business Value | Preserves the trust that underpins the entire value proposition.
Revenue Impact | Failures here drive churn and reputational loss; magnitude Evidence Pending.
Risk if Unmet | Missed/late alerts, spoiled product delivered unknowingly, and collapse of trust in the alert function.
Evidence Source | Standard + Judgment/Assumption — local logging + cadence upload + immediate alerts is based (D2); real-world coverage-gap impact is Evidence Pending.
Supporting Research | GDP continuous-monitoring expectations; WHO/UNICEF last-mile connectivity gaps; store-and-forward telemetry practice.
VOC Reference | VoC theme — "it can't lose the alert just because there's no signal."
Strategic Theme | Never miss the signal.
Product Family | SCM — alert reliability.
Lifecycle Stage | Launch
Traceability | UN-SCM-1802, PRD-SCM-1802; ORCH-018 / EM-ENG-012.
Notes | Local logging must guarantee capture; satellite/NTN (MR-INNO-010) and edge ML (MR-INNO-005) mitigate. Ties to MR-RISK-005 (carrier dependence).
MR-RISK-011 — False-Alarm Erosion of Trust
Requirement Statement | The market demands a low false-alarm rate, so that operators keep trusting and acting on alerts rather than becoming desensitized and ignoring a genuine excursion — a failure mode as dangerous as a missed alert.
Business Objective | Protect the credibility of the alerting system, on which adoption and renewal depend.
Customer Segment | All operational segments managing devices at scale.
Geographic Applicability | Global.
Market Priority | High
Importance | Critical
Competitive Driver | Alarm fatigue is a leading complaint against connected loggers; trustworthy alerting differentiates.
Business Value | Sustains adoption, action-on-alert behaviour, and renewals.
Risk if Unmet | Operators ignore alerts; a real excursion is missed amid noise; trust and renewals collapse.
Revenue Impact | Drives churn and undermines premium positioning; magnitude Evidence Pending.
Evidence Source | Judgment/Assumption — alarm-fatigue phenomenon is well documented in analogous domains (Evidence Pending for SCM rates).
Supporting Research | Alarm-management literature; MR-INNO-004/005 (AI/edge false-alarm suppression).
VOC Reference | VOC-181 (pending) — "too many false alarms and people just stop looking."
Strategic Theme | Trustworthy alerting.
Product Family | SCM — alert quality.
Lifecycle Stage | Growth
Traceability | UN-SCM-1811, PRD-SCM-1811; ORCH-018 / EM-ENG-012.
Notes | Severity high because it can cause a genuine excursion to be ignored. Mitigated by MR-INNO-004/005; validated by MR-1900 alarm-efficacy validation.
6.3 Market-Access Risk
MR-RISK-003 — WHO PQS Non-Qualification
Requirement Statement | The market demands that the device achieve WHO PQS qualification, because without it the SCM is locked out of the largest immunization-procurement channel regardless of its technical merits.
Business Objective | Secure access to WHO/UNICEF/GAVI-linked procurement by achieving and maintaining PQS status.
Customer Segment | National immunization programs, WHO/UNICEF/GAVI procurement, ministries of health.
Geographic Applicability | Global public-health procurement, weighted to LMICs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | PQS is a hard procurement gate; non-qualified products are simply not eligible for the highest-volume channel.
Business Value | Unlocks the largest addressable volume and the mission narrative.
Revenue Impact | Non-qualification forfeits the primary public-health channel; magnitude Evidence Pending.
Risk if Unmet | Exclusion from donor-funded and PQS-gated tenders; loss of the highest-volume market.
Evidence Source | Standard — WHO PQS prequalification framework (category code TBC).
Supporting Research | WHO PQS device categories and qualification process; GAVI/UNICEF procurement linkage.
VOC Reference | VoC theme — "no PQS, no tender."
Strategic Theme | Earn the gate.
Product Family | SCM — PQS qualification.
Lifecycle Stage | Launch
Traceability | UN-SCM-1803, PRD-SCM-1803; ORCH-018 / EM-ENG-012.
Notes | PQS lab qualification is a named validation activity (MR-1900). Confirm correct monitor/logger category early. Ties to MR-RISK-010 (donor funding).
MR-RISK-008 — Regulatory Divergence Across Markets
Requirement Statement | The market demands a single platform that satisfies WHO PQS, EU GDP, US FDA/Part 11, and air/ground transport regimes at once, because divergent and evolving requirements across markets threaten to fragment the product or strand it in one region.
Business Objective | Preserve the one-platform-many-markets thesis against regulatory fragmentation.
Customer Segment | Cross-border manufacturers, distributors, and 3PLs operating in multiple regimes.
Geographic Applicability | Multi-market (EU, US, transport, LMIC).
Market Priority | High
Importance | High
Competitive Driver | Divergent rules push competitors into region-specific products; a genuinely multi-regime device is the differentiator (MR-EXEC-004).
Business Value | Protects the platform-consolidation advantage and avoids costly per-market variants.
Revenue Impact | Divergence handled well protects multi-market revenue; handled poorly forces fragmentation cost; magnitude Evidence Pending.
Risk if Unmet | Product forks per market, eroding margin, or is excluded from a regime by a missed requirement.
Evidence Source | Standard — WHO PQS, EU GDP, FDA/21 CFR Part 11, IATA DG (D4).
Supporting Research | Comparative WHO/EU/US/transport requirement mapping; RED/FCC radio conformity.
VOC Reference | VoC theme — "one device that clears every regulator we ship under."
Strategic Theme | Multi-regime by design.
Product Family | SCM — regulatory strategy.
Lifecycle Stage | Launch
Traceability | UN-SCM-1808, PRD-SCM-1808; ORCH-018 / EM-ENG-012.
Notes | Radio/EMC conformity (FCC/RED/CISPR) also affects market access. Coordinate with regulatory and RF engineers.
6.4 Competitive & Supply Risk
MR-RISK-004 — Competitive Disruption by Incumbents
Requirement Statement | The market demands durable differentiation, because well-capitalized incumbent logger/monitoring vendors and logistics-telemetry platforms could replicate features and out-distribute a new entrant on price and channel.
Business Objective | Build and defend a moat (multi-market compliance, data integrity, reuse economics, analytics) that incumbents cannot quickly copy.
Customer Segment | All commercial segments; especially enterprise accounts incumbents already hold.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Incumbents have installed base, channel, and capital; differentiation must be defensible, not feature-parity.
Business Value | Sustains pricing power and share against fast-following incumbents.
Revenue Impact | Weak differentiation compresses price and share; magnitude Evidence Pending.
Risk if Unmet | Commoditization; incumbents absorb the category and the SCM competes only on price.
Evidence Source | Judgment/Assumption — competitive-landscape assessment pending (Evidence Pending). Multi-market/data-integrity/reuse moat is based (D1, D4).
Supporting Research | Cold-chain monitoring competitive landscape (to be compiled); MR-EXEC-004/005 differentiation.
VOC Reference | VoC theme — "why you and not the vendor we already use?"
Strategic Theme | Defensible moat.
Product Family | SCM — competitive strategy.
Lifecycle Stage | Concept
Traceability | UN-SCM-1804, PRD-SCM-1804; ORCH-018 / EM-ENG-012.
Notes | Moat rests on the hardest-to-copy elements: proven cold endurance (MR-RISK-001), PQS+GDP+FDA span, and data integrity. Validate via MR-1900 competitive/WTP work.
MR-RISK-005 — Cellular Carrier / Roaming / eSIM Dependency
Requirement Statement | The market demands connectivity that works reliably across borders and carriers without being hostage to a single network's coverage, roaming policy, or eSIM/roaming regulatory changes, so cross-border shipments stay connected end to end.
Business Objective | De-risk the connectivity ecosystem dependence that could silently degrade the alert function across regions.
Customer Segment | Cross-border distributors, 3PLs, and multi-country immunization programs.
Geographic Applicability | Global, acute on cross-border and roaming-restricted routes.
Market Priority | High
Importance | High
Competitive Driver | Roaming/eSIM policy shifts and permanent-roaming restrictions can strand cellular-only devices; multi-network resilience differentiates.
Business Value | Protects the real-time-alert promise across the global routes the SCM targets.
Revenue Impact | Connectivity failures drive churn and support cost; magnitude Evidence Pending.
Risk if Unmet | Devices lose connectivity across borders; alerts delayed/lost; support burden rises.
Evidence Source | Judgment/Assumption — carrier/roaming/eSIM dependence is a known IoT risk (Evidence Pending for SCM impact). LTE-M primary with Wi-Fi/BLE is based.
Supporting Research | Global IoT connectivity/roaming and eSIM (GSMA) practice; permanent-roaming regulatory constraints.
VOC Reference | VoC theme — "it can't drop off the network when it crosses a border."
Strategic Theme | Connectivity resilience.
Product Family | SCM — connectivity risk.
Lifecycle Stage | Launch
Traceability | UN-SCM-1805, PRD-SCM-1805; ORCH-018 / EM-ENG-012.
Notes | Multi-IMSI/eSIM and store-and-forward mitigate; satellite/NTN (MR-INNO-010) is the deeper hedge. Ties to MR-RISK-002.
MR-RISK-006 — Single-Source Low-Temperature Cell Supply
Requirement Statement | The market demands supply-chain resilience for the specialty low-temperature cells the device depends on, because single-sourcing a critical, hard-to-substitute component exposes the program to shortage, price shocks, and stop-ship.
Business Objective | Avoid a single point of failure in the supply of the component most central to the D3 performance and safety.
Customer Segment | Fleet buyers relying on assured supply and lead times; internal supply/operations.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Specialty low-temperature cells are less commoditized; concentration risk is real and can halt production.
Business Value | Assured supply protects delivery commitments and pricing stability.
Revenue Impact | A cell shortage can stop shipments and forfeit tenders; magnitude Evidence Pending.
Risk if Unmet | Production stoppage, price shock, or forced redesign under duress.
Evidence Source | Judgment/Assumption — sourcing landscape for low-temperature cells pending (Evidence Pending). Dependence on low-temperature cells is based (D3).
Supporting Research | Battery supply-chain concentration analyses; second-source qualification practice.
VOC Reference | VoC theme — "can you guarantee supply and lead time at fleet scale?"
Strategic Theme | Supply resilience.
Product Family | SCM — supply chain.
Lifecycle Stage | Launch
Traceability | UN-SCM-1806, PRD-SCM-1806; ORCH-018 / EM-ENG-012.
Notes | Second-source qualification interacts with UN 38.3 re-qualification and MR-INNO-008. Coordinate with supply-chain and battery engineers.
6.5 Data & Security Risk
MR-RISK-007 — Data-Integrity / Cybersecurity Breach
Requirement Statement | The market demands that monitoring data and the connected platform be secure and tamper-resistant end to end, because a breach or data-integrity failure destroys the audit-grade-evidence promise that is the product's core differentiator.
Business Objective | Protect the confidentiality, integrity, and availability of data on which disposition decisions and the brand depend.
Customer Segment | Manufacturers, regulators, QA/QP release functions, and all connected-platform users.
Geographic Applicability | Global, weighted to Part 11 / GDP-regulated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Regulated buyers reject data that cannot be trusted or secured; a breach is an existential trust event.
Business Value | Preserves the data-integrity moat (MR-EXEC-005) and regulatory standing.
Revenue Impact | A breach forfeits trust, contracts, and regulated-market access; magnitude Evidence Pending.
Risk if Unmet | Tampered/leaked data, failed audits, regulatory action, and loss of the core value proposition.
Evidence Source | Standard — 21 CFR Part 11 data integrity; EU GDP; general IoT security practice.
Supporting Research | FDA data-integrity guidance; IoT security frameworks; MR-INNO-006 chain-of-custody assurance.
VOC Reference | VoC theme — "if the data can be tampered with, it's worthless to us."
Strategic Theme | Trust and security.
Product Family | SCM — data security.
Lifecycle Stage | Launch
Traceability | UN-SCM-1807, PRD-SCM-1807; ORCH-018 / EM-ENG-012.
Notes | Coordinate with cybersecurity engineer; ties to MR-INNO-006 and the OTA/BLE-provisioning attack surface.
6.6 Commercial Model Risk
MR-RISK-009 — Reverse-Logistics Loss / Device Attrition
Requirement Statement | The market demands that reusable devices be reliably recovered after each trip, because attrition, loss, and non-return in reverse logistics directly destroys the reuse count on which the entire reusable economic and sustainability model depends.
Business Objective | Protect the reuse-count that underpins TCO, margin, and the sustainability thesis.
Customer Segment | Fleet owners, 3PLs, and multi-party chains managing returns.
Geographic Applicability | Global, acute on last-mile/LMIC routes where recovery is hardest.
Market Priority | High
Importance | High
Competitive Driver | Reusable economics only beat single-use if recovery is high; poor recovery quietly turns a reusable device into an expensive single-use one.
Business Value | High recovery preserves the business model's core assumption.
Revenue Impact | Attrition erodes margin and inflates effective cost-per-trip; magnitude Evidence Pending.
Risk if Unmet | The reusable model collapses economically and environmentally.
Evidence Source | Judgment/Assumption — attrition rates unknown and must be measured in pilots (Evidence Pending). Reverse logistics is a evidence-backed consequence of reusability.
Supporting Research | Returnable-asset recovery-rate literature; reverse-logistics loss studies.
VOC Reference | VoC theme — "we won't get half of them back — then what's the math?"
Strategic Theme | Recovery = viability.
Product Family | SCM — reverse logistics.
Lifecycle Stage | Growth
Traceability | UN-SCM-1809, PRD-SCM-1809; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-EXEC-008 (investment case), MR-SUS-001/008 (sustainability), and MR-1900 pilot validation of recovery rates.
MR-RISK-010 — Economic / Tender-Funding Dependence on Donors
Requirement Statement | The market demands a business model that is not fatally dependent on a single donor or funding cycle, because concentration of demand in GAVI/donor-funded tenders exposes the program to funding-cycle and policy shifts outside its control.
Business Objective | Diversify revenue across donor, public, and commercial channels to withstand funding volatility.
Customer Segment | Immunization programs (donor-funded) plus commercial pharma, distributors, and 3PLs.
Geographic Applicability | Global; donor concentration highest in LMIC public-health procurement.
Market Priority | High
Importance | High
Competitive Driver | Over-reliance on donor tenders makes revenue hostage to funding cycles; diversified demand is more resilient.
Business Value | Revenue diversification stabilizes the program across funding cycles.
Revenue Impact | Donor-cycle shifts can swing volume sharply; diversification dampens the swing; magnitude Evidence Pending.
Risk if Unmet | A single funding-cycle or policy change materially threatens program revenue and continuity.
Evidence Source | Judgment/Assumption — donor-funding concentration is a recognized dynamic in immunization supply (Evidence Pending for SCM exposure).
Supporting Research | GAVI/UNICEF/WHO procurement-funding structures; immunization-financing literature.
VOC Reference | VoC theme — "what happens to you if the donor budget moves?"
Strategic Theme | Revenue resilience.
Product Family | SCM — commercial model.
Lifecycle Stage | Growth
Traceability | UN-SCM-1810, PRD-SCM-1810; ORCH-018 / EM-ENG-012.
Notes | Balances the mission channel (MR-EXEC-003) with commercial diversification. Ties to MR-1900 tender/procurement validation.
7. Evidence & Validation Notes
• Anchor facts D1–D4 and the device specification are given. MR-RISK-001 (low-temperature battery endurance and charge safety) is the TOP STRATEGIC RISK and is explicitly UNPROVEN — it must be closed by test (IEC 60068 environmental testing, cold-charge safety, and reuse-cycle-life validation) per MR-1900.
• Standards (PQS, GDP, Part 11, UN 38.3, IEC 62368-1/60068/60529, FCC/RED/CISPR) define hard gates; conformance is Evidence Pending until demonstrated.
• All rates and magnitudes (attrition, false-alarm rate, coverage-gap impact, competitive share, donor exposure) are marked Evidence Pending pending pilot, VOC, and market-analysis data. No precise number is asserted as fact.
Severity-first framing | MR-RISK-001, -002, -003, -007, and -011 carry Critical importance because they threaten safety or loss of the core function/market. Detailed S/O/D and Action Priority scoring live in the HAL/ERFKL; this module states the market-level demand and severity emphasis.
Validation path | MR-1900 (endurance/charge-safety validation, PQS lab qualification, alarm-efficacy, field pilots, WTP/tender validation) converts these Judgment/Assumption risks into cited evidence before scale-up.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-RISK-001 | UN-SCM-1801 | PRD-SCM-1801 | ORCH-018 / EM-ENG-012
MR-RISK-002 | UN-SCM-1802 | PRD-SCM-1802 | ORCH-018 / EM-ENG-012
MR-RISK-003 | UN-SCM-1803 | PRD-SCM-1803 | ORCH-018 / EM-ENG-012
MR-RISK-004 | UN-SCM-1804 | PRD-SCM-1804 | ORCH-018 / EM-ENG-012
MR-RISK-005 | UN-SCM-1805 | PRD-SCM-1805 | ORCH-018 / EM-ENG-012
MR-RISK-006 | UN-SCM-1806 | PRD-SCM-1806 | ORCH-018 / EM-ENG-012
MR-RISK-007 | UN-SCM-1807 | PRD-SCM-1807 | ORCH-018 / EM-ENG-012
MR-RISK-008 | UN-SCM-1808 | PRD-SCM-1808 | ORCH-018 / EM-ENG-012
MR-RISK-009 | UN-SCM-1809 | PRD-SCM-1809 | ORCH-018 / EM-ENG-012
MR-RISK-010 | UN-SCM-1810 | PRD-SCM-1810 | ORCH-018 / EM-ENG-012
MR-RISK-011 | UN-SCM-1811 | PRD-SCM-1811 | ORCH-018 / EM-ENG-012
Related modules: MR-000 (EXEC) sets the goals these risks threaten; MR-1600 (INNO) and MR-1700 (SUS) carry risks governed here; MR-1900 (VAL) is the primary vehicle for closing them; HAL/ERFKL carry the detailed safety and failure analysis.
MR-1900 — Market Validation (VAL)
Traceability: this module rolls up to the SCM Master Market Requirements Library and traces forward to PRD-SCM-19xx (Product Requirements) and UN-SCM-19xx (User Needs), and back to ORCH-018 / EM-ENG-012 (orchestration + originating engineering mandate). It defines the market's demand for evidence — the validation activities that must convert the library's assumptions into proof before scale-up.
1. Purpose
This module states what evidence the market — and the program's own investment discipline — requires before the SCM is scaled, and how that evidence will be obtained. Much of the Master Market Requirements Library rests on evidence-backed device facts plus explicit assumptions; this module is the mechanism that retires those assumptions with real data. It is the demand-side counterpart to MR-1800 (Strategic Risks): where RISK names what could go wrong, VAL names how the program will prove it will not. It carries a particular obligation to validate, by test, the D3 low-temperature battery endurance and charge-safety KEY RISK. It is written for product strategy, program leadership, the investment committee, and the pilot/quality owners who run the studies.
2. Scope
In scope: market demand for customer discovery interviews, field pilots with distributors and EPI/immunization programs, WHO PQS lab qualification testing, beta programs with 3PLs, LMIC last-mile field studies, sensor accuracy/calibration validation, endurance validation at -30 °C (and charge-safety/cycle-life), alarm-efficacy validation, usability studies with frontline health workers, willingness-to-pay/pricing validation, tender/procurement validation, and closed-loop sales-and-support feedback. Each is expressed as a market demand for evidence.
Out of scope: the detailed test protocols, statistical designs, sample-size calculations, and acceptance thresholds — those live in the Verification/Validation libraries (PVL/PVaL) and discipline modules. This module states which evidence the market requires and why; it does not itself execute the studies or set numeric pass/fail criteria (those are Evidence Pending until designed with data owners).
3. Applicable Standards & References
• WHO PQS — prequalification lab qualification and field-evaluation practice; a primary validation gate (category code TBC).
• EU GDP, US FDA / 21 CFR Part 11 — monitoring performance and data-integrity expectations that validation must demonstrate.
• IEC 60068 (environmental testing), IEC 60529 (IP), IEC 62368-1 (safety) — environmental/endurance/safety validation methods, including cold-temperature testing.
• UN 38.3 / IATA DG — battery transport qualification tied to endurance/charge-safety validation.
• ISO 17025 / metrology calibration practice — reference-standard traceability for accuracy/calibration validation (method reference).
• WHO / GAVI / UNICEF Supply Division — pilot, field-evaluation, and procurement-validation context (cited by name).
• Cross-references: MR-000 (EXEC), MR-1600 (INNO), MR-1700 (SUS), MR-1800 (RISK); PVL/PVaL, ROSTER.md, ORCHESTRATION.md, and the SCM engineering discipline modules.
4. Background
The SCM library asserts a compelling thesis — reusable, multi-market, data-integrity-grade monitoring that prevents vaccine loss — but a thesis is not proof. Institutional buyers (WHO/UNICEF/GAVI, ministries of health) and enterprise buyers (manufacturers, distributors, 3PLs) both de-risk platform decisions by demanding evidence: a clean pilot, PQS qualification, demonstrated cold endurance, proven accuracy, and credible economics. The program's own capital discipline demands the same. The most consequential unknown is the D3 low-temperature battery behaviour: endurance at -30 °C, safe handling of the below-0 °C charge inhibition, and capacity fade across reuse cycles are the KEY RISKS (MR-1800 RISK-001) and must be closed by physical test before any scale commitment.
Validation is sequenced from cheap-and-early (customer discovery, usability) to expensive-and-gating (PQS lab qualification, field pilots, endurance testing). It runs as a loop: each study either confirms an assumption (upgrading its Evidence Source from Judgment/Assumption to Cited) or forces a design/strategy change. This module names the studies the market requires; it deliberately does not pre-declare their numeric outcomes, which would violate the evidence-honesty rule.
5. Requirement Categories
• Customer Discovery & VOC — interviews and voice-of-customer capture.
• Field & Pilot Validation — distributor/EPI pilots, 3PL beta, LMIC last-mile studies.
• Technical Validation — accuracy/calibration, -30 °C endurance & charge safety, alarm efficacy, PQS lab qualification.
• Usability Validation — frontline health-worker studies.
• Commercial Validation — willingness-to-pay/pricing, tender/procurement, sales-and-support feedback.
6. Detailed Market Requirements
6.1 Customer Discovery & VOC
MR-VAL-001 — Customer Discovery Interviews
Requirement Statement | The market demands that the program validate the pain, priorities, and buying criteria of each target segment through structured customer interviews before committing scope, so requirements reflect real demand rather than internal assumption.
Business Objective | Ground the requirement library in evidenced voice-of-customer and prioritize the roadmap by real demand.
Customer Segment | Manufacturers, distributors, 3PLs, immunization-program managers, procurement bodies.
Geographic Applicability | Global, sampled across regulated and LMIC markets.
Market Priority | Must-have
Importance | High
Competitive Driver | Products built on evidenced VOC out-position those built on assumption; buyers notice.
Business Value | De-risks scope and roadmap; converts VoC themes into cited VOC references across the library.
Revenue Impact | Improves product-market fit and win rate; magnitude Evidence Pending.
Risk if Unmet | Requirements drift from real demand; wrong features built.
Evidence Source | Judgment/Assumption — VOC program pending execution (Evidence Pending); it is the mechanism that grounds other modules.
Supporting Research | Customer-discovery / VOC methodology; segment definitions in MR-200.
VOC Reference | VOC-190 (pending) — establishes the VOC-### series the library references.
Strategic Theme | Evidence-led scope.
Product Family | SCM (Smart Cold Chain Monitor) — discovery.
Lifecycle Stage | Concept
Traceability | UN-SCM-1901, PRD-SCM-1901; ORCH-018 / EM-ENG-012.
Notes | Outputs populate the pending VOC-### references throughout MR-1600/1700/1800. Run before scope lock.
6.2 Field & Pilot Validation
MR-VAL-002 — Field Pilots with Distributors & EPI Programs
Requirement Statement | The market demands controlled field pilots with real distributors and national immunization (EPI) programs, so the device's end-to-end performance and value are demonstrated in operational conditions before standardization.
Business Objective | Prove real-world efficacy and build the reference deployments buyers require before scaling.
Customer Segment | Distributors, national immunization/EPI programs, ministries of health.
Geographic Applicability | Global, weighted to representative LMIC and regulated deployments.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Buyers gate standardization on a clean pilot; reference deployments are decisive in procurement.
Business Value | Generates the operational evidence and references that unlock scale-up.
Revenue Impact | Gates enterprise/institutional standardization; magnitude Evidence Pending.
Risk if Unmet | No operational proof; adoption stalls despite technical readiness (MR-EXEC-007).
Evidence Source | Judgment/Assumption — pilots pending (Evidence Pending); aligns with PQS field-evaluation practice.
Supporting Research | WHO PQS field-evaluation practice; pilot-design references; GDP monitoring expectations.
VOC Reference | VoC theme — "show us a clean pilot in our chain first."
Strategic Theme | Prove in the field.
Product Family | SCM — field pilot.
Lifecycle Stage | Launch
Traceability | UN-SCM-1902, PRD-SCM-1902; ORCH-018 / EM-ENG-012.
Notes | Pilots also measure reverse-logistics recovery (MR-1800 RISK-009) and real coverage-gap behaviour (RISK-002).
MR-VAL-003 — Beta Program with 3PLs
Requirement Statement | The market demands a beta program with third-party logistics providers, so the device is validated in high-volume commercial logistics workflows and integrations distinct from public-health deployments.
Business Objective | Validate commercial-channel fit, integration, and scale behaviour with sophisticated logistics operators.
Customer Segment | 3PLs and commercial logistics/control-tower operators.
Geographic Applicability | Global commercial logistics corridors.
Market Priority | High
Importance | High
Competitive Driver | 3PLs demand proof of integration and reliability at volume before adopting a new monitor.
Business Value | Validates the commercial channel and integration/API requirements (ties to MR-INNO-003).
Revenue Impact | Unlocks commercial-channel revenue diversification (MR-1800 RISK-010); magnitude Evidence Pending.
Risk if Unmet | Commercial channel unproven; over-reliance on donor channel persists.
Evidence Source | Judgment/Assumption — beta pending (Evidence Pending).
Supporting Research | 3PL telemetry-integration practice; enterprise beta-program design.
VOC Reference | VoC theme — "prove it plugs into our systems at scale."
Strategic Theme | Commercial-channel proof.
Product Family | SCM — 3PL beta.
Lifecycle Stage | Launch
Traceability | UN-SCM-1903, PRD-SCM-1903; ORCH-018 / EM-ENG-012.
Notes | Complements EPI pilots (MR-VAL-002); tests integration and roaming/connectivity resilience (RISK-005).
MR-VAL-004 — LMIC Last-Mile Field Studies
Requirement Statement | The market demands field studies in genuine LMIC last-mile conditions — intermittent connectivity, high ambient heat, constrained charging, and difficult recovery — so performance is proven where cold-chain risk and mission value are highest.
Business Objective | Prove the device delivers in the hardest, highest-mission-value environments, not just benign ones.
Customer Segment | LMIC immunization programs, donor-funded deployments, last-mile health workers.
Geographic Applicability | LMIC last-mile settings across regions.
Market Priority | High
Importance | Critical
Competitive Driver | Many monitors perform in benign conditions but fail in the last mile; proof here is a strong differentiator and mission credential.
Business Value | Validates the mission thesis (MR-EXEC-003) and the toughest technical assumptions at once.
Revenue Impact | Underpins donor/PQS-channel credibility; magnitude Evidence Pending.
Risk if Unmet | Device fails or under-performs in the exact conditions it is sold to address.
Evidence Source | Judgment/Assumption — LMIC field studies pending (Evidence Pending). Connectivity/heat/charging constraints are evidence-backed market context.
Supporting Research | WHO/UNICEF last-mile cold-chain field-evaluation practice; LMIC connectivity/charging constraints.
VOC Reference | VoC theme — "prove it in our worst conditions, not a lab."
Strategic Theme | Prove the hardest case.
Product Family | SCM — LMIC field study.
Lifecycle Stage | Launch
Traceability | UN-SCM-1904, PRD-SCM-1904; ORCH-018 / EM-ENG-012.
Notes | Directly stresses RISK-002 (connectivity), RISK-009 (recovery), and D3 endurance in heat as well as cold.
6.3 Technical Validation
MR-VAL-005 — Sensor Accuracy & Calibration Validation
Requirement Statement | The market demands demonstrated, traceable accuracy and calibration of the temperature (internal + external probe), humidity, and shock sensing, so the data can serve as audit-grade evidence for disposition and regulatory purposes.
Business Objective | Prove the measurement integrity on which the entire data-integrity value proposition rests.
Customer Segment | Manufacturers, QA/QP release, regulators, and all data-dependent users.
Geographic Applicability | Global, binding in regulated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Data that cannot demonstrate traceable accuracy is rejected in audits and disputes.
Business Value | Substantiates the audit-grade-evidence claim (MR-EXEC-005).
Revenue Impact | Gates regulated-market and disposition use; magnitude Evidence Pending.
Risk if Unmet | Data challenged in audit/dispute; core promise fails (ties to RISK-007).
Evidence Source | Standard — metrology/calibration practice (ISO 17025 lineage); GDP/Part 11 data expectations.
Supporting Research | Reference-standard calibration and uncertainty practice; GDP monitoring accuracy expectations.
VOC Reference | VoC theme — "prove the numbers are accurate and traceable."
Strategic Theme | Trustworthy measurement.
Product Family | SCM — metrology validation.
Lifecycle Stage | Launch
Traceability | UN-SCM-1905, PRD-SCM-1905; ORCH-018 / EM-ENG-012.
Notes | Accuracy must hold across the -30 °C to +60 °C range; ultra-cold (MR-INNO-009) needs separate validation. Acceptance thresholds live in PVL.
MR-VAL-006 — Endurance & Charge-Safety Validation at Low Temperature (KEY)
Requirement Statement | The market demands physical validation of endurance across the full -30 °C to +60 °C range, safe behaviour of the below-0 °C charge inhibition, and capacity retention across many reuse cycles, so the program's defining KEY RISK is closed by test before scale-up rather than assumed.
Business Objective | Retire the top strategic risk (MR-1800 RISK-001) with hard environmental-test evidence.
Customer Segment | All segments; critically frozen/cold-climate operators and high-reuse fleets.
Geographic Applicability | Global, acute for cold/frozen operations.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Cold endurance and charge safety are the credibility test for any in-shipper battery monitor; proof is decisive.
Business Value | Converts the program's biggest liability into a proven reliability claim.
Revenue Impact | Existential gate — no credible scale without it; magnitude Evidence Pending.
Risk if Unmet | Scaling on unproven cold performance risks field-wide failure, safety incident, and loss of trust.
Evidence Source | Judgment/Assumption — this validation is PENDING and is the mechanism that grounds RISK-001 (Evidence Pending until tested). Methods based on IEC 60068 / UN 38.3; D3 constraints evidence-backed.
Supporting Research | IEC 60068 environmental/thermal test methods; UN 38.3; low-temperature Li-ion charge-safety and cycle-life references.
VOC Reference | VOC-191 (pending) — "prove -30 °C endurance and safe cold charging, trip after trip."
Strategic Theme | Prove the physics (KEY).
Product Family | SCM — endurance/charge-safety validation.
Lifecycle Stage | Concept
Traceability | UN-SCM-1906, PRD-SCM-1906; ORCH-018 / EM-ENG-012.
Notes | HIGHEST-PRIORITY validation. Directly closes MR-1800 RISK-001; informs MR-SUS-003 (cycle life) and MR-INNO-007/008. Coordinate with battery/power, thermal, and safety engineers; detailed protocol in PVL.
MR-VAL-007 — Alarm-Efficacy Validation
Requirement Statement | The market demands demonstrated alarm efficacy — high detection of true excursions with a low false-alarm rate and reliable delivery — so buyers can trust that alerts are both complete and credible before they rely on them operationally.
Business Objective | Prove the alerting system is both sensitive (no missed excursions) and specific (few false alarms).
Customer Segment | All operational segments relying on excursion alerts.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Alert reliability and low false-alarm rate are the direct antidotes to the market's alarm-fatigue distrust.
Business Value | Substantiates the core alert promise and de-risks RISK-002 and RISK-011 together.
Revenue Impact | Underpins adoption and renewal; magnitude Evidence Pending.
Risk if Unmet | Missed or noisy alerts undermine the product's central function and trust.
Evidence Source | Judgment/Assumption — alarm-efficacy study pending (Evidence Pending). Immediate-excursion-alert function evidence-backed (D2).
Supporting Research | Alarm-management/detection-theory practice; GDP alert-handling expectations.
VOC Reference | VoC theme — "prove it catches the real ones and doesn't cry wolf."
Strategic Theme | Trustworthy alerting.
Product Family | SCM — alarm validation.
Lifecycle Stage | Launch
Traceability | UN-SCM-1907, PRD-SCM-1907; ORCH-018 / EM-ENG-012.
Notes | Validates the payoff of MR-INNO-004/005; closes RISK-011. Missed-alert (sensitivity) severity dominates over false-alarm rate.
MR-VAL-008 — WHO PQS Lab Qualification Testing
Requirement Statement | The market demands completion of WHO PQS laboratory qualification testing, because PQS status is the gate to the largest immunization-procurement channel and cannot be substituted by commercial evidence.
Business Objective | Achieve the PQS qualification that unlocks the primary public-health market (closes RISK-003).
Customer Segment | WHO/UNICEF/GAVI procurement, national immunization programs, ministries of health.
Geographic Applicability | Global public-health procurement, weighted to LMICs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | PQS is a hard, non-substitutable gate; qualified competitors are eligible and unqualified ones are not.
Business Value | Opens the highest-volume channel and anchors the mission narrative.
Revenue Impact | Gates the primary public-health channel; magnitude Evidence Pending.
Risk if Unmet | Exclusion from PQS-gated and donor-funded tenders (RISK-003, RISK-010).
Evidence Source | Standard — WHO PQS prequalification and lab-qualification framework (category code TBC).
Supporting Research | WHO PQS test/qualification procedures; device-category confirmation.
VOC Reference | VoC theme — "no PQS certificate, no conversation."
Strategic Theme | Earn the gate.
Product Family | SCM — PQS qualification.
Lifecycle Stage | Launch
Traceability | UN-SCM-1908, PRD-SCM-1908; ORCH-018 / EM-ENG-012.
Notes | Confirm correct monitor/logger category early; PQS testing overlaps environmental/endurance evidence (MR-VAL-006).
6.4 Usability Validation
MR-VAL-009 — Usability Studies with Frontline Health Workers
Requirement Statement | The market demands usability validation with frontline health workers and logistics operators — including provisioning, reading the OLED/button, responding to alerts, and recharge/return handling — so the device is usable by non-specialists under real field pressure.
Business Objective | Ensure adoption is not blocked by usability failures at the point of use.
Customer Segment | Frontline health workers, warehouse/logistics operators, LMIC last-mile staff.
Geographic Applicability | Global, weighted to LMIC last-mile and high-turnover operations.
Market Priority | High
Importance | High
Competitive Driver | Devices that confuse frontline users are abandoned regardless of technical merit; usability drives real-world efficacy.
Business Value | Protects field efficacy, correct alert response, and recovery behaviour.
Revenue Impact | Poor usability drives misuse, attrition, and support cost; magnitude Evidence Pending.
Risk if Unmet | Misprovisioning, ignored/misread alerts, and device loss erode both efficacy and economics.
Evidence Source | Judgment/Assumption — usability studies pending (Evidence Pending). OLED+button and BLE/mobile provisioning are evidence-backed features.
Supporting Research | Human-factors/usability-engineering practice; WHO field-usability considerations for cold-chain equipment.
VOC Reference | VoC theme — "our field staff must get it right with no training."
Strategic Theme | Usable at the frontline.
Product Family | SCM — usability validation.
Lifecycle Stage | Launch
Traceability | UN-SCM-1909, PRD-SCM-1909; ORCH-018 / EM-ENG-012.
Notes | Usability failures also worsen RISK-009 (recovery) and RISK-011 (alert response). Coordinate with human-factors engineer.
6.5 Commercial Validation
MR-VAL-010 — Willingness-to-Pay / Pricing Validation
Requirement Statement | The market demands validation of willingness-to-pay and pricing structure (device, connectivity/SaaS, per-trip, or pooling), so the reusable business model is proven economically viable across both commercial and donor channels.
Business Objective | Confirm the pricing and revenue model that underpins the investment case (MR-EXEC-008).
Customer Segment | Fleet buyers across commercial and institutional/donor channels; internal investment committee.
Geographic Applicability | Global, differentiated by channel and region.
Market Priority | High
Importance | High
Competitive Driver | Pricing must beat single-use TCO and clear donor budget realities; unvalidated pricing risks mispricing the platform.
Business Value | De-risks the revenue model and pricing strategy before scale.
Revenue Impact | Directly determines margin and viability; magnitude Evidence Pending.
Risk if Unmet | Mispriced offering; model unviable or leaves value on the table.
Evidence Source | Judgment/Assumption — WTP/pricing research pending (Evidence Pending); feeds the business-case workbook.
Supporting Research | Pricing-research methods (e.g., Van Westendorp/conjoint); reusable-vs-single-use TCO framing; GAVI budget context.
VOC Reference | VoC theme — "make the reusable price beat our single-use spend."
Strategic Theme | Prove the economics.
Product Family | SCM — pricing validation.
Lifecycle Stage | Concept
Traceability | UN-SCM-1910, PRD-SCM-1910; ORCH-018 / EM-ENG-012.
Notes | Must incorporate reverse-logistics/attrition (RISK-009) in the per-trip economics. Feeds the business-case workbook (magnitudes deferred there).
MR-VAL-011 — Tender / Procurement Validation & Support Feedback Loop
Requirement Statement | The market demands validation against real tender and procurement criteria — plus a closed-loop sales-and-support feedback mechanism — so the offering matches how buyers actually evaluate and buy, and so field learning continuously refines the product and library.
Business Objective | Ensure the product clears real procurement gates and that post-sale learning feeds back into requirements.
Customer Segment | Institutional/donor procurement, enterprise procurement, and internal sales/support functions.
Geographic Applicability | Global, differentiated by procurement regime.
Market Priority | High
Importance | High
Competitive Driver | Winning depends on matching tender scoring criteria; a feedback loop keeps the product ahead of evolving criteria.
Business Value | Improves win rate and continuously grounds the library in field reality.
Revenue Impact | Directly affects tender win rate and retention; magnitude Evidence Pending.
Risk if Unmet | Product misaligned with procurement scoring; field learning lost; requirements stagnate.
Evidence Source | Judgment/Assumption — tender analysis and feedback loop pending (Evidence Pending). PQS/GDP/FDA tender criteria based on standards.
Supporting Research | GAVI/UNICEF/WHO tender-criteria structures; enterprise-procurement scoring; closed-loop VOC practice.
VOC Reference | VoC theme — "score well on the tender, then keep listening."
Strategic Theme | Buy-the-way-they-buy + continuous learning.
Product Family | SCM — procurement validation.
Lifecycle Stage | Growth
Traceability | UN-SCM-1911, PRD-SCM-1911; ORCH-018 / EM-ENG-012.
Notes | Closes the loop back to MR-VAL-001 (VOC); feeds continuous library updates. Ties to RISK-010 (donor dependence) diversification.
7. Evidence & Validation Notes
• This module is the evidence engine for the library. Each activity here is designed to upgrade specific Judgment/Assumption items in MR-1600/1700/1800 to Cited evidence. Until executed, the activities themselves are Evidence Pending.
• MR-VAL-006 (low-temperature endurance & charge-safety validation) is the highest-priority study, because it closes the program's TOP STRATEGIC RISK (MR-1800 RISK-001). It must precede scale commitment and uses evidence-backed methods (IEC 60068, UN 38.3).
• No study outcome is pre-asserted. Acceptance thresholds, sample sizes, and numeric results are deferred to the PVL/PVaL and business-case workbook; asserting them here would breach the evidence-honesty rule.
Sequencing | discovery/usability early and cheap; PQS lab qualification, endurance testing, and field pilots are the gating, capital-intensive studies before scale-up. Results feed back through MR-VAL-011 into continuous library revision.
• Standards (PQS, GDP, Part 11, IEC 60068/60529/62368-1, UN 38.3, ISO 17025 lineage) define the methods and gates; conformance/pass is Evidence Pending until demonstrated.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-VAL-001 | UN-SCM-1901 | PRD-SCM-1901 | ORCH-018 / EM-ENG-012
MR-VAL-002 | UN-SCM-1902 | PRD-SCM-1902 | ORCH-018 / EM-ENG-012
MR-VAL-003 | UN-SCM-1903 | PRD-SCM-1903 | ORCH-018 / EM-ENG-012
MR-VAL-004 | UN-SCM-1904 | PRD-SCM-1904 | ORCH-018 / EM-ENG-012
MR-VAL-005 | UN-SCM-1905 | PRD-SCM-1905 | ORCH-018 / EM-ENG-012
MR-VAL-006 | UN-SCM-1906 | PRD-SCM-1906 | ORCH-018 / EM-ENG-012
MR-VAL-007 | UN-SCM-1907 | PRD-SCM-1907 | ORCH-018 / EM-ENG-012
MR-VAL-008 | UN-SCM-1908 | PRD-SCM-1908 | ORCH-018 / EM-ENG-012
MR-VAL-009 | UN-SCM-1909 | PRD-SCM-1909 | ORCH-018 / EM-ENG-012
MR-VAL-010 | UN-SCM-1910 | PRD-SCM-1910 | ORCH-018 / EM-ENG-012
MR-VAL-011 | UN-SCM-1911 | PRD-SCM-1911 | ORCH-018 / EM-ENG-012
Related modules: MR-1900 (VAL) is the evidence-generating counterpart to MR-1800 (RISK); it grounds the assumptions in MR-1600 (INNO) and MR-1700 (SUS) and feeds the success criteria in MR-000 (EXEC). Detailed protocols and thresholds live in the PVL/PVaL libraries.
MR-200 — Customer Segmentation (SEG)
Traceability: this module operationalizes the customer segments referenced in MR-EXEC and MR-OPP, and traces forward to PRD-SCM-2xx and UN-SCM-2xx, and back to ORCH-018 / EM-ENG-012.
1. Purpose
This module defines the customer and buyer segments for the Smart Cold Chain Monitor (SCM), their distinct needs, buying behaviours, and priority to the program. It gives product and go-to-market leadership a clear map of who buys, who uses, and who funds — so that requirements, pricing, and channel strategy are matched to real segment economics rather than treated as one homogeneous market.
2. Scope
In scope: pharma manufacturers, distributors/wholesalers, 3PLs/logistics providers, national immunization programs / Ministries of Health, global procurement bodies (GAVI/UNICEF/WHO), hospitals/pharmacies, CDMOs, cold-chain equipment OEMs, channel partners, and regional-market segmentation. Out of scope: market sizing (MR-100), specific problem definitions (MR-300), and pricing models (business-case workbook). Segment volumes and shares are marked Judgment/Assumption or Evidence Pending where not evidence-backed.
3. Applicable Standards & References
• WHO PQS — market-access gate for immunization-program and procurement-body purchases (category code TBC).
• EU GDP; US FDA 21 CFR Part 11 — compliance drivers shaping manufacturer, distributor, and 3PL requirements.
• GAVI, UNICEF Supply Division, WHO — institutional procurement channels and specifications.
• IATA DG / UN 38.3 — transport-of-battery constraints relevant to air-freight 3PLs and forwarders.
• IQVIA / pharma-logistics reporting — segment structure and buyer behaviour context (cited by name).
4. Background
The vaccine cold chain is a multi-actor network: a vaccine is manufactured, sold and released by a manufacturer, moved by distributors/wholesalers and 3PLs, procured for public programs by GAVI/UNICEF/WHO and delivered by national immunization programs/MoH, and finally administered at hospitals and pharmacies. CDMOs manufacture on behalf of others, and cold-chain equipment OEMs and channel partners shape how monitoring is specified, bundled, and sold.
These actors do not share one buying logic. Regulated-market manufacturers and distributors buy on data integrity and audit-defensibility; 3PLs buy on SLA performance and integration; immunization programs and procurement bodies buy on PQS eligibility, ruggedness, and total cost within donor budgets. Effective segmentation prevents the classic error of designing and pricing for an "average" customer who does not exist.
5. Requirement Categories
• Commercial Buyers — manufacturers, distributors/wholesalers, CDMOs.
• Logistics Buyers — 3PLs and freight forwarders.
• Public-Health Buyers — immunization programs/MoH and global procurement bodies.
• Point-of-Use Buyers — hospitals and pharmacies.
• Ecosystem Partners — cold-chain equipment OEMs and channel partners.
• Regional Segmentation — geographic buying-behaviour differences.
6. Detailed Market Requirements
6.1 Commercial Buyers
MR-SEG-001 — Pharmaceutical / Vaccine Manufacturers
Requirement Statement | The market — vaccine manufacturers — demands monitoring that provides audit-grade release-supporting evidence and protects brand and product liability across every regulated market they ship into.
Business Objective | Win manufacturers as anchor buyers who can mandate SCM across their distribution.
Customer Segment | Multinational and regional vaccine/biologic manufacturers.
Geographic Applicability | Global, weighted to EU/US HQ'd manufacturers.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Manufacturers can standardize monitoring downstream, making them high-leverage design partners.
Business Value | Anchor accounts that pull volume through the chain.
Revenue Impact | High-value, standard-setting accounts; magnitude Evidence Pending.
Risk if Unmet | Lose the actor most able to mandate a standard; adoption fragments.
Evidence Source | Standard — FDA 21 CFR Part 11 / EU GDP release expectations; segment sizing Evidence Pending.
Supporting Research | FDA data-integrity guidance; GDP; IQVIA manufacturer-logistics reporting.
VOC Reference | VoC theme — "our name is on the vial; the data must be defensible."
Strategic Theme | Anchor accounts.
Product Family | SCM — commercial buyer.
Lifecycle Stage | Launch
Traceability | UN-SCM-201, PRD-SCM-201; ORCH-018 / EM-ENG-012.
Notes | Device supports disposition; it does not perform batch release.
MR-SEG-002 — Distributors / Wholesalers
Requirement Statement | The market — pharma distributors and wholesalers — demands monitoring that protects GDP compliance, reduces spoilage write-offs, and settles temperature disputes with upstream and downstream partners.
Business Objective | Convert distributors into reusable-fleet buyers with recurring service.
Customer Segment | National and regional pharma distributors/wholesalers.
Geographic Applicability | Global, strongest in GDP-regulated EU and comparable markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Distributors bear direct spoilage and dispute cost; reusable economics fit their volume.
Business Value | Core recurring-revenue segment with high shipment frequency.
Revenue Impact | Large reusable-fleet and connectivity revenue; magnitude Evidence Pending.
Risk if Unmet | Miss the highest-frequency commercial buyer of monitoring.
Evidence Source | Standard — EU GDP distribution-integrity duties; sizing Evidence Pending.
Supporting Research | EU GDP; distributor spoilage/dispute reporting (industry).
VOC Reference | VoC theme — "we eat the write-off when we can't prove who broke the chain."
Strategic Theme | Recurring-revenue core.
Product Family | SCM — commercial buyer.
Lifecycle Stage | Launch
Traceability | UN-SCM-202, PRD-SCM-202; ORCH-018 / EM-ENG-012.
Notes | High shipment frequency favours reusable device economics.
MR-SEG-003 — Contract Development & Manufacturing Organizations (CDMOs)
Requirement Statement | The market — CDMOs — demands monitoring that documents cold-chain integrity for the products they make and ship on behalf of clients, protecting client relationships and contractual quality obligations.
Business Objective | Reach CDMOs as both buyers and referrers into their manufacturer clients.
Customer Segment | Biologic/vaccine CDMOs and fill-finish providers.
Geographic Applicability | Global, concentrated in regulated manufacturing hubs.
Market Priority | Medium
Importance | Medium
Competitive Driver | CDMOs must meet client quality specs; monitoring is a contractual differentiator.
Business Value | Access to multiple downstream clients through one relationship.
Revenue Impact | Moderate direct, high referral leverage; magnitude Evidence Pending.
Risk if Unmet | Miss a channel into multiple manufacturer clients.
Evidence Source | Judgment/Assumption — CDMO buying behaviour inferred; sizing Evidence Pending.
Supporting Research | IQVIA CDMO-market reporting; GDP/FDA quality-agreement practice.
VOC Reference | VoC theme — "our clients audit our cold chain, not just their own."
Strategic Theme | Channel leverage.
Product Family | SCM — commercial buyer.
Lifecycle Stage | Growth
Traceability | UN-SCM-203, PRD-SCM-203; ORCH-018 / EM-ENG-012.
Notes | Often specify to client quality agreements rather than buy autonomously.
6.2 Logistics Buyers
MR-SEG-004 — Third-Party Logistics (3PLs) & Freight Forwarders
Requirement Statement | The market — 3PLs and cold-chain forwarders — demands monitoring that proves SLA performance, integrates with their tracking systems, and limits their liability for excursions in transit, including air freight.
Business Objective | Embed SCM as the monitoring layer within 3PL cold-chain services.
Customer Segment | Pharma-specialized 3PLs, air/ground cold-chain forwarders.
Geographic Applicability | Global, along major air and ground lanes.
Market Priority | Must-have
Importance | High
Competitive Driver | 3PLs need real-time telemetry to defend SLAs and win pharma contracts.
Business Value | High-volume, integration-driven recurring revenue.
Revenue Impact | Fleet + connectivity + API revenue; magnitude Evidence Pending.
Risk if Unmet | 3PLs adopt a competing telemetry standard and lock SCM out.
Evidence Source | Standard — IATA DG / UN 38.3 for air transport of the Li-ion device; sizing Evidence Pending.
Supporting Research | IATA cold-chain/DG guidance; 3PL telemetry-adoption reporting.
VOC Reference | VoC theme — "we get blamed for excursions we didn't cause."
Strategic Theme | Logistics integration.
Product Family | SCM — logistics buyer.
Lifecycle Stage | Launch
Traceability | UN-SCM-204, PRD-SCM-204; ORCH-018 / EM-ENG-012.
Notes | Air transport imposes battery-shipment and in-flight-radio constraints (IATA DG, UN 38.3).
6.3 Public-Health Buyers
MR-SEG-005 — National Immunization Programs / Ministries of Health
Requirement Statement | The market — national immunization programs and MoH — demands PQS-eligible, rugged, low-maintenance monitoring that survives last-mile conditions and gives program managers actionable visibility within constrained budgets.
Business Objective | Win national-program standardization and recurring device replenishment.
Customer Segment | EPI/immunization programs, Ministries of Health.
Geographic Applicability | Global, weighted to LMICs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | PQS eligibility and ruggedness are gating; few devices qualify.
Business Value | Large-volume, mission-critical public-health accounts.
Revenue Impact | High unit volumes via program budgets/donors; magnitude Evidence Pending.
Risk if Unmet | Excluded from national-program deployments and their volumes.
Evidence Source | Standard — WHO PQS eligibility (code TBC); volumes Evidence Pending.
Supporting Research | WHO EPI/PQS documentation; UNICEF/GAVI program reporting.
VOC Reference | VoC theme — "it has to work unattended in tough field conditions."
Strategic Theme | Public-health scale.
Product Family | SCM — public-health buyer.
Lifecycle Stage | Growth
Traceability | UN-SCM-205, PRD-SCM-205; ORCH-018 / EM-ENG-012.
Notes | PQS prequalification is a hard market-access gate.
MR-SEG-006 — Global Procurement Bodies (GAVI / UNICEF / WHO)
Requirement Statement | The market — global procurement and funding bodies — demands devices that meet PQS specifications, offer strong total-cost-of-ownership, and demonstrate measurable public-health impact to justify donor-funded procurement at scale.
Business Objective | Achieve inclusion on institutional procurement catalogues and tenders.
Customer Segment | GAVI, UNICEF Supply Division, WHO procurement.
Geographic Applicability | Global, disbursing to LMICs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Catalogue inclusion unlocks the largest aggregated volumes; PQS + TCO + impact are the gates.
Business Value | Single-relationship access to multi-country volume.
Revenue Impact | Largest aggregated volume channel; magnitude Evidence Pending.
Risk if Unmet | Locked out of the dominant LMIC procurement pathway.
Evidence Source | Standard — WHO PQS; procurement-spec practice; volumes Evidence Pending.
Supporting Research | UNICEF Supply Division catalogues; GAVI CCEOP; WHO PQS.
VOC Reference | VoC theme — "if it's not PQS-listed and cost-effective, we can't buy it."
Strategic Theme | Institutional procurement.
Product Family | SCM — public-health buyer.
Lifecycle Stage | Growth
Traceability | UN-SCM-206, PRD-SCM-206; ORCH-018 / EM-ENG-012.
Notes | Procurement cycles are long; plan PQS listing early.
6.4 Point-of-Use Buyers
MR-SEG-007 — Hospitals & Pharmacies
Requirement Statement | The market — hospitals and pharmacies receiving vaccines — demands simple confirmation at receipt that a shipment stayed in range, so staff can accept, quarantine, or reject stock quickly and defensibly.
Business Objective | Extend SCM value to the receiving end to reinforce chain-wide adoption.
Customer Segment | Hospital pharmacies, retail/community pharmacies, clinics.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Receivers want a fast, unambiguous accept/reject signal, not raw data.
Business Value | Reinforces end-to-end value and receiver confidence.
Revenue Impact | Lower direct spend; strengthens whole-chain case; magnitude Evidence Pending.
Risk if Unmet | Receiver friction undermines the chain-of-custody promise.
Evidence Source | Judgment/Assumption — receiver behaviour inferred; sizing Evidence Pending.
Supporting Research | GDP receipt-check practice; immunization-clinic workflow studies.
VOC Reference | VoC theme — "tell me at the door if I can use this stock."
Strategic Theme | End-to-end confidence.
Product Family | SCM — point-of-use buyer.
Lifecycle Stage | Growth
Traceability | UN-SCM-207, PRD-SCM-207; ORCH-018 / EM-ENG-012.
Notes | OLED + button + BLE support a quick at-receipt status check.
6.5 Ecosystem Partners
MR-SEG-008 — Cold-Chain Equipment OEMs
Requirement Statement | The market — cold-chain packaging and equipment OEMs — demands embeddable, interoperable monitoring they can integrate into insulated shippers and containers as a value-added, differentiating feature.
Business Objective | Establish OEM embedding partnerships that scale distribution.
Customer Segment | Insulated-shipper, container, and refrigeration OEMs.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | OEMs seek differentiation; embedded monitoring adds value to their hardware.
Business Value | Scalable, designed-in distribution channel.
Revenue Impact | OEM-embedded volume + licensing; magnitude Evidence Pending.
Risk if Unmet | OEMs embed a competitor's monitor as the default.
Evidence Source | Judgment/Assumption — OEM-partnership model; sizing Evidence Pending.
Supporting Research | Cold-chain packaging market reporting; OEM integration case studies.
VOC Reference | VoC theme — "we want monitoring built into our shipper, not bolted on."
Strategic Theme | Designed-in distribution.
Product Family | SCM — ecosystem partner.
Lifecycle Stage | Growth
Traceability | UN-SCM-208, PRD-SCM-208; ORCH-018 / EM-ENG-012.
Notes | Fit-inside-shipper form factor (D-facts) supports embedding.
MR-SEG-009 — Channel Partners / Resellers / Systems Integrators
Requirement Statement | The market — channel partners, resellers, and systems integrators — demands a device and data platform they can resell, integrate, and support within regional cold-chain and pharma-logistics offerings.
Business Objective | Build indirect reach into fragmented regional markets.
Customer Segment | Regional resellers, cold-chain SIs, IoT platform partners.
Geographic Applicability | Global, especially markets hard to serve directly.
Market Priority | Medium
Importance | Medium
Competitive Driver | Local partners provide reach, language, and support the vendor cannot supply directly.
Business Value | Cost-efficient regional expansion.
Revenue Impact | Incremental indirect volume; magnitude Evidence Pending.
Risk if Unmet | Slow, expensive regional expansion; gaps competitors fill locally.
Evidence Source | Judgment/Assumption — channel model; sizing Evidence Pending.
Supporting Research | IoT/cold-chain channel-model literature.
VOC Reference | VoC theme — "we need a local partner who can integrate and support it."
Strategic Theme | Indirect reach.
Product Family | SCM — ecosystem partner.
Lifecycle Stage | Growth
Traceability | UN-SCM-209, PRD-SCM-209; ORCH-018 / EM-ENG-012.
Notes | Open APIs and BLE/mobile provisioning enable integrator onboarding.
6.6 Regional Segmentation
MR-SEG-010 — Regulated Mature Markets (EU / US / comparable)
Requirement Statement | The market — buyers in mature regulated regions — demands audit-grade, integration-ready monitoring and will pay a premium for data integrity and compliance assurance.
Business Objective | Serve mature markets as the margin and reference-account base.
Customer Segment | Manufacturers, distributors, 3PLs, hospitals in EU/US and comparable regions.
Geographic Applicability | EU, US, and comparable regulated markets.
Market Priority | Must-have
Importance | High
Competitive Driver | High compliance bar and product value support premium positioning.
Business Value | Margin anchor and credibility reference for other regions.
Revenue Impact | High per-unit and recurring revenue; magnitude Evidence Pending.
Risk if Unmet | Lose margin base and reference accounts.
Evidence Source | Standard — EU GDP, FDA Part 11; sizing Evidence Pending.
Supporting Research | GDP; FDA guidance; IQVIA regulated-market reporting.
VOC Reference | VoC theme — "compliance first, price second."
Strategic Theme | Premium mature markets.
Product Family | SCM — regional segment.
Lifecycle Stage | Launch
Traceability | UN-SCM-210, PRD-SCM-210; ORCH-018 / EM-ENG-012.
Notes | Mirrors MR-OPP-008 regional opportunity.
MR-SEG-011 — LMIC / Last-Mile Markets
Requirement Statement | The market — buyers in LMIC and last-mile settings — demands rugged, affordable, low-maintenance, PQS-eligible monitoring that functions reliably despite intermittent connectivity, power scarcity, and harsh field conditions.
Business Objective | Serve LMIC markets as the mission-impact and high-volume base.
Customer Segment | Immunization programs/MoH, procurement bodies, local partners.
Geographic Applicability | LMICs across multiple regions.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Ruggedness, affordability, low-temperature battery capability, and PQS eligibility together are rare.
Business Value | Highest mission impact and largest unit volumes.
Revenue Impact | High-volume, budget-constrained; magnitude Evidence Pending.
Risk if Unmet | Cede the highest-impact, highest-volume segment.
Evidence Source | Cited — WHO/UNICEF/GAVI last-mile cold-chain reporting (reported/estimated); sizing Evidence Pending.
Supporting Research | WHO EPI assessments; UNICEF Supply Division; GAVI CCEOP.
VOC Reference | VoC theme — "it must survive the field and not need babysitting."
Strategic Theme | Mission-led LMIC.
Product Family | SCM — regional segment.
Lifecycle Stage | Growth
Traceability | UN-SCM-211, PRD-SCM-211; ORCH-018 / EM-ENG-012.
Notes | LTE-M primary with BLE/Wi-Fi store-and-forward matters where coverage is intermittent.
7. Evidence & Validation Notes
• Regulatory and procurement gates (WHO PQS, EU GDP, FDA Part 11, IATA DG) are stated as Standard and are the firmest basis for segment prioritization.
• Segment volumes, shares, and revenue are uniformly Evidence Pending pending the business-case workbook and validated pipeline; buyer-behaviour characterizations not backed by a standard are labelled Judgment/Assumption.
• LMIC/last-mile loss context is Cited to WHO/UNICEF/GAVI and labelled reported/estimated.
Validation path | VoC interviews per segment (to replace pending VOC IDs), pipeline data, and PQS/procurement engagement will confirm prioritization.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-SEG-001 | UN-SCM-201 | PRD-SCM-201 | ORCH-018 / EM-ENG-012
MR-SEG-002 | UN-SCM-202 | PRD-SCM-202 | ORCH-018 / EM-ENG-012
MR-SEG-003 | UN-SCM-203 | PRD-SCM-203 | ORCH-018 / EM-ENG-012
MR-SEG-004 | UN-SCM-204 | PRD-SCM-204 | ORCH-018 / EM-ENG-012
MR-SEG-005 | UN-SCM-205 | PRD-SCM-205 | ORCH-018 / EM-ENG-012
MR-SEG-006 | UN-SCM-206 | PRD-SCM-206 | ORCH-018 / EM-ENG-012
MR-SEG-007 | UN-SCM-207 | PRD-SCM-207 | ORCH-018 / EM-ENG-012
MR-SEG-008 | UN-SCM-208 | PRD-SCM-208 | ORCH-018 / EM-ENG-012
MR-SEG-009 | UN-SCM-209 | PRD-SCM-209 | ORCH-018 / EM-ENG-012
MR-SEG-010 | UN-SCM-210 | PRD-SCM-210 | ORCH-018 / EM-ENG-012
MR-SEG-011 | UN-SCM-211 | PRD-SCM-211 | ORCH-018 / EM-ENG-012
Upstream: MR-000 (EXEC), MR-100 (OPP). Lateral: MR-300 (PROB).
MR-300 — Customer Problems (PROB)
Traceability: this module documents the market problems the SCM must resolve, traces upstream to MR-EXEC/MR-OPP/MR-SEG, forward to PRD-SCM-3xx and UN-SCM-3xx, and back to ORCH-018 / EM-ENG-012.
1. Purpose
This module catalogues the concrete, product-independent problems that vaccine cold-chain customers experience today, expressed as market demands for relief. It is the empirical backbone of the requirements library: every product feature must trace to a problem here, and every problem must trace to a segment (MR-200) and an opportunity (MR-100). It is written so that a competent generalist can see exactly why the SCM should exist and what pain it removes.
2. Scope
In scope: excursion-detection gaps, manual-logging burden, VVM limitations, data-integrity/audit pain, last-mile/LMIC visibility gaps, false alarms/alarm fatigue, device battery failure in cold, wasted-vaccine cost, dispute/liability exposure, and reverse-logistics losses in reusable fleets. Out of scope: product implementation (PRD), segment definitions (MR-200), and market sizing (MR-100). Problems are stated as demands for relief; loss magnitudes are Cited (reported/estimated) or Evidence Pending — never invented as exact facts.
3. Applicable Standards & References
• WHO / UNICEF Supply Division / GAVI — cold-chain loss, wastage, freezing/heat-exposure, and last-mile reporting (reported/estimated).
• WHO PQS — monitoring-device expectations (VVM and electronic monitors; code TBC).
• EU GDP; US FDA 21 CFR Part 11 — data-integrity and audit requirements underlying audit-pain problems.
• IEC 62368-1; UN 38.3; IATA DG — battery-safety context for cold-weather battery-failure problems.
• IEC 60068 / 60529 — environmental/ingress test context for field-robustness problems.
4. Background
The vaccine cold chain fails in characteristic, well-documented ways. WHO and immunization-supply bodies have reported for years that a substantial share of vaccines are exposed to temperatures outside recommended ranges, that inadvertent freezing is as damaging as heat, and that visibility collapses in the last mile (reported/estimated). Legacy tools each solve part of the problem and leave gaps: chemical VVMs give a coarse cumulative heat signal but do not detect freezing, do not quantify, and do not alert; USB dataloggers quantify but must be recovered and manually offloaded, revealing excursions too late; and manual logging is labour-intensive and error-prone.
Layered on top are operational problems: false alarms that breed alarm fatigue, connected devices that die in the cold, disputes over who broke the chain, and — for any reusable fleet — the risk of losing devices in reverse logistics. Each of these is a market demand for relief that the SCM is chartered to address.
5. Requirement Categories
• Detection Gaps — excursions and freezing missed or found too late.
• Operational Burden — manual logging and legacy-tool limitations.
• Data & Compliance Pain — audit, integrity, and dispute exposure.
• Field Reliability — cold-weather battery failure and last-mile visibility.
• Alarm Quality — false alarms and alarm fatigue.
• Economic Loss — wasted vaccine and reverse-logistics losses.
6. Detailed Market Requirements
6.1 Detection Gaps
MR-PROB-001 — Undetected / Late-Detected Excursions
Requirement Statement | The market demands that temperature excursions be detected and communicated in near-real-time, because today excursions are frequently discovered only after delivery, when the vaccine is already compromised.
Business Objective | Eliminate the "found out too late" failure mode across the chain.
Customer Segment | Manufacturers, distributors, 3PLs, immunization programs.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Offline loggers reveal excursions only on recovery; real-time alerting is a decisive advantage.
Business Value | Directly enables intervention and loss avoidance — the core value proposition.
Revenue Impact | Primary purchase driver; magnitude Evidence Pending.
Risk if Unmet | Product fails its central promise; no reason to switch from cheaper loggers.
Evidence Source | Cited — WHO/UNICEF reporting that vaccines are frequently exposed to out-of-range temperatures (reported/estimated).
Supporting Research | WHO immunization-supply-chain publications; UNICEF cold-chain reporting.
VOC Reference | VoC theme — "we find out at delivery, and then it's too late."
Strategic Theme | Real-time detection.
Product Family | SCM — detection gap.
Lifecycle Stage | Concept
Traceability | UN-SCM-301, PRD-SCM-301; ORCH-018 / EM-ENG-012.
Notes | Based on D2 (immediate excursion alerts).
MR-PROB-002 — Undetected Freezing (Not Just Heat)
Requirement Statement | The market demands detection of inadvertent freezing as rigorously as heat excursions, because many vaccines are freeze-sensitive and freezing damage is common, silent, and often missed by heat-biased indicators.
Business Objective | Close the freezing blind spot left by heat-only tools like VVMs.
Customer Segment | Immunization programs, distributors, manufacturers.
Geographic Applicability | Global, notably cold-climate and over-refrigerated last-mile settings.
Market Priority | Must-have
Importance | Critical
Competitive Driver | VVMs and many indicators do not signal freezing; quantitative two-sided monitoring wins.
Business Value | Addresses a large, under-served category of cold-chain loss.
Revenue Impact | Expands value in freeze-sensitive vaccine handling; magnitude Evidence Pending.
Risk if Unmet | Silent freezing losses persist; product seen as no better than VVMs on this axis.
Evidence Source | Cited — WHO/UNICEF reporting that inadvertent freezing of freeze-sensitive vaccines is common (reported/estimated).
Supporting Research | WHO/UNICEF freezing-exposure studies; PQS temperature-monitoring guidance.
VOC Reference | VoC theme — "freezing quietly ruins stock and nobody sees it."
Strategic Theme | Two-sided detection.
Product Family | SCM — detection gap.
Lifecycle Stage | Concept
Traceability | UN-SCM-302, PRD-SCM-302; ORCH-018 / EM-ENG-012.
Notes | Internal + external-probe temperature (device spec) supports freeze detection.
6.2 Operational Burden & Legacy-Tool Limitations
MR-PROB-003 — Manual Logging Burden and Error
Requirement Statement | The market demands elimination of manual temperature logging, because manual checks and paper/USB records are labour-intensive, error-prone, and frequently incomplete or retrospectively "corrected."
Business Objective | Replace manual logging with automated, continuous, trustworthy records.
Customer Segment | Distributors, immunization programs, hospitals/pharmacies, 3PLs.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | High
Competitive Driver | Automation removes labour cost and human error that plague manual methods.
Business Value | Reduces operating cost and improves record completeness.
Revenue Impact | Labour-savings ROI driver; magnitude Evidence Pending.
Risk if Unmet | Customers keep cheap manual methods; automation value unproven.
Evidence Source | Judgment/Assumption — manual-logging burden widely observed; quantification Evidence Pending.
Supporting Research | GDP monitoring practice; immunization-program workflow studies.
VOC Reference | VoC theme — "our staff spend hours logging temperatures by hand."
Strategic Theme | Automation.
Product Family | SCM — operational burden.
Lifecycle Stage | Concept
Traceability | UN-SCM-303, PRD-SCM-303; ORCH-018 / EM-ENG-012.
Notes | Automated local logging (D2) directly addresses this.
MR-PROB-004 — VVM and Chemical-Indicator Limitations
Requirement Statement | The market demands quantitative, alerting, two-sided monitoring to overcome the limits of vaccine vial monitors and chemical indicators, which give only coarse cumulative-heat signals, cannot quantify, cannot detect freezing, and cannot alert.
Business Objective | Position the SCM as the quantitative successor to coarse chemical indicators.
Customer Segment | Immunization programs, procurement bodies, manufacturers.
Geographic Applicability | Global, weighted to LMIC programs relying on VVMs.
Market Priority | High
Importance | High
Competitive Driver | VVMs are cheap and ubiquitous but fundamentally limited; a quantitative device complements/upgrades them.
Business Value | Converts a known limitation into a differentiated value story.
Revenue Impact | Upgrade/complement opportunity at scale; magnitude Evidence Pending.
Risk if Unmet | Programs stay with "good enough" VVMs; upgrade case not made.
Evidence Source | Cited/Standard — WHO PQS VVM function and known limitations (reported).
Supporting Research | WHO PQS VVM specifications; UNICEF VVM guidance.
VOC Reference | VoC theme — "the VVM only tells us so much, and never about freezing."
Strategic Theme | Quantitative upgrade.
Product Family | SCM — legacy-tool limitation.
Lifecycle Stage | Concept
Traceability | UN-SCM-304, PRD-SCM-304; ORCH-018 / EM-ENG-012.
Notes | SCM complements VVMs at consignment/shipper level; VVMs remain at vial level.
6.3 Data & Compliance Pain
MR-PROB-005 — Data-Integrity and Audit Pain
Requirement Statement | The market demands tamper-evident, audit-ready monitoring records, because current records are often incomplete, editable, or hard to reconstruct, exposing customers to GDP/FDA audit findings and data-integrity citations.
Business Objective | Make SCM data the clean, defensible record that passes audits.
Customer Segment | Manufacturers, distributors, QA/QP release, 3PLs.
Geographic Applicability | EU, US, and PQS-gated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Editable or gap-ridden records fail audits; tamper-evident data is a hard requirement.
Business Value | Turns compliance risk into a purchase trigger.
Revenue Impact | Gate to regulated-market revenue; magnitude Evidence Pending.
Risk if Unmet | Product data rejected in audits; disqualified in regulated markets.
Evidence Source | Standard — FDA 21 CFR Part 11; EU GDP data-integrity expectations.
Supporting Research | FDA data-integrity guidance; GDP; MHRA/EU inspection findings (industry).
VOC Reference | VoC theme — "we dread the auditor asking for the temperature history."
Strategic Theme | Audit-grade data.
Product Family | SCM — compliance pain.
Lifecycle Stage | Concept
Traceability | UN-SCM-305, PRD-SCM-305; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-EXEC-005 data-integrity strategy.
MR-PROB-006 — Excursion Disputes and Liability Exposure
Requirement Statement | The market demands objective, attributable evidence of when and where an excursion occurred, because without it, partners dispute responsibility for spoiled shipments and absorb unfair liability and write-offs.
Business Objective | Give customers unimpeachable evidence to resolve who-broke-the-chain disputes.
Customer Segment | Distributors, 3PLs, manufacturers, insurers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Time/location-stamped, tamper-evident data settles disputes that coarse tools cannot.
Business Value | Reduces write-offs and litigation; strengthens partner trust.
Revenue Impact | Loss/dispute-avoidance ROI; magnitude Evidence Pending.
Risk if Unmet | Customers keep absorbing disputed losses; a key ROI lever is missed.
Evidence Source | Judgment/Assumption — dispute burden widely reported by distributors/3PLs; quantification Evidence Pending.
Supporting Research | GDP responsibility allocation; cold-chain claims/insurance practice (industry).
VOC Reference | VoC theme — "we get blamed for excursions we didn't cause."
Strategic Theme | Attributable evidence.
Product Family | SCM — liability pain.
Lifecycle Stage | Concept
Traceability | UN-SCM-306, PRD-SCM-306; ORCH-018 / EM-ENG-012.
Notes | GPS + timestamped logs support attribution.
6.4 Field Reliability
MR-PROB-007 — Device Battery Failure in Cold
Requirement Statement | The market demands monitoring that keeps working through long, cold journeys, because prior connected devices have failed prematurely in low temperatures — losing data and alerts exactly when needed most.
Business Objective | Deliver reliable long-endurance operation across the full cold-chain temperature range.
Customer Segment | All segments shipping through cold environments.
Geographic Applicability | Global, notably cold lanes and deep-cold storage.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Cold-weather reliability is a recurring failure of incumbent connected loggers.
Business Value | Reliability is table stakes for trust; failures destroy the value proposition.
Revenue Impact | Retention-critical; magnitude Evidence Pending.
Risk if Unmet | Data/alert gaps in the cold; loss of customer trust and churn.
Evidence Source | Standard/Judgment — low-temperature Li-ion behaviour and charge-below-0°C hazard are well established (IEC 62368-1 context); field-failure rates Evidence Pending.
Supporting Research | Low-temperature Li-ion literature; IEC 62368-1; UN 38.3.
VOC Reference | VoC theme — "the last logger gave out in the cold and we lost the record."
Strategic Theme | Cold reliability.
Product Family | SCM — field reliability.
Lifecycle Stage | Concept
Traceability | UN-SCM-307, PRD-SCM-307; ORCH-018 / EM-ENG-012.
Notes | Based on D3 (low-temp Li-ion, duty-cycled, charging inhibited below 0°C) and 90d/-30-+60°C targets.
MR-PROB-008 — Last-Mile / LMIC Visibility Gaps
Requirement Statement | The market demands reliable monitoring and reporting through the last mile in LMICs, because connectivity, power, and infrastructure gaps there leave the final delivery leg — where losses concentrate — effectively unmonitored.
Business Objective | Extend dependable visibility into the hardest, highest-loss delivery leg.
Customer Segment | Immunization programs/MoH, procurement bodies, local partners.
Geographic Applicability | LMICs, last-mile.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Few devices maintain useful visibility under intermittent connectivity and power scarcity.
Business Value | Addresses the highest-impact loss zone and unlocks donor channels.
Revenue Impact | Large mission-volume potential; magnitude Evidence Pending.
Risk if Unmet | Last-mile stays a black box; mission promise unfulfilled.
Evidence Source | Cited — WHO/UNICEF/GAVI reporting on last-mile visibility and loss concentration (reported/estimated).
Supporting Research | WHO EPI assessments; UNICEF Supply Division; GAVI CCEOP.
VOC Reference | VoC theme — "the last mile is a black box."
Strategic Theme | Last-mile visibility.
Product Family | SCM — field reliability.
Lifecycle Stage | Concept
Traceability | UN-SCM-308, PRD-SCM-308; ORCH-018 / EM-ENG-012.
Notes | Store-and-forward local logging with LTE-M primary and BLE/Wi-Fi fallback addresses intermittency.
6.5 Alarm Quality
MR-PROB-009 — False Alarms and Alarm Fatigue
Requirement Statement | The market demands accurate, well-calibrated alerting that distinguishes genuine, product-relevant excursions from benign transients, because false and excessive alarms cause alarm fatigue and lead staff to ignore or disable alerts.
Business Objective | Deliver trustworthy alerts that staff act on rather than tune out.
Customer Segment | Distributors, 3PLs, immunization programs, hospitals/pharmacies.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Poorly-tuned alerting is a common failing; credible, low-noise alerting differentiates.
Business Value | Preserves the operational value of alerts and user trust.
Revenue Impact | Retention and satisfaction driver; magnitude Evidence Pending.
Risk if Unmet | Alerts ignored/disabled; real excursions missed despite detection capability.
Evidence Source | Judgment/Assumption — alarm fatigue well documented across monitoring domains; SCM-specific data Evidence Pending.
Supporting Research | Alarm-fatigue literature (clinical/industrial monitoring); GDP alarm-management practice.
VOC Reference | VoC theme — "so many false alerts that we stopped looking."
Strategic Theme | Trustworthy alerting.
Product Family | SCM — alarm quality.
Lifecycle Stage | Concept
Traceability | UN-SCM-309, PRD-SCM-309; ORCH-018 / EM-ENG-012.
Notes | Alert thresholds should reflect product stability profiles, not raw excursions.
6.6 Economic Loss
MR-PROB-010 — Cost of Wasted Vaccine
Requirement Statement | The market demands monitoring that reduces avoidable vaccine wastage, because temperature-related spoilage destroys valuable stock, forces re-shipments, and wastes immunization opportunities.
Business Objective | Anchor SCM ROI in the avoided cost of wasted vaccine.
Customer Segment | Manufacturers, distributors, immunization programs, procurement bodies.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Loss avoidance is a stronger buying argument than feature parity.
Business Value | Converts spoilage cost into the primary ROI justification.
Revenue Impact | Underpins value-based pricing; magnitude Evidence Pending.
Risk if Unmet | ROI case rests on soft benefits; harder to justify spend.
Evidence Source | Cited — WHO/UNICEF reporting on vaccine wastage and temperature-related loss (reported/estimated); exact per-shipment cost Evidence Pending.
Supporting Research | WHO vaccine-management assessments; UNICEF cold-chain cost reporting.
VOC Reference | VoC theme — "we destroy stock we can't vouch for."
Strategic Theme | Avoided-loss ROI.
Product Family | SCM — economic loss.
Lifecycle Stage | Concept
Traceability | UN-SCM-310, PRD-SCM-310; ORCH-018 / EM-ENG-012.
Notes | Mirrors MR-OPP-011; do not assert exact spoilage cost.
MR-PROB-011 — Reverse-Logistics Device Losses (Reusable Fleet)
Requirement Statement | The market demands that reusable monitoring devices be reliably recovered, tracked, and returned, because a reusable fleet loses its economic advantage if devices are not returned after each trip.
Business Objective | Protect reusable economics by minimizing device loss in reverse logistics.
Customer Segment | Distributors, 3PLs, immunization programs operating reusable fleets.
Geographic Applicability | Global, harder in last-mile/LMIC settings.
Market Priority | High
Importance | High
Competitive Driver | Reusable-model ROI (MR-EXEC-008) depends entirely on high return rates.
Business Value | Preserves the per-trip cost advantage central to the business case.
Revenue Impact | Directly affects reusable-model margin; magnitude Evidence Pending.
Risk if Unmet | Device attrition erodes reusable economics; model reverts toward single-use costs.
Evidence Source | Judgment/Assumption — reverse-logistics loss is a known reusable-asset risk; SCM-specific rates Evidence Pending.
Supporting Research | Reusable-container / asset-pooling return-rate literature.
VOC Reference | VoC theme — "reusable only pays if we actually get them back."
Strategic Theme | Reusable-fleet integrity.
Product Family | SCM — economic loss.
Lifecycle Stage | Growth
Traceability | UN-SCM-311, PRD-SCM-311; ORCH-018 / EM-ENG-012.
Notes | GPS/BLE tracking and clear return workflows mitigate; ties to MR-EXEC-008 investment case.
7. Evidence & Validation Notes
• Cited loss context (excursion exposure, freezing, wastage, last-mile losses) is drawn from WHO/UNICEF/GAVI reporting and labelled reported/estimated — no exact figures are asserted.
• Standard-based problems (data integrity, cold-battery hazard) rest on FDA Part 11, EU GDP, IEC 62368-1, and UN 38.3.
• All per-customer loss magnitudes and failure rates are Evidence Pending until VoC interviews, field data, and pilot measurement exist; pending VOC themes will be replaced by VOC-### IDs.
Validation path | segment VoC studies, pilot loss/dispute measurement, and PQS field evaluation will convert Judgment/Assumption items into cited evidence.
8. Traceability Placeholders
MR ID | Forward → User Needs | Forward → PRD | Orchestration / Mandate
MR-PROB-001 | UN-SCM-301 | PRD-SCM-301 | ORCH-018 / EM-ENG-012
MR-PROB-002 | UN-SCM-302 | PRD-SCM-302 | ORCH-018 / EM-ENG-012
MR-PROB-003 | UN-SCM-303 | PRD-SCM-303 | ORCH-018 / EM-ENG-012
MR-PROB-004 | UN-SCM-304 | PRD-SCM-304 | ORCH-018 / EM-ENG-012
MR-PROB-005 | UN-SCM-305 | PRD-SCM-305 | ORCH-018 / EM-ENG-012
MR-PROB-006 | UN-SCM-306 | PRD-SCM-306 | ORCH-018 / EM-ENG-012
MR-PROB-007 | UN-SCM-307 | PRD-SCM-307 | ORCH-018 / EM-ENG-012
MR-PROB-008 | UN-SCM-308 | PRD-SCM-308 | ORCH-018 / EM-ENG-012
MR-PROB-009 | UN-SCM-309 | PRD-SCM-309 | ORCH-018 / EM-ENG-012
MR-PROB-010 | UN-SCM-310 | PRD-SCM-310 | ORCH-018 / EM-ENG-012
MR-PROB-011 | UN-SCM-311 | PRD-SCM-311 | ORCH-018 / EM-ENG-012
Upstream: MR-000 (EXEC), MR-100 (OPP), MR-200 (SEG). This module closes the market-requirements set for the SCM.
MR-400 — Jobs To Be Done (JTBD)
Traceability: ORCH-009 (Market Requirements) · ORCH-018 / EM-ENG-012 · based on 00-concept/concept-and-decisions.md (D1–D4). Module code: JTBD.
1. Purpose
This module states, in Jobs-To-Be-Done terms, the underlying progress that cold-chain stakeholders are trying to make when they place a Smart Cold Chain Monitor (SCM) inside a vaccine shipment. It captures the demand-side jobs — functional, emotional, social, operational, and lifecycle — that the market hires a cold-chain monitor to accomplish, independent of any particular technical implementation. It is the demand-space anchor from which Voice of Customer (VOC), Stakeholder Needs (STN), and downstream User Needs (UN) requirements are derived.
2. Scope
In scope: the jobs of the parties who fund, deploy, carry, receive, audit, service, and act on data from a reusable in-container vaccine monitor across the global cold chain — vaccine programs and their funders (e.g., WHO/UNICEF/Gavi-aligned procurers), national/regional immunization stores, pharmaceutical manufacturers and distributors, third-party logistics providers (3PLs) and carriers, last-mile health workers in low- and middle-income countries (LMICs), quality-assurance and regulatory-affairs teams, and fleet/asset administrators. Out of scope: internal engineering functions of the device (covered in Product Requirements), and the therapeutic performance of the vaccine itself (the SCM is a monitoring/information device, not a temperature-controlling or therapeutic product per the concept intended-use statement).
3. Applicable Standards & References
• WHO PQS — Performance, Quality and Safety prequalification scheme; cold-chain temperature-monitoring device/logger performance specifications (exact category code Evidence Pending).
• EU GDP — Good Distribution Practice for medicinal products (2013/C 343/01); transport temperature control and record expectations.
• US FDA — data-integrity expectations for distribution monitoring; 21 CFR Part 11 (electronic records/signatures). Device classification for the SCM Evidence Pending.
• UN 38.3 / IATA DGR — lithium-battery transport (governs the reverse-logistics and air-transport jobs).
• IEC 60068 / IEC 60529 — environmental and ingress conditions the device is used within.
• Program references (Cited where quantitative): WHO/UNICEF Effective Vaccine Management (EVM) framework; WHO/PATH cold-chain and controlled-temperature-chain guidance. Specific figures flagged per field.
4. Background
The core job is enduring and stable: a stakeholder must be able to know, and prove, that a vaccine shipment stayed within its approved temperature (and, increasingly, humidity/shock) limits from origin to point of use — and act quickly when it did not. Vaccines are potency-sensitive biologics; both heat excursions and, critically, freezing of freeze-sensitive vaccines can silently destroy efficacy while leaving the product visually unchanged. Historically this job has been "hired out" to substitute solutions — vaccine vial monitors (VVMs), freeze indicators, chemical/threshold indicators, single-use USB loggers, and manual chart records — each of which does part of the job but leaves gaps in real-time visibility, reuse economics, auditability, and end-to-end evidence. The SCM concept (reusable, connected, in-container, 90-day, −30 °C to +60 °C) is a bet that the market will "fire" fragmented point solutions in favor of one device that does the whole job across trips and markets. This module expresses those jobs as market demands, not features. It is deliberately implementation-neutral: where the concept fixes a decision (e.g., D2 store-and-forward with immediate excursion alerts), the job is stated as the outcome the market wants, and the implementation is left to Product Requirements.
5. Requirement Categories
1. Core functional job — knowing and proving in-range end-to-end.
2. Emotional / assurance jobs — confidence, reduced anxiety, defensible decisions.
3. Social / accountability jobs — trust between shipper, receiver, regulator, donor.
4. Operational jobs — decide fast on receipt; act on alerts in transit.
5. Emergency / excursion-response jobs — timely warning and triage.
6. Installation / provisioning jobs — arm and place the device quickly, correctly.
7. Maintenance / readiness jobs — keep the device charged, calibrated, trip-ready.
8. Administration / audit jobs — retrieve complete, tamper-evident records for release and inspection.
9. Service / reverse-logistics jobs — recover, clean, and re-deploy the reusable device.
6. Detailed Market Requirements
6.1 Core Functional Job
MR-JTBD-001 — Prove Vaccines Stayed In Range, End To End
Requirement Statement | The market demands a way to know and prove that a vaccine shipment remained within its approved temperature limits (and detect freezing and heat excursions) continuously from packing at origin to arrival at the point of use, with an unbroken record.
Business Objective | Enable confident, defensible product-disposition decisions and reduce wastage of potent vaccine.
Customer Segment | Immunization program managers, distributors, QA/regulatory, receiving health facilities.
Geographic Applicability | Global; heightened value on multi-leg LMIC last-mile routes.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Substitute solutions (VVMs, single-use loggers) each cover only part of the origin-to-use span; the market rewards uninterrupted, provable coverage.
Business Value | Protects vaccine value and program credibility; anchors the entire product's reason to exist.
Revenue Impact | Foundational — this job justifies device purchase; failure to serve it collapses all other value. Precise willingness-to-pay Evidence Pending (VOC).
Risk if Unmet | Undetected excursions release compromised vaccine or, conversely, unnecessary discard of good vaccine — both erode trust and program budgets.
Evidence Source | Standard — WHO PQS / EU GDP cold-chain record intent; based on concept intended-use.
Supporting Research | WHO/UNICEF EVM framework emphasis on continuous temperature evidence; freeze sensitivity of vaccines widely documented (cite in VOC as Cited).
VOC Reference | VOC-### (pending) — theme "prove it stayed cold."
Strategic Theme | End-to-end cold-chain assurance.
Product Family | SCM (Smart Cold Chain Monitor) — core monitoring platform.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-001 / PRD-FUN-001 (placeholders); ORCH-018 / EM-ENG-012.
Notes | This is the "core job." All other JTBD entries are contextual jobs around it.
6.2 Emotional / Assurance Jobs
MR-JTBD-002 — Give Decision-Makers Confidence, Not Anxiety
Requirement Statement | The market demands that the party accepting or releasing a shipment feels confident the cold-chain evidence is complete and trustworthy, removing the anxiety of "signing off blind" on product they cannot inspect for potency.
Business Objective | Reduce decision friction and defensive over-discard driven by uncertainty.
Customer Segment | Receiving pharmacists, store managers, QA release officers, last-mile health workers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Ambiguous or partial records (a lone freeze indicator, a logger someone forgot to start) leave receivers uneasy; trustworthy end-to-end evidence differentiates.
Business Value | Faster, calmer acceptance; fewer disputes; less wasted good stock.
Revenue Impact | Indirect but strong retention driver; supports premium positioning (see MR-POS).
Risk if Unmet | Users distrust the device and revert to substitutes or over-discard, undermining ROI.
Evidence Source | Judgment/Assumption — based on cold-chain decision behaviour; confirm via VOC.
Supporting Research | Behavioural pattern of "when in doubt, discard/accept" reported in EVM assessments (verify as Cited in VOC).
VOC Reference | VOC-### (pending) — theme "confidence at hand-off."
Strategic Theme | Trusted assurance.
Product Family | SCM — HMI/dashboard evidence presentation.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-002 / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Emotional job; realized largely through clarity of status and record, not new sensing.
6.3 Social / Accountability Jobs
MR-JTBD-003 — Establish Shared, Tamper-Evident Truth Between Parties
Requirement Statement | The market demands a single, tamper-evident cold-chain record that shipper, carrier, receiver, regulator, and donor can all trust as the shared source of truth, so accountability is clear when an excursion occurs.
Business Objective | De-risk multi-party hand-offs and dispute resolution across the chain.
Customer Segment | Distributors, 3PLs/carriers, QA/regulatory, donor/funder oversight.
Geographic Applicability | Global; critical where custody changes hands multiple times.
Market Priority | High
Importance | High
Competitive Driver | Paper/threshold indicators are disputable and easily reset; a defensible digital record with data integrity (Part 11-aligned) is a recognized differentiator.
Business Value | Reduces liability disputes, speeds claims, protects donor/program reputation.
Revenue Impact | Enables enterprise/institutional contracts where auditability is a purchase gate.
Risk if Unmet | Finger-pointing at hand-off, unresolved liability, loss of institutional buyers.
Evidence Source | Standard — EU GDP / FDA 21 CFR Part 11 data-integrity expectations.
Supporting Research | GDP transport-record and data-integrity guidance (Standard).
VOC Reference | VOC-### (pending) — theme "whose fault, provable."
Strategic Theme | Accountable custody.
Product Family | SCM — data integrity & audit trail.
Lifecycle Stage | Concept
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Social job; interacts with cybersecurity (data integrity) — flag EM cyber persona in PRD stage.
6.4 Operational Jobs
MR-JTBD-004 — Decide Accept/Reject In Seconds At Receipt
Requirement Statement | The market demands that a receiver can determine within seconds, at the point of delivery, whether a shipment is within limits and safe to accept — without specialist tools, connectivity, or data expertise.
Business Objective | Keep receiving throughput high and disposition decisions immediate and correct.
Customer Segment | Last-mile health workers, store clerks, receiving docks in LMICs.
Geographic Applicability | Global; acute where staff are non-technical and time-pressured.
Market Priority | Must-have
Importance | Critical
Competitive Driver | VVMs and go/no-go indicators are prized precisely for instant at-a-glance disposition; any richer device must match that immediacy or it loses the last-mile job.
Business Value | Prevents bottlenecks and wrong calls at the most error-prone step.
Revenue Impact | Decisive for last-mile/LMIC adoption; gate for volume programs.
Risk if Unmet | Workers ignore the device, mis-accept excursed stock, or discard good stock.
Evidence Source | Judgment/Assumption — based on last-mile workflow reality; confirm via VOC.
Supporting Research | VVM at-a-glance model as the usability benchmark (industry-known).
VOC Reference | VOC-### (pending) — theme "green/red at the door."
Strategic Theme | Instant disposition.
Product Family | SCM — on-device status (OLED/indicator) + offline verdict.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Must work with no connectivity — aligns with D2 (local logging) as an outcome, not a feature.
6.5 Emergency / Excursion-Response Jobs
MR-JTBD-005 — Warn The Right People In Time To Save The Shipment
Requirement Statement | The market demands to be alerted to a temperature excursion (or freeze/shock event) quickly enough, and to the right responder, that intervention can still protect the shipment or divert it before value is lost.
Business Objective | Convert monitoring into loss prevention, not just post-mortem record-keeping.
Customer Segment | Logistics control towers, distributors, program cold-chain officers, 3PLs.
Geographic Applicability | Global; highest value on long-haul and long-dwell legs.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Real-time IoT trackers (Controlant, Tive, Roambee-class) compete primarily on timely alerting; single-use loggers cannot serve this job at all.
Business Value | Enables in-transit rescue and rerouting; directly avoids write-offs.
Revenue Impact | Primary differentiator vs single-use loggers; supports higher price point.
Risk if Unmet | Excursions discovered only on arrival; shipment already lost — device reduced to a "black box recorder."
Evidence Source | Judgment/Assumption + Standard — based on concept D2 (immediate excursion alert); alert latency depends on connectivity (carried as hazard in concept).
Supporting Research | Real-time-alert value proposition of reusable IoT trackers (industry-known).
VOC Reference | VOC-### (pending) — theme "tell me before it's too late."
Strategic Theme | In-transit loss prevention.
Product Family | SCM — excursion detection + alerting/escalation.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-FUN-0xx; HA-### (delayed/missing alert hazard); ORCH-018 / EM-ENG-012.
Notes | Latency is connectivity-bound (LTE-M primary) — state as market demand, resolve in PRD/hazard.
6.6 Installation / Provisioning Jobs
MR-JTBD-006 — Arm And Place A Device Correctly In Under A Minute
Requirement Statement | The market demands that a shipment can be configured, armed, and correctly placed inside the container by ordinary warehouse staff in well under a minute, with the trip's profile (limits, cadence, route) set without error.
Business Objective | Make correct setup the path of least resistance at high-throughput pack-out.
Customer Segment | Manufacturer/distributor pack-out staff, 3PL warehouse operators.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | "Forgot to start the logger," wrong threshold, or bad placement are classic failure modes of legacy loggers; frictionless provisioning is a real differentiator.
Business Value | Fewer void trips and setup errors; higher usable-data yield per shipment.
Revenue Impact | Protects data yield that underpins the value proposition; reduces support cost.
Risk if Unmet | Mis-armed/void devices produce no usable evidence — silent value loss and churn.
Evidence Source | Judgment/Assumption — based on known logger failure modes; confirm via VOC.
Supporting Research | Reported "logger not started/misconfigured" failures in field programs (verify as Cited).
VOC Reference | VOC-### (pending) — theme "set it and forget it, correctly."
Strategic Theme | Frictionless deployment.
Product Family | SCM — mobile/BLE provisioning + on-device confirmation.
Lifecycle Stage | Concept
Traceability | forward to UN-PER-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Placement guidance (int vs ext probe) matters — resolve in User Needs/PRD.
6.7 Maintenance / Readiness Jobs
MR-JTBD-007 — Keep A Reusable Fleet Charged, Calibrated, And Trip-Ready
Requirement Statement | The market demands assurance that any device drawn from the reusable pool is charged, in calibration, and firmware-current — i.e., trustworthy for the next 90-day trip — without manual per-device checking.
Business Objective | Sustain data trustworthiness and endurance across many reuse cycles (D1).
Customer Segment | Fleet/asset administrators, distributor operations, 3PL depots.
Geographic Applicability | Global; concentrated at hubs and return depots.
Market Priority | High
Importance | High
Competitive Driver | Reuse economics are the reusable-device value story; hidden drift or dead batteries destroy that story and hand advantage back to single-use.
Business Value | Protects the reuse ROI and record credibility; avoids deploying stale/uncalibrated units.
Revenue Impact | Underpins the total-cost-of-ownership argument central to SCM's business case.
Risk if Unmet | Fleet drifts out of calibration or endurance; disputed records; ROI evaporates.
Evidence Source | Judgment/Assumption + Standard — based on D1 reusability and calibration norms.
Supporting Research | Metrology re-calibration expectations for monitoring devices (Standard; interval Evidence Pending).
VOC Reference | VOC-### (pending) — theme "is this unit good to send?"
Strategic Theme | Fleet readiness & reuse economics.
Product Family | SCM — battery/charge health, calibration status, OTA.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-PERF-0xx; ORCH-018 / EM-ENG-012.
Notes | Low-temp charging inhibited below 0 °C (D3) affects readiness workflow — carry to PRD.
6.8 Administration / Audit Jobs
MR-JTBD-008 — Retrieve Complete, Inspection-Ready Records On Demand
Requirement Statement | The market demands that a complete, tamper-evident cold-chain record for any shipment or device can be retrieved on demand in a form acceptable to GDP/FDA inspectors and program auditors, without reconstructing data from fragments.
Business Objective | Pass inspections and audits efficiently and defend releases retrospectively.
Customer Segment | QA/regulatory affairs, quality directors, donor/program M&E teams.
Geographic Applicability | Global; regulatory framing varies (WHO PQS / EU GDP / US FDA).
Market Priority | High
Importance | High
Competitive Driver | Enterprise pharma buyers gate purchases on data integrity and audit-readiness; platform players win here, threshold indicators cannot.
Business Value | Lower audit cost, faster inspections, stronger release defensibility.
Revenue Impact | Purchase gate for regulated distributors and manufacturers.
Risk if Unmet | Failed inspections, unusable evidence, exclusion from regulated tenders.
Evidence Source | Standard — EU GDP records; FDA 21 CFR Part 11.
Supporting Research | Part 11 audit-trail and record-retention expectations (Standard).
VOC Reference | VOC-### (pending) — theme "get me the record the auditor wants."
Strategic Theme | Audit-ready by default.
Product Family | SCM — cloud record store, export, audit trail.
Lifecycle Stage | Concept
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Retention period and export format Evidence Pending — set in regulatory stage.
6.9 Service / Reverse-Logistics Jobs
MR-JTBD-009 — Recover, Clean, And Redeploy The Device Safely
Requirement Statement | The market demands a practical way to return, clean/decontaminate, inspect, and re-provision a used device for its next trip — including safe/compliant transport of its lithium battery — so reuse is operationally and legally viable.
Business Objective | Make the reusable model logistically real, not just a spec claim.
Customer Segment | Reverse-logistics operators, 3PL depots, distributor operations, sustainability leads.
Geographic Applicability | Global; return economics hardest on remote LMIC last-mile.
Market Priority | Medium
Importance | High
Competitive Driver | Single-use loggers avoid returns entirely; the reusable value case only holds if reverse logistics is cheap and compliant, or advantage reverts to disposables.
Business Value | Realizes reuse ROI and sustainability benefit; controls per-trip cost.
Revenue Impact | Determines whether the total-cost-of-ownership story survives contact with logistics.
Risk if Unmet | Devices not returned/reused; stranded assets; TCO case fails.
Evidence Source | Standard + Judgment/Assumption — UN 38.3 / IATA govern battery return; return-rate economics Evidence Pending.
Supporting Research | Lithium-battery transport constraints (Standard); reusable-tracker return programs (industry-known).
VOC Reference | VOC-### (pending) — theme "getting the box back."
Strategic Theme | Circular reuse & reverse logistics.
Product Family | SCM — cleanability, return workflow, battery-transport compliance.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-MECH-0xx; ORCH-018 / EM-ENG-012.
Notes | Cleaning between trips flagged as residual in concept D1 — carry to User/Service requirements.
MR-JTBD-010 — Manage The Whole Fleet As One Visible Asset Base
Requirement Statement | The market demands a way to see, locate, and manage the entire deployed and idle fleet as a single asset base — utilization, location, health, and status — so operators can allocate devices, chase strays, and plan capacity.
Business Objective | Maximize utilization and minimize lost/idle assets across a reusable fleet.
Customer Segment | Fleet/asset administrators, operations planners, finance (asset accounting).
Geographic Applicability | Global multi-site networks.
Market Priority | Medium
Importance | Medium
Competitive Driver | Reusable IoT platforms differentiate on fleet analytics and utilization; this is a recognized enterprise expectation.
Business Value | Higher asset utilization, fewer losses, better capital planning.
Revenue Impact | Supports platform/subscription revenue and expansion within accounts.
Risk if Unmet | Fleets leak devices, sit idle, or can't scale — eroding the platform value story.
Evidence Source | Judgment/Assumption — based on fleet-model logic of D1; confirm via VOC.
Supporting Research | Fleet/asset-management dashboards of reusable-tracker vendors (industry-known).
VOC Reference | VOC-### (pending) — theme "where are my devices?"
Strategic Theme | Fleet visibility & utilization.
Product Family | SCM — cloud fleet management + GPS/location.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Location relies on GPS/connectivity; performance inside metallized containers is an open item (concept §7).
7. Evidence & Validation Notes
Evidence posture | Core-job and regulatory-driven jobs (MR-JTBD-001/003/005/008) are anchored in Standards and concept decisions. Behavioural/operational jobs (002/004/006/007/010) are marked Judgment/Assumption and must be confirmed by primary VOC before promotion to User Needs.
Validation method | Contextual inquiry and interviews across each segment (pack-out, carrier, last-mile, QA, fleet admin, reverse-logistics); job-importance/satisfaction survey to size opportunity.
Quantification gap | No market-size, willingness-to-pay, wastage-rate, or return-rate figures are asserted here as fact; all such numbers are Evidence Pending until sourced from WHO/UNICEF/Gavi or named industry analysts in VOC/COMP.
Open concept items carried | WHO PQS category code; US device classification; low-temp battery endurance at −30 °C; antenna performance inside metallized containers (concept §7).
8. Traceability Placeholders
Upstream | 00-concept/concept-and-decisions.md (D1–D4); ORCH-009 market-requirements standard.
Peer modules | MR-500 (COMP), MR-600 (SOTA), MR-700 (POS); MR-100/200/300 (EXEC/OPP/SEG — pending).
Downstream (forward) | VOC-### → STN-### → UN-JTBD-### → PRD-FUN/PERF/MECH-### → SYS-### → HA-### / FA-### → PV-###/VAL-###.
Governance | ORCH-018 (traceability) / EM-ENG-012 (traceability engineer) maintain the end-to-end links.
MR-500 — Competitive Landscape (COMP)
Traceability: ORCH-009 (Market Requirements) · ORCH-018 / EM-ENG-012 · based on 00-concept/concept-and-decisions.md (D1–D4). Module code: COMP.
1. Purpose
This module maps the competitive and substitute landscape the Smart Cold Chain Monitor (SCM) must win in, and derives the market requirements that the competitive field imposes — the capabilities, proof points, and positioning the market will expect because rivals or substitutes already set the bar. It is a demand-side reading of competition: what the market demands given who else serves the job (see MR-400 JTBD), not an internal feature comparison.
2. Scope
In scope: direct competitors (single-use USB temperature loggers; reusable real-time IoT trackers), substitute/adjacent solutions (WHO-PQS-prequalified electronic shipping indicators, freeze indicators, vaccine vial monitors [VVMs], chemical indicators, manual chart recorders), integrated pharma serialization/track-and-trace platforms, and emerging startups. Also in scope: competitive strengths, weaknesses, and a patent-landscape posture note. Out of scope: detailed pricing teardowns and named contract values (Evidence Pending), and internal SCM design responses (Product Requirements).
3. Applicable Standards & References
• WHO PQS — prequalification of cold-chain temperature-monitoring devices / 30-day electronic shipping indicators / freeze indicators / VVMs (specific PQS codes Evidence Pending; do not fabricate).
• EU GDP — Good Distribution Practice; the baseline record/transport expectation competitors meet.
• US FDA — device listing/clearance and data-integrity (21 CFR Part 11); competitor FDA records Evidence Pending.
• UN 38.3 / IATA DGR — lithium-battery transport (constrains reusable/real-time competitors alike).
• FCC Part 15 / CE RED, IEC 62368-1, CISPR/IEC 61000 — radio/EMC/safety regimes competitors certify to.
• Landscape sources: competitor and category facts cited as reported/industry-known; specific regulatory record IDs and market shares flagged Evidence Pending.
4. Background
The cold-chain monitoring market is layered. At the low end, single-use USB temperature loggers (reported classes: Sensitech TempTale, ELPRO LIBERO, DeltaTrak) are cheap, ubiquitous, and trusted for a per-trip PDF record, but offer no real-time alerting and no reuse. At the high end, reusable real-time IoT trackers (reported classes: Controlant, Roambee, Tive, Berlinger/Q-tag, Monnit, LogTag) offer live cellular visibility, alerting, and fleet dashboards — the segment the SCM competes in most directly. In parallel, the vaccine-specific world relies on WHO-PQS-prequalified indicators and VVMs (e.g., Temptime/Zebra-associated technologies): 30-day electronic shipping indicators, freeze indicators, and vial-level VVMs that are inexpensive, PQS-credentialed, and dominant in public-sector immunization — strong substitutes for parts of the SCM's job at the last mile. Integrated serialization/track-and-trace platforms frame monitoring as one feed within broader supply-chain visibility, competing for the buyer's budget and integration attention. Finally, emerging startups push NB-IoT/LTE-M sensor tags, AI excursion prediction, and low-cost reusable form factors. The SCM's wedge is the combination the concept locks: reusable + in-container + 90-day + −30 °C-to-+60 °C + multi-market (D1–D4) + immediate alerts (D2). No single incumbent obviously owns that exact intersection, but each does part of it well; the competitive requirements below capture what the SCM must therefore deliver and prove to the market.
4.1 Competitive Analysis Table (evidence-backed)
Facts are reported/industry-known category characterizations; specific regulatory record IDs and share figures are Evidence Pending and must not be invented.
Company | Product | Description | Regulatory/Clearance records | Associated Regulation/Standard | Comments
Sensitech (Carrier) | TempTale (single-use USB logger class) | Disposable per-trip temperature logger; on-arrival PDF/CSV record; widely used in pharma distribution | Evidence Pending (FDA listing / declarations) | EU GDP; FDA data integrity; IEC 62368-1; UN 38.3 (battery) | Trusted, cheap per unit; no real-time alerting; no reuse; not the in-container 90-day reusable job
ELPRO | LIBERO (single-use & data-logger family) | Temperature/humidity loggers with validated software reports; strong pharma-GxP positioning | Evidence Pending | EU GDP; 21 CFR Part 11-aligned reporting | Strong GxP/reporting credibility; limited real-time; reuse limited by model
DeltaTrak | Single-use / Flash Loggers & real-time class | Broad logger range incl. some connected models; food + pharma cold chain | Evidence Pending | EU GDP; FDA; radio certs where connected | Broad, value-priced; mixed real-time maturity; brand more food-chain weighted
Controlant | Reusable real-time IoT logger + platform | Cellular reusable loggers with cloud platform; used by major pharma for global distribution | Evidence Pending | EU GDP; 21 CFR Part 11; UN 38.3; radio/EMC | Closest platform analog; enterprise-pharma entrenched; managed-service model
Roambee | Reusable IoT sensor + BeeBeacon/analytics | Real-time multi-sensor (temp/humidity/shock/location) reusable trackers as a service | Evidence Pending | EU GDP; radio certs; UN 38.3 | Strong analytics/as-a-service; broader-than-vaccine focus
Tive | Solo 5G/real-time tracker | Single-use-leaning real-time trackers (temp/light/shock/location) with live dashboard | Evidence Pending | FDA data integrity; radio certs; UN 38.3 | Strong live visibility & UX; positioning less LMIC/PQS-oriented
Berlinger / Q-tag | Q-tag / cold-chain indicators & loggers | PDF loggers and indicators; long vaccine cold-chain heritage incl. WHO programs | Evidence Pending (some WHO PQS-associated) | WHO PQS; EU GDP | Deep vaccine/WHO credibility; more indicator/logger than connected fleet
Monnit / LogTag | Low-cost sensors / loggers | Value loggers and wireless sensors; LogTag widely used in vaccine fridges/transport | Evidence Pending (LogTag WHO PQS-associated for some models) | WHO PQS (select); EU GDP | Very cost-competitive; last-mile familiar; limited in-container real-time cellular
Temptime (Zebra) | VVMs, freeze indicators, 30-day electronic shipping indicators | Vaccine-specific indicators incl. WHO-PQS 30-day electronic shipping indicators & VVMs | Evidence Pending (WHO PQS prequalified — codes TBC) | WHO PQS; UNICEF/Gavi procurement specs | Dominant vaccine-substitute at vial/shipment level; ultra-low-cost; no reuse/real-time fleet
VVMs / freeze / chemical indicators (category) | Vial monitors, freeze/threshold indicators | Passive, at-a-glance go/no-go potency/freeze proxies on vials or in cartons | WHO PQS prequalified (category); codes Evidence Pending | WHO PQS; program procurement | The last-mile usability benchmark; passive, disposable, non-auditable digitally
Serialization / track-&-trace platforms (category) | e.g., enterprise supply-chain visibility suites | Integrated serialization + condition-monitoring feeds within broader logistics visibility | Evidence Pending | EU FMD/US DSCSA (serialization); GDP | Compete for buyer budget/integration; monitoring is one feed, not core device
Emerging startups (category) | NB-IoT/LTE-M tags, AI excursion prediction | Low-cost reusable sensor tags, predictive analytics, novel form factors | Evidence Pending | Regional radio/EMC; GDP | Fast-moving; price/analytics pressure; variable regulatory maturity
5. Requirement Categories
1. Parity-to-compete — table stakes competitors already set (record quality, GDP/Part 11).
2. Real-time differentiation — alerting/visibility vs single-use loggers.
3. Substitute displacement — winning the last-mile job from VVMs/indicators.
4. PQS/public-sector credibility — competing where WHO PQS gates procurement.
5. Platform & integration — competing with integrated track-and-trace suites.
6. Reuse economics — the total-cost-of-ownership battleground.
7. Emerging-threat resilience — startups, AI, NB-IoT pressure.
8. Positioning & IP — strengths/weaknesses framing and patent-landscape posture.
6. Detailed Market Requirements
6.1 Parity To Compete
MR-COMP-001 — Match The Trusted Per-Trip Record Incumbents Set
Requirement Statement | The market demands that any new monitor produce a per-trip cold-chain record at least as complete, exportable, and trusted as the PDF/CSV records single-use loggers (TempTale/LIBERO/ DeltaTrak class) have made the norm.
Business Objective | Remove "switching risk" so buyers can adopt without losing familiar evidence.
Customer Segment | Distributors, QA/regulatory, 3PLs migrating from single-use loggers.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Incumbent loggers have trained the market to expect a clean, auditable per-trip report; falling short of that baseline is disqualifying.
Business Value | Enables one-for-one displacement of single-use loggers.
Revenue Impact | Gate to the largest replaceable installed base.
Risk if Unmet | Buyers see the SCM as unproven and stay with familiar disposables.
Evidence Source | Standard + reported — EU GDP records; incumbent report formats (industry-known).
Supporting Research | Ubiquity of logger PDF reports in pharma distribution (industry-known).
VOC Reference | VOC-### (pending) — theme "give me the report I already trust."
Strategic Theme | Frictionless displacement.
Product Family | SCM — record generation & export.
Lifecycle Stage | Launch
Traceability | forward to UN-JTBD-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Parity is necessary, not sufficient; differentiation comes in 6.2–6.6.
MR-COMP-002 — Meet The GDP / Part 11 Data-Integrity Bar Buyers Already Enforce
Requirement Statement | The market demands data-integrity, audit-trail, and validated-reporting credibility equal to the GxP-strong incumbents (e.g., ELPRO/Controlant class), because regulated buyers will not adopt a monitor that weakens their compliance posture.
Business Objective | Clear the compliance gate that governs regulated-pharma procurement.
Customer Segment | Pharma manufacturers, regulated distributors, QA/regulatory affairs.
Geographic Applicability | Global (EU GDP, US FDA especially).
Market Priority | Must-have
Importance | Critical
Competitive Driver | GxP-credible vendors win enterprise pharma partly on validated software and data-integrity documentation; this is a known purchase gate.
Business Value | Access to high-value regulated accounts.
Revenue Impact | Gate to enterprise/pharma revenue tier.
Risk if Unmet | Excluded from regulated tenders regardless of hardware merit.
Evidence Source | Standard — EU GDP; FDA 21 CFR Part 11.
Supporting Research | Vendor emphasis on validated/Part 11 reporting (industry-known).
VOC Reference | VOC-### (pending) — theme "won't break my audit."
Strategic Theme | Compliance credibility.
Product Family | SCM — data integrity, audit trail, validation package.
Lifecycle Stage | Launch
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Coordinates with cybersecurity + regulatory stages.
6.2 Real-Time Differentiation
MR-COMP-003 — Beat Single-Use Loggers On Real-Time Visibility And Alerting
Requirement Statement | The market demands in-transit visibility and immediate excursion alerting that single-use loggers structurally cannot offer, making live intervention — not just post-hoc discovery — the reason to switch.
Business Objective | Establish the primary "why change" over the disposable installed base.
Customer Segment | Distributors, control towers, 3PLs, high-value-shipment owners.
Geographic Applicability | Global; strongest on long-haul/high-value lanes.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Real-time IoT trackers (Controlant/Tive/Roambee class) have proven the market values live alerting; single-use loggers are exposed here.
Business Value | In-transit loss avoidance; premium justification.
Revenue Impact | Core differentiator enabling higher price vs disposables.
Risk if Unmet | SCM reduced to an expensive reusable logger with no live edge.
Evidence Source | reported + Judgment/Assumption — based on concept D2 (immediate alerts).
Supporting Research | Real-time-tracker value narrative (industry-known).
VOC Reference | VOC-### (pending) — theme "know now, not on arrival."
Strategic Theme | Live loss prevention.
Product Family | SCM — alerting/escalation + connectivity (LTE-M).
Lifecycle Stage | Launch
Traceability | forward to UN-JTBD-005 / PRD-FUN-0xx; HA-### (alert latency); ORCH-018 / EM-ENG-012.
Notes | Alert latency is connectivity-bound — market demand stated as outcome.
6.3 Substitute Displacement (Last Mile)
MR-COMP-004 — Match VVM/Indicator At-A-Glance Simplicity To Win The Last Mile
Requirement Statement | The market demands that, at the point of use, a receiver gets a go/no-go verdict as instantly and simply as a VVM or freeze indicator — otherwise the SCM cannot displace the substitutes that dominate LMIC immunization.
Business Objective | Compete against passive indicators on their strongest ground: instant usability.
Customer Segment | Last-mile health workers, immunization programs, LMIC stores.
Geographic Applicability | Global; decisive in LMIC public-sector.
Market Priority | High
Importance | High
Competitive Driver | VVMs/indicators are ultra-cheap and universally understood; a connected device must equal their immediacy to displace them at the vial/carton edge.
Business Value | Opens the large public-sector immunization opportunity.
Revenue Impact | Access to WHO/UNICEF/Gavi-aligned volume (sizing Evidence Pending).
Risk if Unmet | Substitutes retain the last mile; SCM confined to commercial pharma lanes.
Evidence Source | reported + Judgment/Assumption — VVM/indicator dominance (industry-known).
Supporting Research | WHO-PQS indicator/VVM prevalence in immunization (industry-known; codes Evidence Pending).
VOC Reference | VOC-### (pending) — theme "as simple as a VVM."
Strategic Theme | Last-mile displacement.
Product Family | SCM — on-device verdict / offline status.
Lifecycle Stage | Growth
Traceability | forward to UN-JTBD-004 / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | SCM complements (not replaces) vial-level VVMs; position as shipment/container layer.
6.4 PQS / Public-Sector Credibility
MR-COMP-005 — Earn WHO-PQS-Grade Credibility Where Procurement Requires It
Requirement Statement | The market demands WHO-PQS-level performance credibility (accuracy, freeze detection, robustness) so the SCM can compete in public-sector immunization procurement where PQS status, which incumbents like Temptime/Berlinger/LogTag hold for various products, gates the buy.
Business Objective | Qualify for PQS-gated tenders and program procurement.
Customer Segment | Ministries of health, UNICEF/Gavi-aligned procurers, program managers.
Geographic Applicability | Global public-sector; concentrated in LMICs.
Market Priority | High
Importance | Critical
Competitive Driver | PQS prequalification is a hard gate; competitors with PQS-listed products have a structural advantage in these channels.
Business Value | Unlocks donor-funded and government procurement volume.
Revenue Impact | Large but access-gated; sizing Evidence Pending.
Risk if Unmet | Locked out of the biggest vaccine-specific channel.
Evidence Source | Standard — WHO PQS scheme; competitor PQS records Evidence Pending.
Supporting Research | PQS gating role in immunization procurement (Standard/industry-known).
VOC Reference | VOC-### (pending) — theme "is it PQS?"
Strategic Theme | Public-sector qualification.
Product Family | SCM — performance + PQS submission.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-PERF-0xx; regulatory stage; ORCH-018 / EM-ENG-012.
Notes | Exact PQS category code is an open concept item (§7) — do not assert.
6.5 Platform & Integration
MR-COMP-006 — Integrate, Rather Than Compete Head-On, With Track-And-Trace Platforms
Requirement Statement | The market demands that a cold-chain monitor feed cleanly into the integrated serialization/track-and-trace and supply-chain-visibility platforms buyers already run, so the SCM adds a best-in-class condition-monitoring feed instead of forcing a rip-and-replace.
Business Objective | Convert platform incumbents from competitors into distribution partners.
Customer Segment | Enterprise pharma IT, distributors, 3PLs with existing visibility suites.
Geographic Applicability | Global enterprise.
Market Priority | High
Importance | High
Competitive Driver | Integrated platforms compete for the same budget; interoperability neutralizes the "single-pane-of-glass" objection and rides existing adoption.
Business Value | Faster enterprise adoption; ecosystem leverage.
Revenue Impact | Enables partner/channel-led revenue and stickiness.
Risk if Unmet | Seen as a data silo; loses to bundled platform monitoring.
Evidence Source | Judgment/Assumption + reported — platform integration expectations (industry-known).
Supporting Research | DSCSA/FMD serialization + visibility convergence (Standard/industry-known).
VOC Reference | VOC-### (pending) — theme "must fit our stack."
Strategic Theme | Interoperable-by-design.
Product Family | SCM — cloud APIs/integration.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Integration scope set with cybersecurity/architecture stages.
6.6 Reuse Economics
MR-COMP-007 — Win The Total-Cost-Of-Ownership Argument Against Disposables
Requirement Statement | The market demands a demonstrable per-trip total-cost-of-ownership advantage over single-use loggers across a realistic reuse life, or reusable devices lose to the low unit cost and zero-return simplicity of disposables.
Business Objective | Make reuse economics, not just features, the winning argument.
Customer Segment | Procurement/finance, operations, sustainability leads.
Geographic Applicability | Global; hardest to prove on high-loss LMIC last-mile.
Market Priority | High
Importance | High
Competitive Driver | Disposables win on low unit price and no reverse logistics; the reusable case must quantify break-even reuse count and net savings.
Business Value | Justifies capital purchase / subscription over per-trip disposables.
Revenue Impact | Central to the SCM business model's viability.
Risk if Unmet | Buyers default to disposables on headline unit price.
Evidence Source | Judgment/Assumption — TCO model inputs (return rate, reuse cycles) Evidence Pending.
Supporting Research | Reusable-vs-disposable TCO debates in cold chain (industry-known).
VOC Reference | VOC-### (pending) — theme "does reuse actually pay?"
Strategic Theme | Reuse economics.
Product Family | SCM — durability, reuse-cycle life, reverse logistics.
Lifecycle Stage | Growth
Traceability | forward to UN-JTBD-009 / PRD-MECH-0xx; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-JTBD-009 (reverse logistics) and MR-POS pricing.
6.7 Emerging-Threat Resilience
MR-COMP-008 — Stay Ahead Of Low-Cost NB-IoT/LTE-M Tags And AI Prediction Startups
Requirement Statement | The market demands that the SCM keep pace with fast-moving startups pushing cheaper NB-IoT/LTE-M sensor tags and AI-driven excursion prediction, so it is not undercut on price or out-innovated on analytics before it scales.
Business Objective | Protect differentiation and price against rapid new entrants.
Customer Segment | All segments, especially price-sensitive and analytics-forward buyers.
Geographic Applicability | Global.
Market Priority | Medium
Importance | High
Competitive Driver | Low barriers to basic connected sensing mean startups can erode price and claim analytics leadership quickly.
Business Value | Sustains margin and technical lead over the product lifecycle.
Revenue Impact | Protects pricing power and market share as the category commoditizes.
Risk if Unmet | Rapid commoditization; SCM caught between low-cost tags and entrenched platforms.
Evidence Source | Judgment/Assumption — startup activity (industry-known); see MR-600 SOTA.
Supporting Research | NB-IoT/LTE-M and predictive-analytics trends (industry-known; detailed in SOTA).
VOC Reference | VOC-### (pending) — theme "will this be obsolete fast?"
Strategic Theme | Defensible innovation.
Product Family | SCM — roadmap, OTA extensibility, analytics.
Lifecycle Stage | Sustaining
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; MR-600 SOTA; ORCH-018 / EM-ENG-012.
Notes | OTA (concept) is a key hedge — enables field feature/analytics upgrades.
6.8 Positioning & IP
MR-COMP-009 — Own A Defensible Position At The Reusable / In-Container / Cold-Capable / Multi-Market Intersection
Requirement Statement | The market demands a clearly articulated, defensible position at the specific intersection no incumbent fully owns — reusable + in-container + −30 °C-capable + 90-day + multi-market with immediate alerts — supported by an IP posture that protects it.
Business Objective | Convert a genuine white-space into a durable, communicable, and protected position.
Customer Segment | Buyers, partners, investors, and (via IP) the competitive field.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Each rival owns part of the job; the SCM's advantage is the combination, which must be both messaged (see MR-POS) and, where possible, protected.
Business Value | Durable differentiation and defensibility; investor/partner narrative.
Revenue Impact | Supports premium and reduces fast-follow erosion.
Risk if Unmet | Value proposition copied; position blurs into a crowded field.
Evidence Source | Judgment/Assumption — white-space read; patent landscape is Evidence Pending (formal FTO/landscape search required before any IP claim).
Supporting Research | Category segmentation above (industry-known); IP landscape not yet searched.
VOC Reference | VOC-### (pending) — theme "what makes it different, durably."
Strategic Theme | Defensible white-space.
Product Family | SCM — whole-product positioning + IP.
Lifecycle Stage | Concept
Traceability | forward to MR-POS-### / UN-VIS-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Patent landscape note: no freedom-to-operate or patentability search has been performed; treat all IP defensibility as an open action, not a claim. Coordinate with legal/IP function.
7. Evidence & Validation Notes
Evidence posture | All competitor and category descriptions are reported/industry-known; no regulatory record IDs, PQS codes, FDA listings, market shares, or contract values are asserted as fact — each is marked Evidence Pending. The table's "Regulatory/Clearance records" column is deliberately Evidence Pending pending records verification.
Validation method | Confirm competitor regulatory records via WHO PQS database, FDA listings, and vendor documentation; commission a formal patent/freedom-to-operate search (MR-COMP-009); size substitute installed base via WHO/UNICEF/Gavi and named analysts.
Key open items | exact WHO PQS category and competitor PQS codes; competitor FDA classifications; TCO model inputs (reuse-cycle life, return rates); patent landscape.
Do-not-fabricate reminder | per project hallucination policy, record numbers and precise shares remain gaps until sourced.
8. Traceability Placeholders
Upstream | 00-concept/concept-and-decisions.md (D1–D4); MR-400 (JTBD).
Peer modules | MR-600 (SOTA), MR-700 (POS); MR-100/200/300 (pending).
Downstream (forward) | VOC-### → STN-### → UN-VIS/STK/JTBD-### → PRD-FUN/PERF/MECH-### → SYS-### → PV-###/VAL-###.
Governance | ORCH-018 / EM-ENG-012 maintain competitive-claim traceability and evidence status.
MR-600 — State of the Art (SOTA)
Traceability: ORCH-009 (Market Requirements) · ORCH-018 / EM-ENG-012 · based on 00-concept/concept-and-decisions.md (D1–D4). Module code: SOTA.
1. Purpose
This module records the state of the art in cold-chain monitoring technology and derives the market requirements that current technical maturity imposes on the Smart Cold Chain Monitor (SCM). It answers: given what is now technically normal, expected, or emerging in loggers, real-time IoT, connectivity, low-temperature power, cloud/AI analytics, indicators/VVMs, and data integrity — what will the market expect the SCM to embody, avoid, or exceed? It is the technology-baseline anchor for downstream requirements.
2. Scope
In scope: the current and near-term technical state of single-use and reusable logging; real-time IoT architectures; LTE-M/NB-IoT and multi-bearer connectivity; low-temperature rechargeable battery technology and duty-cycling; cloud/analytics/AI adoption; WHO PQS baseline device technology; VVM and indicator state of the art; and data-integrity/blockchain and automation trends. Out of scope: SCM-specific design choices (Product/System Requirements) and vendor-by-vendor comparison (MR-500 COMP).
3. Applicable Standards & References
• WHO PQS — baseline performance technology for cold-chain monitors/indicators/VVMs (codes Evidence Pending).
• EU GDP / FDA 21 CFR Part 11 — data-integrity state of practice.
• 3GPP LTE-M (Cat-M1) / NB-IoT — low-power wide-area cellular standards underpinning connectivity SOTA.
• UN 38.3 / IATA DGR — lithium-battery constraints bounding battery/energy SOTA.
• IEC 62368-1 / CISPR / IEC 61000 / IEC 60529 / IEC 60068 — safety/EMC/ingress/environmental baselines.
• Technology sources cited as reported/industry-known; quantitative claims flagged per field, precise figures Evidence Pending.
4. Background
Cold-chain monitoring technology has moved through three overlapping generations. Generation one — passive indicators (VVMs, freeze/threshold indicators, chemical labels) and chart recorders — remains dominant at the vaccine vial/last-mile edge because it is cheap, robust, and instantly readable; VVM chemistry in particular is a mature, WHO-PQS-credentialed technology. Generation two — electronic single-use USB/Bluetooth data loggers — brought accurate digital records, configurable thresholds, and software reports, and is now the workhorse of commercial pharma distribution. Generation three — reusable, cellular, real-time IoT loggers with cloud platforms — is the current frontier: LTE-M/NB-IoT low-power WAN makes always-connected multi-year sensing feasible, cloud analytics turn raw temperature traces into excursion insight, and fleets are managed as connected assets. The SCM sits squarely in generation three but pushes several parameters to their current limits simultaneously: 90-day endurance at −30 °C stresses low-temperature battery state of the art (standard Li-ion loses usable capacity and cannot charge below 0 °C — concept D3); in-container operation stresses RF propagation through insulated/metallized walls (concept open item §7); and multi-market data integrity stresses the Part-11/GDP state of practice. Emerging trends — predictive/AI excursion analytics, blockchain-style immutable custody records, and warehouse/pack-out automation — are shaping what buyers will soon consider standard. The requirements below translate this technology baseline into market expectations.
5. Requirement Categories
1. Logging technology baseline — accuracy, memory, reporting expectations.
2. Real-time IoT maturity — always-connected architecture expectations.
3. Connectivity SOTA — LTE-M/NB-IoT, multi-bearer, GPS.
4. Low-temperature power SOTA — cold-capable rechargeable energy + duty-cycling.
5. Cloud / analytics / AI — insight, prediction, dashboards.
6. WHO PQS baseline tech — public-sector device performance norms.
7. Indicator / VVM SOTA — passive substitute technology.
8. Data integrity / immutability — Part 11, tamper-evidence, blockchain trends.
9. Automation trends — provisioning, fleet, integration automation.
6. Detailed Market Requirements
6.1 Logging Technology Baseline
MR-SOTA-001 — Deliver Accuracy And Record Depth Equal To Modern Electronic Loggers
Requirement Statement | The market demands temperature-measurement accuracy, sampling depth, and non-volatile record integrity at least equal to current electronic loggers, since second-generation loggers have made calibrated, high-resolution digital records the accepted norm.
Business Objective | Match the digital-record quality the market already treats as standard.
Customer Segment | Distributors, QA/regulatory, all record-consuming stakeholders.
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Modern loggers set expectations for calibrated accuracy and complete, gap-free traces; anything less reads as a regression.
Business Value | Credible, defensible records; disposition confidence.
Revenue Impact | Baseline requirement to be considered at all.
Risk if Unmet | Records distrusted; product dismissed as sub-standard.
Evidence Source | Standard + reported — WHO PQS accuracy classes / logger norms (specific tolerances Evidence Pending).
Supporting Research | Electronic-logger accuracy/calibration practice (industry-known).
VOC Reference | VOC-### (pending) — theme "accurate as my current logger."
Strategic Theme | Measurement credibility.
Product Family | SCM — temperature (int + ext probe) sensing & storage.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-PERF-0xx; ORCH-018 / EM-ENG-012.
Notes | Accuracy class targets to be set against PQS in regulatory stage.
6.2 Real-Time IoT Maturity
MR-SOTA-002 — Embody Always-Connected, Store-And-Forward IoT Architecture
Requirement Statement | The market demands the always-connected, store-and-forward reliability that reusable IoT loggers have made state of the art — buffering offline and uploading opportunistically so no data is lost across connectivity gaps.
Business Objective | Meet the reliability expectation set by mature real-time platforms.
Customer Segment | Distributors, control towers, 3PLs, fleet operators.
Geographic Applicability | Global; critical on intermittent-coverage lanes.
Market Priority | Must-have
Importance | High
Competitive Driver | Gen-three trackers have normalized "never lose data, sync when you can"; gaps read as unreliability.
Business Value | Complete records + live visibility despite coverage gaps.
Revenue Impact | Enables the real-time value proposition end-to-end.
Risk if Unmet | Data gaps undermine both alerts and audit records.
Evidence Source | reported + Standard — based on concept D2 (store-and-forward).
Supporting Research | Store-and-forward as IoT-logger norm (industry-known).
VOC Reference | VOC-### (pending) — theme "never loses my data."
Strategic Theme | Reliable connectivity.
Product Family | SCM — buffering, cadence upload, opportunistic sync.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Directly instantiates concept D2 as a market expectation.
6.3 Connectivity State Of The Art
MR-SOTA-003 — Adopt LTE-M/NB-IoT-Class Low-Power WAN With Multi-Bearer Fallback
Requirement Statement | The market demands connectivity built on the current low-power WAN state of the art (LTE-M primary, with Wi-Fi/BLE offload and GPS), because LTE-M/NB-IoT have become the expected bearer for long-life connected sensing with global-ish roaming.
Business Objective | Achieve long-life connectivity that works across international lanes.
Customer Segment | Global distributors, 3PLs, program logistics.
Geographic Applicability | Global; coverage varies by region/carrier.
Market Priority | Must-have
Importance | High
Competitive Driver | LTE-M/NB-IoT are now the default low-power cellular bearers for this class; multi-bearer fallback is increasingly expected for coverage resilience.
Business Value | Reliable low-power connectivity and location across markets.
Risk if Unmet | Coverage gaps or high power draw undermine endurance and alerts.
Revenue Impact | Enables the multi-market (D4) reach central to the value proposition.
Evidence Source | Standard + reported — 3GPP LTE-M/NB-IoT; concept connectivity set (D4).
Supporting Research | LTE-M/NB-IoT adoption for logistics IoT (industry-known); roaming/coverage specifics Evidence Pending.
VOC Reference | VOC-### (pending) — theme "works wherever it ships."
Strategic Theme | Resilient global connectivity.
Product Family | SCM — LTE-M/Wi-Fi/BLE + GPS.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-0xx / PRD-ELEC-0xx; ORCH-018 / EM-ENG-012.
Notes | In-container RF attenuation through metallized walls is an open concept item (§7) — antenna performance to be validated.
6.4 Low-Temperature Power State Of The Art
MR-SOTA-004 — Meet Endurance At Cold Using Current Low-Temp Battery + Duty-Cycling SOTA
Requirement Statement | The market demands multi-week endurance in genuine cold using the current state of the art in low-temperature-capable rechargeable chemistry and aggressive duty-cycling, recognizing that standard Li-ion loses usable capacity and cannot be charged below 0 °C.
Business Objective | Deliver the endurance/cold combination that defines the SCM's technical wedge.
Customer Segment | All segments shipping frozen/ultra-cold or long-dwell vaccine.
Geographic Applicability | Global cold and frozen chains.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Few reusable devices credibly combine long endurance with true low-temperature operation; this is where the SCM must exceed, not just match, the field.
Business Value | Enables the −30 °C-to-+60 °C, 90-day claim that competitors struggle to make.
Revenue Impact | Core differentiator; also the top technical risk to the business case.
Risk if Unmet | Endurance/charge-safety shortfalls at cold invalidate the headline claim.
Evidence Source | Judgment/Assumption + Standard — based on concept D3 (KEY RISK); UN 38.3 bounds chemistry; low-temp capacity/charge behaviour Evidence Pending (validate by test).
Supporting Research | Low-temp Li-ion / LTO-class behaviour (industry-known; datasheets Evidence Pending).
VOC Reference | VOC-### (pending) — theme "lasts the whole frozen trip."
Strategic Theme | Cold-capable endurance.
Product Family | SCM — low-temp battery, BMS, charge-inhibit <0 °C, duty-cycling.
Lifecycle Stage | Concept
Traceability | forward to UN-JTBD-007 / PRD-ELEC/PERF-0xx; HA-###; power-budget ADR; ORCH-018 / EM-ENG-012.
Notes | This is the concept's identified top risk (D3) — carried to power budget, reliability, hazards.
6.5 Cloud / Analytics / AI
MR-SOTA-005 — Provide Cloud Analytics And Emerging AI-Grade Excursion Insight
Requirement Statement | The market increasingly demands that raw temperature data be turned into insight — excursion characterization, mean-kinetic-temperature-style analysis, and emerging predictive/AI early-warning — because cloud analytics are now the expected layer on top of connected sensing.
Business Objective | Move from data capture to decision support and predictive value.
Customer Segment | QA/regulatory, logistics analytics, program M&E, control towers.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Platform players and startups increasingly compete on analytics and prediction; a raw-trace-only device looks dated.
Business Value | Better, faster disposition decisions; predictive loss avoidance.
Revenue Impact | Supports platform/subscription revenue and differentiation.
Risk if Unmet | Perceived as a commodity logger; loses analytics-forward buyers.
Evidence Source | reported + Judgment/Assumption — analytics/AI adoption trend (industry-known).
Supporting Research | Cloud dashboards, MKT analysis, predictive-excursion trend (industry-known).
VOC Reference | VOC-### (pending) — theme "tell me what the data means."
Strategic Theme | Insight over raw data.
Product Family | SCM — cloud dashboard + analytics (OTA-extensible).
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-CLD-0xx; MR-COMP-008; ORCH-018 / EM-ENG-012.
Notes | AI claims must be evidence-backed and validated; avoid over-claiming predictive accuracy.
6.6 WHO PQS Baseline Technology
MR-SOTA-006 — Meet Or Exceed The WHO PQS Device Performance Baseline
Requirement Statement | The market demands performance at or above the WHO PQS baseline for cold-chain monitoring devices (accuracy, freeze/heat detection, robustness, battery life expectations) that defines what a credible vaccine-monitoring device is in the public sector.
Business Objective | Be recognized as a legitimate vaccine-grade monitoring technology.
Customer Segment | Immunization programs, ministries, UNICEF/Gavi-aligned procurers.
Geographic Applicability | Global public sector; LMIC-weighted.
Market Priority | High
Importance | Critical
Competitive Driver | PQS defines the technical floor for vaccine monitoring; devices below it are not credible in this domain.
Business Value | Vaccine-grade credibility and procurement eligibility.
Revenue Impact | Gate to public-sector volume (sizing Evidence Pending).
Risk if Unmet | Not taken seriously as a vaccine device; confined to commercial lanes.
Evidence Source | Standard — WHO PQS device specs (exact category/thresholds Evidence Pending).
Supporting Research | PQS performance-spec role (Standard/industry-known).
VOC Reference | VOC-### (pending) — theme "vaccine-grade, not generic."
Strategic Theme | Vaccine-grade baseline.
Product Family | SCM — performance vs PQS.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-PERF-0xx; MR-COMP-005; regulatory stage; ORCH-018 / EM-ENG-012.
Notes | Exact PQS thresholds are an open concept item — do not assert specific numbers.
6.7 Indicator / VVM State Of The Art
MR-SOTA-007 — Respect Mature VVM/Indicator Tech As Complement, Not Ignore It
Requirement Statement | The market demands that a connected monitor acknowledge the mature, trusted state of the art in VVMs and freeze/threshold indicators — complementing vial-level potency/freeze proxies rather than pretending to replace them — so the SCM fits how vaccines are actually judged at the edge.
Business Objective | Position the SCM at the shipment/container layer above the vial-level indicator layer.
Customer Segment | Immunization programs, last-mile workers, QA.
Geographic Applicability | Global; LMIC-weighted.
Market Priority | Medium
Importance | Medium
Competitive Driver | VVMs are entrenched, WHO-PQS-credentialed, and near-free; competing head-on at the vial level is futile — complementary positioning is the mature technical stance.
Business Value | Coherent layered story; avoids a losing battle against near-free indicators.
Revenue Impact | Protects last-mile credibility and adoption path.
Risk if Unmet | Overreach vs VVMs alienates the immunization community.
Evidence Source | reported + Judgment/Assumption — VVM/indicator maturity (industry-known).
Supporting Research | VVM chemistry maturity and PQS status (industry-known; codes Evidence Pending).
VOC Reference | VOC-### (pending) — theme "works with our VVMs."
Strategic Theme | Layered complementarity.
Product Family | SCM — shipment/container-level monitoring.
Lifecycle Stage | Growth
Traceability | forward to UN-JTBD-004 / MR-POS-###; ORCH-018 / EM-ENG-012.
Notes | Aligns with MR-COMP-004 (last-mile) and MR-POS positioning.
6.8 Data Integrity / Immutability
MR-SOTA-008 — Match The State Of The Art In Tamper-Evident, Part-11-Grade Data Integrity
Requirement Statement | The market demands data integrity at the current state of the art — Part-11-grade audit trails, secure time-stamping, and tamper-evidence, with immutable/blockchain-style custody records an emerging expectation for high-value/regulated shipments.
Business Objective | Make the cold-chain record legally and technically unimpeachable.
Customer Segment | QA/regulatory, enterprise pharma, donors requiring accountability.
Geographic Applicability | Global (EU GDP, US FDA especially).
Market Priority | High
Importance | High
Competitive Driver | Data integrity is the current battleground for regulated buyers; immutable custody is an emerging differentiator among platform players.
Business Value | Defensible records; access to regulated/high-value shipments.
Revenue Impact | Gate to regulated enterprise revenue; supports premium.
Risk if Unmet | Records challengeable; excluded from regulated use.
Evidence Source | Standard + reported — FDA 21 CFR Part 11; EU GDP; blockchain-custody trend (industry-known).
Supporting Research | Part 11 audit-trail practice; immutable-ledger pilots in pharma supply chain (industry-known).
VOC Reference | VOC-### (pending) — theme "record no one can dispute."
Strategic Theme | Unimpeachable records.
Product Family | SCM — data integrity, secure time, audit trail.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-CLD-0xx; cybersecurity stage; ORCH-018 / EM-ENG-012.
Notes | Blockchain framed as emerging option, not a requirement to over-engineer at launch.
6.9 Automation Trends
MR-SOTA-009 — Embrace Automated Provisioning, Fleet, And Integration Workflows
Requirement Statement | The market increasingly demands automation of the surrounding workflow — bulk/zero-touch provisioning, automated fleet health and calibration tracking, and API-driven integration — because manual per-device handling does not scale to fleet operations that are now state of the art.
Business Objective | Make the SCM operable at fleet scale with minimal manual effort.
Customer Segment | Fleet/asset admins, distributor/3PL operations, enterprise IT.
Geographic Applicability | Global enterprise networks.
Market Priority | Medium
Importance | Medium
Competitive Driver | Leading platforms automate provisioning, fleet health, and integration; manual workflows are a competitive liability at scale.
Business Value | Lower operating cost; scalable deployments; stickier accounts.
Revenue Impact | Enables large-fleet expansion and subscription value.
Risk if Unmet | Operational friction caps deployment size and retention.
Evidence Source | reported + Judgment/Assumption — automation trend in IoT fleet ops (industry-known).
Supporting Research | Zero-touch provisioning and fleet-automation practice (industry-known).
VOC Reference | VOC-### (pending) — theme "runs itself at scale."
Strategic Theme | Operational automation.
Product Family | SCM — provisioning, fleet management, APIs, OTA.
Lifecycle Stage | Sustaining
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; MR-COMP-006; ORCH-018 / EM-ENG-012.
Notes | Ties to MR-JTBD-006/010 and MR-COMP-006 (integration).
7. Evidence & Validation Notes
Evidence posture | Technology trends are reported/industry-known; standards references are Standard; the cold-endurance items are Judgment/Assumption anchored in concept D3 and require test validation. No specific accuracy tolerances, battery capacities at −30 °C, PQS thresholds, or coverage figures are asserted as fact — all Evidence Pending.
Validation method | Technology scan against 3GPP/LTE-M documentation, low-temp battery datasheets, WHO PQS specs, and vendor technical literature; bench and environmental testing (IEC 60068) for the cold-endurance and in-container RF open items.
Key open items | low-temp battery endurance/charge behaviour at −30 °C (D3, top risk); antenna/RF performance inside metallized containers (§7); exact PQS thresholds; realistic AI-prediction accuracy.
8. Traceability Placeholders
Upstream | 00-concept/concept-and-decisions.md (D1–D4); MR-400 (JTBD); MR-500 (COMP).
Peer modules | MR-700 (POS); MR-100/200/300 (pending).
Downstream (forward) | VOC-### → STN-### → UN-JTBD/STK-### → PRD-FUN/PERF/ELEC/CLD-### → SYS-### → HA-### / FA-### → PV-###/VAL-###; power-budget ADR.
Governance | ORCH-018 / EM-ENG-012 maintain technology-claim traceability and evidence status.
MR-700 — Product Positioning (POS)
Traceability: ORCH-009 (Market Requirements) · ORCH-018 / EM-ENG-012 · based on 00-concept/concept-and-decisions.md (D1–D4). Module code: POS.
1. Purpose
This module defines how the Smart Cold Chain Monitor (SCM) must be positioned in the market — its unique value proposition, differentiators, reusable-vs-single-use stance, premium-vs-value posture, feature and price positioning, brand/trust posture, and go-to-market assumptions. It expresses positioning as market requirements: what the market must be able to understand, believe, and pay for, given the jobs (MR-400), the competitive field (MR-500), and the technology baseline (MR-600).
2. Scope
In scope: the demand-side positioning requirements — value proposition, differentiation, segment/price posture, brand and trust, and go-to-market assumptions across public-sector immunization and commercial pharma distribution. Out of scope: detailed pricing models, channel contracts, and campaign execution (commercial planning), and internal design (Product Requirements). Specific price points, market sizes, and share targets are Evidence Pending.
3. Applicable Standards & References
• WHO PQS — credibility credential shaping public-sector positioning (codes Evidence Pending).
• EU GDP / FDA 21 CFR Part 11 — compliance credibility underpinning trust positioning.
• UN 38.3 / IATA DGR — reusable/reverse-logistics positioning constraints (battery transport).
• MR-400 (JTBD), MR-500 (COMP), MR-600 (SOTA) — the demand, competitive, and technology inputs to positioning.
• Market/economic figures cited as Evidence Pending unless sourced from WHO/UNICEF/Gavi or named analysts.
4. Background
Positioning must resolve a genuine tension. On one axis, the market is anchored by cheap, trusted single-use loggers and near-free VVMs/indicators; on the other, by entrenched enterprise real-time platforms (e.g., Controlant-class). The SCM's defensible ground, established in MR-500, is a white-space intersection no incumbent fully owns: reusable + in-container + −30 °C-to-+60 °C + 90-day + multi-market + immediate excursion alerts (concept D1–D4). Positioning therefore cannot be "cheapest logger" (loses to disposables and VVMs) nor "generic IoT tracker" (loses to entrenched platforms). It must be the vaccine-grade, cold-capable, reusable assurance platform — a device credible enough for WHO-PQS public-sector immunization and GDP/Part-11 commercial pharma, that pays for itself through reuse and prevents loss through real-time alerting. The requirements below specify what the market must believe for that position to hold, and flag every economic assumption as Evidence Pending. The device is consistently positioned as a monitoring/information product that supports disposition decisions — not a temperature- controlling or therapeutic product (concept intended-use), to keep claims truthful and regulatorily safe.
5. Requirement Categories
1. Unique value proposition — the one-line reason to choose the SCM.
2. Differentiators — the defensible, communicable points of difference.
3. Reusable-vs-single-use positioning — the model choice narrative.
4. Premium-vs-value posture — where on the value curve the SCM sits.
5. Feature positioning — which capabilities lead the story.
6. Price positioning — how price is framed (TCO/outcome, not unit).
7. Brand / trust positioning — vaccine-grade credibility and integrity.
8. Go-to-market assumptions — channels, segments, entry sequence.
6. Detailed Market Requirements
6.1 Unique Value Proposition
MR-POS-001 — Own "Vaccine-Grade, Cold-Capable, Reusable Cold-Chain Assurance"
Requirement Statement | The market demands a single, unmistakable value proposition — that the SCM is the reusable, vaccine-grade monitor that proves vaccines stayed in range end-to-end, even in deep cold, and warns in time to save the shipment.
Business Objective | Give buyers one clear reason to choose the SCM over loggers, platforms, and indicators.
Customer Segment | All target segments (public-sector immunization + commercial pharma).
Geographic Applicability | Global.
Market Priority | Must-have
Importance | Critical
Competitive Driver | No incumbent occupies the reusable + cold-capable + vaccine-grade + real-time intersection; a crisp UVP claims it before fast-followers blur it.
Business Value | Clear, memorable market identity anchoring all messaging and pricing.
Revenue Impact | Foundational to conversion and premium; enables the whole commercial story.
Risk if Unmet | Positioned as "just another tracker/logger"; commoditized.
Evidence Source | Judgment/Assumption — synthesis of MR-400/500/600; confirm resonance via VOC.
Supporting Research | JTBD core job (MR-JTBD-001) + white-space (MR-COMP-009).
VOC Reference | VOC-### (pending) — theme "the one reason to buy."
Strategic Theme | Own the white-space.
Product Family | SCM — whole-product value proposition.
Lifecycle Stage | Concept
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; MR-COMP-009; ORCH-018 / EM-ENG-012.
Notes | Keep claims to monitoring/assurance — never imply the device preserves or treats vaccine.
6.2 Differentiators
MR-POS-002 — Lead With The Cold-Capable, In-Container, Multi-Market Combination
Requirement Statement | The market demands that the SCM's differentiation be communicated as the combination — deep-cold capability, in-container operation, 90-day endurance, and multi-market compliance in one reusable device — since each element alone is matched somewhere, but the combination is not.
Business Objective | Make the differentiator hard to copy and easy to grasp.
Customer Segment | Buyers evaluating against loggers and platforms; partners; investors.
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Point-feature comparisons favor specialists; the combination is the SCM's defensible, communicable edge.
Business Value | Durable differentiation resistant to single-feature undercutting.
Revenue Impact | Sustains premium and reduces feature-price erosion.
Risk if Unmet | Reduced to a feature checklist where specialists win each row.
Evidence Source | Judgment/Assumption — based on concept D1–D4 and MR-COMP-009.
Supporting Research | Competitive segmentation (MR-500); technology limits (MR-600).
VOC Reference | VOC-### (pending) — theme "why not just a cheaper logger?"
Strategic Theme | Combination advantage.
Product Family | SCM — integrated whole product.
Lifecycle Stage | Launch
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; MR-COMP-009; ORCH-018 / EM-ENG-012.
Notes | Cold-endurance claim depends on the D3 top risk being validated — don't message ahead of proof.
6.3 Reusable-vs-Single-Use Positioning
MR-POS-003 — Position Reuse As Both Economic And Sustainability Advantage
Requirement Statement | The market demands a clear reusable-vs-disposable narrative that frames reuse as both a total-cost-of-ownership win and a sustainability/waste-reduction win over single-use loggers — while honestly addressing the reverse-logistics effort reuse requires.
Business Objective | Convert the reusable model from a cost objection into a purchase reason.
Customer Segment | Procurement/finance, sustainability leads, operations.
Geographic Applicability | Global; nuanced on high-loss LMIC last-mile.
Market Priority | High
Importance | High
Competitive Driver | Disposables win on unit price and no returns; the reusable story must reframe the comparison to TCO + sustainability to win.
Business Value | Justifies the model against the disposable default.
Revenue Impact | Underpins the capital/subscription business case.
Risk if Unmet | Buyers anchor on unit price and choose disposables.
Evidence Source | Judgment/Assumption — TCO and return-rate inputs Evidence Pending; based on D1.
Supporting Research | MR-COMP-007 (TCO) and MR-JTBD-009 (reverse logistics).
VOC Reference | VOC-### (pending) — theme "reuse vs disposable."
Strategic Theme | Reuse as advantage.
Product Family | SCM — reusable device + reverse-logistics program.
Lifecycle Stage | Growth
Traceability | forward to UN-JTBD-009 / PRD-MECH-0xx; MR-COMP-007; ORCH-018 / EM-ENG-012.
Notes | Sustainability claims must be substantiated (avoid greenwashing); quantify at VOC/commercial stage.
6.4 Premium-vs-Value Posture
MR-POS-004 — Position As Premium-Justified-By-Outcome, Not Cheapest-Per-Unit
Requirement Statement | The market demands that the SCM be positioned as a premium, outcome-justified device — priced against the value of vaccine saved and audits passed — rather than competing on lowest unit cost, which cedes ground to disposables and VVMs.
Business Objective | Anchor value on outcomes (loss avoided, compliance) not hardware unit price.
Customer Segment | Enterprise pharma, high-value-shipment owners, program QA.
Geographic Applicability | Global; adapt framing for price-sensitive public sector.
Market Priority | High
Importance | High
Competitive Driver | Competing on unit price is unwinnable against near-free indicators/disposables; outcome-based premium is defensible.
Business Value | Protects margin and signals quality/credibility.
Revenue Impact | Supports higher ASP and healthier unit economics (targets Evidence Pending).
Risk if Unmet | Margin erosion; race to the bottom against disposables.
Evidence Source | Judgment/Assumption — willingness-to-pay Evidence Pending (VOC).
Supporting Research | Value-of-vaccine-saved framing (Evidence Pending quantification).
VOC Reference | VOC-### (pending) — theme "worth the premium?"
Strategic Theme | Outcome-based premium.
Product Family | SCM — premium positioning tier.
Lifecycle Stage | Launch
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; MR-POS-006; ORCH-018 / EM-ENG-012.
Notes | Public-sector may need a value-tier / program-priced variant — flag for GTM.
6.5 Feature Positioning
MR-POS-005 — Lead With Assurance And Alerting, Not A Sensor Spec Sheet
Requirement Statement | The market demands that positioning lead with the outcomes buyers care about — provable end-to-end assurance and timely excursion alerting — with sensors, connectivity, and battery positioned as the means, not the headline.
Business Objective | Sell the job done (MR-400), not the component list.
Customer Segment | Decision-makers (QA, program, operations) vs technical evaluators.
Geographic Applicability | Global.
Market Priority | Medium
Importance | High
Competitive Driver | Spec-sheet wars favor commoditized loggers; outcome-led framing differentiates and resists comparison-shopping.
Business Value | Higher-value conversations; less feature-price haggling.
Revenue Impact | Supports premium and shortens the "why" of the sale.
Risk if Unmet | Buyers compare on specs and default to cheaper equivalents.
Evidence Source | Judgment/Assumption — based on JTBD outcomes (MR-400).
Supporting Research | MR-JTBD-001/005 (assurance + alerting as core jobs).
VOC Reference | VOC-### (pending) — theme "what problem does it solve."
Strategic Theme | Outcome-led messaging.
Product Family | SCM — messaging hierarchy.
Lifecycle Stage | Launch
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; ORCH-018 / EM-ENG-012.
Notes | Provide a spec sheet for technical buyers, but keep it beneath the outcome story.
6.6 Price Positioning
MR-POS-006 — Frame Price As Cost-Per-Trip / Outcome, Enabling Flexible Commercial Models
Requirement Statement | The market demands that price be framed as cost-per-trip or cost-per-outcome (with device-as-a-service / subscription options), so buyers compare against per-trip disposable cost and vaccine value, not against a one-time hardware sticker price.
Business Objective | Make the SCM's price legible and favorable in the buyer's own terms.
Customer Segment | Procurement/finance, program budgeting, enterprise sourcing.
Geographic Applicability | Global; model mix varies public vs commercial.
Market Priority | High
Importance | High
Competitive Driver | Reusable platforms increasingly sell as-a-service; per-trip framing directly counters the disposable unit-price anchor.
Business Value | Lowers adoption barrier; enables recurring revenue.
Revenue Impact | Enables subscription/recurring revenue streams (pricing Evidence Pending).
Risk if Unmet | Sticker shock at hardware price stalls adoption.
Evidence Source | Judgment/Assumption — pricing/commercial-model inputs Evidence Pending.
Supporting Research | As-a-service trend among reusable trackers (industry-known); MR-COMP-007.
VOC Reference | VOC-### (pending) — theme "how do we buy this?"
Strategic Theme | Legible, flexible pricing.
Product Family | SCM — hardware + subscription/service.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; MR-POS-004; ORCH-018 / EM-ENG-012.
Notes | Connectivity (LTE-M data) and cloud are recurring-cost drivers — reflect in model.
6.7 Brand / Trust Positioning
MR-POS-007 — Build A Vaccine-Grade Trust Brand On Credentials And Data Integrity
Requirement Statement | The market demands that the SCM brand stand for vaccine-grade trust — evidenced by WHO-PQS/GDP/Part-11 credentials and unimpeachable data integrity — because in cold-chain assurance, trust is the product and credentials are its proof.
Business Objective | Make trust the brand's defining, defensible attribute.
Customer Segment | QA/regulatory, programs, donors, enterprise pharma.
Geographic Applicability | Global.
Market Priority | High
Importance | Critical
Competitive Driver | Buyers stake vaccine value and compliance on the record; credential-backed trust beats feature novelty in this domain.
Business Value | Preference, retention, and premium rooted in credibility.
Revenue Impact | Enables access to regulated/public-sector revenue and premium pricing.
Risk if Unmet | Any trust failure (bad record, security lapse) is existential in this category.
Evidence Source | Standard + Judgment/Assumption — WHO PQS / GDP / Part 11 credentials (records Evidence Pending).
Supporting Research | MR-COMP-002/005, MR-SOTA-006/008 (compliance + integrity).
VOC Reference | VOC-### (pending) — theme "can we trust it with our vaccines?"
Strategic Theme | Trust as brand.
Product Family | SCM — brand + compliance credentials.
Lifecycle Stage | Launch
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; regulatory + cybersecurity stages; ORCH-018 / EM-ENG-012.
Notes | Do not claim credentials before they are earned — mark PQS/FDA status Evidence Pending until granted.
6.8 Go-To-Market Assumptions
MR-POS-008 — Sequence Entry Via Commercial Pharma, Then Scale Into Public-Sector Immunization
Requirement Statement | The market demands a go-to-market posture that assumes early traction in commercial pharma/enterprise distribution (faster procurement, willingness to pay for real-time), then scales into WHO-PQS-gated public-sector immunization once credentials and unit economics mature.
Business Objective | Enter where value is realized fastest, then broaden to the largest volume channel.
Customer Segment | Commercial pharma/3PLs first; ministries/UNICEF/Gavi-aligned procurers next.
Geographic Applicability | Global, sequenced by channel readiness.
Market Priority | High
Importance | High
Competitive Driver | PQS qualification and public-sector procurement cycles are slow; commercial lanes reward real-time value sooner and fund the path to PQS.
Business Value | Faster revenue, credential-building, and economics before the harder channel.
Revenue Impact | Shapes revenue timing and investment sequencing (figures Evidence Pending).
Risk if Unmet | Chasing slow PQS procurement first starves the business of early revenue.
Evidence Source | Judgment/Assumption — channel-dynamics read; validate with commercial + VOC.
Supporting Research | MR-COMP-005 (PQS gating), MR-POS-004/006 (premium/pricing).
VOC Reference | VOC-### (pending) — theme "who buys first."
Strategic Theme | Sequenced go-to-market.
Product Family | SCM — commercial-first, public-sector-scale.
Lifecycle Stage | Launch
Traceability | forward to UN-VIS-0xx / PRD-FUN-0xx; MR-100/200 (EXEC/OPP, pending); ORCH-018 / EM-ENG-012.
Notes | A value-tier variant may be needed for public sector (see MR-POS-004) — carry to product-family strategy.
MR-POS-009 — Position As Partner-Friendly, Not A Rip-And-Replace, To Ride Existing Ecosystems
Requirement Statement | The market demands that the SCM be positioned as complementary and integration-friendly toward existing platforms, indicators, and logistics stacks, so buyers and partners see it as additive value rather than a disruptive replacement they must fight to justify.
Business Objective | Lower adoption friction and unlock channel/partner-led growth.
Customer Segment | Enterprise pharma, 3PLs, platform partners, immunization programs.
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Rip-and-replace positioning invites resistance from incumbents and IT; additive positioning rides existing adoption (see MR-COMP-006, MR-SOTA-007).
Business Value | Faster adoption, partner leverage, reduced competitive friction.
Revenue Impact | Enables channel/partner revenue and account expansion.
Risk if Unmet | Framed as a threat; blocked by incumbents and integration inertia.
Evidence Source | Judgment/Assumption — based on MR-COMP-006 and MR-SOTA-007.
Supporting Research | Integration + complementarity requirements (MR-500/600).
VOC Reference | VOC-### (pending) — theme "fits what we already run."
Strategic Theme | Additive, partner-friendly.
Product Family | SCM — interoperable positioning.
Lifecycle Stage | Growth
Traceability | forward to UN-STK-0xx / PRD-FUN-0xx; MR-COMP-006; ORCH-018 / EM-ENG-012.
Notes | Complements vial-level VVMs (MR-SOTA-007) and existing platforms (MR-COMP-006).
7. Evidence & Validation Notes
Evidence posture | Positioning is a synthesis of MR-400/500/600 and is therefore largely Judgment/Assumption, explicitly flagged. Credential-based trust claims are Standard in intent but the actual records (WHO PQS, FDA, GDP certification) are Evidence Pending until granted. No price points, market sizes, willingness-to-pay, or share targets are asserted as fact.
Validation method | Test value proposition, differentiators, and price framing through VOC (message testing, willingness-to-pay research, segment interviews) and early commercial pilots; validate cold-endurance-dependent claims only after the D3 risk is closed by test.
Truthfulness guardrail | All positioning keeps the SCM a monitoring/information device (concept intended-use) — no claim implies it preserves, controls, or treats vaccine.
Key open items | willingness-to-pay and TCO inputs; PQS/FDA credential status; public-sector value-tier need; partner/channel structure.
8. Traceability Placeholders
Upstream | 00-concept/concept-and-decisions.md (D1–D4); MR-400 (JTBD), MR-500 (COMP), MR-600 (SOTA).
Peer/adjacent modules | MR-100/200/300 (EXEC/OPP/SEG — pending) feed and consume this positioning.
Downstream (forward) | VOC-### → STN-### → UN-VIS/STK-### → PRD-FUN/PERF-### → SYS-### → PV-###/VAL-###; commercial/GTM planning.
Governance | ORCH-018 / EM-ENG-012 maintain positioning-claim traceability and evidence status.
MR-800 — Commercial Requirements (COMM)
Traceability note: This module (code COMM) forms part of the Smart Cold Chain Monitor (SCM) Master Market Requirements Library. Each requirement traces forward to User Needs (UN-…) and Product Requirements (PRD-…) placeholders and is governed under ORCH-018 (Traceability) with market-requirements authorship under ORCH-009 / EM-ENG-012.
1. Purpose
This module captures the commercial market requirements for the Smart Cold Chain Monitor (SCM): the pricing, cost, margin, subscription, service, and business-model demands that the global vaccine cold-chain market places on a reusable, battery-operated, in-container temperature/humidity/shock monitor. It expresses what the market will pay for and how it expects to buy, own, and be served — deliberately separated from how the product is engineered.
The purpose is to give downstream engineering, product, and go-to-market teams a evidence-backed commercial envelope so that device cost targets, the SaaS/data-plan model, the reverse-logistics/pooling economics, and the accessory and service ecosystem are all aligned to demonstrable buyer behaviour in the cold-chain sector (WHO PQS procurement, EU GDP-regulated distributors, US pharma 3PLs, and donor-funded LMIC programs).
2. Scope
In scope: target selling-price positioning; device unit cost and gross-margin targets; recurring SaaS/data-plan subscription; reverse-logistics and device-pooling business model; accessory ecosystem (external probes, mounts, cradles, charging); warranty; calibration-as-a-service; customer support model; aftermarket/service revenue; and volume/tender pricing for institutional buyers (e.g., GAVI, UNICEF Supply Division tenders).
Out of scope: detailed engineering implementation (owned by the System and Subsystem Requirements modules), regulatory approval mechanics (see MR-1100 REG), success KPIs and targets (see MR-1000 MET), and internal budget/capability constraints (see MR-900 CON). Commercial requirements here state market demand; the business's ability to meet them is captured as constraints in MR-900.
3. Applicable Standards & References
• WHO PQS (Performance, Quality and Safety) prequalification scheme — the procurement gate for cold-chain equipment bought by WHO/UNICEF/GAVI-funded programs; commercial acceptance in donor markets is effectively conditioned on PQS listing (exact device category code TBC; see MR-1100 REG).
• EU Good Distribution Practice (GDP) — Guidelines 2013/C 343/01 — frames the commercial expectations of EU-regulated pharma distributors and 3PLs who are the paying customers.
• US FDA 21 CFR Part 11 (electronic records/signatures) and FDA data-integrity guidance — shapes what US pharma customers require of a monitoring/records product before they will buy.
• UN 38.3 and IATA Dangerous Goods Regulations (Li battery provisions) — govern the transportability of the reusable device and therefore its reverse-logistics cost structure.
• UNICEF Supply Division / GAVI cold-chain equipment procurement practice — the institutional tender mechanism through which large LMIC volumes are bought (referenced as market practice; specific tender terms Evidence Pending).
• Commercial figures (prices, unit costs, margins, subscription levels) in this module are Judgment/Assumption unless explicitly cited; they define target envelopes to be validated by Voice-of-Customer (VOC) and pricing research, not asserted facts.
4. Background
The vaccine cold chain loses product to temperature excursions across the last mile, particularly in LMIC distribution and in air/ground transport where shipments leave the controlled environment of a facility. Buyers span three commercially distinct groups: (1) donor-funded public health programs (WHO/UNICEF/GAVI and national EPI programs) that buy at scale through tenders and are highly price- and total-cost-sensitive; (2) EU/US regulated pharma distributors and 3PLs who buy on compliance, data integrity, and operational reliability and will pay premium for GDP/Part-11 defensibility; and (3) commercial pharma manufacturers shipping high-value biologics who value loss avoidance per shipment.
The SCM's defining commercial characteristic is reusability. A reusable, rechargeable, multi-trip device changes the business model from a per-shipment consumable to a durable asset with recurring SaaS/data-plan revenue plus a reverse-logistics obligation. This creates both the opportunity (recurring revenue, pooling/rental) and the cost driver (retrieval, cleaning, recalibration, redeployment). The commercial requirements below reflect that the market increasingly expects a service, not just a device — dashboards, alerts, calibration certificates, and support — sold on subscription.
5. Requirement Categories
1. Pricing & Positioning — target selling price, tender/volume pricing.
2. Cost & Margin — device unit-cost and gross-margin envelope.
3. Recurring Revenue — SaaS/data-plan subscription model.
4. Business Model — reverse-logistics, pooling/rental, device ownership models.
5. Ecosystem & Aftermarket — accessories, warranty, calibration-as-a-service, support, service revenue.
6. Detailed Market Requirements
6.1 Pricing & Positioning
MR-COMM-001 — Reusable Total-Cost-of-Ownership Advantage
Requirement Statement | The market demands that a reusable in-container monitor deliver a demonstrably lower total cost of ownership per monitored trip than single-use/disposable data loggers over the device's service life.
Business Objective | Win switching decisions from incumbent single-use loggers by making the reusable economic case unambiguous.
Customer Segment | Pharma distributors, 3PLs, and donor-funded programs currently using disposable loggers.
Geographic Applicability | Global; strongest where shipment volumes are high and retrieval is feasible (EU, US, regional LMIC hubs).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Single-use loggers have low unit price but high cumulative cost at volume; reusability is the primary differentiator.
Business Value | Enables premium device pricing justified by per-trip economics; anchors the entire value proposition.
Revenue Impact | High — determines device ASP ceiling and the credibility of the recurring model.
Risk if Unmet | Buyers default to cheap disposables; reusable premium cannot be sustained.
Evidence Source | Judgment/Assumption — per-trip TCO advantage is the reusability thesis; magnitude Evidence Pending (requires validated cost model + VOC).
Supporting Research | Reusability anchor fact D1; cold-chain logger market practice. Quantified TCO delta pending pricing study.
VOC Reference | VOC-### (pending) — theme: "disposable loggers are cheap per unit but add up."
Strategic Theme | Reusability economics.
Product Family | SCM (Smart Cold Chain Monitor) — core device + service.
Lifecycle Stage | Concept
Traceability | forward to UN-… (cost-of-ownership need), PRD-… (service-life / trips-per-life target); ORCH-018 / EM-ENG-012.
Notes | TCO model must include retrieval, recalibration, and loss-rate assumptions from MR-900 CON.
MR-COMM-002 — Target Selling-Price Positioning
Requirement Statement | The market demands a device selling price positioned as a durable capital/asset purchase (or embedded in subscription), materially above disposable loggers but justified by multi-trip reuse.
Business Objective | Establish a price band that supports margin while remaining winnable against disposables on TCO.
Customer Segment | All paying segments; most price-elastic in donor/LMIC tenders.
Geographic Applicability | Global, with tender-driven price sensitivity in LMIC markets.
Market Priority | Must-have
Importance | High
Competitive Driver | Reusable multi-trip loggers and pooled-asset trackers set the competitive price envelope.
Business Value | Correct positioning maximises adoption × margin.
Revenue Impact | High — directly sets device revenue line.
Risk if Unmet | Overpricing stalls adoption; underpricing erodes margin and signals low quality.
Evidence Source | Judgment/Assumption — specific price band Evidence Pending (pricing research + competitor teardown).
Supporting Research | Reusable-logger and pooled-tracker price points to be benchmarked. No verified figure asserted.
VOC Reference | VOC-### (pending) — theme: procurement price ceilings.
Strategic Theme | Value-based pricing.
Product Family | SCM — device SKU tiers.
Lifecycle Stage | Concept
Traceability | forward to PRD-… (BOM cost budget), UN-…; ORCH-018 / EM-ENG-012.
Notes | Price band to be expressed as tiers (standard vs extended-probe/GPS) once BOM cost is known.
MR-COMM-003 — Institutional Tender / Volume Pricing (GAVI / UNICEF)
Requirement Statement | The market demands volume-based and tender-compliant pricing structures compatible with GAVI/UNICEF Supply Division and national EPI procurement, including multi-year framework pricing.
Business Objective | Access large donor-funded volumes that require tender-format, transparent, volume-tiered pricing.
Customer Segment | Donor-funded public-health programs; UNICEF Supply Division; national immunization programs.
Geographic Applicability | LMICs and donor-market geographies globally.
Market Priority | High
Importance | High
Competitive Driver | WHO PQS listing plus tender-ready pricing is the gate to institutional volume.
Business Value | Unlocks step-change unit volume and reference credibility.
Revenue Impact | High volume / lower margin — critical to scale.
Risk if Unmet | Excluded from the largest cold-chain buyers; lose scale economics.
Evidence Source | Cited — UNICEF/GAVI cold-chain equipment procured via tender is established market practice; specific terms/prices Evidence Pending.
Supporting Research | UNICEF Supply Division cold-chain procurement model; PQS-conditioned eligibility (see MR-1100 REG).
VOC Reference | VOC-### (pending) — theme: tender eligibility and price ceilings.
Strategic Theme | Institutional access.
Product Family | SCM — tender SKU + support package.
Lifecycle Stage | Growth
Traceability | forward to MR-1100 REG (PQS listing), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Tender participation is contingent on PQS listing (MR-REG-001).
6.2 Cost & Margin
MR-COMM-004 — Device Unit-Cost Target Envelope
Requirement Statement | The market's price sensitivity demands that landed device unit cost sit within an envelope that preserves target gross margin at the achievable selling price across segments.
Business Objective | Constrain BOM + manufacturing + logistics cost so margin targets are met at market-clearing prices.
Customer Segment | Internal commercial (derived from buyer price ceilings across all segments).
Geographic Applicability | Global (manufacturing/logistics cost varies by region).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Disposable-logger price points compress the acceptable cost envelope.
Business Value | Protects profitability and pricing flexibility for tenders.
Revenue Impact | High — governs margin on every unit.
Risk if Unmet | Product is unprofitable at winnable prices, especially in tenders.
Evidence Source | Judgment/Assumption — specific cost target Evidence Pending (BOM cost-down + DFM from MR-900 CON).
Supporting Research | Low-temp Li-ion, LTE-M module, and sensor cost structure to be costed; no figure asserted.
VOC Reference | n/a (internal-derived from pricing VOC).
Strategic Theme | Cost leadership within reusables.
Product Family | SCM — device BOM.
Lifecycle Stage | Concept
Traceability | forward to PRD-… (cost budget), MR-900 CON (supply constraints); ORCH-018 / EM-ENG-012.
Notes | Cost envelope must absorb reverse-logistics and recalibration amortised over trips-per-life.
MR-COMM-005 — Gross-Margin Target
Requirement Statement | The business model demands a blended gross-margin target across device + subscription that sustains the reusable/service model and funds reverse logistics and support.
Business Objective | Ensure the combined hardware-plus-SaaS economics are viable and scalable.
Customer Segment | Internal commercial / investor stakeholders.
Geographic Applicability | Global blended.
Market Priority | Must-have
Importance | Critical
Competitive Driver | SaaS-attached hardware businesses are valued on recurring margin, not one-time device margin.
Business Value | Recurring high-margin SaaS offsets thinner device margin; supports valuation.
Revenue Impact | High — defines overall profitability profile.
Risk if Unmet | Reverse-logistics and support costs erode margin below viability.
Evidence Source | Judgment/Assumption — target margin Evidence Pending (financial model).
Supporting Research | Standard hardware+SaaS margin logic; specific numbers pending.
VOC Reference | n/a (internal).
Strategic Theme | Recurring-revenue economics.
Product Family | SCM — device + SaaS blended.
Lifecycle Stage | Concept
Traceability | forward to MR-1000 MET (margin KPI), MR-900 CON; ORCH-018 / EM-ENG-012.
Notes | Ties directly to MR-MET margin metric.
6.3 Recurring Revenue
MR-COMM-006 — SaaS / Data-Plan Subscription
Requirement Statement | The market demands that ongoing monitoring — cloud dashboard, alerting, reporting, and connectivity — be available as a recurring subscription rather than solely a one-time purchase.
Business Objective | Establish predictable recurring revenue and fund the connectivity + cloud cost base.
Customer Segment | All segments; strongest with distributors/3PLs valuing continuous compliance reporting.
Geographic Applicability | Global (LTE-M/roaming availability dependent — see MR-900 CON).
Market Priority | Must-have
Importance | Critical
Competitive Driver | Cold-chain visibility competitors sell platform subscriptions; the market expects a data service.
Business Value | Recurring, high-margin revenue; increases switching costs and lifetime value.
Revenue Impact | High and recurring — the strategic revenue engine.
Risk if Unmet | Revenue is one-time only; connectivity/cloud costs become a loss; valuation multiple falls.
Evidence Source | Judgment/Assumption — subscription level/tiers Evidence Pending; recurring model based on D2 (cadence upload) and stated business-model implication.
Supporting Research | Anchor fact: cadence upload + alerts + dashboard (D2); reusability implies recurring SaaS/data-plan revenue.
VOC Reference | VOC-### (pending) — theme: "we need continuous visibility, not just a file at the end."
Strategic Theme | Data-as-a-service.
Product Family | SCM — cloud platform + data plan.
Lifecycle Stage | Launch
Traceability | forward to PRD-… (connectivity/cloud), MR-1000 MET (renewal/retention); ORCH-018 / EM-ENG-012.
Notes | Data-plan cost (cellular) is a key COGS input; see MR-900 CON carrier constraints.
6.4 Business Model
MR-COMM-007 — Reverse-Logistics / Device-Pooling Model
Requirement Statement | The market demands a workable reverse-logistics and device-pooling/return model so reusable devices can be retrieved, refurbished, and redeployed without imposing unacceptable handling burden on the buyer.
Business Objective | Make reuse operationally real by owning or enabling the return loop; protect per-trip economics.
Customer Segment | Distributors, 3PLs, and programs unwilling to manage device retrieval themselves.
Geographic Applicability | Global; hardest in dispersed last-mile LMIC routes.
Market Priority | High
Importance | Critical
Competitive Driver | Pooled-asset logistics providers set the expectation that the vendor manages the loop.
Business Value | Enables reusability at scale; can itself be a priced service (pooling/rental).
Revenue Impact | Medium/High — enables device amortisation and optional pooling revenue.
Risk if Unmet | Devices are not returned; reuse economics collapse; effectively a costly disposable.
Evidence Source | Judgment/Assumption — return rates and loop cost Evidence Pending; reverse-logistics obligation is a evidence-backed implication of reusability.
Supporting Research | Stated business-model implication: reusability requires reverse logistics; UN 38.3 / IATA DG affect return shipping cost (MR-1100 REG).
VOC Reference | VOC-### (pending) — theme: "who is responsible for getting the device back?"
Strategic Theme | Reverse-logistics as enabler.
Product Family | SCM — pooling/rental service option.
Lifecycle Stage | Growth
Traceability | forward to MR-1000 MET (utilization/trips-per-year, return rate), MR-900 CON (logistics); ORCH-018 / EM-ENG-012.
Notes | Return-loop economics drive MR-COMM-001 TCO and MR-COMM-005 margin.
MR-COMM-008 — Flexible Ownership Models (Buy vs Rent/Subscribe-to-Asset)
Requirement Statement | The market demands flexible commercial models — outright purchase, rental/leasing, or monitoring-as-a-service — to match diverse buyer capital and procurement constraints.
Business Objective | Remove capital-budget barriers and match donor, distributor, and commercial procurement norms.
Customer Segment | Capital-constrained programs (favour OpEx/rental); large distributors (may favour purchase).
Geographic Applicability | Global; OpEx models attractive in budget-constrained LMIC and 3PL contexts.
Market Priority | Medium
Importance | High
Competitive Driver | Device-as-a-service is an emerging expectation in IoT asset markets.
Business Value | Broadens addressable market; smooths revenue.
Revenue Impact | Medium — mix effect and improved win rate.
Risk if Unmet | Lose buyers whose procurement cannot accommodate the offered model.
Evidence Source | Judgment/Assumption — model mix Evidence Pending.
Supporting Research | IoT device-as-a-service commercial patterns; specific uptake pending VOC.
VOC Reference | VOC-### (pending) — theme: CapEx vs OpEx procurement preference.
Strategic Theme | Commercial flexibility.
Product Family | SCM — commercial packaging.
Lifecycle Stage | Growth
Traceability | forward to PRD-…, MR-1000 MET; ORCH-018 / EM-ENG-012.
Notes | Rental model depends on reverse logistics (MR-COMM-007).
6.5 Ecosystem & Aftermarket
MR-COMM-009 — Accessory Ecosystem (Probes, Mounts, Charging)
Requirement Statement | The market demands a supporting accessory ecosystem — external temperature probes, container mounts/cradles, and charging accessories — to fit diverse container types and use cases.
Business Objective | Increase fit-for-use across shipment types and generate attach/accessory revenue.
Customer Segment | All device buyers; especially those with varied packaging (cartons, cold boxes, pallets).
Geographic Applicability | Global.
Market Priority | Medium
Importance | Medium
Competitive Driver | Configurable probe/mount ecosystems differentiate on fit and reduce install friction.
Business Value | Attach revenue plus improved deployment success and stickiness.
Revenue Impact | Medium — accessory attach revenue.
Risk if Unmet | Poor physical fit reduces adoption; lost attach revenue.
Evidence Source | Judgment/Assumption — attach rates Evidence Pending; external probe capability evidence-backed (external probe in device scope).
Supporting Research | Anchor fact: internal + external temperature probe supported; mounting varies by container.
VOC Reference | VOC-### (pending) — theme: "it has to fit our cold boxes/pallets."
Strategic Theme | Ecosystem attach.
Product Family | SCM — accessories line.
Lifecycle Stage | Launch
Traceability | forward to PRD-… (probe/mount interfaces), MR-1000 MET (attach rate); ORCH-018 / EM-ENG-012.
Notes | External probe accuracy/calibration ties to MR-COMM-011 calibration service.
MR-COMM-010 — Warranty & Customer Support Model
Requirement Statement | The market demands a defined warranty and responsive support model appropriate to a compliance-critical, multi-trip asset used across global time zones.
Business Objective | Provide the assurance and serviceability that regulated buyers require to standardise on the product.
Customer Segment | Regulated distributors/3PLs and institutional programs.
Geographic Applicability | Global; multi-language/time-zone support expectation.
Market Priority | High
Importance | High
Competitive Driver | Enterprise cold-chain buyers require SLAs and warranty before standardising.
Business Value | Enables enterprise/institutional adoption; reduces churn.
Revenue Impact | Medium — supports retention and can be a paid premium-support tier.
Risk if Unmet | Buyers perceive operational risk; blocked from enterprise/tender wins.
Evidence Source | Judgment/Assumption — SLA/warranty terms Evidence Pending.
Supporting Research | Enterprise IoT and GDP-regulated procurement expectations; specifics pending.
VOC Reference | VOC-### (pending) — theme: "what happens when one fails in the field?"
Strategic Theme | Trust & serviceability.
Product Family | SCM — support/warranty program.
Lifecycle Stage | Launch
Traceability | forward to MR-1000 MET (support cost/device, NPS), MR-900 CON; ORCH-018 / EM-ENG-012.
Notes | Support cost per device is a tracked KPI (MR-1000 MET).
MR-COMM-011 — Calibration-as-a-Service & Recertification
Requirement Statement | The market demands ongoing sensor calibration/recertification (with traceable certificates) across the reusable device's service life, because monitoring data is only trusted if the sensors are demonstrably in calibration.
Business Objective | Preserve data trust across reuse and create a recurring service-revenue stream.
Customer Segment | GDP/Part-11-regulated distributors, pharma, and PQS-procured programs.
Geographic Applicability | Global; certificate expectations strongest in EU/US regulated markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Regulated cold-chain requires calibrated, traceable measurement; recalibration is a known reusable-asset need.
Business Value | Recurring service revenue plus a compliance moat; keeps reused devices trustworthy.
Revenue Impact | Medium/High recurring — calibration-as-a-service.
Risk if Unmet | Data rejected by regulated buyers; devices unusable after drift; reuse thesis undermined.
Evidence Source | Standard — GDP and data-integrity practice require calibrated, traceable measurement (see MR-1100 REG); pricing Judgment/Assumption (Evidence Pending).
Supporting Research | EU GDP and FDA data-integrity expectations for measurement traceability; recalibration inherent to multi-trip reuse.
VOC Reference | VOC-### (pending) — theme: "we need a calibration certificate we can show an auditor."
Strategic Theme | Compliance-grade reuse.
Product Family | SCM — calibration service.
Lifecycle Stage | Growth
Traceability | forward to MR-1100 REG (GDP/Part-11), MR-1000 MET (service revenue), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Calibration loop should integrate with reverse-logistics refurbishment (MR-COMM-007).
7. Evidence & Validation Notes
• All price, unit-cost, margin, subscription, and attach-rate figures are Judgment/Assumption placeholders and must be validated by (a) competitor pricing/teardown analysis, (b) a bottom-up BOM + landed-cost model, (c) a reverse-logistics cost model, and (d) structured VOC/willingness-to-pay research across the three buyer segments.
• Institutional procurement claims (GAVI/UNICEF tender) are based on known market practice but specific tender terms and eligibility are Evidence Pending and interlock with PQS listing (MR-1100 REG).
• Calibration/data-trust requirements are Standard-backed (EU GDP, FDA data-integrity) — see MR-1100 REG for citations.
• Recommended validations: pricing conjoint/Van Westendorp study; TCO model peer review; reverse-logistics pilot to measure real return rates feeding MR-1000 MET.
8. Traceability Placeholders
MR ID | Forward to User Needs | Forward to Product Requirements | Governance
MR-COMM-001 | UN-… (TCO/reuse) | PRD-… (service life/trips) | ORCH-018 / EM-ENG-012
MR-COMM-002 | UN-… | PRD-… (BOM budget) | ORCH-018 / EM-ENG-012
MR-COMM-003 | UN-… | PRD-… (tender SKU) | ORCH-009 / ORCH-018
MR-COMM-004 | — | PRD-… (cost budget) | ORCH-018 / EM-ENG-012
MR-COMM-005 | — | PRD-… (financial model) | ORCH-018 / EM-ENG-012
MR-COMM-006 | UN-… (continuous visibility) | PRD-… (connectivity/cloud) | ORCH-018 / EM-ENG-012
MR-COMM-007 | UN-… (return loop) | PRD-… (pooling/refurb) | ORCH-018 / EM-ENG-012
MR-COMM-008 | UN-… | PRD-… | ORCH-018 / EM-ENG-012
MR-COMM-009 | UN-… (fit/mount) | PRD-… (probe/mount IF) | ORCH-018 / EM-ENG-012
MR-COMM-010 | UN-… (support) | PRD-… (serviceability) | ORCH-018 / EM-ENG-012
MR-COMM-011 | UN-… (calibration cert) | PRD-… (calibration/traceability) | ORCH-018 / EM-ENG-012
Cross-module links: MR-COMM ↔ MR-900 CON (cost/logistics constraints), MR-1000 MET (revenue/margin/retention/service KPIs), MR-1100 REG (PQS listing, GDP/Part-11 calibration & data trust).
MR-900 — Business Constraints (CON)
Traceability note: This module (code CON) forms part of the Smart Cold Chain Monitor (SCM) Master Market Requirements Library. Constraints traced forward to User Needs (UN-…) and Product Requirements (PRD-…) placeholders; governed under ORCH-018 (Traceability), authored under ORCH-009 / EM-ENG-012.
1. Purpose
This module captures the business and market constraints that bound what the Smart Cold Chain Monitor (SCM) can commercially and technically be: development budget and time-to-market, manufacturing capability, the low-temperature battery and semiconductor/cellular supply chain, global reverse-logistics realities, cellular carrier/roaming limitations, cross-market regulatory limits, and business/IP risk.
Constraints differ from commercial demands (MR-800): they express limits the market and operating environment impose on the business, which every downstream requirement must respect. Stating them explicitly prevents engineering and go-to-market from committing to requirements the supply chain, regulatory landscape, or capital plan cannot support.
2. Scope
In scope: development budget and schedule/time-to-market; manufacturing and test capability (including calibration at volume); low-temp rechargeable Li-ion battery supply and availability; semiconductor and cellular-module (LTE-M) availability; global logistics for reverse flows; cellular carrier/roaming/eSIM constraints; regulatory limitations that constrain the commercial rollout sequence; and business, competitive, and IP risk.
Out of scope: the commercial offer itself (MR-800 COMM), success targets (MR-1000 MET), and the detailed regulatory register (MR-1100 REG) — though constraints here reference REG where a regulatory limit acts as a business constraint on market entry timing or cost.
3. Applicable Standards & References
• UN 38.3 (Transport of Dangerous Goods — Li battery testing) and IATA Dangerous Goods Regulations (Section II / lithium provisions) — constrain how the reusable, battery-containing device can be shipped and returned globally.
• GSMA / 3GPP LTE-M (Cat-M1) specifications and regional carrier deployment — constrain where always-on cellular monitoring is actually available; roaming/eSIM policy constrains global operability.
• WHO PQS prequalification and country type-approval / import regimes — constrain the sequence and timing of market entry (see MR-1100 REG).
• IEC 62368-1, UN 38.3, and battery-cell industry supply data — constrain low-temperature rechargeable Li-ion sourcing (charge inhibited below 0 °C is a evidence-backed design fact, D3).
• Semiconductor and cellular-module lead-time/allocation conditions are referenced as market practice / Judgment/Assumption; no specific lead-time number is asserted as fact (Evidence Pending).
• Budget, schedule, and capacity figures are Judgment/Assumption unless cited; they are planning envelopes, not committed facts.
4. Background
The SCM is a globally-sold, battery-operated, cellular-connected reusable asset. Each of those attributes carries a hard constraint. "Battery-operated" and "shipped globally" collide with UN 38.3 / IATA DG rules that govern lithium-cell transport and therefore reverse logistics. "Low-temperature operation to −30 °C" with rechargeable Li-ion is a specialised cell requirement (charging must be inhibited below 0 °C — anchor fact D3), narrowing the qualified supplier pool. "Cellular-connected" depends on LTE-M coverage and roaming agreements that are uneven across exactly the LMIC last-mile geographies the product targets. "Multi-market regulated" (WHO PQS + EU GDP + US FDA + air/ground transport — anchor fact D4) means market entry is gated and sequenced by multiple independent approval processes.
Layered on top are the ordinary constraints of a hardware+SaaS venture: finite development budget and time-to-market pressure, the need to build or contract manufacturing and volume calibration capability, exposure to semiconductor/module allocation cycles, and competitive/IP risk in an active cold-chain-visibility market. These constraints are the guardrails within which MR-800, MR-1000, and MR-1100 must be satisfied.
5. Requirement Categories
1. Program Constraints — budget, schedule/time-to-market.
2. Manufacturing & Test — production and volume-calibration capability.
3. Supply Chain — low-temp battery, semiconductor/cellular module.
4. Connectivity — carrier coverage, roaming/eSIM.
5. Logistics — global reverse flows and DG transport.
6. Regulatory & Business Risk — cross-market approval limits, competitive/IP risk.
6. Detailed Market Requirements
6.1 Program Constraints
MR-CON-001 — Development Budget Ceiling
Requirement Statement | The business is constrained to deliver the SCM within a finite development budget envelope, so the market requirement set must be prioritised to a fundable scope.
Business Objective | Deliver a market-viable device+platform without over-investing ahead of validated demand.
Customer Segment | Internal (investors / program sponsors); indirectly all customers via scope.
Geographic Applicability | Global program.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Capital efficiency versus well-funded cold-chain-visibility competitors.
Business Value | Protects runway; forces MVP discipline aligned to Must-have market requirements.
Revenue Impact | Indirect — governs time-to-first-revenue.
Risk if Unmet | Scope creep exhausts capital before market entry.
Evidence Source | Judgment/Assumption — specific budget figure Evidence Pending (program financial plan).
Supporting Research | Standard hardware+SaaS NPD budgeting; no figure asserted.
VOC Reference | n/a (internal constraint).
Strategic Theme | Capital discipline.
Product Family | SCM — program.
Lifecycle Stage | Concept
Traceability | forward to PRD-… (MVP scope), MR-800 COMM (margin); ORCH-018 / EM-ENG-012.
Notes | Drives prioritisation of Must-have MRs across all four modules.
MR-CON-002 — Time-to-Market Window
Requirement Statement | The market imposes a time-to-market window driven by procurement cycles (donor tenders, distributor budget years) and competitor moves, constraining acceptable development schedule.
Business Objective | Enter the market in time to capture upcoming tender/budget cycles.
Customer Segment | Institutional buyers (tender calendars); distributors (annual budgets).
Geographic Applicability | Global; tender calendars are region/donor specific.
Market Priority | High
Importance | High
Competitive Driver | Missing a tender cycle can cost a multi-year window; first-mover credibility matters.
Business Value | Aligns launch to buying windows, maximising early revenue capture.
Revenue Impact | High — a missed cycle defers large volume by years.
Risk if Unmet | Product misses procurement windows; competitors entrench.
Evidence Source | Cited — donor/tender procurement operates on defined cycles (market practice); exact dates Evidence Pending.
Supporting Research | UNICEF/GAVI and national EPI procurement calendars; PQS listing lead time (MR-1100 REG).
VOC Reference | VOC-### (pending) — theme: procurement timing.
Strategic Theme | Time-to-market.
Product Family | SCM — program schedule.
Lifecycle Stage | Concept
Traceability | forward to MR-1100 REG (approval lead times), PRD-…; ORCH-018 / EM-ENG-012.
Notes | PQS + type-approval lead times (MR-1100 REG) are the critical path.
6.2 Manufacturing & Test
MR-CON-003 — Manufacturing & Volume-Calibration Capability
Requirement Statement | The business is constrained by the need to establish manufacturing and volume sensor-calibration/test capability capable of producing GDP/PQS-grade, individually-calibrated devices at target volume.
Business Objective | Ensure devices can be built and calibrated at scale with traceable quality records.
Customer Segment | Regulated buyers requiring calibrated, traceable devices.
Geographic Applicability | Global supply to all markets.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Volume calibration with traceable certificates is a barrier and a differentiator.
Business Value | Enables compliant supply and calibration-as-a-service (MR-COMM-011).
Revenue Impact | High — gates ability to fulfil regulated orders.
Risk if Unmet | Cannot scale compliant supply; calibration bottleneck caps volume.
Evidence Source | Standard — GDP/data-integrity require traceable calibration; capacity figures Judgment/Assumption (Evidence Pending).
Supporting Research | EU GDP measurement traceability; contract-manufacturing + metrology capability (see MR-1100 REG).
VOC Reference | n/a (internal constraint from regulated demand).
Strategic Theme | Compliant manufacturability.
Product Family | SCM — manufacturing/test.
Lifecycle Stage | Launch
Traceability | forward to MR-COMM-011 (calibration service), MR-1100 REG, PRD-…; ORCH-018 / EM-ENG-012.
Notes | Volume-calibration throughput is a likely capacity constraint.
6.3 Supply Chain
MR-CON-004 — Low-Temperature Rechargeable Battery Supply
Requirement Statement | The business is constrained by the limited supplier pool for rechargeable Li-ion cells qualified for reliable operation to −30 °C with charging safely inhibited below 0 °C, which narrows sourcing and raises supply risk.
Business Objective | Secure a qualified, second-sourced low-temp cell supply that meets endurance and transport requirements.
Customer Segment | All (battery reliability underpins the whole value proposition).
Geographic Applicability | Global sourcing; global cold-environment operation.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Low-temp performance and 90-day endurance depend on cell chemistry availability.
Business Value | Protects the core reliability claim and endurance spec.
Revenue Impact | High — a battery supply failure halts production and field reliability.
Risk if Unmet | Single-source dependency; excursions in the device's own performance in cold; UN 38.3 requalification on cell change.
Evidence Source | Standard/Judgment — charge-inhibit below 0 °C and −30 °C operation are anchor facts (D3, range spec); supplier-pool scarcity is Judgment/Assumption (Evidence Pending).
Supporting Research | Anchor facts: low-temp Li-ion, duty-cycled, charge inhibited <0 °C (D3); UN 38.3 requalification on cell change.
VOC Reference | n/a (internal supply constraint).
Strategic Theme | Supply resilience.
Product Family | SCM — battery subsystem.
Lifecycle Stage | Concept
Traceability | forward to MR-1100 REG (UN 38.3/IATA), PRD-… (battery spec); ORCH-018 / EM-ENG-012.
Notes | Cell changes trigger UN 38.3 retest — a schedule and cost constraint.
MR-CON-005 — Semiconductor & Cellular-Module Availability
Requirement Statement | The business is constrained by semiconductor and LTE-M cellular-module availability, lead times, and allocation cycles, which can limit production volume and design longevity.
Business Objective | Choose components and second sources that de-risk allocation and long-term availability.
Customer Segment | All (affects supply continuity).
Geographic Applicability | Global.
Market Priority | High
Importance | High
Competitive Driver | Component availability determines who can actually ship at volume.
Business Value | Supply continuity; avoids costly redesigns and recertification.
Revenue Impact | High — allocation shortfalls cap revenue.
Risk if Unmet | Production stalls; forced module change triggers re-certification (FCC/CE/PQS).
Evidence Source | Judgment/Assumption — allocation/lead-time conditions are market-practice; specific figures Evidence Pending.
Supporting Research | LTE-M module ecosystem; module change implies EMC/radio recert (see MR-1100 REG).
VOC Reference | n/a (internal supply constraint).
Strategic Theme | Supply resilience.
Product Family | SCM — connectivity/compute BOM.
Lifecycle Stage | Concept
Traceability | forward to MR-1100 REG (radio/EMC recert), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Design for module second-sourcing to protect certification investment.
6.4 Connectivity
MR-CON-006 — Cellular Coverage, Roaming & eSIM Constraints
Requirement Statement | The business is constrained by uneven LTE-M coverage and cross-border roaming/eSIM policy, which limits where continuous cellular monitoring is available and complicates a single global connectivity offer.
Business Objective | Deliver dependable connectivity and cost across target geographies despite fragmented carrier availability.
Customer Segment | Buyers shipping through LMIC last-mile and across borders.
Geographic Applicability | Global; most acute in LMIC and cross-border transport corridors.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Reliable multi-country connectivity is a core differentiator for in-transit monitoring.
Business Value | Enables the SaaS/data-plan revenue model to function globally.
Revenue Impact | High — connectivity gaps break the alerting value proposition and data-plan revenue.
Risk if Unmet | Devices go dark on key routes; excursion alerts (D2) fail to deliver in real time.
Evidence Source | Cited/Judgment — LTE-M coverage and roaming policy vary by country (market practice); specific coverage maps Evidence Pending.
Supporting Research | Evidence: LTE-M primary with Wi-Fi/BLE fallbacks and store-and-forward cadence upload (D2); roaming/eSIM commercial terms pending.
VOC Reference | VOC-### (pending) — theme: "does it work on our routes?"
Strategic Theme | Global connectivity.
Product Family | SCM — connectivity + data plan.
Lifecycle Stage | Launch
Traceability | forward to MR-COMM-006 (SaaS/data plan), PRD-… (store-and-forward), MR-1000 MET; ORCH-018 / EM-ENG-012.
Notes | Store-and-forward + multi-bearer partly mitigates coverage gaps; roaming cost affects data-plan COGS.
6.5 Logistics
MR-CON-007 — Global Reverse Logistics & DG Transport of Batteries
Requirement Statement | The business is constrained by the cost, complexity, and dangerous-goods rules of moving battery-containing devices globally in reverse (returns), which the reusable model depends upon.
Business Objective | Establish a reverse-logistics network that is DG-compliant and economically viable per trip.
Customer Segment | All reusable-device buyers relying on return/pooling.
Geographic Applicability | Global; hardest from dispersed LMIC last-mile.
Market Priority | High
Importance | Critical
Competitive Driver | Viable reverse logistics is what makes reusability real versus disposables.
Business Value | Preserves per-trip economics and enables pooling/rental (MR-COMM-007/008).
Revenue Impact | High — return-loop cost directly determines device amortisation and margin.
Risk if Unmet | Devices not recovered or non-compliantly shipped; reuse economics and legal compliance fail.
Evidence Source | Standard — UN 38.3 / IATA DG govern Li-battery transport; return-rate/cost figures Judgment/Assumption (Evidence Pending).
Supporting Research | UN 38.3, IATA DGR lithium provisions; reverse-logistics obligation from reusability (MR-800 COMM).
VOC Reference | VOC-### (pending) — theme: return handling burden.
Strategic Theme | Reverse-logistics viability.
Product Family | SCM — reverse logistics/pooling.
Lifecycle Stage | Growth
Traceability | forward to MR-COMM-007, MR-1100 REG (UN 38.3/IATA), MR-1000 MET (utilization/returns); ORCH-018 / EM-ENG-012.
Notes | DG classification/packaging drives return shipping cost and options.
6.6 Regulatory & Business Risk
MR-CON-008 — Cross-Market Regulatory Sequencing Limits
Requirement Statement | The business is constrained by multiple independent, time-consuming approval processes (WHO PQS, EU GDP alignment, US FDA data-integrity posture, radio/EMC, per-country type-approval/import), which limit the pace and sequence of global market entry.
Business Objective | Sequence approvals to reach priority markets first without over-committing resources.
Customer Segment | All regulated/donor buyers (approval gates purchase eligibility).
Geographic Applicability | Global; requirements differ per market.
Market Priority | Must-have
Importance | Critical
Competitive Driver | Approvals gate market access; PQS listing gates donor tenders.
Business Value | A realistic rollout plan aligned to achievable approvals.
Revenue Impact | High — each unattained approval closes a market.
Risk if Unmet | Over-promised launch dates; blocked market entry; wasted certification spend.
Evidence Source | Standard — PQS/GDP/FDA/radio/type-approval are established regimes (see MR-1100 REG); timelines Evidence Pending.
Supporting Research | Evidence-backed multi-market fact D4; detailed register in MR-1100 REG.
VOC Reference | n/a (regulatory constraint).
Strategic Theme | Regulated market access.
Product Family | SCM — regulatory program.
Lifecycle Stage | Concept
Traceability | forward to MR-1100 REG (all approvals), MR-CON-002 (time-to-market); ORCH-018 / EM-ENG-012.
Notes | Detailed obligations and Evidence-Pending classifications live in MR-1100 REG.
MR-CON-009 — Data Sovereignty & Privacy Constraints on the SaaS Platform
Requirement Statement | The business is constrained by cross-border data-transfer, privacy (e.g., GDPR and regional regimes), and data-residency requirements that limit how monitoring/telemetry data can be stored and moved across markets.
Business Objective | Operate the cloud platform lawfully across all served jurisdictions.
Customer Segment | EU/regulated buyers; government/donor programs with data-residency requirements.
Geographic Applicability | Global; strictest in EU and data-localisation jurisdictions.
Market Priority | High
Importance | High
Competitive Driver | Enterprise/government buyers require demonstrable data-protection compliance.
Business Value | Enables lawful SaaS operation and unlocks privacy-sensitive buyers.
Revenue Impact | Medium/High — non-compliance blocks EU/government revenue and risks fines.
Risk if Unmet | Regulatory penalties; loss of enterprise/government contracts.
Evidence Source | Standard — GDPR and regional privacy/residency regimes apply (see MR-1100 REG); specific residency needs Evidence Pending.
Supporting Research | GDPR; regional data-localisation trends; cross-referenced in MR-1100 REG.
VOC Reference | VOC-### (pending) — theme: "where does our data live?"
Strategic Theme | Trust & compliance.
Product Family | SCM — cloud platform.
Lifecycle Stage | Launch
Traceability | forward to MR-1100 REG (GDPR/residency), MR-COMM-006 (SaaS); ORCH-018 / EM-ENG-012.
Notes | May require regional cloud hosting — a cost/architecture constraint.
MR-CON-010 — Competitive & Intellectual-Property Risk
Requirement Statement | The business is constrained by an active competitive landscape and IP risk (existing cold-chain-monitoring patents and fast-following competitors), which limits freedom-to-operate and defensible differentiation.
Business Objective | Secure freedom-to-operate and defensible IP while differentiating on reuse + data.
Customer Segment | Internal (affects all customers via product viability).
Geographic Applicability | Global (patent regimes are territorial).
Market Priority | Medium
Importance | High
Competitive Driver | Cold-chain visibility is a crowded, patent-active market.
Business Value | Protects margin and market position; avoids costly litigation/redesign.
Revenue Impact | Medium/High — IP disputes can block sales or force redesign.
Risk if Unmet | Infringement exposure or easy competitor replication erodes advantage.
Evidence Source | Judgment/Assumption — FTO/IP landscape Evidence Pending (formal FTO search required).
Supporting Research | Active cold-chain-monitoring competitive field; formal patent landscape pending.
VOC Reference | n/a (business risk).
Strategic Theme | Defensibility.
Product Family | SCM — IP/strategy.
Lifecycle Stage | Concept
Traceability | forward to MR-800 COMM (positioning), PRD-…; ORCH-018 / EM-ENG-012.
Notes | Commission FTO search before design freeze.
7. Evidence & Validation Notes
• Budget, schedule, capacity, lead-time, return-rate, and coverage figures are Judgment/Assumption / Evidence Pending and require: a program financial plan, a manufacturing/metrology capacity study, supplier qualification and second-sourcing analysis, carrier coverage/roaming assessment per corridor, a reverse-logistics cost/return-rate pilot, and a formal freedom-to-operate search.
• Regulatory-sequencing and DG-transport constraints are Standard-backed; specific timelines, PQS category, and per-country type-approval details are Evidence Pending and detailed in MR-1100 REG.
• Battery low-temp and charge-inhibit constraints are anchor facts (D3 and range spec); supplier-pool scarcity remains an assumption to be validated by sourcing.
8. Traceability Placeholders
MR ID | Forward to Product Requirements | Cross-module | Governance
MR-CON-001 | PRD-… (MVP scope) | MR-800 COMM | ORCH-018 / EM-ENG-012
MR-CON-002 | PRD-… (schedule) | MR-1100 REG | ORCH-018 / EM-ENG-012
MR-CON-003 | PRD-… (mfg/calibration) | MR-COMM-011, MR-1100 REG | ORCH-018 / EM-ENG-012
MR-CON-004 | PRD-… (battery) | MR-1100 REG (UN 38.3) | ORCH-018 / EM-ENG-012
MR-CON-005 | PRD-… (module 2nd source) | MR-1100 REG (recert) | ORCH-018 / EM-ENG-012
MR-CON-006 | PRD-… (multi-bearer/S&F) | MR-COMM-006, MR-1000 MET | ORCH-018 / EM-ENG-012
MR-CON-007 | PRD-… (reverse logistics) | MR-COMM-007, MR-1100 REG | ORCH-018 / EM-ENG-012
MR-CON-008 | PRD-… (regulatory plan) | MR-1100 REG | ORCH-018 / EM-ENG-012
MR-CON-009 | PRD-… (data residency) | MR-1100 REG (GDPR) | ORCH-018 / EM-ENG-012
MR-CON-010 | PRD-… (IP/FTO) | MR-800 COMM | ORCH-018 / EM-ENG-012
Cross-module links: MR-900 CON bounds MR-800 COMM (what the business can offer), gates MR-1000 MET (achievable targets), and is detailed by MR-1100 REG (regulatory obligations).

## 2 Voice of Customer
Voice of Customer
What customers, partners and front-line staff actually told us, in their own words.
↩ Contents
In this section:
• Voice of Customer (VoC) Library — Smart Cold Chain Monitor (SCM)
• VoC — Field Observations, Complaints & Support, Distributor/Installer/Service/Clinical Feedback
• VoC — Depth-Interview Findings (VOC-001 … VOC-024)
• VoC — Survey Themes & Lost-Deal / No-Decision Reasons
Voice of Customer (VoC) Library — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 02 — Voice of Customer · Owner: Principal Market Research Director (with EM-ENG-001 Chief Systems Engineer as integrator; Human-Factors contributor for field/point-of-use themes; EM-ENG-006 Quality/Regulatory for audit/compliance themes) · Convening: feeds Stakeholder Needs (STN-###, folder 03) and User Needs (Stage 4, ORCH-010) · Traceability anchors: ORCH-018 / EM-ENG-012 · Status: Authored — representative, product-specific, honestly labelled.
---
1. Purpose
This library captures the Voice of the Customer for the SCM: what the people who fund, deploy, carry, receive, audit, service, and act on the data actually say they need, in their own terms. It converts the market baseline (Stage 1, 01-market-requirements/) into concrete stakeholder statements, findings, and signals that can be traced forward into Stakeholder Needs (STN-###) and User Needs (UN-###).
It is based on the single source of truth, 00-concept/concept-and-decisions.md (foundational decisions D1–D4), and aligned to the Master Market Requirements Library (01-market-requirements/index.md), especially MR-200 (Segmentation, SEG), MR-300 (Problems, PROB), MR-400 (Jobs To Be Done, JTBD), and MR-1200 (Personas, PER).
2. Honesty / evidence posture (READ THIS FIRST)
Per 00-governance/evidence-standard.md and hallucination-policy.md, this stage is authored before a primary VoC field program has been executed. Therefore:
• No verbatim below is presented as a real, attributed citation from a named individual or organization.
• Every statement is representative synthesized VoC — a plausible paraphrase constructed from the evidence-backed market problems (MR-PROB), jobs (MR-JTBD), personas (MR-PER), cited public-health reporting (WHO/UNICEF/GAVI, reported/estimated), and applicable standards (WHO PQS, EU GDP, FDA 21 CFR Part 11, UN 38.3 / IATA DGR, IEC 62368-1 / 60068 / 60529). These are explicitly labelled Judgment/Assumption or Evidence Pending.
• Where a statement rests directly on a published standard or widely reported public-health fact, it is labelled Standard or Cited (reported/estimated) in the Confidence field.
• Each item carries a Confidence rating and a Notes field naming the study that would confirm it. These synthesized items are hypotheses to be validated by the recommended VoC program (Section 6), and the placeholder VOC-### (pending) references already seeded in the MR library will be back-filled with the real IDs from this library.
Bottom line: treat every quoted line as representative of a theme we must go and confirm, not as an audited quotation. This is the discipline the org requires (CLAUDE.md rule 1: ground everything; separate fact from inference; state the test that would confirm).
3. ID scheme
Prefix | Meaning | Where
VOC-### | Primary Voice-of-Customer item (interview, survey theme, synthesized verbatim) | all files
FLD-### | Field observation / contextual-inquiry finding | VOC-field-and-support.md
CFB-### | Customer feedback: complaint / support ticket / distributor-installer-service feedback | VOC-field-and-support.md
LD-### | Lost-deal / no-decision reason | VOC-survey-and-lostdeals.md
VoC items trace backward to Market Requirements (MR-…) and forward to Stakeholder Needs (STN-###, folder 03) and, subsequently, User Needs (UN-###, Stage 4).
4. Library contents
File | Content | Items
[VOC-interviews.md](VOC-interviews.md) | Depth-interview findings across every stakeholder segment (synthesized) | VOC-001 … VOC-024
[VOC-field-and-support.md](VOC-field-and-support.md) | Field / contextual-inquiry observations, complaints & support issues, distributor / installer / service / clinical-operational feedback | FLD-001 … FLD-012; CFB-001 … CFB-016
[VOC-survey-and-lostdeals.md](VOC-survey-and-lostdeals.md) | Survey themes (importance/satisfaction), and lost-deal / no-decision reasons | VOC-050 … VOC-064; LD-001 … LD-010
5. Theme roll-up (what customers keep telling us)
The synthesized VoC converges on a small number of durable themes. Each maps to evidence-backed market problems and jobs; each is carried forward to Stakeholder Needs.
Theme | One-line customer voice | Primary MR anchor | Carries to STN
T1 Real-time excursion warning | "Tell me before it's too late, not at delivery." | MR-PROB-001, MR-JTBD-005 | STN-001, STN-018
T2 Two-sided (freeze + heat) detection | "Freezing quietly ruins stock and nobody sees it." | MR-PROB-002 | STN-002
T3 Cold-endurance / battery survives the trip | "The last logger gave out in the cold and we lost the record." | MR-PROB-007, D3 | STN-003
T4 Instant, offline accept/reject at the door | "Green/red at the door, no phone, no manual." | MR-PROB-004, MR-JTBD-004, MR-PER-007 | STN-004
T5 Audit-grade, tamper-evident data | "I have to defend every excursion to an inspector." | MR-PROB-005, MR-JTBD-008, MR-PER-009 | STN-005
T6 Attributable evidence (who broke the chain) | "We get blamed for excursions we didn't cause." | MR-PROB-006 | STN-006
T7 Alert quality / no alarm fatigue | "So many false alerts that we stopped looking." | MR-PROB-009 | STN-007
T8 Frictionless provisioning / arming | "Set it and forget it — correctly — in seconds." | MR-JTBD-006, MR-PER-005 | STN-008
T9 Last-mile / intermittent-connectivity resilience | "The last mile is a black box." | MR-PROB-008 | STN-009
T10 Reuse economics: recover, clean, recal, redeploy | "Reusable only pays if we actually get them back." | MR-PROB-011, MR-JTBD-007/009, MR-PER-011 | STN-010, STN-016
T11 PQS eligibility + total cost of ownership | "If it's not PQS-listed and cost-effective, we can't buy it." | MR-SEG-005/006, MR-PER-008 | STN-011
T12 Transport / battery compliance (air & ground) | "It can't trip a battery-shipping rule or slow the dock." | MR-SEG-004, MR-PER-004, D4 | STN-012
T13 Enterprise data governance / integration | "It must pass our CSV and SSO/access policy." | MR-PER-003 | STN-013
T14 Fleet visibility & ROI for the executive | "Show me network-wide loss reduction and cost-per-trip." | MR-JTBD-010, MR-PER-010 | STN-014
6. Recommended VoC validation program (to convert synthesized → cited)
To promote these items from Judgment/Assumption to Cited, the org recommends a structured primary study before User Needs (Stage 4) freeze:
1. Contextual inquiry at 2–3 distributor dispatch docks and cold stores (provisioning, gloved handling).
2. Ride-along / observation with ≥2 LMIC last-mile immunization teams (offline read, accept/reject).
3. Depth interviews across each persona class (MR-PER-001…011), min. one per class, spanning EU, US, and one LMIC corridor.
4. Audit simulation with QA/regulatory personas (MR-PER-009) to test record defensibility.
5. Service-depot time-and-motion study for reverse logistics / recalibration (MR-PER-011).
6. Importance–satisfaction survey across segments to quantify theme priority and size opportunity gaps.
Quantitative gaps that remain Evidence Pending until this program runs: per-shipment spoilage cost, willingness-to-pay, excursion/freeze frequency by lane, false-alarm rate tolerance, device return rate, service throughput, and OLED low-temperature legibility thresholds.
7. Traceability (forward / backward)
Backward: each VoC item cites its MR-… anchor (and D1–D4 where a foundational decision drives it). Forward: each theme and item is carried into Stakeholder Needs (STN-###, 03-stakeholder-needs/), then into User Needs (UN-###, Stage 4, ORCH-010). The end-to-end trace matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18).
Flow: MR (Stage 1) → VOC-### (this stage) → STN-### (Stage 3) → UN-### (Stage 4) → PRD-### → … → Verification / Validation.
VoC — Field Observations, Complaints & Support, Distributor/Installer/Service/Clinical Feedback
Evidence note (read index.md §2): the SCM is pre-launch, so there is no real deployed-fleet
support history yet. Every item below is anticipated / representative synthesized VoC — field
observations projected from analogous cold-chain deployments and incumbent-logger behaviour, and support
issues anticipated from the concept's known tensions (D1–D4) and evidence-backed market problems (MR-PROB).
They are hypotheses for the field/pilot program, not audited incident records. Backward trace to MR-…;
forward trace to STN-### / UN-###. Governance: ORCH-018 / EM-ENG-012.
FLD record fields: ID · Observation · Source type · Stakeholder/segment · Context · Interpreted need theme · Priority/frequency signal · Linked MR · Confidence · Notes. CFB record fields: ID · Reported issue (paraphrased) · Source type · Stakeholder/segment · Context · Interpreted need theme · Priority/frequency signal · Linked MR · Confidence · Notes.
---
Part 1 — Field / contextual-inquiry observations (FLD-001 … FLD-012)
FLD-001 — Devices arrive un-armed ("forgot to start the logger")
Observation | At high-throughput pack-out, observers project that a meaningful fraction of devices would ship without being correctly armed/configured unless arming is a single, confirmed, hard-to-skip action.
Source type | Contextual inquiry (anticipated, synthesized) — distributor pack-out dock.
Stakeholder/segment | Pack-out / warehouse operators (MR-PER-002/005).
Context | Time-pressured dispatch; legacy loggers' classic "not started" failure.
Interpreted need theme | T8 — provisioning must be frictionless and error-proof; on-device armed confirmation.
Priority/frequency signal | High; silent value loss (void trips) if unmanaged.
Linked MR | MR-JTBD-006; personas MR-PER-002/005.
Confidence | Judgment/Assumption — based on reported logger failure modes; SCM rate Evidence Pending.
Notes | OLED + BLE provisioning with a clear "armed/placed" confirmation directly mitigates.
FLD-002 — Wrong probe placement inside the shipper
Observation | Operators place the device/probe inconveniently (against a coolant pack, at the lid) rather than in the product zone, biasing readings toward false freeze/heat events.
Source type | Contextual inquiry (anticipated) — cold-store staging.
Stakeholder/segment | Warehouse cold-store operators (MR-PER-005).
Context | Internal vs. external-probe placement; freeze-detection accuracy.
Interpreted need theme | T2 + T8 — placement guidance so freeze/heat detection reflects the product.
Priority/frequency signal | Medium-High; drives false alarms and disputed records.
Linked MR | MR-PROB-002/009, MR-JTBD-006; MR-PER-005.
Confidence | Judgment/Assumption — placement error common in field programs; confirm via observation.
Notes | Resolve placement guidance (int vs ext probe) in User Needs / PRD; ties to VOC-011.
FLD-003 — Gloved, cold-numbed interaction with the OLED/button
Observation | In +2/+8 °C rooms and −20 °C freezers, staff operate gloved with low dexterity; multi-step or small-target interactions are skipped or fumbled; OLED legibility at low temperature is uncertain.
Source type | Contextual inquiry (anticipated) — cold-store / freezer.
Stakeholder/segment | Warehouse cold-store operators (MR-PER-005).
Context | Human factors at temperature; single-action requirement.
Interpreted need theme | T8 — glove-friendly, cold-legible, single-action interaction.
Priority/frequency signal | Medium-High; handling errors corrupt trip records.
Linked MR | MR-PER-005; MR-1300 ENV; IEC 60068 (low-temp).
Confidence | Judgment/Assumption — cold-store ergonomics; OLED low-temp legibility thresholds Evidence Pending.
Notes | Flag for HMI/display-engineering review (OLED range to −30 °C).
FLD-004 — Deep-cold lane: connectivity and battery both degrade together
Observation | On long deep-cold legs, LTE-M coverage is patchy and battery capacity is depressed; worst-case is a device that must buffer for days then upload a backlog when it warms and reconnects.
Source type | Field observation (anticipated) — cold-lane transit.
Stakeholder/segment | 3PL / distributor cold-lane (MR-SEG-002/004).
Context | D2 store-and-forward + D3 low-temp energy budget interacting.
Interpreted need theme | T1 + T3 + T9 — buffer offline, endure the cold, upload opportunistically.
Priority/frequency signal | Critical; worst-case for both alert latency and endurance.
Linked MR | MR-PROB-007/008, MR-JTBD-005; D2/D3; MR-RISK-001.
Confidence | Standard/Judgment — based on D2/D3; quantified endurance Evidence Pending (test).
Notes | Non-volatile buffering must not lose data across the cold dwell; carry to hazard (data-loss) + test.
FLD-005 — Antenna performance inside metallized / insulated containers
Observation | Placing a cellular/GPS device inside an insulated, sometimes foil-lined shipper degrades RF; uploads and location fixes may fail from within the box even where coverage exists outside.
Source type | Field observation (anticipated) — in-container placement.
Stakeholder/segment | All shipping segments; fleet admins (MR-JTBD-010).
Context | Concept §7 open item — antenna performance inside metallized containers.
Interpreted need theme | T1 + T9 + T14 — reliable upload/location from inside the shipper.
Priority/frequency signal | High; a named open technical risk that undercuts real-time/location value.
Linked MR | MR-JTBD-005/010, MR-PROB-008; concept §7.
Confidence | Judgment/Assumption — flagged as open item in concept; must be characterized by test.
Notes | Carry to system/architecture + verification (RF-in-container test); may drive external-antenna option.
FLD-006 — Last-mile: no phone, no network, staff rotation
Observation | At rural posts, the receiving worker frequently has neither a smartphone nor connectivity, and may be seeing the device for the first time; anything requiring an app or account fails in practice.
Source type | Ride-along (anticipated) — LMIC last-mile.
Stakeholder/segment | Last-mile health workers (MR-PER-007).
Context | Offline accept/reject at point of administration.
Interpreted need theme | T4 — offline, self-evident, on-device verdict.
Priority/frequency signal | Critical for last-mile adoption.
Linked MR | MR-JTBD-004, MR-PROB-008; MR-PER-007; MR-1300 ENV.
Confidence | Standard/Judgment — WHO last-mile reality; confirm via ride-along.
Notes | OLED + button must carry the full accept/reject decision with zero connectivity.
FLD-007 — Reverse logistics: devices stranded at the edge
Observation | After delivery, devices accumulate at remote posts with no easy, funded return path; a fraction never come back, directly eroding reuse economics.
Source type | Field observation (anticipated) — LMIC reverse logistics.
Stakeholder/segment | Programs, distributors, reverse-logistics operators (MR-PER-011).
Context | D1 reusable model; return-rate economics on hard routes.
Interpreted need theme | T10 — recoverable fleet; return workflow; location to chase strays.
Priority/frequency signal | High; determines whether TCO story holds on export/last-mile lanes.
Linked MR | MR-PROB-011, MR-JTBD-009/010; MR-EXEC-008.
Confidence | Judgment/Assumption — return-rate Evidence Pending.
Notes | Consider region-specific single-use vs. reusable strategy; GPS/BLE to locate; ties to VOC-006.
FLD-008 — Charge-inhibit-below-0 °C surprises the depot
Observation | Devices returning frozen refuse to charge until warmed; without clear status the depot reads this as a "non-working device," delaying turnaround and inflating apparent failure rate.
Source type | Field observation (anticipated) — return depot.
Stakeholder/segment | Service technicians (MR-PER-011).
Context | D3 charge-below-0 °C inhibit.
Interpreted need theme | T10 + T3 — clear charge-state/readiness indication at temperature.
Priority/frequency signal | Medium-High; operational friction and false failure reports.
Linked MR | D3; MR-JTBD-007; MR-PER-011; IEC 62368-1.
Confidence | Standard/Judgment — charge-below-0 hazard established; workflow impact Evidence Pending.
Notes | Depot dock or device must show "warming / will charge"; self-heat-before-charge under evaluation (D3).
FLD-009 — Excursion clustered at custody hand-offs (tarmac / transfer)
Observation | Excursions concentrate at hand-offs — tarmac dwell, cross-docks, transfers — precisely where custody and accountability are ambiguous.
Source type | Field observation (anticipated) — multi-leg transit.
Stakeholder/segment | Distributors, 3PLs, manufacturers (MR-SEG-002/004/001).
Context | Multi-party custody; attribution.
Interpreted need theme | T6 — time/location-stamped, attributable evidence at hand-offs.
Priority/frequency signal | High; the moments that generate disputes.
Linked MR | MR-PROB-006, MR-JTBD-003; MR-PER-004.
Confidence | Judgment/Assumption — hand-off risk widely reported; confirm via lane data.
Notes | GPS + timestamp at event enables who/where/when attribution.
FLD-010 — Alarm fatigue from benign transients
Observation | Brief door-opens, truck-to-dock transfers, and tarmac spikes generate alerts that are not product-relevant; without smart logic staff mute alerts and later miss a real excursion.
Source type | Field observation (anticipated) — distribution operations.
Stakeholder/segment | Distributors, 3PLs, programs (MR-SEG-002/004/005).
Context | Threshold vs. product-stability-aware alerting.
Interpreted need theme | T7 — alerts reflect cumulative/stability-relevant exposure, not raw transients.
Priority/frequency signal | High; determines whether alerts remain trusted.
Linked MR | MR-PROB-009, MR-JTBD-005.
Confidence | Judgment/Assumption — alarm-fatigue literature; SCM tuning Evidence Pending.
Notes | Time-above/below-threshold and mean-kinetic-temperature-style logic candidates; resolve in PRD.
FLD-011 — Cleaning / decontamination between trips
Observation | Reusable devices returning from clinical/field settings need practical cleaning; crevices, ports, and seals must tolerate wipe-down without ingress or label loss.
Source type | Field observation (anticipated) — service depot.
Stakeholder/segment | Service technicians, sustainability leads (MR-PER-011).
Context | D1 reuse; cleanability; IP integrity (IEC 60529).
Interpreted need theme | T10 — cleanable, sealed, durable-labelled enclosure for many cycles.
Priority/frequency signal | Medium-High; hygiene + durability across reuse cycles.
Linked MR | MR-JTBD-009; MR-PER-011; D1; IEC 60529.
Confidence | Judgment/Assumption — cleaning is a stated D1 residual; method Evidence Pending.
Notes | Flag materials + mechanical engineers (chemical compatibility, seal life, USB-C port sealing).
FLD-012 — GPS/location reliability inside the box vs. cost of always-on
Observation | Fleet/asset teams want "where are my devices?" but continuous GPS is energy-expensive and fixes fail inside shippers; teams accept periodic/opportunistic location as a compromise.
Source type | Field observation (anticipated) — fleet operations.
Stakeholder/segment | Fleet/asset administrators (MR-PER-010; MR-JTBD-010).
Context | D3 energy budget vs. location value; concept §7 RF-in-container.
Interpreted need theme | T14 — useful location within the 90-day energy budget.
Priority/frequency signal | Medium; enables fleet visibility and stray recovery.
Linked MR | MR-JTBD-010, MR-PROB-011; D3; concept §7.
Confidence | Judgment/Assumption — energy/location tradeoff; quantify in power budget + test.
Notes | A conflict (continuous location vs. 90-day endurance); carry to STN-015 tradeoff.
---
Part 2 — Complaints, support issues & channel/service feedback (CFB-001 … CFB-016)
(Anticipated support taxonomy: what the support desk and channel would most likely hear once deployed, derived from the concept's tensions and incumbent-logger complaint patterns. Pre-launch — Evidence Pending.)
CFB-001 — "The device stopped working before the trip ended"
Reported issue | Battery depleted before the 90-day / end-of-trip target, especially on cold lanes.
Source type | Anticipated support ticket theme.
Stakeholder/segment | All shipping segments (MR-SEG-002/004/005).
Context | D3 low-temp endurance; the program's top risk.
Interpreted need theme | T3 — endurance across full range/duration.
Priority/frequency signal | Critical (severity-dominant); trust-destroying.
Linked MR | MR-PROB-007; D3; MR-RISK-001; MR-VAL-006.
Confidence | Standard/Judgment — evidence-backed tension; rate Evidence Pending (test).
Notes | Provide battery/health telemetry and end-of-life warning; validate endurance at −30 °C.
CFB-002 — "We got no alert even though there was an excursion"
Reported issue | Delayed or missing excursion alert due to no connectivity at the time.
Source type | Anticipated support/complaint theme.
Stakeholder/segment | Distributors, 3PLs, programs (MR-SEG-002/004/005).
Context | D2 alert-latency-connectivity dependence (concept-carried hazard).
Interpreted need theme | T1 — timely alerting; set expectations for offline behaviour.
Priority/frequency signal | Critical (safety-relevant: delayed/missing alert).
Linked MR | MR-PROB-001, MR-JTBD-005; D2; HA-### (delayed/missing alert).
Confidence | Standard/Judgment — latency is connectivity-bound by design.
Notes | Carry as hazard; alert must still be captured locally and surfaced on reconnection/at receipt.
CFB-003 — "Too many false alarms — we turned them off"
Reported issue | Excessive/benign alerts leading to muting and missed real events.
Source type | Anticipated support/complaint theme.
Stakeholder/segment | Distributors, 3PLs, programs, pharmacies.
Context | Alarm tuning.
Interpreted need theme | T7 — trustworthy, low-noise alerting.
Priority/frequency signal | High.
Linked MR | MR-PROB-009.
Confidence | Judgment/Assumption.
Notes | Ties to VOC-022, FLD-010; product-stability-aware thresholds.
CFB-004 — "The record has a gap / looks like it could be edited"
Reported issue | Perceived data gaps or lack of tamper-evidence undermining audit defensibility.
Source type | Anticipated support/QA-escalation theme.
Stakeholder/segment | QA, auditors (MR-PER-001/009).
Context | Part 11 / GDP data integrity.
Interpreted need theme | T5 — complete, tamper-evident, exportable records.
Priority/frequency signal | Critical (compliance gate).
Linked MR | MR-PROB-005, MR-JTBD-008; MR-PER-009.
Confidence | Standard.
Notes | Continuous local logging + signed records + audit trail; gaps must be explicitly flagged, never hidden.
CFB-005 — "Provisioning failed / device wouldn't arm at the dock"
Reported issue | BLE/app provisioning failures or unclear armed state at pack-out.
Source type | Anticipated support theme.
Stakeholder/segment | Pack-out staff, coordinators (MR-PER-002/005).
Context | Provisioning UX and reliability.
Interpreted need theme | T8 — reliable, confirmed provisioning.
Priority/frequency signal | High.
Linked MR | MR-JTBD-006; MR-PER-002/005.
Confidence | Judgment/Assumption.
Notes | On-device fallback provisioning + clear confirmation; ties to FLD-001.
CFB-006 — "Held at customs/airline as a battery/dangerous-goods issue"
Reported issue | Shipment delayed because the Li-ion device tripped a transport-compliance question.
Source type | Anticipated channel/3PL escalation.
Stakeholder/segment | 3PLs, forwarders (MR-PER-004).
Context | D4/D3 — UN 38.3 / IATA DGR.
Interpreted need theme | T12 — clear transport compliance & documentation.
Priority/frequency signal | High (channel-scale blocker).
Linked MR | MR-SEG-004, MR-PER-004; D3/D4.
Confidence | Standard.
Notes | Provide UN 38.3 test summary + transport guidance in the box/onboarding.
CFB-007 — "OLED unreadable / button unresponsive in the freezer"
Reported issue | Display or control degraded at low temperature.
Source type | Anticipated support theme.
Stakeholder/segment | Cold-store operators, last-mile workers (MR-PER-005/007).
Context | Low-temp HMI.
Interpreted need theme | T4/T8 — legible, operable HMI across range.
Priority/frequency signal | Medium-High.
Linked MR | MR-PER-005/007; IEC 60068.
Confidence | Judgment/Assumption.
Notes | OLED low-temp spec to −30 °C; flag HMI engineer.
CFB-008 — "Device didn't come back / can't find it"
Reported issue | Reusable device lost in reverse logistics; location unknown.
Source type | Anticipated fleet-ops theme.
Stakeholder/segment | Fleet admins, distributors (MR-PER-010/011).
Context | D1 reuse; return-rate.
Interpreted need theme | T10/T14 — locate & recover fleet.
Priority/frequency signal | High (erodes reuse ROI).
Linked MR | MR-PROB-011, MR-JTBD-009/010.
Confidence | Judgment/Assumption.
Notes | Last-known-location + return workflow; ties to FLD-07, VOC-006.
CFB-009 — "Backlog dump after reconnect confused our dashboard"
Reported issue | After a long offline dwell, a burst of buffered data uploads out of order / late.
Source type | Anticipated support/integration theme.
Stakeholder/segment | Coordinators, IT (MR-PER-002/003).
Context | D2 store-and-forward backlog handling.
Interpreted need theme | T1/T13 — coherent, time-ordered data after offline dwell.
Priority/frequency signal | Medium.
Linked MR | MR-JTBD-005; MR-PER-003; D2.
Confidence | Judgment/Assumption.
Notes | Timestamp integrity + ordered ingestion; ties to FLD-004.
CFB-010 — "Couldn't integrate with our TMS/WMS / no API"
Reported issue | Data siloed; no integration into logistics/quality systems.
Source type | Anticipated enterprise/IT theme.
Stakeholder/segment | IT admins, 3PLs (MR-PER-003; MR-SEG-004).
Context | Enterprise integration.
Interpreted need theme | T13 — open APIs / integration.
Priority/frequency signal | High for enterprise/3PL.
Linked MR | MR-PER-003; MR-SEG-004; MR-1500 ECO.
Confidence | Judgment/Assumption.
Notes | Open API + export; ties to VOC-018.
CFB-011 — "Firmware update failed / device offline after OTA"
Reported issue | OTA update interrupted (connectivity/power), leaving device in a bad state.
Source type | Anticipated support theme.
Stakeholder/segment | IT, service (MR-PER-003/011).
Context | D2 OTA robustness + security.
Interpreted need theme | T13/T5 — safe, signed, resumable OTA.
Priority/frequency signal | Medium-High.
Linked MR | MR-PER-003; MR-PROB-005; D2.
Confidence | Judgment/Assumption.
Notes | Signed, fail-safe, resumable OTA; flag cybersecurity persona.
CFB-012 — "Calibration expired / can we trust this reading?"
Reported issue | Device out of calibration in circulation; readings questioned.
Source type | Anticipated QA/service theme.
Stakeholder/segment | QA, auditors, service (MR-PER-009/011).
Context | D1 reuse; metrology.
Interpreted need theme | T10/T5 — calibration currency management.
Priority/frequency signal | High (record credibility).
Linked MR | MR-JTBD-007; MR-PER-009/011; ISO/IEC 17025 concepts.
Confidence | Standard/Judgment — interval Evidence Pending.
Notes | Calibration status visible to depot and audit; block deploy if expired.
CFB-013 — "Installer/OEM: doesn't fit our shipper / mounting"
Reported issue (installer/OEM) | Form factor or mounting doesn't integrate into insulated shippers.
Source type | Anticipated OEM/installer feedback.
Stakeholder/segment | Cold-chain equipment OEMs, installers (MR-SEG-008).
Context | Designed-in / embedded monitoring (MR-SEG-008).
Interpreted need theme | Embeddable, standard form factor & mounting.
Priority/frequency signal | Medium-High for OEM channel.
Linked MR | MR-SEG-008; MR-JTBD-006.
Confidence | Judgment/Assumption.
Notes | Fit-inside-shipper form factor; mounting/retention spec; carry to STN (installer).
CFB-014 — "Reseller/SI: hard to onboard/support across regions"
Reported issue (channel) | Provisioning, language, and support tooling insufficient for local partners.
Source type | Anticipated channel-partner feedback.
Stakeholder/segment | Resellers, systems integrators (MR-SEG-009).
Context | Indirect regional reach.
Interpreted need theme | Partner enablement, localization, integration.
Priority/frequency signal | Medium.
Linked MR | MR-SEG-009; MR-1500 ECO.
Confidence | Judgment/Assumption.
Notes | Open APIs + BLE/mobile provisioning + localized docs enable integrators.
CFB-015 — "Clinical/operational: accept/reject rule didn't match our policy"
Reported issue | Device verdict didn't align with the program's product-specific stability policy.
Source type | Anticipated clinical/operational feedback.
Stakeholder/segment | EPI managers, QA (MR-PER-006/001).
Context | Configurable, product-specific thresholds vs. fixed rules.
Interpreted need theme | T4/T7 — configurable, product-relevant accept/reject logic.
Priority/frequency signal | High.
Linked MR | MR-JTBD-004; MR-PROB-009; MR-PER-006.
Confidence | Judgment/Assumption.
Notes | Thresholds per vaccine/stability profile; PQS-consistent; coordinate regulatory.
CFB-016 — "USB-C charging behaviour unclear / slow to ready after cold"
Reported issue | Charging state / readiness ambiguous, especially post-cold (charge-inhibit).
Source type | Anticipated service/support theme.
Stakeholder/segment | Service technicians, ops (MR-PER-011).
Context | D3 charge management; USB-C charging.
Interpreted need theme | T10/T3 — clear charge/readiness state.
Priority/frequency signal | Medium-High.
Linked MR | D3; MR-JTBD-007; MR-PER-011.
Confidence | Standard/Judgment.
Notes | Ties to FLD-008/CFB-001; readiness indication at depot.
---
Backward-trace summary (this file)
Item | Theme | MR anchors | → STN
FLD-001 | T8 | MR-JTBD-006, MR-PER-002/005 | STN-008
FLD-002 | T2/T8 | MR-PROB-002/009, MR-PER-005 | STN-002, STN-008
FLD-003 | T8 | MR-PER-005, IEC 60068 | STN-008
FLD-004 | T1/T3/T9 | MR-PROB-007/008, D2/D3, MR-RISK-001 | STN-001, STN-003, STN-009
FLD-005 | T1/T9/T14 | MR-JTBD-005/010, concept §7 | STN-001, STN-009, STN-014
FLD-006 | T4 | MR-JTBD-004, MR-PER-007 | STN-004
FLD-007 | T10 | MR-PROB-011, MR-JTBD-009/010 | STN-010, STN-016
FLD-008 | T10/T3 | D3, MR-JTBD-007, MR-PER-011 | STN-010, STN-003
FLD-009 | T6 | MR-PROB-006, MR-JTBD-003 | STN-006
FLD-010 | T7 | MR-PROB-009 | STN-007
FLD-011 | T10 | MR-JTBD-009, D1, IEC 60529 | STN-010, STN-016
FLD-012 | T14 | MR-JTBD-010, D3, concept §7 | STN-014, STN-015 (conflict)
CFB-001 | T3 | MR-PROB-007, D3, MR-RISK-001 | STN-003
CFB-002 | T1 | MR-PROB-001, MR-JTBD-005, D2 | STN-001, STN-018
CFB-003 | T7 | MR-PROB-009 | STN-007
CFB-004 | T5 | MR-PROB-005, MR-JTBD-008 | STN-005
CFB-005 | T8 | MR-JTBD-006 | STN-008
CFB-006 | T12 | MR-SEG-004, D3/D4 | STN-012
CFB-007 | T4/T8 | MR-PER-005/007, IEC 60068 | STN-004, STN-008
CFB-008 | T10/T14 | MR-PROB-011, MR-JTBD-010 | STN-010, STN-014
CFB-009 | T1/T13 | MR-JTBD-005, MR-PER-003, D2 | STN-001, STN-013
CFB-010 | T13 | MR-PER-003, MR-SEG-004 | STN-013
CFB-011 | T13/T5 | MR-PER-003, MR-PROB-005, D2 | STN-013, STN-005
CFB-012 | T10/T5 | MR-JTBD-007, MR-PER-009/011 | STN-010, STN-005
CFB-013 | (OEM fit) | MR-SEG-008, MR-JTBD-006 | STN-017
CFB-014 | (channel) | MR-SEG-009 | STN-017
CFB-015 | T4/T7 | MR-JTBD-004, MR-PER-006 | STN-004, STN-007
CFB-016 | T10/T3 | D3, MR-JTBD-007 | STN-010, STN-003
VoC — Depth-Interview Findings (VOC-001 … VOC-024)
Evidence note (read index.md §2): every "verbatim" below is representative synthesized VoC, not
an attributed real quotation. Confidence and Notes state the basis and the study that would confirm it.
Backward trace to MR-…; forward trace to STN-### (folder 03) and UN-### (Stage 4). Governance: ORCH-018 / EM-ENG-012.
Record fields: ID · Verbatim/paraphrased · Source type · Stakeholder/segment · Context · Interpreted need theme · Priority/frequency signal · Linked Market Requirement · Confidence · Notes.
---
A. Pharma manufacturers & QA / regulatory affairs (MR-SEG-001, MR-PER-001/003/009)
VOC-001 — "Our name is on the vial; the data has to survive an inspector"
Verbatim (synthesized) | "When there's an excursion, I'm the one who has to stand in front of an FDA or MHRA inspector and defend the disposition. If the temperature record has gaps, or looks editable, the whole batch decision is in question. I need a record I can hand over and not flinch."
Source type | Depth interview (synthesized) — QA / Responsible Person.
Stakeholder/segment | Pharma / vaccine manufacturer QA (MR-SEG-001; persona MR-PER-001, MR-PER-009).
Context | Batch-disposition / release defense in a regulated (EU GDP / FDA Part 11) market.
Interpreted need theme | T5 — audit-grade, tamper-evident, complete records.
Priority/frequency signal | Very high; raised by every regulated-market quality persona. Purchase gate.
Linked Market Requirement | MR-PROB-005, MR-JTBD-008, MR-JTBD-003; personas MR-PER-001/009.
Confidence | Standard (rests on EU GDP + 21 CFR Part 11) — theme High; exact acceptance workflow Evidence Pending.
Notes | Confirm via audit-simulation with QA personas. Ties to MR-EXEC-005 data-integrity strategy.
VOC-002 — "Provisioning cannot become the bottleneck at dispatch"
Verbatim (synthesized) | "I love the idea of better data, but if arming the device adds thirty seconds per box at pack-out, my dispatch team will quietly go back to the old loggers — or not start it at all."
Source type | Depth interview (synthesized) — QA manager / dispatch owner.
Stakeholder/segment | Manufacturer / distributor QA & dispatch (MR-PER-001, MR-PER-002).
Context | High-throughput pack-out; audit rigor vs. dispatch speed tension.
Interpreted need theme | T8 — frictionless provisioning; covert requirement that rigor not slow dispatch.
Priority/frequency signal | High; recurring covert requirement behind adoption.
Linked Market Requirement | MR-JTBD-006; personas MR-PER-001/002.
Confidence | Judgment/Assumption — based on known logger failure modes ("forgot to start"); confirm via contextual inquiry.
Notes | Provisioning-time target is Evidence Pending; MR-JTBD-006 states "well under a minute."
VOC-003 — "It has to work in every market we ship into, or it's useless to me"
Verbatim (synthesized) | "We ship the same product to the EU, the US, and through GAVI channels. I'm not going to run three different monitoring systems. One device, defensible everywhere."
Source type | Depth interview (synthesized) — global QA / supply-chain quality.
Stakeholder/segment | Multinational manufacturer (MR-SEG-001, MR-SEG-010).
Context | Multi-market distribution; D4 multi-market regulatory posture.
Interpreted need theme | T5 + T11 — one audit-grade, multi-market-compliant record set.
Priority/frequency signal | High among multinationals; standard-setting leverage (anchor accounts).
Linked Market Requirement | MR-SEG-001, MR-SEG-010; D4; MR-REG (WHO PQS / EU GDP / US FDA).
Confidence | Standard/Judgment — multi-market intent based on D4; exact per-market claims Evidence Pending.
Notes | Reinforces anchor-account strategy; manufacturers can mandate SCM downstream.
B. Distributors / wholesalers (MR-SEG-002, MR-PER-002)
VOC-004 — "We absorb the loss when we can't show where the chain broke"
Verbatim (synthesized) | "A shipment arrives warm, the customer refuses it, and the argument starts. Was it us, the airline, the last handler? Without hard evidence we usually just absorb the loss to keep the relationship. That's real money every month."
Source type | Depth interview (synthesized) — distributor operations / QA.
Stakeholder/segment | Pharma distributor / wholesaler (MR-SEG-002).
Context | Multi-party custody hand-offs; dispute and liability allocation.
Interpreted need theme | T6 — attributable, time/location-stamped evidence.
Priority/frequency signal | High; strong ROI lever cited by distributors and 3PLs alike.
Linked Market Requirement | MR-PROB-006, MR-JTBD-003; segment MR-SEG-002.
Confidence | Judgment/Assumption — dispute burden widely reported; per-firm magnitude Evidence Pending.
Notes | GPS + tamper-evident timestamped logs support attribution. Quantify write-off cost in survey.
VOC-005 — "I manage 200 boxes a day; tell me only what needs action"
Verbatim (synthesized) | "I don't want to read temperature graphs. I want a list of the shipments that have a problem right now, and everything else to stay quiet."
Source type | Depth interview (synthesized) — distributor logistics coordinator.
Stakeholder/segment | Distributor logistics coordinator (MR-PER-002).
Context | Concurrent high-volume shipment management; dashboard/alert design.
Interpreted need theme | T7 — action-only, role-appropriate alerting; no alarm fatigue.
Priority/frequency signal | High; daily active user whose adoption drives utilization/ROI.
Linked Market Requirement | MR-PROB-009; persona MR-PER-002; MR-JTBD-005.
Confidence | Judgment/Assumption — based on GDP dispatch workflow; confirm via contextual inquiry.
Notes | Alert fatigue is the failure mode; needs configurable, product-relevant thresholds.
VOC-006 — "Reusable only pays if we actually get them back"
Verbatim (synthesized) | "The cost-per-trip story only works if the devices come home. On domestic lanes we can manage returns; on some export and last-mile routes, honestly, we'd write half of them off."
Source type | Depth interview (synthesized) — distributor operations / finance.
Stakeholder/segment | Distributor / 3PL running a reusable fleet (MR-SEG-002/004).
Context | D1 reusable model; reverse-logistics economics.
Interpreted need theme | T10 — reverse-logistics recovery & fleet integrity.
Priority/frequency signal | High; determines whether the TCO/business case survives.
Linked Market Requirement | MR-PROB-011, MR-JTBD-009/010; D1; MR-EXEC-008 investment case.
Confidence | Judgment/Assumption — return-rate economics Evidence Pending.
Notes | GPS/BLE location + clear return workflow mitigate; measure return rate in pilot.
C. 3PL / logistics / air-cargo (MR-SEG-004, MR-PER-004)
VOC-007 — "It can't slow the dock or trip a battery-shipping rule"
Verbatim (synthesized) | "We move thousands of pieces an hour. The monitor has to be invisible to my handlers, safe to fly, and never the reason a shipment gets held for dangerous-goods paperwork."
Source type | Depth interview (synthesized) — 3PL operations lead.
Stakeholder/segment | Pharma 3PL / air-cargo forwarder (MR-SEG-004; persona MR-PER-004).
Context | High-throughput handling; air transport of a Li-ion device.
Interpreted need theme | T12 — transport/battery compliance + handling neutrality.
Priority/frequency signal | High; table-stakes gate for 3PL SOP inclusion; channel-scale lever.
Linked Market Requirement | MR-SEG-004; persona MR-PER-004; D3/D4 (UN 38.3, IATA DGR).
Confidence | Standard — UN 38.3 / IATA DGR govern; specific handling steps Evidence Pending.
Notes | Li-ion transport classification (D3) is decisive; coordinate battery + regulatory engineers.
VOC-008 — "Real-time is why we'd choose you over a USB logger"
Verbatim (synthesized) | "A logger someone downloads after the fact tells me what I already lost. If your device pings me mid-flight or on a long dwell so I can reroute or re-ice, that's worth paying for."
Source type | Depth interview (synthesized) — 3PL control-tower / cold-chain manager.
Stakeholder/segment | 3PL logistics control tower (MR-SEG-004; MR-JTBD-005).
Context | Long-haul / long-dwell legs; in-transit intervention.
Interpreted need theme | T1 — real-time excursion warning in time to act.
Priority/frequency signal | High; primary differentiator vs. single-use loggers; supports premium price.
Linked Market Requirement | MR-PROB-001, MR-JTBD-005; competitors MR-COMP (Controlant/Tive/Roambee-class).
Confidence | Judgment/Assumption + Standard (D2 immediate excursion alert); latency is connectivity-bound.
Notes | Carry delayed/missing-alert as a hazard (HA-###). Alert latency depends on LTE-M coverage.
D. Immunization / EPI program managers & MoH (MR-SEG-005, MR-PER-006)
VOC-009 — "I need a clear go/no-go, not a data dump"
Verbatim (synthesized) | "When a consignment lands, my staff must decide accept or reject fast. A graph doesn't help them. I need an unambiguous verdict that lines up with what PQS expects."
Source type | Depth interview (synthesized) — national EPI / immunization program manager.
Stakeholder/segment | National immunization program / MoH (MR-SEG-005; persona MR-PER-006).
Context | Consignment arrival decision in a PQS-governed program.
Interpreted need theme | T4 — unambiguous, PQS-aligned accept/reject verdict.
Priority/frequency signal | Critical; decisive in program procurement.
Linked Market Requirement | MR-JTBD-004, MR-PROB-004; persona MR-PER-006; MR-REG (WHO PQS E006, code TBC).
Confidence | Standard — WHO PQS decision intent; exact rule-set / category code Evidence Pending.
Notes | Accept/reject thresholds must be configurable and PQS-consistent; coordinate with regulatory.
VOC-010 — "It has to work unattended in tough field conditions"
Verbatim (synthesized) | "Our conditions are harsh — heat, dust, no reliable power, staff who rotate. Whatever we deploy has to keep working for weeks without anyone babysitting it."
Source type | Depth interview (synthesized) — EPI logistics / cold-chain officer.
Stakeholder/segment | National immunization program (MR-SEG-005, MR-SEG-011).
Context | LMIC field deployment; low maintenance, ruggedness.
Interpreted need theme | T3 + T9 — endurance and field robustness with minimal intervention.
Priority/frequency signal | Critical for LMIC adoption; ruggedness + low-temp + PQS is a rare combination.
Linked Market Requirement | MR-SEG-005/011, MR-PROB-007/008; D3; IEC 60068 / 60529.
Confidence | Cited (WHO/UNICEF/GAVI last-mile reporting, reported/estimated) + Standard.
Notes | 90-day endurance at −30…+60 °C is the top engineering tension (D3, MR-RISK-001). Validate by test.
VOC-011 — "Freezing quietly ruins stock and nobody sees it"
Verbatim (synthesized) | "Everyone worries about heat, but our bigger silent problem is freezing. A vial can look perfect and already be spoiled. Our indicators don't catch it."
Source type | Depth interview (synthesized) — EPI cold-chain manager.
Stakeholder/segment | Immunization program / distributor (MR-SEG-005, MR-SEG-002).
Context | Freeze-sensitive vaccines; over-refrigeration in last-mile fridges.
Interpreted need theme | T2 — two-sided (freeze + heat) detection.
Priority/frequency signal | Critical; large under-served loss category; VVMs do not signal freezing.
Linked Market Requirement | MR-PROB-002, MR-PROB-004; MR-JTBD-001.
Confidence | Cited (WHO/UNICEF freezing-exposure reporting, reported/estimated).
Notes | Internal + external-probe temperature supports freeze detection; confirm probe placement in UN/PRD.
E. Global procurement / GAVI / UNICEF (MR-SEG-006, MR-PER-008)
VOC-012 — "If it's not PQS-listed and cost-effective, we can't buy it"
Verbatim (synthesized) | "Our hands are tied by procurement rules. Show me the PQS listing and the cost-per-trip over the device's life, and demonstrate the public-health impact, or you're not on the tender."
Source type | Depth interview (synthesized) — pooled-procurement officer.
Stakeholder/segment | GAVI / UNICEF Supply Division / WHO procurement (MR-SEG-006; persona MR-PER-008).
Context | Pooled, tendered, donor-funded procurement for LMICs.
Interpreted need theme | T11 — PQS eligibility + TCO/cost-per-trip + demonstrable impact.
Priority/frequency signal | Critical; catalogue inclusion unlocks the largest aggregated volume channel.
Linked Market Requirement | MR-SEG-006, MR-PER-008; MR-REG (WHO PQS); MR-EXEC-008 (reuse TCO).
Confidence | Standard/Cited — PQS + UNICEF/GAVI frameworks; exact category & TCO figures Evidence Pending.
Notes | Long procurement cycles; plan PQS listing early. Document reuse economics as cost-per-trip.
VOC-013 — "Total cost, not sticker price — but our budget is still a hard wall"
Verbatim (synthesized) | "I understand reusable can be cheaper per trip. But donor budgets are fixed, and a rugged low-temperature device that lasts 90 days at −30 will not be cheap. Prove the lifetime math to me."
Source type | Depth interview (synthesized) — procurement / donor programs.
Stakeholder/segment | Global procurement body (MR-SEG-006).
Context | Budget-constrained volume buying; cost vs. capability tension.
Interpreted need theme | T11 — TCO justification against a hard affordability ceiling.
Priority/frequency signal | Critical; a core stakeholder conflict (low cost vs. −30 °C / 90-day endurance).
Linked Market Requirement | MR-SEG-006/011, MR-PROB-010; MR-800 (COMM); D3.
Confidence | Judgment/Assumption — price/TCO figures Evidence Pending.
Notes | Carry as STN conflict (STN-011 vs STN-003). Feeds MR-RISK-001 cost/endurance tension.
F. Clinic / last-mile point-of-use (MR-SEG-007, MR-PER-007)
VOC-014 — "No signal, no phone — the box itself must tell me"
Verbatim (synthesized) | "Out here I don't have a smartphone or a network. When the vaccine arrives I look at the box. If it can tell me green or red right on the screen, I can do my job. If I need an app, I can't."
Source type | Depth interview / ride-along (synthesized) — last-mile nurse / CHW.
Stakeholder/segment | Last-mile health worker (MR-SEG-007; persona MR-PER-007).
Context | Rural health post, intermittent/no connectivity, minimal device training.
Interpreted need theme | T4 — offline, glanceable on-device accept/reject status.
Priority/frequency signal | Critical for last-mile trust and program scale.
Linked Market Requirement | MR-JTBD-004, MR-PROB-004/008; persona MR-PER-007; MR-1300 ENV (rural post).
Confidence | Standard/Judgment — WHO last-mile guidance; literacy/training assumptions Evidence Pending.
Notes | OLED + button is the sole interface here; low-literacy, offline-legible design essential.
VOC-015 — "Tell me at the door if I can use this stock"
Verbatim (synthesized) | "I'm not trying to become a data analyst. I just need to know, at receipt, whether to put this in the fridge, quarantine it, or send it back."
Source type | Depth interview (synthesized) — hospital / community pharmacy receiver.
Stakeholder/segment | Hospital / pharmacy point-of-use (MR-SEG-007).
Context | Receipt-check workflow at a health facility.
Interpreted need theme | T4 — fast, unambiguous accept/quarantine/reject at receipt.
Priority/frequency signal | Medium-High; reinforces end-to-end chain-of-custody confidence.
Linked Market Requirement | MR-SEG-007, MR-JTBD-002/004.
Confidence | Judgment/Assumption — receiver behaviour inferred; confirm via receipt-workflow observation.
Notes | OLED + button + BLE support a quick at-receipt status check.
G. Cold-chain technicians / field service (MR-PER-011)
VOC-016 — "I need to clean, recharge, recal, and redeploy fast and verifiably"
Verbatim (synthesized) | "When a device comes back, I have to turn it around quickly. Is it charged? Is it still in calibration? Is the firmware current? If checking each one is manual, I can't keep up with the fleet."
Source type | Depth interview (synthesized) — field-service / calibration technician.
Stakeholder/segment | Service depot / pooling operator (persona MR-PER-011).
Context | Reverse-logistics turnaround; reuse sustainment (D1).
Interpreted need theme | T10 — serviceable, calibration-manageable, fast-turnaround fleet.
Priority/frequency signal | High; turnaround directly sets device utilization and margin.
Linked Market Requirement | MR-JTBD-007/009, MR-PROB-011; persona MR-PER-011; ISO/IEC 17025 concepts.
Confidence | Standard/Judgment — calibration traceability evidence-backed; throughput figures Evidence Pending.
Notes | Charge-inhibit-below-0 °C (D3) and cleaning-between-trips intersect battery + materials engineering.
VOC-017 — "Charging inhibited in the cold surprises the depot"
Verbatim (synthesized) | "If devices come back frozen and won't take a charge until they warm up, my turnaround plan breaks. I need the device — or the dock — to tell me clearly what's happening."
Source type | Depth interview (synthesized) — service depot lead.
Stakeholder/segment | Service / reverse logistics (persona MR-PER-011).
Context | D3 charge-below-0 °C inhibit; depot workflow.
Interpreted need theme | T10 + T3 — predictable charge/readiness behaviour at temperature.
Priority/frequency signal | Medium-High; operational surprise that erodes reuse economics if unmanaged.
Linked Market Requirement | D3; MR-JTBD-007; persona MR-PER-011.
Confidence | Standard/Judgment — low-temp charge hazard well established (IEC 62368-1 context).
Notes | Self-heat-before-charge is under evaluation (D3 residual). Carry to User/Service requirements.
H. IT / security & data governance (MR-PER-003)
VOC-018 — "It must pass our CSV and SSO/access policy"
Verbatim (synthesized) | "Before this touches our network it goes through computerized-system validation, single sign-on, role-based access, and data-retention review. If it can't clear those, it never deploys, no matter how good the sensor is."
Source type | Depth interview (synthesized) — enterprise IT / quality-systems admin.
Stakeholder/segment | Enterprise pharma/distributor IT (persona MR-PER-003).
Context | Platform deployment gate; Part 11 / GAMP 5 / data governance.
Interpreted need theme | T13 — enterprise data governance, validation, and integration.
Priority/frequency signal | High; hard gate for enterprise rollout and platform expansion.
Linked Market Requirement | persona MR-PER-003; MR-PROB-005; MR-1500 ECO; Standard (21 CFR Part 11, GAMP 5, ALCOA+).
Confidence | Standard — validation/data-governance obligations evidence-backed; integration specifics Evidence Pending.
Notes | Coordinate with regulatory auditor (MR-PER-009) and cybersecurity persona in PRD stage.
VOC-019 — "If it's connected, prove it can't be tampered with"
Verbatim (synthesized) | "A cellular device on our shipments is an attack surface and a data-integrity question. Show me the record can't be altered, the firmware updates are signed, and access is controlled."
Source type | Depth interview (synthesized) — IT security / data-integrity owner.
Stakeholder/segment | IT / security admin (persona MR-PER-003).
Context | OTA + cellular connectivity security; tamper-evidence.
Interpreted need theme | T5 + T13 — tamper-evident data and secure OTA/access.
Priority/frequency signal | High in regulated enterprise accounts.
Linked Market Requirement | MR-PROB-005, MR-JTBD-003; persona MR-PER-003; D2 (OTA).
Confidence | Standard/Judgment — data-integrity + secure-update expectations evidence-backed; threat model Evidence Pending.
Notes | Flag EM cybersecurity persona for PRD/hazard stages; ties to signed-OTA and audit-trail requirements.
I. Executive / supply-chain leadership (MR-PER-010)
VOC-020 — "Show me network-wide loss reduction and cost-per-trip"
Verbatim (synthesized) | "I don't care about one device. I care about how much cold-chain loss we cut across the network, how our compliance exposure drops, and what each trip costs. Give me that and I'll fund standardization."
Source type | Depth interview (synthesized) — VP/Director Supply Chain.
Stakeholder/segment | Executive / supply-chain director (persona MR-PER-010).
Context | Fleet-scale, multi-year budget and standardization decision.
Interpreted need theme | T14 — fleet-level ROI, loss reduction, and cost-per-trip visibility.
Priority/frequency signal | High; executive sponsorship anchors large multi-year contracts.
Linked Market Requirement | MR-JTBD-010; persona MR-PER-010; MR-1000 (MET); MR-EXEC-008.
Confidence | Judgment/Assumption — ROI figures Evidence Pending.
Notes | Consumes fleet-management analytics (MR-1500 ECO); distinct from QA-manager needs.
J. Cross-cutting / whole-chain themes
VOC-021 — "We find out at delivery, and then it's too late"
Verbatim (synthesized) | "By the time we open the box and read the logger, the damage is done. We're documenting a loss, not preventing one."
Source type | Depth interview (synthesized) — distributor / program (multi-segment).
Stakeholder/segment | Distributors, 3PLs, immunization programs (MR-SEG-002/004/005).
Context | Legacy offline-logger workflow; the core pain the SCM exists to remove.
Interpreted need theme | T1 — real-time detection replacing after-the-fact discovery.
Priority/frequency signal | Critical; the central value proposition; near-universal.
Linked Market Requirement | MR-PROB-001, MR-JTBD-001/005; D2.
Confidence | Cited (WHO/UNICEF out-of-range exposure reporting, reported/estimated).
Notes | If unmet, no reason to switch from cheaper loggers (MR-PROB-001 risk-if-unmet).
VOC-022 — "So many false alerts that we stopped looking"
Verbatim (synthesized) | "Our last system alerted on everything — every door-open, every truck transfer. After a while everyone muted it, and an important one slipped past."
Source type | Depth interview (synthesized) — distributor / program operations.
Stakeholder/segment | Distributors, 3PLs, programs, pharmacies (MR-SEG-002/004/005/007).
Context | Alarm-management; benign transients vs. product-relevant excursions.
Interpreted need theme | T7 — accurate, product-relevant alerting; no alarm fatigue.
Priority/frequency signal | High; directly determines whether alerts are acted on.
Linked Market Requirement | MR-PROB-009, MR-JTBD-005.
Confidence | Judgment/Assumption — alarm fatigue well documented across monitoring domains; SCM-specific data Evidence Pending.
Notes | Thresholds should reflect product stability profiles, not raw excursions.
VOC-023 — "The last mile is a black box"
Verbatim (synthesized) | "We have good visibility until the regional store. After that — motorbikes, cold boxes, remote posts — we lose visibility, and that's exactly where product is most at risk."
Source type | Depth interview (synthesized) — program / procurement.
Stakeholder/segment | Immunization programs, procurement bodies (MR-SEG-005/006/011).
Context | LMIC last-mile; connectivity/power/infrastructure gaps.
Interpreted need theme | T9 — last-mile visibility under intermittent connectivity.
Priority/frequency signal | Critical; highest-loss, lowest-visibility leg; unlocks donor channels.
Linked Market Requirement | MR-PROB-008, MR-JTBD-001; D2 (store-and-forward).
Confidence | Cited (WHO/UNICEF/GAVI last-mile reporting, reported/estimated).
Notes | Store-and-forward with LTE-M primary + BLE/Wi-Fi offload addresses intermittency.
VOC-024 — "The last logger gave out in the cold and we lost the record"
Verbatim (synthesized) | "On a deep-cold lane last winter the device stopped working. No data, no alert, and no way to prove the shipment was fine. We'd like to avoid that going forward."
Source type | Depth interview (synthesized) — 3PL / distributor cold-lane operator.
Stakeholder/segment | All segments shipping cold lanes (MR-SEG-002/004/005).
Context | Cold-weather battery failure of incumbent connected loggers.
Interpreted need theme | T3 — cold endurance; battery survives the full range and duration.
Priority/frequency signal | Critical; retention-critical reliability; recurring incumbent failure.
Linked Market Requirement | MR-PROB-007; D3 (low-temp Li-ion, duty-cycled); MR-RISK-001.
Confidence | Standard/Judgment — low-temp Li-ion behaviour established; field-failure rates Evidence Pending.
Notes | Top engineering tension; closure-by-test in MR-VAL-006. Validate 90-day at −30 °C endurance.
---
Backward-trace summary (this file)
VOC | Theme | MR anchors | → STN
VOC-001 | T5 | MR-PROB-005, MR-JTBD-008/003, MR-PER-001/009 | STN-005
VOC-002 | T8 | MR-JTBD-006, MR-PER-001/002 | STN-008
VOC-003 | T5/T11 | MR-SEG-001/010, D4, MR-REG | STN-005, STN-011
VOC-004 | T6 | MR-PROB-006, MR-JTBD-003 | STN-006
VOC-005 | T7 | MR-PROB-009, MR-PER-002 | STN-007
VOC-006 | T10 | MR-PROB-011, MR-JTBD-009/010, D1 | STN-010, STN-016
VOC-007 | T12 | MR-SEG-004, MR-PER-004, D3/D4 | STN-012
VOC-008 | T1 | MR-PROB-001, MR-JTBD-005 | STN-001
VOC-009 | T4 | MR-JTBD-004, MR-PER-006, MR-REG | STN-004
VOC-010 | T3/T9 | MR-SEG-005/011, MR-PROB-007/008, D3 | STN-003, STN-009
VOC-011 | T2 | MR-PROB-002/004 | STN-002
VOC-012 | T11 | MR-SEG-006, MR-PER-008, MR-REG | STN-011
VOC-013 | T11 | MR-SEG-006/011, MR-800, D3 | STN-011 (conflict w/ STN-003)
VOC-014 | T4 | MR-JTBD-004, MR-PER-007 | STN-004
VOC-015 | T4 | MR-SEG-007, MR-JTBD-004 | STN-004
VOC-016 | T10 | MR-JTBD-007/009, MR-PER-011 | STN-010, STN-016
VOC-017 | T10/T3 | D3, MR-JTBD-007, MR-PER-011 | STN-010, STN-003
VOC-018 | T13 | MR-PER-003, MR-PROB-005 | STN-013
VOC-019 | T5/T13 | MR-PROB-005, MR-PER-003, D2 | STN-005, STN-013
VOC-020 | T14 | MR-JTBD-010, MR-PER-010 | STN-014
VOC-021 | T1 | MR-PROB-001, MR-JTBD-001/005 | STN-001
VOC-022 | T7 | MR-PROB-009 | STN-007
VOC-023 | T9 | MR-PROB-008, D2 | STN-009
VOC-024 | T3 | MR-PROB-007, D3, MR-RISK-001 | STN-003
VoC — Survey Themes & Lost-Deal / No-Decision Reasons
Evidence note (read index.md §2): the survey themes (VOC-050…) are projected importance/satisfaction
hypotheses, not results from a fielded survey; the lost-deal reasons (LD-001…) are anticipated loss
patterns synthesized from the concept's tensions and competitive landscape (MR-COMP: Controlant / Tive /
Roambee-class real-time trackers, plus incumbent single-use loggers and VVMs). All are Evidence Pending
until the recommended VoC/survey program and real pipeline data exist. Backward trace to MR-…; forward to
STN-### / UN-###. Governance: ORCH-018 / EM-ENG-012.
---
Part 1 — Survey themes (importance vs. satisfaction with today's solutions) — VOC-050 … VOC-064
Method (planned): importance–satisfaction survey across segments (manufacturers, distributors, 3PLs, immunization programs, procurement, point-of-use, service). Opportunity = Importance + max(Importance − Satisfaction, 0) with today's tools. Scores below are hypothesized 1–5 (Judgment/Assumption) to prioritize the field program — not measured data.
Record fields: ID · Survey theme (attribute) · Hypothesized importance · Hypothesized satisfaction w/ today · Interpreted opportunity · Linked MR · Confidence · Notes.
VOC-050 — Real-time excursion alerting (in time to act)
Importance (hyp.) | 5 (Critical) · Satisfaction w/ today (hyp.): 2 (single-use loggers can't) · Opportunity: Very high (underserved).
Linked MR | MR-PROB-001, MR-JTBD-005. · Confidence: Judgment/Assumption. · → STN-001.
Notes | Core differentiator vs. offline loggers; the single strongest expected gap.
VOC-051 — Freeze (two-sided) detection
Importance | 5 · Satisfaction: 2 (VVMs heat-only) · Opportunity: Very high.
Linked MR | MR-PROB-002/004. · Confidence: Cited theme (WHO/UNICEF freezing). · → STN-002.
VOC-052 — Cold-weather endurance / battery survives the trip
Importance | 5 · Satisfaction: 2 (incumbents fail in cold) · Opportunity: Very high.
Linked MR | MR-PROB-007; D3; MR-RISK-001. · Confidence: Judgment/Assumption. · → STN-003.
VOC-053 — Instant offline accept/reject at the door
Importance | 5 (last-mile) · Satisfaction: 3 (VVM good but coarse/heat-only) · Opportunity: High.
Linked MR | MR-JTBD-004; MR-PER-006/007. · Confidence: Standard/Judgment. · → STN-004.
VOC-054 — Audit-grade tamper-evident data
Importance | 5 (regulated) · Satisfaction: 3 (loggers exportable but disputable/gappy) · Opportunity: High.
Linked MR | MR-PROB-005; MR-JTBD-008; MR-PER-009. · Confidence: Standard. · → STN-005.
VOC-055 — Attributable evidence (who broke the chain)
Importance | 4 · Satisfaction: 2 · Opportunity: High.
Linked MR | MR-PROB-006; MR-JTBD-003. · Confidence: Judgment/Assumption. · → STN-006.
VOC-056 — Alert accuracy / low false-alarm rate
Importance | 4 · Satisfaction: 2 (alarm fatigue common) · Opportunity: High.
Linked MR | MR-PROB-009. · Confidence: Judgment/Assumption. · → STN-007.
VOC-057 — Fast, error-proof provisioning
Importance | 4 · Satisfaction: 2 ("forgot to start") · Opportunity: High.
Linked MR | MR-JTBD-006; MR-PER-002/005. · Confidence: Judgment/Assumption. · → STN-008.
VOC-058 — Last-mile / intermittent-connectivity resilience
Importance | 5 (LMIC) · Satisfaction: 2 · Opportunity: Very high.
Linked MR | MR-PROB-008; D2. · Confidence: Cited theme. · → STN-009.
VOC-059 — Reuse economics (recover, clean, recal, redeploy)
Importance | 4 · Satisfaction: 2 (single-use avoids it; reusable hard) · Opportunity: High.
Linked MR | MR-PROB-011; MR-JTBD-007/009; D1. · Confidence: Judgment/Assumption. · → STN-010, STN-016.
VOC-060 — PQS eligibility + total cost of ownership
Importance | 5 (procurement) · Satisfaction: 3 (few PQS-listed reusables) · Opportunity: High.
Linked MR | MR-SEG-005/006; MR-PER-008; MR-REG. · Confidence: Standard/Cited. · → STN-011.
VOC-061 — Transport / battery compliance (air & ground)
Importance | 4 (3PL gate) · Satisfaction: 3 · Opportunity: Medium-High.
Linked MR | MR-SEG-004; MR-PER-004; D3/D4. · Confidence: Standard. · → STN-012.
VOC-062 — Enterprise data governance & integration (CSV/SSO/API)
Importance | 4 (enterprise) · Satisfaction: 3 · Opportunity: Medium-High.
Linked MR | MR-PER-003; MR-1500 ECO. · Confidence: Standard/Judgment. · → STN-013.
VOC-063 — Fleet visibility & ROI reporting (executive)
Importance | 4 · Satisfaction: 2 · Opportunity: High.
Linked MR | MR-JTBD-010; MR-PER-010; MR-1000. · Confidence: Judgment/Assumption. · → STN-014.
VOC-064 — Affordable purchase price within donor budgets
Importance | 5 (LMIC/procurement) · Satisfaction: 3 (VVMs/loggers cheap) · Opportunity: High — and a conflict.
Linked MR | MR-SEG-006/011; MR-800; D3. · Confidence: Judgment/Assumption (price Evidence Pending).
Notes | Directly conflicts with VOC-052 (−30 °C/90-day endurance) — carry to STN conflict register (STN-011 vs STN-003).
Survey opportunity roll-up (hypothesized): the largest expected gaps (Importance high, Satisfaction low) are VOC-050 (real-time alerting), VOC-051 (freeze detection), VOC-052 (cold endurance), and VOC-058 (last-mile resilience) — i.e., the SCM's core reason to exist. These become the top-priority Stakeholder Needs and must be the first to be validated and, later, verified.
---
Part 2 — Lost-deal / no-decision reasons (LD-001 … LD-010)
Method (planned): capture reasons deals are lost or stall, from pipeline reviews and win/loss interviews. Below are anticipated loss patterns (Judgment/Assumption) to pre-empt in design and go-to-market.
Record fields: ID · Loss reason (paraphrased) · Segment · Competing alternative · Interpreted need/gap · Linked MR · Confidence · Notes.
LD-001 — "Not PQS-listed yet, so we couldn't tender it"
Segment | Procurement / immunization programs (MR-SEG-005/006). · Competing alt: PQS-listed devices / status quo.
Interpreted gap | PQS prequalification is a hard market-access gate (T11).
Linked MR | MR-SEG-006; MR-PER-008; MR-REG. · Confidence: Standard. · → STN-011.
Notes | Plan PQS listing early; long cycles. Losing this excludes the largest LMIC volume channel.
LD-002 — "Too expensive for our budget"
Segment | LMIC programs / procurement (MR-SEG-011/006). · Competing alt: VVMs, single-use loggers.
Interpreted gap | Purchase-price ceiling vs. rugged/low-temp cost (T11 vs T3 conflict).
Linked MR | MR-800; MR-SEG-011; D3. · Confidence: Judgment/Assumption (price Evidence Pending). · → STN-011 (conflict STN-003).
Notes | TCO/cost-per-trip story + possible value/entry SKU (MR-1400). Core conflict.
LD-003 — "We didn't trust it would survive our cold lanes"
Segment | Distributors / 3PLs (MR-SEG-002/004). · Competing alt: Proven incumbent (even if offline).
Interpreted gap | Unproven cold endurance / reliability (T3).
Linked MR | MR-PROB-007; D3; MR-RISK-001; MR-VAL-006. · Confidence: Judgment/Assumption. · → STN-003.
Notes | Needs demonstrable −30 °C / 90-day test evidence to win trust.
LD-004 — "The incumbent real-time tracker was already embedded"
Segment | 3PLs / enterprise (MR-SEG-004). · Competing alt: Controlant / Tive / Roambee-class (MR-COMP).
Interpreted gap | Switching cost + integration lock-in (T13/T14).
Linked MR | MR-COMP; MR-PER-003/010; MR-1500 ECO. · Confidence: Judgment/Assumption. · → STN-013, STN-014.
Notes | Compete on cold endurance, freeze detection, offline verdict, TCO, and open integration.
LD-005 — "Failed our IT/CSV/security review"
Segment | Enterprise pharma/distributor (MR-PER-003). · Competing alt: Validated incumbent platform / no-decision.
Interpreted gap | Enterprise data governance/validation not cleared (T13).
Linked MR | MR-PER-003; MR-PROB-005; Part 11/GAMP 5. · Confidence: Standard. · → STN-013.
Notes | Validation pack, SSO/RBAC, retention, signed OTA must be ready pre-sales.
LD-006 — "No decision — the pilot data got lost / device wasn't started"
Segment | Any (MR-SEG-002/004/005). · Competing alt: Status quo (pilot failed to prove value).
Interpreted gap | Provisioning/void-trip failure destroyed the pilot's evidence (T8).
Linked MR | MR-JTBD-006; MR-PER-002/005. · Confidence: Judgment/Assumption. · → STN-008.
Notes | Error-proof arming + armed-confirmation is a sales-enabling requirement, not just UX.
LD-007 — "3PL wouldn't accept it as air cargo without clear battery paperwork"
Segment | 3PL / forwarder (MR-SEG-004; MR-PER-004). · Competing alt: Ground-only / incumbent with known DG status.
Interpreted gap | Transport-compliance clarity (T12).
Linked MR | MR-SEG-004; D3/D4 (UN 38.3, IATA). · Confidence: Standard. · → STN-012.
Notes | UN 38.3 summary + IATA guidance shipped with the product removes this objection.
LD-008 — "Reverse logistics didn't pencil out on our routes"
Segment | Distributors / programs on hard lanes (MR-SEG-002/011). · Competing alt: Single-use loggers.
Interpreted gap | Return-rate/TCO on export/last-mile routes (T10).
Linked MR | MR-PROB-011; MR-JTBD-009; D1; MR-EXEC-008. · Confidence: Judgment/Assumption. · → STN-010, STN-016.
Notes | Region-specific reusable-vs-single-use strategy; location + return workflow; ties to FLD-007.
LD-009 — "The verdict didn't match our vaccine's stability policy"
Segment | Programs / QA (MR-PER-006/001). · Competing alt: Manual assessment / incumbent with configurable rules.
Interpreted gap | Inflexible accept/reject logic (T4/T7).
Linked MR | MR-JTBD-004; MR-PROB-009; MR-PER-006. · Confidence: Judgment/Assumption. · → STN-004, STN-007.
Notes | Per-product configurable, PQS-consistent thresholds; coordinate regulatory. Ties to CFB-015.
LD-010 — "Connectivity in our region was too poor to rely on live data"
Segment | LMIC programs / remote lanes (MR-SEG-011). · Competing alt: Offline loggers / status quo.
Interpreted gap | Perceived over-reliance on live connectivity (T9).
Linked MR | MR-PROB-008; D2 (store-and-forward). · Confidence: Cited theme. · → STN-009.
Notes | Emphasize local logging + offline verdict + opportunistic upload; live data is a bonus, not a dependency.
---
Backward-trace summary (this file)
Item | Theme | MR anchors | → STN
VOC-050…064 | T1–T14 (+ affordability) | MR-PROB / MR-JTBD / MR-SEG / MR-PER / MR-REG / D1–D4 | STN-001…016
LD-001 | T11 | MR-SEG-006, MR-PER-008, MR-REG | STN-011
LD-002 | T11/T3 | MR-800, MR-SEG-011, D3 | STN-011 (conflict STN-003)
LD-003 | T3 | MR-PROB-007, D3, MR-VAL-006 | STN-003
LD-004 | T13/T14 | MR-COMP, MR-PER-003/010 | STN-013, STN-014
LD-005 | T13 | MR-PER-003, Part 11/GAMP 5 | STN-013
LD-006 | T8 | MR-JTBD-006, MR-PER-002/005 | STN-008
LD-007 | T12 | MR-SEG-004, D3/D4 | STN-012
LD-008 | T10 | MR-PROB-011, MR-JTBD-009, D1 | STN-010, STN-016
LD-009 | T4/T7 | MR-JTBD-004, MR-PER-006 | STN-004, STN-007
LD-010 | T9 | MR-PROB-008, D2 | STN-009

## 3 Stakeholder Needs
Stakeholder Needs
What every group touched by the product needs to succeed — and where those needs pull in different directions.
↩ Contents
In this section:
• Stakeholder Needs Library — Smart Cold Chain Monitor (SCM)
• Stakeholder-Need Conflicts & Trade-offs — SCM
• Stakeholder Needs — Detailed Records (STN-001 … STN-027), by Stakeholder Class
• Stakeholder Register & Influence/Interest Map — SCM
Stakeholder Needs Library — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 03 — Stakeholder Needs · Owner: Principal Systems Engineer + Market Research Director (EM-ENG-001 integrator; EM-ENG-006 Quality/Regulatory for compliance needs; Human-Factors contributor for point-of-use needs) · Convening: consumes Voice of Customer (VOC-###, folder 02) and Market Requirements (MR-…, folder 01); feeds User Needs (Stage 4, ORCH-010) · Traceability anchors: ORCH-018 / EM-ENG-012 · Status: Authored — product-specific, honestly labelled.
---
1. Purpose
This library states, in each stakeholder's own terms, what that stakeholder needs from the SCM program. It sits between the demand-space evidence (Voice of Customer, folder 02) and the solution-neutral User Needs (Stage 4). A stakeholder need answers "who needs what, and why" — it is not yet a product requirement (that is Stage 5, PRD). Needs are written so a competent generalist can follow, jargon defined on first use (per CLAUDE.md rule 4).
It is based on 00-concept/concept-and-decisions.md (decisions D1–D4), the Master Market Requirements Library (01-market-requirements/), and the Voice-of-Customer library (02-voice-of-customer/).
2. Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, each need is tagged by its evidentiary basis. Because the primary VoC field program is not yet executed (see folder 02, §2), needs that rest on synthesized VoC are labelled Judgment/Assumption or Evidence Pending; needs that rest on a published standard or a widely reported public-health fact are labelled Standard or Cited (reported/estimated). No success measures are asserted as achieved; targets marked (target — Evidence Pending) are to be set/ confirmed in Product Requirements and closed by Verification/Validation.
3. ID scheme & record structure
Needs use STN-###. Each need record carries: ID · Stakeholder · Need statement (stakeholder terms) · Rationale · Priority · Source (trace to VOC-### and/or MR-…) · Success measure · Conflicts-with · Notes.
Priority uses a severity-first posture consistent with CLAUDE.md rule 3 (safety and loss-of-function dominate): Critical > High > Medium.
4. Library contents
File | Content
[STN-stakeholder-register.md](STN-stakeholder-register.md) | Stakeholder register & influence/interest map: every stakeholder class, role in the reusable multi-trip lifecycle, personas/segments, key needs, engagement priority
[STN-needs-by-stakeholder.md](STN-needs-by-stakeholder.md) | The detailed STN-### needs, grouped by stakeholder class, each with full record + trace
[STN-conflicts-and-tradeoffs.md](STN-conflicts-and-tradeoffs.md) | Explicit stakeholder-need conflicts (e.g., low cost vs −30 °C/90-day endurance; live data vs battery life) carried forward for resolution
5. Stakeholder classes (summary)
ID | Stakeholder class | Primary personas (MR-PER) | Primary segment (MR-SEG)
SH-01 | Immunization / EPI program managers & MoH | MR-PER-006 | MR-SEG-005
SH-02 | Pharma manufacturers (QA / regulatory affairs) | MR-PER-001, MR-PER-009 | MR-SEG-001
SH-03 | Distributors / wholesalers | MR-PER-002 | MR-SEG-002
SH-04 | 3PL / logistics / air-cargo | MR-PER-004 | MR-SEG-004
SH-05 | Clinic / health-facility staff (point of use) | MR-PER-007 | MR-SEG-007
SH-06 | Cold-chain technicians / field service | MR-PER-011 | MR-SEG-002/004 (depots)
SH-07 | Regulators / QA / auditors | MR-PER-009 | cross-cutting
SH-08 | Procurement / GAVI / UNICEF | MR-PER-008 | MR-SEG-006
SH-09 | IT / security administrators | MR-PER-003 | MR-SEG-001/002 (enterprise)
SH-10 | Finance / asset accounting | (buying committee) | cross-cutting
SH-11 | End beneficiaries / patients (public-health outcome) | — (not a device user) | outcome
SH-12 | Installers (OEM-embed / field install) | MR-PER-005 | MR-SEG-008
SH-13 | Service / reverse-logistics operations | MR-PER-011 | MR-SEG-002/004
SH-14 | Ecosystem partners (equipment OEMs, resellers/SIs) | — | MR-SEG-008/009
SH-15 | Executive / supply-chain leadership (buyer) | MR-PER-010 | cross-cutting
6. Needs index (STN-001 … STN-027)
STN | Need (short) | Primary stakeholder | Theme | Source
STN-001 | Real-time excursion warning in time to act | SH-04/SH-03 | T1 | VOC-008/021, CFB-002
STN-002 | Detect freezing as well as heat | SH-01 | T2 | VOC-011, FLD-002
STN-003 | Endure the full cold trip (−30 °C, 90 days) | SH-03/SH-04 | T3 | VOC-010/024, CFB-001
STN-004 | Instant, offline accept/reject verdict | SH-05/SH-01 | T4 | VOC-009/014/015, FLD-006
STN-005 | Audit-grade, tamper-evident records | SH-07/SH-02 | T5 | VOC-001, CFB-004
STN-006 | Attributable evidence (who/where/when) | SH-03/SH-04 | T6 | VOC-004, FLD-009
STN-007 | Accurate alerts, no alarm fatigue | SH-03 | T7 | VOC-005/022, FLD-010
STN-008 | Frictionless, error-proof provisioning | SH-03 | T8 | VOC-002, FLD-001, LD-006
STN-009 | Work through intermittent connectivity | SH-01 | T9 | VOC-023, FLD-004, LD-010
STN-010 | Reusable readiness (charge/cal/firmware) | SH-06 | T10 | VOC-016/017, CFB-012
STN-011 | PQS-eligible & affordable / low TCO | SH-08 | T11 | VOC-012/013/064, LD-001/002
STN-012 | Safe, compliant air/ground transport | SH-04 | T12 | VOC-007, CFB-006, LD-007
STN-013 | Enterprise data governance & integration | SH-09 | T13 | VOC-018, CFB-010, LD-005
STN-014 | Fleet visibility & ROI evidence | SH-15 | T14 | VOC-020, FLD-012
STN-015 | Useful location within the energy budget | SH-15/SH-09 | T14 | FLD-005/012
STN-016 | Recover/return the reusable fleet | SH-13 | T10 | VOC-006, FLD-007, LD-008
STN-017 | Embeddable form + channel enablement | SH-14 | (OEM) | CFB-013/014
STN-018 | Alerts routed/escalated to the right responder | SH-03/SH-04 | T1 | VOC-008, CFB-002
STN-019 | Cost-per-trip & asset accounting evidence | SH-10 | T14/T11 | VOC-020, VOC-060
STN-020 | No compromised vaccine given; no safe dose wasted | SH-11 | (outcome) | VOC-011/021, MR-PROB-010
STN-021 | Calibration currency & traceability | SH-07 | T5/T10 | CFB-012, VOC-016
STN-022 | Regulatory classification / market-access clarity | SH-07 | T11 | MR-REG, LD-001
STN-023 | Standardize monitoring; protect brand | SH-02 | T5 | VOC-001/003
STN-024 | Physical fit, mounting & placement guidance | SH-12 | T8 | CFB-013, FLD-002
STN-025 | Secure connectivity & signed OTA | SH-09 | T13/T5 | VOC-019, CFB-011
STN-026 | End-of-life, battery disposal & circularity | SH-13 | (sustainability) | MR-1700, D1/D3
STN-027 | Fast, unambiguous receipt-check workflow | SH-05 | T4 | VOC-015, CFB-015
7. Conflicts carried forward (see STN-conflicts-and-tradeoffs.md)
The most consequential stakeholder-need conflicts are recorded and carried forward, not silently resolved:
• C1 Low purchase cost (STN-011/019) vs −30 °C / 90-day endurance (STN-003) — the D3 tension / MR-RISK-001.
• C2 Live/real-time data (STN-001) vs battery life & endurance (STN-003) — resolved in direction by D2.
• C3 Continuous GPS location (STN-015) vs 90-day endurance (STN-003).
• C4 Audit rigor / complete provisioning (STN-005/008) vs dispatch speed (STN-008 covert requirement).
• C5 Rich data for QA/executives (STN-005/014) vs dead-simple offline verdict for last-mile (STN-004).
• C6 Reusable TCO (STN-010/019) vs reverse-logistics cost on hard lanes (STN-016).
• C7 Real-time connectivity value (STN-001/014) vs last-mile no-connectivity reality (STN-009).
• C8 Alert sensitivity/immediacy (STN-001/018) vs false-alarm suppression (STN-007).
• C9 Enterprise security lockdown (STN-013/025) vs frictionless field use & open integration (STN-004/017).
• C10 Sustainability/reuse (STN-026) vs Li-battery transport/disposal constraints & single-use fallback.
8. Traceability (forward / backward)
Backward: each STN cites its VOC-### and/or MR-… source (and D1–D4 where a foundational decision drives it). Forward: each STN flows to User Needs (UN-###, Stage 4, ORCH-010) and thence to Product Requirements (PRD-###); conflicts are carried into design trade studies and, where safety-relevant, into Hazard Analysis (HA-###, ORCH-017). The end-to-end trace matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18).
Flow: MR (Stage 1) → VOC-### (Stage 2) → STN-### (this stage) → UN-### (Stage 4) → PRD-### → … → Verification / Validation.
Stakeholder-Need Conflicts & Trade-offs — SCM
Purpose: Per CLAUDE.md rule 6 (design for the whole life) and the org's honesty posture, stakeholder needs
that pull against each other are recorded explicitly and carried forward — never silently reconciled. Each
conflict names the opposing STN needs and stakeholders, the physical/commercial root cause (based on D1–D4),
the interim direction (where a foundational decision already sets one), and where it is resolved downstream.
Governance: ORCH-018 / EM-ENG-012. Forward to User Needs (Stage 4), design trade studies, and — where
safety-relevant — Hazard Analysis (HA-###, ORCH-017).
Severity-first tie-breaker (CLAUDE.md rule 3): where a conflict touches safe disposition of vaccine, STN-020 (no compromised dose administered; no safe dose wasted — the beneficiary/patient outcome, SH-11) dominates.
---
Conflict register
Fields: ID · Conflict · Opposing needs (stakeholders) · Root cause · Severity/priority · Interim direction · Resolved / carried to · Confidence.
C1 — Low cost vs −30 °C / 90-day endurance (the D3 tension — top program risk)
Opposing needs | STN-011 / STN-019 affordability & TCO (SH-08 procurement, SH-10 finance, SH-01 programs) vs STN-003 cold endurance/ruggedness (SH-03 distributors, SH-04 3PLs).
Root cause | A low-temperature-capable chemistry (low-temp Li-ion / LTO-class) sized for 90 days at −30 °C, plus ruggedization and PQS-grade build, raises unit cost — but donor budgets are hard ceilings (D3; VOC-013/064; LD-002).
Severity/priority | Critical. This is MR-RISK-001 and the program's headline engineering/commercial tension.
Interim direction | D3 — low-temp chemistry + aggressive duty-cycling + charge-inhibit-below-0 °C; pursue a possible value/entry SKU (MR-1400) to meet price ceilings while a professional SKU carries full endurance.
Resolved / carried to | Power-budget + business-case trade study; MR-VAL-006 closure-by-test; PRD SKU strategy; HA-### (dead-in-cold hazard).
Confidence | Standard/Judgment — tension is based; cost/endurance numbers Evidence Pending.
C2 — Live/real-time data vs battery life
Opposing needs | STN-001 real-time excursion warning (SH-03/SH-04) vs STN-003 endurance (all shippers).
Root cause | Continuous live cellular every few minutes for 90 days is energy-prohibitive at the endurance target (D2 rationale; confirmed in power budget).
Severity/priority | Critical (safety-relevant: delayed alert vs dead device).
Interim direction | D2 resolves the direction — local logging + scheduled cadence upload + immediate excursion alert when connectivity is available; buffer offline and upload opportunistically.
Resolved / carried to | D2 (decided); alert-latency hazard HA-### (delayed/missing alert); PRD alerting/energy requirements.
Confidence | Standard/Judgment — direction set by D2; residual latency risk carried.
C3 — Continuous GPS location vs 90-day endurance
Opposing needs | STN-015 useful location / STN-014 fleet visibility (SH-15, SH-09) vs STN-003 endurance.
Root cause | Continuous GPS is energy-expensive and fixes often fail inside metallized/insulated shippers (D3; concept §7 open item; FLD-005/012).
Severity/priority | Medium.
Interim direction | Opportunistic/periodic location within the energy budget; evaluate external-antenna option and event-triggered fixes rather than always-on GPS.
Resolved / carried to | Power-budget trade study; RF-in-container verification test; PRD location cadence requirement.
Confidence | Judgment/Assumption — quantify in power budget + test.
C4 — Audit rigor / complete provisioning vs dispatch speed
Opposing needs | STN-005 complete tamper-evident record + STN-008 error-proof provisioning completeness (SH-02, SH-07) vs STN-008 covert "must not slow dispatch" (SH-02/SH-03 dispatch).
Root cause | Rigorous setup and complete records add steps; high-throughput pack-out punishes any added time, driving skipped/void trips (VOC-002; FLD-001; LD-006).
Severity/priority | High.
Interim direction | Make the compliant path the fastest path — single confirmed arming action, sane defaults, on-device "armed/placed" confirmation; no trade of completeness for speed, but remove friction.
Resolved / carried to | Human-factors design; provisioning-time target in PRD; contextual-inquiry validation.
Confidence | Judgment/Assumption — provisioning-time target Evidence Pending.
C5 — Rich data (QA/executive) vs dead-simple offline verdict (last-mile)
Opposing needs | STN-005 audit-grade data + STN-014 fleet KPIs (SH-02, SH-07, SH-15) vs STN-004 instant offline accept/reject for non-technical last-mile users (SH-01, SH-05).
Root cause | One device serves both the most exacting auditor and the least-trained, offline nurse (MR-1200 central tension); a data-rich interface overwhelms the edge, a bare verdict fails the audit.
Severity/priority | High (safety-relevant at the edge).
Interim direction | Layered information — OLED + button gives an unambiguous offline verdict; full ALCOA+ record lives in the log/cloud for QA/audit. Same truth, role-appropriate presentation.
Resolved / carried to | HMI/UX design; PRD (on-device verdict vs cloud record); usability validation across personas.
Confidence | Judgment/Assumption — confirm via persona testing.
C6 — Reusable TCO vs reverse-logistics cost on hard lanes
Opposing needs | STN-010/STN-019 reuse economics (SH-06, SH-08, SH-10) vs STN-016 reverse-logistics cost on export/last-mile routes (SH-13, SH-01).
Root cause | The reusable model (D1) needs high return rates to beat single-use; on remote LMIC last-mile lanes returns are hard/expensive and devices are stranded (VOC-006; FLD-007; LD-008).
Severity/priority | High (determines whether the business case survives).
Interim direction | Region/lane-specific strategy — reusable on manageable lanes, evaluate single-use/entry SKU or pooled-return incentives on hard lanes; location (STN-015) to chase strays.
Resolved / carried to | Business-case + return-rate pilot; SKU/portfolio strategy (MR-1400); MR-EXEC-008.
Confidence | Judgment/Assumption — return-rate/TCO Evidence Pending (pilot).
C7 — Real-time connectivity value vs last-mile no-connectivity reality
Opposing needs | STN-001 real-time warning + STN-014 live fleet KPIs (SH-03/04/15) vs STN-009 offline operation where connectivity/power are unreliable (SH-01, SH-05).
Root cause | Real-time value depends on coverage that does not exist on the highest-loss last-mile legs (MR-PROB-008; LD-010; FLD-004).
Severity/priority | Critical (the black-box last mile is where losses concentrate).
Interim direction | D2 store-and-forward — never depend on live connectivity for the core job; local log + offline verdict always; live alerting/KPIs are an upside where coverage exists.
Resolved / carried to | D2 (decided); PRD offline-first requirements; connectivity-resilience validation.
Confidence | Cited/Standard — based on D2 and last-mile reporting.
C8 — Alert sensitivity/immediacy vs false-alarm suppression
Opposing needs | STN-001/STN-018 immediate, routed alerts (SH-03/SH-04) vs STN-007 accurate, low-noise alerts that avoid fatigue (SH-03 coordinators, SH-05).
Root cause | Maximizing sensitivity catches every transient (door-open, transfer) and breeds alarm fatigue; suppressing transients risks a delayed real alert (VOC-005/022; FLD-010; MR-PROB-009).
Severity/priority | High (safety-relevant: muted alerts miss real excursions).
Interim direction | Product-stability-aware alerting (time-over/under-threshold, cumulative/MKT-style logic), configurable per vaccine (STN-004); escalate genuine excursions, suppress benign transients.
Resolved / carried to | Excursion-logic design; PRD alerting requirements; alarm-management validation; HA-### (missed excursion).
Confidence | Judgment/Assumption — thresholds/false-alarm tolerance Evidence Pending.
C9 — Enterprise security lockdown vs frictionless field use & open integration
Opposing needs | STN-013/STN-025 CSV, SSO/RBAC, signed OTA, controlled access (SH-09, SH-07) vs STN-004 zero-friction offline edge use (SH-05) and STN-017 open APIs / easy partner integration (SH-14).
Root cause | Hardening for enterprise/regulatory trust adds gates and access controls that can obstruct offline field use and open ecosystem integration (VOC-018/019; LD-005; CFB-010/014).
Severity/priority | High.
Interim direction | Separate the trust boundaries — device edge stays offline-simple; platform enforces security/validation; expose secured open APIs (authN/Z, signed) rather than closed silo or open-and-unsafe.
Resolved / carried to | Architecture + cybersecurity design (flag EM cyber persona); PRD security vs usability requirements.
Confidence | Judgment/Assumption — reconcilable by architecture; confirm with IT/security personas.
C10 — Sustainability / reuse vs Li-battery transport & disposal constraints
Opposing needs | STN-026 circularity/reuse + STN-010 long service life (SH-13) vs STN-012 Li-battery transport compliance (SH-04) and end-of-life disposal constraints.
Root cause | Reuse and returns move Li-ion batteries repeatedly (UN 38.3 / IATA), and end-of-life disposal is regulated; single-use would sidestep returns but harms sustainability/TCO (D1/D3; FLD-011; CFB-006).
Severity/priority | Medium.
Interim direction | Design for compliant battery transport and a defined disposal/EOL path; treat returns as a compliance-managed process, not an afterthought.
Resolved / carried to | Sustainability (MR-1700) + regulatory + mechanical stages; PRD battery/EOL requirements.
Confidence | Standard/Judgment — transport/disposal governed by standards; EOL path Evidence Pending.
---
Conflict-to-decision map (how each is expected to close)
ID | Already decided by | To be resolved in | Safety-relevant?
C1 | Direction by D3 (chemistry + duty-cycle + SKU) | Power budget, business case, MR-VAL-006 test, PRD SKU | Yes (dead-in-cold)
C2 | Resolved by D2 (store-and-forward + immediate alert) | Alert-latency hazard HA-###; PRD | Yes
C3 | — | Power budget + RF-in-container test; PRD location cadence | No
C4 | — | Human-factors design; PRD provisioning-time | Indirect (void trips)
C5 | — | HMI/UX; PRD verdict vs record; persona usability test | Yes (edge decision)
C6 | Direction by D1 (reusable) | Return-rate pilot; portfolio/SKU strategy | No (economic)
C7 | Resolved by D2 (offline-first) | PRD offline requirements; resilience test | Yes
C8 | — | Excursion-logic design; PRD; alarm-management test; HA-### | Yes
C9 | — | Architecture + cybersecurity; PRD security-vs-usability | Indirect
C10 | Direction by D1/D3 + standards | Sustainability + regulatory + mechanical; PRD | No
Traceability
Backward: STN-### (STN-needs-by-stakeholder.md); VOC-### (folder 02); MR-… (folder 01); D1–D4 (concept). Forward: User Needs (Stage 4, ORCH-010), design trade studies, and Hazard Analysis (HA-###, ORCH-017) for safety-relevant conflicts. End-to-end matrix assembled by ORCH-018 / EM-ENG-012 (Stage 18).
Stakeholder Needs — Detailed Records (STN-001 … STN-027), by Stakeholder Class
Evidence note (read index.md §2): needs are tagged by evidentiary basis (Standard / Cited /
Judgment/Assumption / Evidence Pending). Success measures marked (target — Evidence Pending) are to be
set in Product Requirements and closed by Verification/Validation. Backward trace to VOC-### / MR-… / D1–D4;
forward trace to UN-### (Stage 4). Governance: ORCH-018 / EM-ENG-012.
Record fields: ID · Stakeholder · Need statement (stakeholder terms) · Rationale · Priority · Source (VOC-### / MR-…) · Success measure · Conflicts-with · Notes.
Where a need is genuinely shared, it is placed under its primary stakeholder and cross-referenced.
---
SH-01 — Immunization / EPI program managers & MoH
STN-002 — "Catch freezing, not just heat"
Stakeholder | EPI program managers (shared: distributors, manufacturers).
Need statement | "I need the monitor to detect and flag inadvertent freezing as reliably as heat, because freeze-sensitive vaccines are silently ruined by over-cold fridges and cold-lane exposure, and our current indicators miss it."
Rationale | Freezing damage is common, silent, and under-detected by heat-biased VVMs; a large share of loss.
Priority | Critical.
Source | VOC-011, FLD-002; MR-PROB-002/004; MR-JTBD-001.
Success measure | Two-sided detection with configurable low/high thresholds; freeze events flagged in verdict and record (accuracy target — Evidence Pending; set in PRD, closed by V&V).
Conflicts-with | — (probe placement dependency, see STN-024).
Notes | Internal + external-probe temperature supports freeze detection; Cited to WHO/UNICEF freezing reporting.
STN-004 — "Give me a clear go/no-go, offline, at the door"
Stakeholder | EPI program managers (shared: clinic staff SH-05).
Need statement | "When a consignment arrives, my staff must decide accept or reject in seconds, from the device itself, with no phone, network, or data expertise — and the verdict must line up with PQS expectations."
Rationale | Last-mile staff are non-technical and time-pressured; a data dump does not support the decision.
Priority | Critical.
Source | VOC-009/014/015, FLD-006, CFB-015, LD-009; MR-JTBD-004; MR-PROB-004; MR-PER-006/007.
Success measure | On-device accept/reject verdict readable offline within seconds; rule-set configurable and PQS-consistent (verdict-match target — Evidence Pending; confirm via audit sim + WHO PQS category, code TBC).
Conflicts-with | STN-005/STN-014 (rich data vs simple verdict — conflict C5).
Notes | OLED + button is the sole interface at the edge; low-literacy, offline-legible design essential.
STN-009 — "Work through poor connectivity and power"
Stakeholder | EPI program managers (shared: 3PLs, distributors on remote lanes).
Need statement | "In our regions connectivity and power are unreliable, so the monitor must keep logging and give a verdict without a live connection, and upload whenever it can — I can't depend on live data."
Rationale | Losses concentrate in the last mile where connectivity/power gaps leave shipments unmonitored.
Priority | Critical.
Source | VOC-023, FLD-004, LD-010; MR-PROB-008; D2 (store-and-forward).
Success measure | Full local log + offline verdict retained; buffered data uploaded opportunistically with intact time order (buffer duration/data-loss target — Evidence Pending; test).
Conflicts-with | STN-001/STN-014 (real-time value vs no-connectivity reality — conflict C7).
Notes | LTE-M primary with BLE/Wi-Fi offload; live data is a bonus, not a dependency. Cited (WHO/UNICEF/GAVI).
STN-011 — "PQS-eligible and affordable within our budget"
Stakeholder | EPI programs (shared/primary with procurement SH-08).
Need statement | "If it isn't PQS-listed and the lifetime cost fits our donor budget, we simply cannot buy it — no matter how good it is."
Rationale | PQS prequalification is a hard market-access gate; donor budgets are fixed ceilings.
Priority | Critical.
Source | VOC-012/013/064, LD-001/002; MR-SEG-005/006/011; MR-PER-008; MR-REG; MR-800; D3.
Success measure | WHO PQS listing achieved (category/code TBC — Evidence Pending); documented cost-per-trip and TCO within target budget band (price/TCO — Evidence Pending).
Conflicts-with | STN-003 (endurance/ruggedness cost vs affordability — conflict C1, the D3 tension / MR-RISK-001).
Notes | Plan PQS listing early (long cycles). Possible value/entry SKU (MR-1400) to meet price ceilings.
STN-020 — "No child gets a dead dose; no good dose is thrown away" (represented for SH-11)
Stakeholder | End beneficiaries / patients — represented by EPI programs, regulators, procurement.
Need statement | "The people we serve need assurance that no compromised vaccine is administered and that potent vaccine is not needlessly discarded — the monitor exists to protect that outcome."
Rationale | The public-health outcome is the ultimate purpose; both undetected excursions and defensive over-discard harm beneficiaries and waste scarce doses.
Priority | Critical (severity anchor — CLAUDE.md rule 3).
Source | VOC-011/021; MR-PROB-010/002/001; MR-JTBD-001; concept intended-use.
Success measure | Correct disposition rate — no false "accept" of excursed stock; no false "reject" of in-range stock (target — Evidence Pending; closed by validation, ties to hazard analysis HA-###).
Conflicts-with | — (this need is the tie-breaker when other needs conflict).
Notes | Device supports, does not make, the disposition; it is not therapeutic. This STN is the top severity reference for risk prioritization across the program.
---
SH-02 — Pharma manufacturers (QA & regulatory affairs)
STN-005 — "A record I can defend to any inspector"
Stakeholder | Manufacturer QA/RA (shared: regulators SH-07, distributors, 3PLs).
Need statement | "When there is an excursion I must defend the disposition to an FDA/MHRA/GDP inspector, so the temperature record must be complete, tamper-evident, attributable, and exportable — no gaps, nothing that looks editable."
Rationale | Editable or gap-ridden records fail audits and invalidate release decisions; brand/liability on the line.
Priority | Critical.
Source | VOC-001/003/019, CFB-004; MR-PROB-005; MR-JTBD-008/003; MR-PER-001/009; MR-EXEC-005.
Success measure | Continuous local logging with tamper-evident, ALCOA+-aligned, exportable records and full audit trail; gaps explicitly flagged, never hidden (Part 11/GDP conformance — Standard; validated in V&V).
Conflicts-with | STN-004 (rich audit data vs simple offline verdict — conflict C5); STN-008 (rigor vs dispatch speed — C4).
Notes | Standard basis: EU GDP + 21 CFR Part 11 + ALCOA+. Coordinate cybersecurity persona for tamper-evidence.
STN-008 — "Arm and place it correctly in seconds — every time"
Stakeholder | Manufacturer/distributor dispatch (shared/primary with distributors SH-03; installers SH-12).
Need statement | "At pack-out my team must configure, arm, and place the device correctly in well under a minute, with no way to forget to start it or set the wrong trip profile — audit rigor must not slow dispatch."
Rationale | "Forgot to start the logger" and misconfiguration are classic void-trip failures that destroy data yield and pilots; provisioning is a covert adoption gate.
Priority | High.
Source | VOC-002, FLD-001/003, CFB-005, LD-006; MR-JTBD-006; MR-PER-001/002/005.
Success measure | Provisioning time (target well under 1 min — Evidence Pending); on-device "armed/placed" confirmation; error-proof profile selection; void-trip rate minimized (target — Evidence Pending).
Conflicts-with | STN-005 (completeness/rigor vs speed — conflict C4).
Notes | Glove-friendly, cold-legible, single-action interaction (FLD-003). On-device fallback if BLE/app fails.
STN-023 — "Standardize monitoring across my whole chain and protect the brand"
Stakeholder | Manufacturer QA/supply-chain.
Need statement | "I want one monitoring approach I can mandate across all markets and partners, so my product's cold-chain story is consistent and my brand is protected wherever it ships."
Rationale | Manufacturers are high-leverage anchors who can standardize monitoring downstream; fragmentation raises risk.
Priority | High.
Source | VOC-001/003; MR-SEG-001/010; MR-EXEC-005; D4.
Success measure | Single device/platform defensible across WHO PQS + EU GDP + US markets (multi-market conformance — Standard/Evidence Pending, confirm in regulatory stage).
Conflicts-with | — (supports STN-005/STN-011 alignment).
Notes | Anchor-account leverage; winning SH-02 pulls volume through SH-03/SH-04.
---
SH-03 — Distributors / wholesalers
STN-001 — "Warn me in time to act, not at delivery"
Stakeholder | Distributors (shared/primary with 3PLs SH-04).
Need statement | "I need to be alerted to an excursion while I can still do something — reroute, re-ice, divert — not discover it when I open the box. Real-time is why I would pay more than for a USB logger."
Rationale | After-the-fact discovery documents loss instead of preventing it; real-time enables in-transit rescue.
Priority | Critical.
Source | VOC-008/021, CFB-002/009; MR-PROB-001; MR-JTBD-005; MR-COMP; D2.
Success measure | Immediate excursion alert pushed when connectivity available; event captured locally and surfaced on reconnection/at receipt if offline (alert-latency target — Evidence Pending; connectivity-bound).
Conflicts-with | STN-003 (live data vs battery life — conflict C2, directionally resolved by D2); STN-007 (sensitivity/immediacy vs false-alarm suppression — conflict C8); STN-009 (C7).
Notes | Delayed/missing alert carried as a hazard (HA-###). Alert latency depends on LTE-M coverage.
STN-003 — "Survive the whole cold trip — device must not die in the cold"
Stakeholder | Distributors (shared: 3PLs SH-04, EPI programs SH-01 — all cold-lane shippers).
Need statement | "The monitor must keep working for the full trip through deep cold — no dead battery, no lost data — because a device that quits in the cold loses the record exactly when we need it."
Rationale | Cold-weather battery failure is a recurring incumbent failure and the program's top engineering risk.
Priority | Critical.
Source | VOC-010/024, CFB-001/016, FLD-004/008, LD-003; MR-PROB-007; D3; MR-RISK-001; MR-VAL-006.
Success measure | 90-day endurance across −30 °C to +60 °C; no data loss over cold dwell; charge inhibited below 0 °C with clear readiness state (endurance — top target, Evidence Pending; closed by test / MR-VAL-006).
Conflicts-with | STN-011/STN-019 (endurance cost vs affordability — C1); STN-001 (live data vs battery — C2); STN-015 (continuous location vs endurance — C3).
Notes | The D3 low-temperature battery tension is MR-RISK-001, carried to hazard/reliability/power budget. Validate by test; may drive chemistry choice (low-temp Li-ion / LTO-class) and aggressive duty-cycling.
STN-006 — "Prove who broke the chain"
Stakeholder | Distributors (shared/primary with 3PLs SH-04, manufacturers SH-02).
Need statement | "When a shipment arrives out of range I need objective, time- and location-stamped evidence of when and where it happened, so we stop absorbing write-offs for excursions we didn't cause."
Rationale | Without attributable evidence, partners dispute responsibility and distributors eat the loss.
Priority | High.
Source | VOC-004, FLD-009; MR-PROB-006; MR-JTBD-003.
Success measure | Excursion events time- and location-stamped and tamper-evident, sufficient to attribute at a custody hand-off (attribution-usefulness target — Evidence Pending; requires GPS-in-container performance, FLD-005).
Conflicts-with | STN-015 (location value vs energy budget — related to C3).
Notes | GPS + timestamped logs support attribution; hand-offs (tarmac/cross-dock) are where disputes cluster.
STN-007 — "Alert me only when it matters — no alarm fatigue"
Stakeholder | Distributors / logistics coordinators (shared: 3PLs, programs, pharmacies).
Need statement | "Managing hundreds of shipments, I need action-only alerts for genuine, product-relevant excursions — not door-opens and transfers — or my team mutes them and misses the real one."
Rationale | Poorly tuned alerting causes alarm fatigue; muted alerts defeat the whole detection capability.
Priority | High.
Source | VOC-005/022, FLD-010, CFB-003, CFB-015, LD-009; MR-PROB-009; MR-PER-002.
Success measure | Alerts reflect product-stability-relevant exposure (e.g., time-over-threshold / cumulative logic), not raw transients; false-alarm rate within tolerance (target — Evidence Pending).
Conflicts-with | STN-001 (immediacy/sensitivity vs false-alarm suppression — conflict C8).
Notes | Thresholds should reflect product stability profiles; configurable per vaccine (ties STN-004).
STN-008 — (primary record under SH-02 above; distributors are a co-primary stakeholder)
Cross-reference | See STN-008 under SH-02. Distributor dispatch (MR-PER-002) shares full ownership; the covert "must not slow dispatch" requirement is strongest here.
STN-016 — "Get the reusable devices back" (shared/primary with service SH-13)
Cross-reference | See STN-016 under SH-13. Distributors running reusable fleets are a co-primary stakeholder; return economics on export/last-mile lanes are decisive for their TCO (VOC-006, LD-008).
STN-018 — "Route the alert to the person who can act"
Stakeholder | Distributors / 3PL control tower (SH-03/SH-04).
Need statement | "An excursion alert is useless unless it reaches the right responder fast — the coordinator, control tower, or program officer who can actually intervene — with escalation if unacknowledged."
Rationale | Detection only prevents loss if the right person is warned in time (MR-JTBD-005 outcome).
Priority | High.
Source | VOC-008, CFB-002; MR-JTBD-005; MR-PROB-001; MR-PER-002.
Success measure | Configurable alert routing/escalation by role/shipment; acknowledgement tracked (routing/escalation behaviour — Evidence Pending; set in PRD).
Conflicts-with | STN-007 (escalation volume vs fatigue — related C8).
Notes | Cloud/dashboard function; depends on connectivity (STN-009) for push.
---
SH-04 — 3PL / logistics / air-cargo
STN-001 / STN-006 / STN-018 — (co-primary; see SH-03 records)
Cross-reference | 3PLs share STN-001 (real-time warning), STN-006 (attributable evidence), and STN-018 (alert routing) as co-primary stakeholders; control-tower operations amplify the real-time and routing needs.
STN-012 — "Don't slow my dock or trip a battery-shipping rule"
Stakeholder | 3PL operations lead (shared: service/reverse-logistics SH-13).
Need statement | "The monitor must be invisible to my handlers, safe to fly, and never the reason a shipment is held for dangerous-goods paperwork — transport and battery compliance is table stakes for us."
Rationale | 3PL SOP inclusion is a channel-scale lever; a device that trips DG rules or slows handling is excluded.
Priority | High.
Source | VOC-007, CFB-006, LD-007; MR-SEG-004; MR-PER-004; D3/D4 (UN 38.3, IATA DGR).
Success measure | UN 38.3 test summary and IATA-compliant transport guidance provided; device handling-neutral (transport-compliance conformance — Standard; documentation Evidence Pending).
Conflicts-with | STN-026 (battery transport/disposal constraints — related C10).
Notes | Li-ion transport classification (D3) decisive; coordinate battery + regulatory engineers.
STN-003 — (co-primary; see SH-03 record — 3PLs bear in-transit cold-endurance failures directly)
---
SH-05 — Clinic / health-facility staff (point of use)
STN-004 — (primary record under SH-01 above; clinic staff are co-primary)
Cross-reference | See STN-004 under SH-01. The last-mile health worker (MR-PER-007) is the archetypal user of the offline, glanceable verdict (VOC-014); zero-connectivity operation is mandatory here (FLD-006).
STN-027 — "Tell me at the door whether to fridge it, quarantine it, or send it back"
Stakeholder | Hospital / pharmacy / clinic receivers.
Need statement | "At receipt I need a fast, unambiguous status so I can accept, quarantine, or reject stock quickly and defensibly — I'm not trying to become a data analyst."
Rationale | Receivers want a fast, unambiguous accept/reject signal; friction undermines chain-of-custody confidence.
Priority | Medium-High.
Source | VOC-015, CFB-015; MR-SEG-007; MR-JTBD-002/004.
Success measure | At-receipt status (accept/quarantine/reject) readable via OLED/button/BLE within seconds (usability target — Evidence Pending; confirm via receipt-workflow observation).
Conflicts-with | — (aligned with STN-004).
Notes | Reinforces end-to-end confidence; supports whole-chain adoption case.
STN-020 — (represented; see SH-01 record) — clinic staff act on behalf of the beneficiary at the point of care.
---
SH-06 — Cold-chain technicians / field service
STN-010 — "Turn a returned device around fast — charged, in-cal, firmware-current — verifiably"
Stakeholder | Field-service / calibration technicians (shared: service ops SH-13).
Need statement | "When a device comes back I must quickly know it's charged, still in calibration, and firmware-current, and re-provision it — without manually checking each one, or I can't keep the fleet moving."
Rationale | Turnaround time directly sets device utilization and reuse margin; hidden drift or dead batteries destroy the reuse ROI and record credibility.
Priority | High.
Source | VOC-016/017, CFB-012/016, FLD-008; MR-JTBD-007/009; MR-PER-011; D1/D3; ISO/IEC 17025 concepts.
Success measure | Device reports charge state, calibration status, and firmware version to the depot; clear "warming/will charge" indication after cold return; deploy blocked if out-of-cal (throughput/interval targets — Evidence Pending; calibration interval set in regulatory/service stage).
Conflicts-with | STN-003 (charge-inhibit-below-0 °C affects readiness timing — related).
Notes | Charge-inhibit-below-0 °C (D3) and self-heat-before-charge under evaluation; cleaning between trips (STN-026).
STN-021 — "Prove the readings can be trusted — calibration is current and traceable"
Stakeholder | Service technicians (shared/primary with regulators/QA SH-07).
Need statement | "Every device in circulation must have current, traceable calibration, and I need to manage and evidence that across the fleet so no out-of-cal device produces a disputed reading."
Rationale | Record credibility and audit defensibility depend on calibration currency/traceability.
Priority | High.
Source | CFB-012, VOC-016; MR-JTBD-007; MR-PER-009/011; ISO/IEC 17025 concepts.
Success measure | Per-device calibration status tracked with traceable records; expiry blocks deployment (calibration interval/traceability — Standard concept; interval Evidence Pending).
Conflicts-with | — (supports STN-005/STN-010).
Notes | Links service (SH-06) and audit (SH-07); metrology traceability to a recognized standard.
---
SH-07 — Regulators / QA / auditors
STN-005 / STN-021 — (co-primary; see SH-02 / SH-06 records)
Cross-reference | Regulators/auditors are the co-primary stakeholders for STN-005 (tamper-evident records) and STN-021 (calibration traceability); their acceptance defines whether records are admissible.
STN-022 — "Clear regulatory classification and market-access basis"
Stakeholder | Regulatory affairs / QA (internal), regulators (external).
Need statement | "We need the device's regulatory status clear in each market — WHO PQS category, US FDA classification/pathway, EU GDP fit — so we know how to place it, register it, and defend its use."
Rationale | Market access is gated by PQS listing and correct classification; ambiguity blocks tenders and sales.
Priority | High.
Source | MR-REG, LD-001; MR-SEG-005/006; D4; concept §7 open items.
Success measure | Confirmed WHO PQS category/code (TBC), US classification, and EU GDP applicability documented (all Evidence Pending — to be confirmed in the Regulatory stage).
Conflicts-with | — (enables STN-011).
Notes | Device is an information/accessory tool, not therapeutic; classification confirmation is a named concept gap.
STN-025 — (co-primary; see SH-09 record) — auditors require evidence of data-integrity/security controls.
---
SH-08 — Procurement / GAVI / UNICEF
STN-011 — (primary record under SH-01 above; procurement is co-primary/primary)
Cross-reference | See STN-011. Pooled procurement (MR-PER-008) makes PQS listing + reuse TCO the gating tender criteria (VOC-012, LD-001); this is the largest aggregated LMIC volume channel.
STN-019 — "Show me cost-per-trip over the device life"
Stakeholder | Procurement (shared/primary with finance SH-10).
Need statement | "I need documented cost-per-trip and total cost of ownership across the device's reusable life, and supply assurance at scale, to justify donor-funded procurement."
Rationale | Reuse economics only justify purchase if the lifetime math is proven; budgets are fixed.
Priority | High.
Source | VOC-012/013/020/060; MR-EXEC-008; MR-800; MR-1000; D1.
Success measure | Documented cost-per-trip and TCO model with return-rate assumptions; supply-assurance plan (figures — Evidence Pending; validated by pilot return-rate and business-case workbook).
Conflicts-with | STN-003 (endurance/ruggedness cost vs affordability — C1); STN-016 (return rate drives TCO).
Notes | Cost-per-trip is the linchpin of the reusable business case (MR-EXEC-008).
STN-020 / STN-022 — (represented / co-primary; see SH-01 / SH-07 records)
---
SH-09 — IT / security administrators
STN-013 — "It must pass our validation, access, and integration bar"
Stakeholder | Enterprise IT / quality-systems admins.
Need statement | "Before this touches our network it must clear computerized-system validation, single sign-on and role-based access, data-retention rules, and integrate with our logistics/quality systems via open APIs — or it never deploys."
Rationale | Enterprise deployment and platform expansion are gated by IT validation and integration; a siloed device with no API stalls at procurement.
Priority | High.
Source | VOC-018, CFB-010, LD-005; MR-PER-003; MR-1500 ECO; Standard (21 CFR Part 11, GAMP 5, ALCOA+).
Success measure | CSV/validation pack, SSO/RBAC, configurable retention, and documented open APIs available (conformance — Standard; integration specifics Evidence Pending).
Conflicts-with | STN-004/STN-017 (security lockdown vs frictionless field use & open integration — conflict C9).
Notes | Coordinate with regulatory auditor (SH-07) and cybersecurity persona in PRD stage.
STN-025 — "Secure the connected device — signed OTA, controlled access, tamper-evident data"
Stakeholder | IT / security (shared: regulators SH-07).
Need statement | "A cellular device on our shipments is an attack surface and a data-integrity risk, so the data must be tamper-evident, firmware updates signed and fail-safe, and access controlled."
Rationale | Connectivity + OTA introduce security and integrity threats that must be controlled for audit and safety.
Priority | High.
Source | VOC-019, CFB-011; MR-PROB-005; MR-JTBD-003; MR-PER-003; D2 (OTA).
Success measure | Signed, resumable, fail-safe OTA; controlled access; tamper-evident records (security controls — Standard/Judgment; threat model Evidence Pending; set in PRD/hazard with cybersecurity persona).
Conflicts-with | STN-004/STN-017 (C9).
Notes | Flag EM cybersecurity persona for PRD/hazard; ties to STN-005 tamper-evidence.
STN-015 — "Give me useful location without killing the battery"
Stakeholder | IT / fleet admins (shared/primary with executive SH-15).
Need statement | "I want to know where devices are for fleet management and stray recovery, but I understand continuous GPS is energy-expensive and fails inside shippers — give me useful, periodic location within the endurance budget."
Rationale | Location enables fleet visibility and recovery, but continuous GPS conflicts with the 90-day energy budget.
Priority | Medium.
Source | FLD-005/012; MR-JTBD-010; MR-PROB-011; D3; concept §7 (RF-in-container).
Success measure | Opportunistic/periodic location adequate for fleet management and recovery, within endurance target (location cadence vs energy — Evidence Pending; explicit tradeoff, closed by power budget + test).
Conflicts-with | STN-003 (continuous location vs 90-day endurance — conflict C3).
Notes | Antenna/GPS performance inside metallized containers is a named concept gap (FLD-005); may drive external-antenna option.
---
SH-10 — Finance / asset accounting
STN-019 — (primary record under SH-08 above; finance is co-primary)
Cross-reference | See STN-019. Finance owns capital-vs-opex, asset utilization/depreciation, and the loss/ write-off reduction case; cost-per-trip and return rate are the levers.
STN-014 — (co-primary; see SH-15 record) — finance consumes the fleet ROI/loss-reduction evidence.
---
SH-11 — End beneficiaries / patients (public-health outcome)
STN-020 — (primary record above, represented via SH-01/07/08)
Cross-reference | See STN-020 under SH-01. This is the program's severity anchor: correct disposition so no compromised dose is given and no safe dose is wasted. It is the tie-breaker when other needs conflict (CLAUDE.md rule 3).
---
SH-12 — Installers (OEM-embed / field install)
STN-024 — "Make it fit, mount securely, and guide correct placement"
Stakeholder | Installers / OEM integrators (shared: warehouse operators, ecosystem partners SH-14).
Need statement | "I need a standard, embeddable form factor that fits and mounts securely in insulated shippers/containers, with clear guidance on where to place the probe so readings reflect the product zone."
Rationale | Poor fit blocks OEM embedding; wrong probe placement biases freeze/heat detection and causes false alarms.
Priority | Medium.
Source | CFB-013, FLD-002; MR-SEG-008; MR-JTBD-006; MR-PER-005.
Success measure | Fit-inside-shipper form factor + defined mounting/retention; placement guidance for int vs ext probe (fit/placement spec — Evidence Pending; set in PRD/mechanical).
Conflicts-with | — (supports STN-002/STN-008).
Notes | Flag mechanical + materials engineers (mounting, seal integrity at USB-C port, label durability).
STN-008 / STN-002 — (co-primary; see SH-02 / SH-01 records) — correct arming and probe placement at install.
---
SH-13 — Service / reverse-logistics operations & sustainability
STN-016 — "Recover and return the reusable devices — reliably and affordably"
Stakeholder | Reverse-logistics / service operations (shared/primary with distributors SH-03, programs SH-01).
Need statement | "The reusable model only pays if devices actually come home, so I need a practical, affordable return workflow and the ability to locate strays — especially on export and last-mile routes where returns are hard."
Rationale | Device attrition in reverse logistics erodes reuse economics; return rate determines whether TCO holds.
Priority | High.
Source | VOC-006, FLD-007, CFB-008, LD-008; MR-PROB-011; MR-JTBD-009/010; D1; MR-EXEC-008.
Success measure | Defined return workflow; last-known-location to chase strays; return rate within target (return-rate/TCO — Evidence Pending; pilot-measured).
Conflicts-with | STN-019 (reuse TCO vs reverse-logistics cost on hard lanes — conflict C6).
Notes | Region-specific reusable-vs-single-use strategy may be warranted; GPS/BLE location (STN-015) supports recovery.
STN-026 — "Clean, dispose, and close the loop safely"
Stakeholder | Service / sustainability operations (shared: 3PLs SH-04).
Need statement | "I need to clean/decontaminate devices between trips without damaging them, and have a safe, compliant path for battery transport, disposal, and end-of-life across many reuse cycles."
Rationale | Reuse/circularity is a stated D1 objective and a sustainability driver; cleaning and Li-battery disposal are practical and regulatory obligations.
Priority | Medium.
Source | FLD-011; MR-JTBD-009; MR-1700 (SUS); D1/D3; IEC 60529; UN 38.3 / IATA (transport).
Success measure | Cleanable, sealed enclosure (IP-rated) durable over reuse cycles; documented battery disposal/end-of-life path (cleaning method / cycle life / EOL — Evidence Pending; set in mechanical/sustainability stages).
Conflicts-with | STN-012 (battery transport constraints — related C10).
Notes | Flag materials (chemical compatibility, seal life) + mechanical (USB-C port sealing) engineers.
STN-010 / STN-012 — (co-primary; see SH-06 / SH-04 records)
---
SH-14 — Ecosystem partners (equipment OEMs, resellers, SIs)
STN-017 — "Make it embeddable and easy to resell, integrate, and support"
Stakeholder | Cold-chain equipment OEMs, resellers, systems integrators.
Need statement | "As an OEM I want monitoring I can build into my shipper as a differentiator; as a reseller/SI I need open APIs, provisioning tooling, and localized support so I can integrate and support it in my region."
Rationale | OEM embedding and channel partners scale distribution into fragmented and hard-to-serve markets.
Priority | Medium-High.
Source | CFB-013/014; MR-SEG-008/009; MR-1500 ECO; MR-JTBD-006.
Success measure | Embeddable form factor + open APIs + BLE/mobile provisioning + localized docs/enablement (partner-enablement scope — Evidence Pending; set in ecosystem/PRD stages).
Conflicts-with | STN-013/STN-025 (open integration vs enterprise security lockdown — conflict C9).
Notes | Designed-in distribution channel; ties installer (STN-024) and IT integration (STN-013).
---
SH-15 — Executive / supply-chain leadership (strategic buyer)
STN-014 — "Show me network-wide loss reduction, cost-per-trip, and standardization evidence"
Stakeholder | Executive / supply-chain director (shared: finance SH-10).
Need statement | "I don't care about one device — I care about how much cold-chain loss we cut across the network, how compliance exposure drops, and what each trip costs, so I can fund fleet-scale standardization."
Rationale | Executive sponsorship secures budget and standardization decisions; anchors multi-year contracts.
Priority | High.
Source | VOC-020, FLD-012; MR-JTBD-010; MR-PER-010; MR-1000 (MET); MR-EXEC-008.
Success measure | Fleet dashboard with network loss-reduction, compliance, utilization, and cost-per-trip KPIs (ROI figures — Evidence Pending; validated by deployment metrics).
Conflicts-with | STN-004 (fleet-level richness vs edge simplicity — related C5); STN-009 (KPI liveness vs connectivity reality — related C7).
Notes | Consumes MR-1500 ECO fleet analytics; distinct from QA-manager needs (STN-005).
STN-015 / STN-019 / STN-011 — (co-primary; see SH-09 / SH-08 records)
---
Forward-trace note
Every STN above flows to User Needs (UN-###, Stage 4, ORCH-010) and thence to Product Requirements (PRD-###). Safety-relevant needs (notably STN-020, STN-001/018, STN-003, STN-002, STN-004) are carried into Hazard Analysis (HA-###, ORCH-017); conflicts (see STN-conflicts-and-tradeoffs.md) are carried into design trade studies. Governance and end-to-end matrix: ORCH-018 / EM-ENG-012.
Stakeholder Register & Influence/Interest Map — SCM
Evidence: 00-concept/concept-and-decisions.md (D1–D4), 01-market-requirements/ (MR-…), and
02-voice-of-customer/ (VOC-###). Governance: ORCH-018 / EM-ENG-012. This register enumerates every
stakeholder class that interacts with, depends on, procures, funds, administers, is protected by, or is
held accountable by the SCM across its reusable multi-trip lifecycle (D1): provision → transit →
receipt → excursion response → audit → reverse logistics → service → re-deploy.
1. Register
Fields: ID · Stakeholder class · What they do in the SCM lifecycle · Interest (what they care about) · Influence on purchase/design · Primary personas (MR-PER) · Primary segment (MR-SEG) · Key needs (STN) · Engagement priority.
SH-01 — Immunization / EPI program managers & Ministries of Health
Lifecycle role | Deploy devices in national/last-mile programs; accept/reject arriving consignments; own program-level cold-chain integrity.
Interest | PQS-eligible, rugged, low-maintenance monitoring; unambiguous accept/reject verdicts; last-mile visibility; program budget fit.
Influence | High (specify + deploy at national scale). · Personas: MR-PER-006. · Segment: MR-SEG-005/011.
Key needs | STN-002, STN-004, STN-009, STN-011, STN-020. · Engagement priority: Critical.
SH-02 — Pharma / vaccine manufacturers (QA & regulatory affairs)
Lifecycle role | Provision and release shipments; own the release-supporting temperature record; can mandate SCM downstream.
Interest | Audit-grade defensible data across every market; brand/liability protection; standardization; provisioning that doesn't slow dispatch.
Influence | Very high (anchor accounts; standard-setting). · Personas: MR-PER-001, MR-PER-009. · Segment: MR-SEG-001.
Key needs | STN-005, STN-008, STN-023, STN-006. · Engagement priority: Critical.
SH-03 — Distributors / wholesalers
Lifecycle role | Highest-frequency commercial shippers; book/track/reconcile many concurrent shipments; run reusable fleets.
Interest | GDP compliance; reduced spoilage write-offs; dispute settlement; action-only alerts; reuse economics.
Influence | High (core recurring-revenue buyer). · Personas: MR-PER-002. · Segment: MR-SEG-002.
Key needs | STN-001, STN-003, STN-006, STN-007, STN-008, STN-016, STN-018. · Engagement priority: Critical.
SH-04 — 3PL / logistics / air-cargo forwarders
Lifecycle role | Move shipments (air & ground) without opening them; run control towers; bear in-transit excursion liability.
Interest | Real-time telemetry to defend SLAs; handling neutrality; transport/battery compliance; integration.
Influence | High (channel-scale standardization). · Personas: MR-PER-004. · Segment: MR-SEG-004.
Key needs | STN-001, STN-003, STN-006, STN-012, STN-018, STN-013. · Engagement priority: High.
SH-05 — Clinic / health-facility staff (point of use)
Lifecycle role | Receive consignments; read status; accept/quarantine/reject; administer vaccine (last mile).
Interest | Instant, offline, self-evident accept/reject at the door; no phone/app/network required.
Influence | Medium (adoption/trust at the edge). · Personas: MR-PER-007. · Segment: MR-SEG-007.
Key needs | STN-004, STN-027, STN-020. · Engagement priority: Critical (safety-relevant).
SH-06 — Cold-chain technicians / field service
Lifecycle role | Clean, recharge, verify calibration, re-provision, and re-pool returned devices.
Interest | Fast, verifiable turnaround; clear charge/readiness state; calibration management; serviceable, sealed, cleanable hardware.
Influence | Medium-High (determines fleet utilization/margin). · Personas: MR-PER-011. · Segment: depots (MR-SEG-002/004).
Key needs | STN-010, STN-016, STN-021, STN-026. · Engagement priority: High.
SH-07 — Regulators / QA / auditors
Lifecycle role | Verify data integrity, calibration currency, excursion handling; grant/deny market access (WHO PQS, EU GDP, FDA).
Interest | Tamper-evident, ALCOA+ traceable, exportable records; calibration traceability; clear device classification.
Influence | Very high (market-access gatekeepers). · Personas: MR-PER-009. · Segment: cross-cutting; MR-REG.
Key needs | STN-005, STN-021, STN-022, STN-025. · Engagement priority: Critical.
SH-08 — Procurement / GAVI / UNICEF / donor bodies
Lifecycle role | Fund and run pooled tenders that determine LMIC market access at volume.
Interest | PQS listing; reuse TCO/cost-per-trip; supply assurance; demonstrable public-health impact; affordability within donor budgets.
Influence | Very high (largest aggregated volume channel). · Personas: MR-PER-008. · Segment: MR-SEG-006.
Key needs | STN-011, STN-019, STN-022, STN-020. · Engagement priority: Critical.
SH-09 — IT / security administrators
Lifecycle role | Validate and integrate the platform; manage access, data integrity, retention, and connected-device security.
Interest | CSV/GAMP validation; SSO/RBAC; data retention; secure signed OTA; open APIs / system integration.
Influence | High (enterprise deployment gate). · Personas: MR-PER-003. · Segment: MR-SEG-001/002 (enterprise).
Key needs | STN-013, STN-025, STN-005, STN-015. · Engagement priority: High.
SH-10 — Finance / asset accounting
Lifecycle role | Fund capital purchase; account for reusable fleet as an asset; own the ROI/TCO case.
Interest | Cost-per-trip evidence; capital vs. opex; asset utilization and depreciation; loss/write-off reduction.
Influence | Medium-High (authorizes spend). · Personas: buying committee (with MR-PER-010). · Segment: cross-cutting.
Key needs | STN-019, STN-014, STN-011. · Engagement priority: Medium-High.
SH-11 — End beneficiaries / patients (public-health outcome)
Lifecycle role | Not a device user — the party whose health outcome depends on potent vaccine being delivered and unspoiled dose not being wasted.
Interest | No compromised vaccine administered; no safe vaccine needlessly discarded; equitable last-mile access.
Influence | Indirect but paramount (the ultimate purpose). · Personas: — (excluded as persona per MR-1200). · Segment: outcome.
Key needs | STN-020 (represented on their behalf by SH-01/07/08). · Engagement priority: Critical (represented interest).
SH-12 — Installers (OEM-embed / field install)
Lifecycle role | Physically integrate/mount the device into insulated shippers/containers, or place it correctly per trip.
Interest | Standard, embeddable form factor; secure mounting/retention; clear placement guidance (product zone vs coolant).
Influence | Medium (OEM channel & data quality). · Personas: MR-PER-005. · Segment: MR-SEG-008.
Key needs | STN-024, STN-008, STN-002. · Engagement priority: Medium.
SH-13 — Service / reverse-logistics operations & sustainability
Lifecycle role | Run device return, cleaning, and re-pooling; own circularity and end-of-life.
Interest | Recoverable fleet; return workflow; cleanability/durability; battery-safe transport, disposal, and end-of-life.
Influence | Medium-High (reuse economics + sustainability compliance). · Personas: MR-PER-011. · Segment: MR-SEG-002/004.
Key needs | STN-016, STN-010, STN-026, STN-012. · Engagement priority: High.
SH-14 — Ecosystem partners (equipment OEMs, resellers, systems integrators)
Lifecycle role | Embed monitoring into shippers (OEMs); resell/integrate/support regionally (SIs/resellers).
Interest | Embeddable, interoperable device; open APIs; provisioning tooling; localization; partner enablement.
Influence | Medium-High (scalable distribution). · Personas: —. · Segment: MR-SEG-008/009.
Key needs | STN-017, STN-013, STN-024. · Engagement priority: Medium-High.
SH-15 — Executive / supply-chain leadership (strategic buyer)
Lifecycle role | Authorize fleet-scale, multi-year standardization; accountable for network loss, compliance exposure, KPIs.
Interest | Fleet-level ROI, network-wide loss reduction, cost-per-trip, standardization evidence.
Influence | Very high (budget & standardization). · Personas: MR-PER-010. · Segment: cross-cutting.
Key needs | STN-014, STN-015, STN-019, STN-011. · Engagement priority: High.
2. Influence / interest map (qualitative)
Placement is Judgment/Assumption pending VoC-quantified weighting (folder 02, §6). Rows are stakeholder classes; the two axes are influence on purchase/design and interest/stake in outcome.
Influence ↓ / Interest → | High interest | Medium interest
Very high influence | SH-02 Manufacturers · SH-07 Regulators · SH-08 Procurement/GAVI/UNICEF · SH-15 Executive | —
High influence | SH-01 EPI programs · SH-03 Distributors · SH-04 3PLs · SH-09 IT/security | SH-14 Ecosystem partners
Medium influence | SH-05 Clinic staff · SH-06 Service techs · SH-13 Service/reverse-logistics · SH-10 Finance | SH-12 Installers
Indirect (paramount stake) | SH-11 Beneficiaries/patients (managed for via SH-01/07/08) | —
Engagement implications:
Manage closely (high influence + high interest) | SH-01, SH-02, SH-03, SH-04, SH-07, SH-08, SH-15 — these set market access (PQS/GDP/FDA), anchor volume, and gate purchase; involve early and continuously.
Keep satisfied / design carefully for | SH-05, SH-06, SH-09, SH-10, SH-13 — decisive for field trust, reuse economics, enterprise deployment, and the ROI case.
Keep informed / enable | SH-12, SH-14 — scale distribution and data quality.
Represent explicitly | SH-11 — the public-health outcome is the reason the program exists; its need (STN-020) is the severity anchor for risk prioritization (CLAUDE.md rule 3).
3. Traceability
Backward: personas MR-PER-001…011; segments MR-SEG-001…011; VOC-### (folder 02); D1–D4. Forward: STN-### (see STN-needs-by-stakeholder.md) → UN-### (Stage 4) → PRD-###. Governance ORCH-018 / EM-ENG-012.

## 4 User Needs
User Needs
What the people who use the product need it to do for them, described as outcomes rather than features.
↩ Contents
In this section:
• Master User Needs Library — Smart Cold Chain Monitor (SCM)
• UN-000 — Product Vision (VIS)
• UN-100 — Stakeholders (STK)
• UN-1000 — Security & Privacy Needs (SEC)
• UN-1100 — Physical Needs (PHY)
• UN-1200 — Service Needs (SVC)
• UN-1300 — Regulatory Needs (REG)
• UN-1400 — Quality Needs (QUAL)
• UN-1500 — Ecosystem Needs (ECO)
• UN-1600 — Business Needs (BIZ)
• UN-1700 — Validation Needs (VAL)
• UN-200 — Personas (PER)
• UN-300 — Jobs To Be Done (JTBD)
• UN-400 — Use Environments (ENV)
• UN-500 — User Workflow (FLOW)
• UN-600 — Functional Needs (FUN)
• UN-700 — User Experience Needs (UX)
• UN-800 — Performance Needs (PERF)
• UN-900 — Safety Needs (SAFE)
Master User Needs Library — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 04 — User Needs (ORCH-010) · Owner: Chief Systems Engineer (EM-ENG-001) as integrator; Human-Factors/Usability contributor leads UX/PER/UN-SAFE use-error; Quality/Regulatory (EM-ENG-006) leads UN-REG · Status: Authored — instantiated, product-specific.
This library is the highest-level customer-driven baseline for the SCM lifecycle. It converts the Market Requirements (Stage 01), Voice of Customer (Stage 02), and Stakeholder Needs (Stage 03) into solution-neutral user needs — what users need the product to do and satisfy, never how it is implemented. It follows 00-governance/user-needs-standard.md and the 18-module map at 14-tools/market-copilot/user-needs-library/references/module-map.md, and is based on 00-concept/concept-and-decisions.md (foundational decisions D1–D4).
Every need uses the semantic ID scheme UN-<CODE>-NNN and carries the mandated 13-field block (ID · Title · Need Statement · Primary Stakeholder · Secondary Stakeholders · Importance · Market Priority · Source · Voice of Customer Reference · User Journey · Hazard Relationship · Traceability · Notes).
---
The cardinal rule (enforced throughout)
User Needs describe what users need, never how the product implements them. Every statement reads "the &lt;stakeholder&gt; needs …", is atomic, testable, and traceable, and carries no technical solution. Good: "The last-mile nurse needs to read the trip verdict offline, without a phone or training." Bad: "The device shall render the verdict on a 1.3-inch OLED via a BLE-provisioned rule engine."
---
Product in one line
The SCM is a reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shipping containers to continuously log temperature/humidity/shock, warn stakeholders of excursions in time to act, give an offline accept/reject verdict at the door, and provide an auditable, tamper-evident cold-chain record for product-disposition decisions. It is a monitoring/information device — it does not control temperature and is not therapeutic.
Anchor facts (D1–D4): reusable/rechargeable multi-trip (D1 → cleaning/return/re-provisioning needs) · local logging + scheduled cadence upload + immediate excursion alerts (D2 → alert-latency-awareness needs) · low-temperature-capable rechargeable Li-ion, duty-cycled, charge-inhibited below 0 °C (D3 → endurance/ charge needs) · multi-market WHO PQS + EU GDP + US FDA + air/ground transport (D4 → regulatory needs). Target endurance 90 days; operating range −30 °C to +60 °C.
---
Intended Use Statement
The Smart Cold Chain Monitor (SCM) is a reusable, cellular-connected environmental data logger intended to be placed inside insulated vaccine shipping containers, cartons, and cold boxes across the global vaccine distribution chain — from manufacturer pack-out, through distributors, 3PLs, and air/ground transport, to national immunization programs and the last-mile point of use. It is intended to continuously measure and record temperature (with humidity and shock), detect and flag both freezing and overheating excursions relative to configurable product-specific thresholds, alert accountable stakeholders of excursions in time to intervene where connectivity permits, present an unambiguous offline accept/reject verdict on the device at receipt, and provide an auditable, tamper-evident cold-chain record to support release, acceptance, and disposition decisions.
The SCM is an information/monitoring device that supports human and organizational decisions; it does not control temperature, is not the primary means of preserving the vaccine, and is not therapeutic. Its intended users span immunization/EPI program staff, clinic and last-mile health workers (including minimally trained users), cold-chain and calibration technicians, distributor and 3PL operators, manufacturer QA/regulatory affairs, regulators/auditors, IT/security administrators, installers/OEM integrators, and service/reverse-logistics operations. Its intended benefit is to protect the potency of vaccines reaching beneficiaries — ensuring no temperature-compromised dose is administered and no potent dose is needlessly wasted — while providing defensible, attributable evidence of cold-chain integrity across a reusable, multi-market fleet. Intended operating range is −30 °C to +60 °C with a 90-day endurance target; the device is intended for compliant use and transport under WHO PQS, EU GDP, US FDA data-integrity expectations, and air/ground dangerous-goods (Li-battery) regimes.
Evidence: this statement is derived from the concept intended-use (00-concept, §2), STN-020 (severity anchor), and D1–D4. The exact WHO PQS device category and US FDA classification are Evidence Pending (concept §7) and confirmed in the Regulatory stage.
---
Module table
Module | Code | File | Needs | Applicability
UN-000 Product Vision | VIS | [UN-000-product-vision.md](UN-000-product-vision.md) | 7 | Applicable
UN-100 Stakeholders | STK | [UN-100-stakeholders.md](UN-100-stakeholders.md) | 15 | Applicable
UN-200 Personas | PER | [UN-200-personas.md](UN-200-personas.md) | 11 | Applicable
UN-300 Jobs To Be Done | JTBD | [UN-300-jobs-to-be-done.md](UN-300-jobs-to-be-done.md) | 12 | Applicable
UN-400 Use Environments | ENV | [UN-400-use-environments.md](UN-400-use-environments.md) | 9 | Applicable
UN-500 User Workflow | FLOW | [UN-500-user-workflow.md](UN-500-user-workflow.md) | 11 | Applicable
UN-600 Functional | FUN | [UN-600-functional.md](UN-600-functional.md) | 18 | Applicable
UN-700 User Experience | UX | [UN-700-user-experience.md](UN-700-user-experience.md) | 12 | Applicable
UN-800 Performance | PERF | [UN-800-performance.md](UN-800-performance.md) | 11 | Applicable
UN-900 Safety | SAFE | [UN-900-safety.md](UN-900-safety.md) | 10 | Applicable
UN-1000 Security & Privacy | SEC | [UN-1000-security-privacy.md](UN-1000-security-privacy.md) | 9 | Applicable
UN-1100 Physical | PHY | [UN-1100-physical.md](UN-1100-physical.md) | 9 | Applicable
UN-1200 Service | SVC | [UN-1200-service.md](UN-1200-service.md) | 10 | Applicable
UN-1300 Regulatory | REG | [UN-1300-regulatory.md](UN-1300-regulatory.md) | 10 | Applicable
UN-1400 Quality | QUAL | [UN-1400-quality.md](UN-1400-quality.md) | 7 | Applicable
UN-1500 Ecosystem | ECO | [UN-1500-ecosystem.md](UN-1500-ecosystem.md) | 9 | Applicable
UN-1600 Business | BIZ | [UN-1600-business.md](UN-1600-business.md) | 8 | Applicable
UN-1700 Validation | VAL | [UN-1700-validation.md](UN-1700-validation.md) | 9 | Applicable
Total |  | 18 modules | 187 | 18 / 18 applicable
---
Applicability Register
Per the user-needs standard, scope is never silently dropped: every module is emitted, and any exclusion is recorded with a rationale. For the SCM, all 18 modules are Applicable — the product is a globally distributed, regulated, reusable, connected, safety-significant IoT device, so every module carries real customer-driven needs.
Module (code) | Status | Rationale
All 18 (VIS … VAL) | Applicable | The SCM is a multi-market (WHO PQS/EU GDP/US/air-transport), reusable, connected, audit-instrument device serving a dozen stakeholder classes across a multi-trip lifecycle. Every module maps to real user needs: vision/stakeholders/personas/JTBD/environments (who and what), functional/UX/performance (the product's job), safety/security/regulatory/quality (trust and compliance), physical/service (reusable hardware and its loop), ecosystem/business (fleet and commercial), and validation (how needs are confirmed).
Borderline topics flagged inside modules (confirm, not excluded):
• UN-REG — exact WHO PQS device category/code and US FDA classification/pathway are Evidence Pending (concept §7); confirmed in the Regulatory stage. Standards are cited by designation; record IDs are not fabricated.
• UN-PERF / UN-PHY — OLED low-temperature legibility thresholds, gloved-dexterity thresholds, cycle life, and RF/GPS-in-container performance are Evidence Pending (Stage 02 §6 field program; MR-VAL-006).
• UN-BIZ / UN-VAL — price, cost-per-trip, TCO, return rate, ROI, and false-alarm tolerance are Judgment/Assumption ("Evidence Pending"); no precise financials or field rates are asserted as fact.
• UN-FUN-016 / UN-PERF-009 (location) — continuous GPS is treated as a boundary topic subordinate to the energy budget (conflict C3), not a committed always-on capability.
---
Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, every need is based and honestly labelled. Needs resting on a published standard (WHO PQS, EU GDP, 21 CFR Part 11 / ALCOA+, UN 38.3 / IATA, IEC 62368-1 / 60068 / 60529, IEC 62366, ISO 14971, ISO/IEC 17025) are marked Standard. Needs resting on widely reported public-health facts (WHO/UNICEF/GAVI freezing and wastage reporting) are Cited (reported/estimated). Needs derived from synthesized VoC / stakeholder inference are Judgment/Assumption, and unresolved quantities are Evidence Pending — never invented. Because the primary VoC field program has not yet executed (Stage 02 §2), success measures are targets to be set in Product Requirements and closed by Verification/Validation. The D3 low-temperature battery tension (90-day endurance + charge safety at −30 °C + capacity fade over reuse) is carried as the top risk and closed by test (UN-VAL-007 / MR-VAL-006).
---
Stakeholder conflicts carried forward (from Stage 3)
The ten Stage-3 conflicts (C1–C10, 03-stakeholder-needs/STN-conflicts-and-tradeoffs.md) are carried into these needs, not silently reconciled — each is named in the relevant needs' Notes/Traceability. The severity-first tie-breaker (CLAUDE.md rule 3) is UN-VIS-001 / UN-STK-014 / UN-SAFE-001 (STN-020): no compromised dose administered, no safe dose wasted — the beneficiary/patient outcome dominates when needs conflict. Key conflicts and where they surface: C1 cost vs cold endurance (UN-BIZ-001, UN-PERF-003, UN-VIS-006); C2 live data vs battery (UN-FUN-006/010, resolved-direction by D2); C3 GPS vs endurance (UN-PERF-009, UN-FUN-016); C4 rigor vs dispatch speed (UN-UX-005/007, UN-SAFE-006); C5 rich data vs simple verdict (UN-VIS-007, UN-UX-011); C6 reuse TCO vs reverse-logistics cost (UN-SVC-001, UN-BIZ-002); C7 real-time vs no-connectivity (UN-FUN-009, UN-ENV-007, resolved-direction by D2); C8 sensitivity vs false-alarm (UN-FUN-005, UN-SAFE-004); C9 security lockdown vs edge/open integration (UN-SEC-003/004, UN-ECO-005); C10 sustainability vs Li-battery constraints (UN-SVC-007, UN-REG-006/008).
---
Traceability note (chain position)
This library is the authoritative customer-driven baseline and the primary input to Product Requirements. Backward, each need traces to STN-### (Stage 03), VOC-### (Stage 02), MR-… (Stage 01), and D1–D4 (concept). Forward, each need seeds one or more Product Requirements (PRD-<CODE>-NNN, ORCH-011, Stage 05) and one or more validations; safety-relevant needs seed the Preliminary Hazard Analysis (HA-###, ORCH-017) and Risk-Control Requirements.
Flow: MR (Stage 01) → VOC-### (Stage 02) → STN-### (Stage 03) → UN-&lt;CODE&gt;-NNN (this library, Stage 04) → PRD-&lt;CODE&gt;-NNN (Stage 05) → System/Subsystem Requirements → Architecture → Design Outputs (ORCH-013) → Hazard Analysis (ORCH-017) / ERFKL (ORCH-016) → Risk-Control Requirements → Verification (ORCH-014) / Validation (ORCH-015).
The end-to-end trace matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18). Each need's Traceability field carries the forward placeholders (PRD-SCM-…, PV-…, VAL-…, HA-…) that the matrix will resolve.
Convening: ORCH-010 (User Needs). Integrator: EM-ENG-001. Human-Factors/Usability contributor: UX/PER/ use-error needs. Quality/Regulatory (EM-ENG-006): UN-REG and standards citations. This baseline feeds ORCH-011 (Product Requirements, Stage 05) and ORCH-018 / EM-ENG-012 for lifecycle traceability.
UN-000 — Product Vision (VIS)
Master User Needs Library — Smart Cold Chain Monitor (SCM). Stage 04 (ORCH-010). Traceability: this module (code VIS) is based on 00-concept/concept-and-decisions.md (D1–D4), the Market Requirements Library (Stage 01), Voice of Customer (Stage 02), and Stakeholder Needs (Stage 03); it feeds forward to Product Requirements (Stage 05, ORCH-011) and the end-to-end trace matrix (ORCH-018 / EM-ENG-012).
1. Purpose
This module states, in solution-neutral user terms, the shared outcome the Smart Cold Chain Monitor (SCM) exists to deliver: that vaccines reach the people who need them potent, and that no one is forced to choose between administering a possibly-spoiled dose and discarding a possibly-good one. It captures the highest-level needs — the "why" — from which every other module is derived. It is not a feature list; it is the statement of what the stakeholder community needs to be true of the world once the SCM is in use.
2. Scope
In scope: vision-level needs that span the whole stakeholder community and the reusable multi-trip lifecycle (D1) — the outcome (potent vaccine delivered, correct disposition), the value proposition (prevention over documentation), the differentiation the market needs, and the success conditions. Out of scope: how any of this is implemented; specific functions (see UN-FUN), performance targets (UN-PERF), and business/commercial mechanics (UN-BIZ).
3. Applicable Standards & References
• WHO PQS E006 — temperature-monitoring device performance category (device eligibility gate; exact category Evidence Pending, concept §7).
• EU GDP (2013/C 343/01) and US FDA 21 CFR Part 11 — the compliance frame the record must satisfy.
• ISO 14971 — the risk frame within which "no compromised dose administered" is the severity anchor.
• WHO/UNICEF/GAVI public-health reporting on cold-chain wastage — cited as reported/estimated.
4. Background
A cold chain is a relay of custodians; a single break, often silent (freezing), can destroy potency without any visible sign. Today the loss is frequently discovered only at delivery — too late to rescue the stock or attribute the fault. The SCM's reason to exist is to move that knowledge forward in time (warn before it is too late) and to make the resulting record trustworthy enough to base a release/accept/reject decision on. Because the concept is reusable, cellular, rechargeable, and multi-market (D1–D4), the vision must also encompass fleet-scale trust, sustainability, and global regulatory acceptance — not just a single trip.
5. Need Categories
• 5.1 Outcome needs — what must be true for the beneficiary/patient and the program.
• 5.2 Value-proposition needs — why a stakeholder chooses this over the status quo.
• 5.3 Differentiation & success needs — what the market needs the product to prove.
6. Detailed User Needs
6.1 Outcome needs
UN-VIS-001 — No compromised dose given; no safe dose wasted
Need Statement | The immunization program and the beneficiary they serve need assurance that no vaccine compromised by a temperature excursion is administered, and that no potent vaccine is needlessly discarded.
Primary Stakeholder | End beneficiaries / patients (represented by EPI programs, SH-11 / SH-01).
Secondary Stakeholders | Regulators/QA (SH-07), procurement (SH-08), manufacturers (SH-02).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Stakeholder Interview (represented) / Regulatory
Voice of Customer Reference | STN-020; VOC-011/021; MR-PROB-010.
User Journey | Spans the whole journey — provisioning → transit → receipt → disposition → point of care.
Hazard Relationship | The program's severity anchor (ISO 14971). Both failure directions cause harm: undetected excursion → spoiled dose administered → patient harm / ineffective immunization; and false "reject" → safe dose discarded → stock-out → unvaccinated patient. This need is the tie-breaker when other needs conflict (CLAUDE.md rule 3).
Traceability | Back: STN-020, MR-PROB-010, concept intended-use. Forward: all safety needs (UN-SAFE-), verdict needs (UN-FUN-, UN-UX-), and validation (UN-VAL-, HA-### via ORCH-017).
Notes | The device supports the disposition decision; it is not therapeutic and does not make the decision. Judgment: this is the north star for every downstream trade-off.
UN-VIS-002 — Know a break is happening in time to act
Need Statement | Shipment-accountable stakeholders need to learn that a cold-chain break is occurring early enough to intervene (reroute, re-ice, divert, escalate), rather than only discovering it after the product is already lost.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | EPI programs (SH-01), manufacturers (SH-02).
Importance | Essential
Market Priority | Must-have
Source | VoC / JTBD
Voice of Customer Reference | STN-001; VOC-008/021; MR-PROB-001; MR-JTBD-005; D2.
User Journey | Transit and hand-off legs of the journey.
Hazard Relationship | A delayed or missing warning (ISO 14971) means a rescuable excursion becomes an unrecoverable loss; alert latency is connectivity-bound and carried as a hazard.
Traceability | Back: STN-001, MR-JTBD-005. Forward: UN-FUN alerting needs, UN-PERF alert-latency need, hazard analysis (HA-###).
Notes | D2 sets direction (immediate excursion alert when connectivity is available); the vision-level need is prevention over documentation.
UN-VIS-003 — Trust the cold-chain record enough to decide on it
Need Statement | Everyone who must release, accept, reject, or defend a consignment needs a cold-chain record they can trust as complete, attributable, and untampered — trustworthy enough to base a real disposition and regulatory defence on.
Primary Stakeholder | Manufacturer QA/RA (SH-02), regulators/auditors (SH-07).
Secondary Stakeholders | Distributors (SH-03), programs (SH-01), IT/security (SH-09).
Importance | Essential
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-005; VOC-001/003; MR-PROB-005; MR-JTBD-008.
User Journey | Receipt, disposition, and audit legs.
Hazard Relationship | An incomplete or editable record (ISO 14971) can cause a wrong disposition (spoiled dose accepted, or good dose rejected) and fails audit, invalidating the release decision.
Traceability | Back: STN-005/023. Forward: UN-SEC data-integrity needs, UN-REG needs, UN-VAL audit-sim.
Notes | Standard basis: EU GDP + 21 CFR Part 11 + ALCOA+. This is what makes the SCM an instrument of record, not merely a sensor.
6.2 Value-proposition needs
UN-VIS-004 — Prevent loss, don't just document it
Need Statement | Cold-chain operators need monitoring that helps them prevent product loss and disputes, not merely produce a post-mortem of a loss that already happened.
Primary Stakeholder | Distributors (SH-03), executive/supply-chain leadership (SH-15).
Secondary Stakeholders | 3PLs (SH-04), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | Market Analysis / VoC
Voice of Customer Reference | STN-001/014; VOC-008/020; MR-COMP; MR-JTBD-005/010.
User Journey | Whole lifecycle; realized in transit (rescue) and at fleet level (loss reduction).
Hazard Relationship | Indirect — prevention reduces the frequency of spoiled-dose exposure across the network; supports the UN-VIS-001 outcome.
Traceability | Back: STN-001/014, MR-COMP. Forward: UN-FUN alerting, UN-ECO fleet analytics, UN-BIZ ROI.
Notes | This is the reason a buyer would pay more than for a passive USB logger (VOC-008).
UN-VIS-005 — One trustworthy monitoring approach across the whole chain and all markets
Need Statement | Manufacturers and programs need a single, consistent monitoring approach they can rely on across every partner, lane, and market, so the cold-chain story is coherent end-to-end and the brand/program is protected wherever product ships.
Primary Stakeholder | Manufacturer QA/supply-chain (SH-02).
Secondary Stakeholders | Executive leadership (SH-15), programs (SH-01), regulators (SH-07).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-023; VOC-001/003; MR-SEG-001/010; D4.
User Journey | Standardization decision (buyer) and every downstream custody leg.
Hazard Relationship | Indirect — fragmentation of monitoring raises the chance of unmonitored legs and inconsistent disposition decisions.
Traceability | Back: STN-023, MR-EXEC-005, D4. Forward: UN-REG multi-market needs, UN-ECO integration.
Notes | Multi-market conformance (WHO PQS + EU GDP + US) is the enabler; anchor-account leverage.
6.3 Differentiation & success needs
UN-VIS-006 — The reusable model must actually pay off over its life
Need Statement | Procurement, finance, and program buyers need confidence that the reusable, multi-trip model delivers a lower, provable lifetime cost per trip than the single-use status quo — otherwise it cannot be chosen no matter how capable it is.
Primary Stakeholder | Procurement / GAVI / UNICEF (SH-08), finance (SH-10).
Secondary Stakeholders | Executive leadership (SH-15), service ops (SH-13), programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-011/019; VOC-012/013/020; MR-EXEC-008; D1.
User Journey | Procurement decision, then the recover→refurbish→redeploy loop across many trips.
Hazard Relationship | Not a direct hazard; economic viability that determines whether the safety benefit is deployed at scale at all.
Traceability | Back: STN-019, MR-EXEC-008, D1. Forward: UN-BIZ TCO needs, UN-SVC turnaround needs, UN-VAL return-rate pilot.
Notes | Cost-per-trip is the linchpin of the reusable business case; return rate (UN-SVC) is decisive. Conflict C1 (cost vs −30 °C/90-day endurance) is the headline tension carried from Stage 3.
UN-VIS-007 — Serve the least-trained edge user and the most exacting auditor with the same truth
Need Statement | The stakeholder community needs one device that gives an untrained, offline last-mile health worker an unambiguous accept/reject answer while giving an auditor a complete, defensible record — both drawn from the same underlying truth.
Primary Stakeholder | Last-mile health worker (SH-05), regulators/QA (SH-07).
Secondary Stakeholders | EPI programs (SH-01), executive leadership (SH-15).
Importance | Essential
Market Priority | Differentiator
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004/005; conflict C5; MR-PER-006/007/009.
User Journey | Point-of-use receipt (edge) and audit (back office) — the two ends of one journey.
Hazard Relationship | A data-rich interface overwhelms the edge user (wrong disposition at point of care); a bare verdict fails the audit — both are ISO 14971 concerns handled by layered, role-appropriate presentation of the same data.
Traceability | Back: STN-004/005, conflict C5. Forward: UN-UX layered-information needs, UN-FUN verdict + record needs, UN-VAL persona usability validation.
Notes | This "one device, many literacies" tension (MR-1200) is a defining vision constraint.
7. Validation Notes
Vision needs are validated indirectly — by the success of the needs derived from them — and directly by market/field validation (UN-VAL): pilot loss-reduction evidence (UN-VIS-002/004), audit-simulation acceptance of records (UN-VIS-003), correct-disposition validation (UN-VIS-001, ties to HA-### / validation), and a return-rate/TCO pilot (UN-VIS-006). Quantitative outcomes remain Evidence Pending until the VoC and pilot programs run (Stage 02 §6).
8. Traceability Placeholders
Upstream | concept D1–D4; MR-EXEC/PROB/JTBD; VOC themes T1–T14; STN-001/005/011/019/020/023.
Downstream | every other UN module derives from this vision; forward to PRD (ORCH-011), Hazard Analysis (ORCH-017), Validation (ORCH-015). End-to-end matrix: ORCH-018 / EM-ENG-012.
UN-100 — Stakeholders (STK)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on the Stakeholder Register (03-stakeholder-needs/STN-stakeholder-register.md), MR-1200 personas, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module records, per stakeholder class, the single most-defining need that class has of the SCM — the need that, if unmet, causes that stakeholder to reject, work around, or fail to adopt the product. It makes the stakeholder map actionable by attaching one anchor need to each class; the detailed, atomic needs live in the topical modules (UN-FUN, UN-UX, UN-SVC, etc.) and are cross-referenced here.
2. Scope
In scope: all fifteen stakeholder classes (SH-01…SH-15) in the register, including the non-user beneficiary (SH-11) represented by proxies, and low-training/vulnerable users flagged where relevant. Out of scope: persona-level detail (UN-PER), and the full topical needs (other modules).
3. Applicable Standards & References
• EU GDP (Responsible Person, distributor/personnel roles), 21 CFR Part 11 (IT/data-admin role), WHO PQS E006 (program/operator role), UN 38.3 / IATA DGR (3PL role), ISO/IEC 17025 concepts (calibration technician role). All cited as Standard; role-specific quantitative attributes are Evidence Pending pending the VoC program (Stage 02 §6).
4. Background
A reusable SCM passes through a dozen hands on one trip and returns for another. Because it is reusable (D1), classes absent from a single-use logger's life — service/reverse-logistics (SH-13), calibration technicians (SH-06), fleet/IT admins (SH-09) — are first-class stakeholders. Literacy and accountability span from a supply-chain director reading fleet KPIs to a last-mile nurse reading a two-line OLED offline.
5. Need Categories
• 5.1 Custody & provisioning (SH-02, SH-03, SH-09) · 5.2 Transit & handling (SH-04) · 5.3 Receipt & point-of-use (SH-01, SH-05) · 5.4 Governance & procurement (SH-07, SH-08, SH-10) · 5.5 Sustainment & ecosystem (SH-06, SH-12, SH-13, SH-14) · 5.6 Outcome & leadership (SH-11, SH-15).
6. Detailed User Needs
6.1 Custody & provisioning
UN-STK-001 — Manufacturer QA/RA: a defensible record without slowing dispatch
Need Statement | The manufacturer QA/regulatory-affairs stakeholder needs to release each shipment with a temperature record they can defend to any inspector, produced without adding time that would slow high-throughput pack-out.
Primary Stakeholder | Pharma manufacturers — QA/RA (SH-02).
Secondary Stakeholders | Regulators/auditors (SH-07), distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-005/008/023; VOC-001/002; MR-PER-001.
User Journey | Provisioning and release-to-ship.
Hazard Relationship | Slow rigor drives skipped/void trips (no record → undetected excursion, ISO 14971); editable records fail audit and invalidate release.
Traceability | Back: STN-005/008. Forward: UN-FUN provisioning + record, UN-SEC integrity, UN-UX speed.
Notes | Provisioning speed is a covert adoption gate (conflict C4). QA manager is the economic approver.
UN-STK-002 — Distributor: shipment status and action-only alerts, no data interpretation
Need Statement | The distributor coordinator needs to know each shipment's status and receive alerts only when action is required, without having to interpret raw sensor data across hundreds of concurrent shipments.
Primary Stakeholder | Distributors / wholesalers (SH-03).
Secondary Stakeholders | 3PLs (SH-04), programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-001/007/018; VOC-005/008; MR-PER-002.
User Journey | Daily dispatch monitoring and excursion response.
Hazard Relationship | Alarm fatigue (ISO 14971) leads to muted alerts and a missed real excursion.
Traceability | Back: STN-007/018. Forward: UN-FUN alerting/routing, UN-UX action-only signaling.
Notes | Daily active user; adoption drives per-trip pull-through. Alert fatigue is the failure mode.
UN-STK-003 — IT/security admin: clear our validation, access, and integration bar
Need Statement | The enterprise IT/security administrator needs the device and its platform to satisfy computerized-system validation, single sign-on and role-based access, data-retention rules, and secure integration with existing systems before it can be allowed onto the network.
Primary Stakeholder | IT / security administrators (SH-09).
Secondary Stakeholders | Regulators/QA (SH-07), ecosystem partners (SH-14).
Importance | Important
Market Priority | Must-have (enterprise segment)
Source | Stakeholder Interview / Regulatory
Voice of Customer Reference | STN-013/025; VOC-018/019; MR-PER-003.
User Journey | Enterprise onboarding and ongoing fleet administration.
Hazard Relationship | Uncontrolled access / unsigned updates (ISO 14971 + security) can corrupt records and disposition decisions.
Traceability | Back: STN-013/025. Forward: UN-SEC access/OTA needs, UN-ECO integration.
Notes | A siloed device with no API stalls at procurement (conflict C9 with edge simplicity).
6.2 Transit & handling
UN-STK-004 — 3PL: invisible to handlers, safe to fly, never a hold reason
Need Statement | The 3PL operations stakeholder needs the device to be handling-neutral, safe for air/ground transport, and never the reason a shipment is delayed for dangerous-goods paperwork.
Primary Stakeholder | 3PL / logistics / air-cargo (SH-04).
Secondary Stakeholders | Service/reverse-logistics (SH-13), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-012; VOC-007; MR-PER-004; D3/D4.
User Journey | High-throughput handling and air/ground transport legs.
Hazard Relationship | Non-compliant Li-battery transport (UN 38.3 / IATA) is a transport-safety hazard and a channel-exclusion risk.
Traceability | Back: STN-012. Forward: UN-REG transport needs, UN-PHY handling-neutral form, UN-SAFE.
Notes | 3PL SOP inclusion multiplies pull across many shippers.
6.3 Receipt & point-of-use
UN-STK-005 — EPI program manager: a clear, PQS-aligned accept/reject verdict
Need Statement | The immunization/EPI program manager needs an unambiguous, WHO-PQS-consistent accept/reject verdict for arriving consignments, so accept/reject decisions are correct and defensible across a national program.
Primary Stakeholder | Immunization / EPI program managers & MoH (SH-01).
Secondary Stakeholders | Clinic staff (SH-05), procurement (SH-08), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | STN-004; VOC-009/014; MR-PER-006.
User Journey | Consignment arrival and disposition.
Hazard Relationship | Wrong verdict (ISO 14971) → spoiled dose accepted or safe dose rejected; ties to the UN-VIS-001 severity anchor.
Traceability | Back: STN-004. Forward: UN-FUN verdict logic, UN-UX offline legibility, UN-REG PQS.
Notes | Verdict rule-sets must be configurable and PQS-consistent (category TBC — Evidence Pending).
UN-STK-006 — Clinic / last-mile staff: decide accept/quarantine/reject in seconds, offline, untrained
Need Statement | The clinic/last-mile health worker — often minimally trained, gloved, and without connectivity — needs to decide whether to fridge, quarantine, or return arriving stock in seconds, reading the device itself with no phone, network, or manual.
Primary Stakeholder | Clinic / health-facility staff (SH-05). (Vulnerable/low-training user class.)
Secondary Stakeholders | EPI programs (SH-01), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004/027; VOC-014/015; FLD-006; MR-PER-007.
User Journey | Point-of-use receipt at rural health post / mobile team.
Hazard Relationship | Misread status by an untrained user (ISO 14971, IEC 62366 use-error) → wrong disposition at the point of care — highest-consequence edge decision.
Traceability | Back: STN-004/027. Forward: UN-UX low-literacy legibility, UN-FUN offline verdict, UN-VAL.
Notes | OLED + single button is the sole interface here; design for the least-trained, most-remote user.
6.4 Governance & procurement
UN-STK-007 — Regulator/auditor: verify integrity, calibration, and excursion handling without ambiguity
Need Statement | The regulator/auditor needs to verify data integrity, calibration currency, and how excursions were handled, from records that are complete, attributable, and unambiguous.
Primary Stakeholder | Regulators / QA / auditors (SH-07).
Secondary Stakeholders | Manufacturers (SH-02), service technicians (SH-06), IT (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-005/021/022; MR-PER-009.
User Journey | Audit / inspection and disposition defence.
Hazard Relationship | Ambiguous or incomplete records (ISO 14971) → inadmissible evidence → wrong or indefensible disposition.
Traceability | Back: STN-005/021. Forward: UN-SEC integrity, UN-SVC calibration, UN-REG.
Notes | Make the SCM record the easiest artifact in the audit — a differentiator, not a liability.
UN-STK-008 — Procurement (GAVI/UNICEF): PQS eligibility and proven reuse TCO at scale
Need Statement | The institutional procurement stakeholder needs the device to be PQS-eligible and to come with documented reuse total-cost-of-ownership evidence and supply assurance, or it cannot be placed into a pooled tender.
Primary Stakeholder | Procurement / GAVI / UNICEF (SH-08).
Secondary Stakeholders | Finance (SH-10), programs (SH-01), executive (SH-15).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-011/019; VOC-012/013; MR-PER-008.
User Journey | Tender qualification and framework award.
Hazard Relationship | Not a direct hazard; gates whether the safety benefit reaches LMIC scale.
Traceability | Back: STN-011/019. Forward: UN-REG PQS, UN-BIZ TCO/supply, UN-VAL market validation.
Notes | Largest aggregated LMIC volume channel; PQS listing has long lead times — plan early.
UN-STK-009 — Finance: cost-per-trip and asset-accounting evidence over device life
Need Statement | The finance/asset-accounting stakeholder needs documented cost-per-trip, utilization, and depreciation evidence across the reusable device life to justify capital outlay against loss reduction.
Primary Stakeholder | Finance / asset accounting (SH-10).
Secondary Stakeholders | Procurement (SH-08), executive (SH-15).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-019; VOC-020/060; MR-EXEC-008.
User Journey | Business-case approval and ongoing asset management.
Hazard Relationship | None direct; economic gate for scale.
Traceability | Back: STN-019. Forward: UN-BIZ cost-per-trip, UN-ECO fleet utilization reporting.
Notes | Return rate (UN-SVC) is the swing variable in the TCO model (conflict C6).
6.5 Sustainment & ecosystem
UN-STK-010 — Calibration/service technician: fast, verifiable turnaround of returned devices
Need Statement | The cold-chain/calibration technician needs to quickly confirm a returned device is charged, in-calibration, and firmware-current, and re-provision it, without hand-checking each unit, to keep the reusable fleet moving.
Primary Stakeholder | Cold-chain technicians / field service (SH-06).
Secondary Stakeholders | Service ops (SH-13), regulators (SH-07).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / Stakeholder Interview
Voice of Customer Reference | STN-010/021; VOC-016/017; MR-PER-011.
User Journey | Return → inspect → refurbish → recalibrate → redeploy.
Hazard Relationship | An out-of-cal or under-charged device redeployed (ISO 14971) → disputed readings or dead-in-cold on the next trip.
Traceability | Back: STN-010/021. Forward: UN-SVC turnaround/calibration, UN-PERF charge readiness.
Notes | Turnaround time sets device utilization and reuse margin; charge-inhibit-below-0 °C (D3) affects readiness timing.
UN-STK-011 — Installer / OEM integrator: it must fit, mount securely, and guide correct placement
Need Statement | The installer/OEM integrator needs a device that fits and mounts securely inside insulated shippers/containers, with clear guidance on probe placement so readings reflect the product zone.
Primary Stakeholder | Installers (OEM-embed / field install) (SH-12).
Secondary Stakeholders | Ecosystem partners (SH-14), warehouse operators (SH-05).
Importance | Important
Market Priority | Differentiator
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-024; CFB-013; FLD-002; MR-PER-005.
User Journey | Installation / OEM embedding and pre-trip placement.
Hazard Relationship | Wrong probe placement (ISO 14971) biases freeze/heat detection → missed excursion or false alarm.
Traceability | Back: STN-024. Forward: UN-PHY form/mounting, UN-FUN placement guidance, UN-UX guidance.
Notes | Poor fit blocks OEM embedding, a key channel; flag mechanical + materials engineers.
UN-STK-012 — Service / reverse-logistics: recover, clean, and close the loop safely
Need Statement | The service/reverse-logistics stakeholder needs a practical way to get reusable devices back, clean them between trips without damage, and dispose of batteries/end-of-life units safely and compliantly across many reuse cycles.
Primary Stakeholder | Service / reverse-logistics operations (SH-13).
Secondary Stakeholders | Distributors (SH-03), 3PLs (SH-04), sustainability.
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / Market Analysis
Voice of Customer Reference | STN-016/026; VOC-006; FLD-007/011; MR-PROB-011.
User Journey | Return logistics, cleaning, refurbishment, end-of-life.
Hazard Relationship | Improper cleaning (ingress → sensor error) or unsafe battery disposal (ISO 14971, UN 38.3) — safety and environmental.
Traceability | Back: STN-016/026. Forward: UN-SVC return/cleaning, UN-PHY sealing, UN-REG battery EOL.
Notes | Return rate determines whether reuse TCO holds (conflict C6/C10).
UN-STK-013 — Ecosystem partner (OEM/reseller/SI): embeddable, integrable, supportable in-region
Need Statement | The ecosystem partner needs an embeddable form factor, secure open integration points, provisioning tooling, and localized enablement so they can build in, resell, integrate, and support the device in their market.
Primary Stakeholder | Ecosystem partners — OEMs/resellers/SIs (SH-14).
Secondary Stakeholders | Installers (SH-12), IT/security (SH-09).
Importance | Desirable
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-017; CFB-013/014; MR-SEG-008/009.
User Journey | Partner onboarding, integration, and regional support.
Hazard Relationship | Open integration vs security lockdown is conflict C9 (indirect data-integrity).
Traceability | Back: STN-017. Forward: UN-ECO API/partner needs, UN-PHY embeddable form.
Notes | Designed-in distribution channel into fragmented markets.
6.6 Outcome & leadership
UN-STK-014 — Beneficiary/patient (represented): the outcome must be protected
Need Statement | The beneficiary/patient — not a device user — needs, through their proxies, assurance that the monitoring exists to protect the potency of the dose they receive and to avoid needless stock-outs.
Primary Stakeholder | End beneficiaries / patients (SH-11), represented by SH-01/07/08.
Secondary Stakeholders | All.
Importance | Essential
Market Priority | Must-have
Source | Stakeholder Interview (represented) / Regulatory
Voice of Customer Reference | STN-020; MR-PROB-010.
User Journey | Point of care (outcome), spanning the whole chain.
Hazard Relationship | The severity anchor (ISO 14971); see UN-VIS-001.
Traceability | Back: STN-020. Forward: all safety needs; the tie-breaker in conflicts.
Notes | Represented, not a direct user; carries the top severity reference for the program.
UN-STK-015 — Executive/supply-chain leadership: network loss reduction, compliance, and cost-per-trip
Need Statement | The executive/supply-chain leader needs fleet-level evidence of cold-chain loss reduction, reduced compliance exposure, and cost-per-trip, to fund and standardize on the program at scale.
Primary Stakeholder | Executive / supply-chain leadership (SH-15).
Secondary Stakeholders | Finance (SH-10), IT (SH-09).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-014; VOC-020; MR-PER-010.
User Journey | Sponsorship, standardization, and multi-year contract decisions.
Hazard Relationship | None direct; funds deployment of the safety benefit at scale.
Traceability | Back: STN-014. Forward: UN-ECO fleet KPIs, UN-BIZ scalability.
Notes | Cares about the network, not one device; distinct from the QA manager's needs.
7. Validation Notes
Each anchor need is validated where its detailed derivatives are validated (topical modules). Class existence and priority are High-confidence (based on the register); quantitative weighting per class is Evidence Pending until the VoC program quantifies frequency and influence (Stage 02 §6).
8. Traceability Placeholders
Upstream | STN-stakeholder-register; MR-1200 personas; D1–D4.
Downstream | each anchor need points into topical UN modules and forward to PRD (ORCH-011). Matrix: ORCH-018 / EM-ENG-012.
UN-1000 — Security & Privacy Needs (SEC)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on STN-013/025, 21 CFR Part 11 / ALCOA+, GAMP 5, and D2 (connectivity/OTA). Forward to PRD (ORCH-011), cybersecurity design, and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users need for the SCM's data and connectivity to be trustworthy and controlled — a cellular, OTA-updatable device on shipments is both an attack surface and a data-integrity dependency for disposition decisions. Needs are solution-neutral ("the user needs updates to be trustworthy," not "use code-signing algorithm X").
2. Scope
In scope: data integrity/tamper-evidence, access control (SSO/RBAC), secure connectivity, trustworthy firmware update, data ownership/retention, auditability, and privacy of any personal/location data. Out of scope: the security mechanisms/architecture (design stage) and the safety framing of security (UN-SAFE-010).
3. Applicable Standards & References
• FDA 21 CFR Part 11 — electronic records/signatures, audit trails, access control.
• ALCOA+ — data-integrity principles (attributable, legible, contemporaneous, original, accurate, +).
• GAMP 5 / ISPE data-integrity guidance — computerized-system validation and governance.
• EU GDP — data-governance expectations for distribution records.
• IoT security good practice (e.g., signed/fail-safe update, least privilege) — Judgment/Standard; a formal threat model is Evidence Pending (set in PRD/hazard with the cybersecurity persona).
4. Background
Because a disposition and an audit rest on the record, its integrity is a first-order need: it must be tamper-evident and attributable. Connectivity and OTA introduce threats — spoofed data, unauthorized access, malicious/failed updates — that must be controlled without destroying the frictionless offline edge use (conflict C9). Enterprise deployment is gated by IT security and CSV (STN-013).
5. Need Categories
• 5.1 Data integrity & tamper-evidence · 5.2 Access control & identity · 5.3 Secure connectivity & update · 5.4 Data ownership, retention & privacy · 5.5 Auditability.
6. Detailed User Needs
6.1 Data integrity & tamper-evidence
UN-SEC-001 — Trust that records have not been altered
Need Statement | Users who rely on the record need it to be tamper-evident — any alteration detectable — so an excursion record cannot be quietly edited and a disposition or audit rests on trustworthy data.
Primary Stakeholder | Regulators/auditors (SH-07), QA/RA (SH-02).
Secondary Stakeholders | IT (SH-09), distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / VoC
Voice of Customer Reference | STN-005/025; VOC-001/019; 21 CFR Part 11; ALCOA+.
User Journey | Release; audit; dispute.
Hazard Relationship | Editable record (ISO 14971 + integrity) → wrong/indefensible disposition.
Traceability | Back: STN-005/025. Forward: PRD tamper-evidence, UN-FUN record, UN-REG Part 11.
Notes | Overlaps UN-FUN-014 (function side); this is the integrity guarantee side.
UN-SEC-002 — Attribute every record and action to a source
Need Statement | Users need each record and change attributable to a device, user, or action (the "attributable" and "original" of ALCOA+), so evidence can be trusted and responsibility assigned.
Primary Stakeholder | Regulators/auditors (SH-07).
Secondary Stakeholders | QA/RA (SH-02), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-005/006; ALCOA+; 21 CFR Part 11.
User Journey | Audit; attribution dispute.
Hazard Relationship | Unattributable data (ISO 14971) → indefensible/contested disposition.
Traceability | Back: STN-005/006. Forward: PRD attribution/audit-trail, UN-FUN attribution.
Notes | Ties UN-FUN-016 (time/location stamping) and audit trail.
6.2 Access control & identity
UN-SEC-003 — Control who can access, configure, and read fleet data
Need Statement | The IT/security stakeholder needs role-based access with enterprise sign-on so only authorized people can configure devices, read data, and change settings, so the platform meets enterprise and Part 11 access obligations.
Primary Stakeholder | IT/security admins (SH-09).
Secondary Stakeholders | QA/RA (SH-02), regulators (SH-07).
Importance | Important
Market Priority | Must-have (enterprise)
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-013/025; VOC-018; GAMP 5; 21 CFR Part 11.
User Journey | Enterprise administration.
Hazard Relationship | Uncontrolled access (ISO 14971 + security) → tampered config/data → wrong verdict.
Traceability | Back: STN-013. Forward: PRD SSO/RBAC, UN-ECO enterprise integration.
Notes | Must not obstruct offline edge use (conflict C9): edge stays offline-simple; platform enforces.
UN-SEC-004 — Keep the device edge usable without a login while keeping data controlled
Need Statement | Last-mile users need to read the on-device verdict without any login or account, while the underlying data remains controlled and attributable, so security does not defeat frictionless offline use.
Primary Stakeholder | Clinic/last-mile staff (SH-05).
Secondary Stakeholders | IT/security (SH-09), ecosystem partners (SH-14).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Stakeholder Interview
Voice of Customer Reference | STN-004/013/017; conflict C9.
User Journey | Point-of-use receipt vs enterprise governance.
Hazard Relationship | Login gate at edge (IEC 62366; ISO 14971) → verdict unreadable → wrong disposition.
Traceability | Back: conflict C9. Forward: PRD trust-boundary split, UN-UX offline verdict.
Notes | Separate trust boundaries: read-only verdict at edge; controlled config/data on platform.
6.3 Secure connectivity & update
UN-SEC-005 — Trust that firmware updates are genuine and cannot brick the fleet
Need Statement | IT/service users need firmware updates to be verifiably genuine, resumable, and fail-safe, so an update cannot be malicious, corrupt a device, or leave it in an unknown state.
Primary Stakeholder | IT/security (SH-09), service (SH-06).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-025; CFB-011; D2 (OTA).
User Journey | OTA/depot update.
Hazard Relationship | Unsigned/failed update (ISO 14971 + security) → compromised/bricked device.
Traceability | Back: STN-025. Forward: PRD signed/fail-safe OTA, UN-SAFE-010, UN-FLOW update.
Notes | Ties safety fail-safe (UN-SAFE-010); flag cybersecurity persona for PRD/hazard.
UN-SEC-006 — Protect data in transit and at rest against interception/spoofing
Need Statement | Users need data protected as it moves from device to cloud and where it is stored, so records cannot be intercepted, spoofed, or exposed in transit or at rest.
Primary Stakeholder | IT/security (SH-09).
Secondary Stakeholders | QA/RA (SH-02), regulators (SH-07).
Importance | Important
Market Priority | Must-have (enterprise)
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-025; VOC-019; GAMP 5.
User Journey | Upload; storage; access.
Hazard Relationship | Spoofed/intercepted data (ISO 14971 + security) → false record → wrong disposition.
Traceability | Back: STN-025. Forward: PRD secure transport/storage, UN-ECO platform.
Notes | Threat model Evidence Pending; set with cybersecurity persona.
6.4 Data ownership, retention & privacy
UN-SEC-007 — Control data ownership and retention to policy
Need Statement | The enterprise/program stakeholder needs to own its data and control how long it is retained and when it is deleted, so records comply with data-retention and governance policy.
Primary Stakeholder | IT/security (SH-09), QA/RA (SH-02).
Secondary Stakeholders | Executive (SH-15), regulators (SH-07).
Importance | Important
Market Priority | Must-have (enterprise)
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-013; VOC-018; GDP/GAMP 5.
User Journey | Governance; retention/deletion.
Hazard Relationship | None direct; compliance/governance.
Traceability | Back: STN-013. Forward: PRD retention config/ownership, UN-REG data governance.
Notes | Configurable retention; data-ownership terms (ties UN-BIZ subscription).
UN-SEC-008 — Handle any location/personal data proportionately and privately
Need Statement | Stakeholders need any location or personal data the device generates handled proportionately and privately, so fleet/location tracking does not create privacy exposure beyond its operational purpose.
Primary Stakeholder | IT/security (SH-09).
Secondary Stakeholders | Ecosystem partners (SH-14), executive (SH-15).
Importance | Desirable
Market Priority | Nice-to-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-015; MR-1100.
User Journey | Fleet management; location use.
Hazard Relationship | None direct; privacy.
Traceability | Back: STN-015. Forward: PRD privacy handling, UN-REG regional privacy.
Notes | Device monitors cargo, not people; location is asset-level, but confirm regional privacy law (Evidence Pending).
6.5 Auditability
UN-SEC-009 — Have a complete audit trail of who did what to the data
Need Statement | Auditors and IT need a complete, tamper-evident audit trail of configuration changes, access, and data actions, so the system's governance can itself be inspected and defended.
Primary Stakeholder | Regulators/auditors (SH-07), IT (SH-09).
Secondary Stakeholders | QA/RA (SH-02).
Importance | Important
Market Priority | Must-have (regulated/enterprise)
Source | Regulatory
Voice of Customer Reference | STN-005/013/025; 21 CFR Part 11; ALCOA+.
User Journey | Audit / inspection of the system.
Hazard Relationship | No audit trail (ISO 14971) → ungovernable system → contested records.
Traceability | Back: STN-005/013. Forward: PRD audit trail, UN-FUN record, UN-REG Part 11.
Notes | Underpins UN-VIS-003 (trust the record) and enterprise CSV.
7. Validation Notes
Security/privacy needs are validated by CSV/validation review (GAMP 5), audit-trail and access-control testing, tamper-evidence verification, OTA signing/fail-safe testing (with UN-SAFE-010), and a security assessment against a threat model. The threat model, penetration scope, and privacy-law mapping are Evidence Pending (set in PRD/hazard with the cybersecurity persona).
8. Traceability Placeholders
Upstream | STN-005/006/013/015/025; 21 CFR Part 11; ALCOA+; GAMP 5; D2; conflict C9.
Downstream | UN-FUN, UN-SAFE, UN-ECO, UN-REG; forward to PRD (ORCH-011), cybersecurity design, hazard analysis. Matrix: ORCH-018 / EM-ENG-012.
UN-1100 — Physical Needs (PHY)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-1300/1400, STN-024/026, concept subsystems, and D1 (reusable). Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the physical needs users have of the SCM — size, form, mounting, controls/display, durability, sealing, and portability — expressed as needs ("must fit inside a shipper and mount securely"), not as design specifications (materials, dimensions). It is the human/physical-fit layer that installers, operators, and service users depend on.
2. Scope
In scope: form factor and fit inside insulated shippers, mounting/retention, probe form, display and control ergonomics, glove-friendly operation, ruggedness, sealing/ingress, cleanability, portability, and labeling durability. Out of scope: enclosure materials/tolerances (design outputs) and environmental survival conditions (UN-ENV).
3. Applicable Standards & References
• IEC 60529 (IP) — ingress protection.
• IEC 60068 — mechanical robustness (shock/vibration).
• WHO PQS E006 — physical/label expectations for cold-chain devices.
• Cold/gloved ergonomics literature — Judgment/Assumption; dexterity thresholds Evidence Pending.
4. Background
The device lives inside insulated shippers of varying size and must be embeddable by OEMs (SH-14) — so a standard, compact, mountable form is a channel requirement (STN-017/024). It is handled gloved in the cold, read at arm's length in poor light, cleaned repeatedly, and reused for years (D1) — so ruggedness, sealing, cleanability, and durable labeling are physical needs, not niceties.
5. Need Categories
• 5.1 Form, fit & mounting · 5.2 Display & controls (physical) · 5.3 Ruggedness & sealing · 5.4 Cleanability & durability over reuse · 5.5 Portability & labeling.
6. Detailed User Needs
6.1 Form, fit & mounting
UN-PHY-001 — Fit inside standard insulated shippers and containers
Need Statement | Installers and operators need the device to fit inside the insulated shippers and containers used across the chain, so it can be placed with the product without displacing payload.
Primary Stakeholder | Installers (SH-12), warehouse operators (SH-05).
Secondary Stakeholders | Ecosystem partners (SH-14), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-024; CFB-013; MR-SEG-008.
User Journey | Placement / OEM embedding.
Hazard Relationship | Poor fit (ISO 14971) → misplacement → biased detection.
Traceability | Back: STN-024. Forward: PRD form factor, UN-ECO embeddable form.
Notes | Fit-inside-shipper spec Evidence Pending; multiple shipper sizes.
UN-PHY-002 — Mount and retain securely so it stays in place and position through transit
Need Statement | Users need the device (and probe) to mount and stay securely in position through handling and vibration, so it does not shift, so readings stay representative and it is not lost in the box.
Primary Stakeholder | Installers (SH-12), 3PLs (SH-04).
Secondary Stakeholders | Warehouse operators (SH-05).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-024; FLD-002; IEC 60068.
User Journey | Placement; transit.
Hazard Relationship | Shifted probe (ISO 14971) → unrepresentative reading → wrong verdict.
Traceability | Back: STN-024. Forward: PRD mounting/retention, UN-FUN placement.
Notes | Defined mounting/retention; internal vs external probe form.
UN-PHY-003 — Be embeddable by OEMs as a standard module
Need Statement | OEM partners need a standard, embeddable physical module so they can build monitoring into their shippers/containers as a differentiator.
Primary Stakeholder | Ecosystem partners / OEMs (SH-14).
Secondary Stakeholders | Installers (SH-12).
Importance | Desirable
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-017; CFB-013/014; MR-SEG-008/009.
User Journey | OEM integration.
Hazard Relationship | None direct; channel enablement.
Traceability | Back: STN-017. Forward: PRD embeddable form, UN-ECO partner enablement.
Notes | Designed-in distribution channel.
6.2 Display & controls (physical)
UN-PHY-004 — Have a display readable at a glance and in the cold
Need Statement | Users need a display physically sized and positioned to be read at a glance at the point of use, remaining legible at low temperature, so the verdict is readable in the field and freezer.
Primary Stakeholder | Clinic/last-mile staff (SH-05), warehouse operators (SH-05).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/008; FLD-003; MR-PER-005/007.
User Journey | Receipt; cold-store staging.
Hazard Relationship | Unreadable display (ISO 14971) → misread verdict.
Traceability | Back: STN-004. Forward: UN-UX/UN-ENV legibility, PRD display.
Notes | OLED low-temp legibility thresholds Evidence Pending (display-engineering review).
UN-PHY-005 — Have controls operable with gloved, cold hands
Need Statement | Users need physical controls large and tactile enough to operate with gloved, cold hands in a single action, so the device is usable in the cold store and field.
Primary Stakeholder | Warehouse operators (SH-05), nurse (SH-05).
Secondary Stakeholders | Service (SH-06), dispatch (SH-02/SH-03).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008; FLD-003; MR-PER-005.
User Journey | Arming; receipt; handling.
Hazard Relationship | Fiddly controls (IEC 62366; ISO 14971) → mis-operation.
Traceability | Back: STN-008. Forward: UN-UX one-action, PRD controls.
Notes | Glove-friendly single button/indicator; dexterity thresholds Evidence Pending.
6.3 Ruggedness & sealing
UN-PHY-006 — Survive drops, handling, and vibration in the field and transit
Need Statement | Users need the device to survive the drops, rough handling, and vibration of real logistics and field use without damage or measurement error, so it lasts its reusable life.
Primary Stakeholder | 3PLs (SH-04), service ops (SH-13).
Secondary Stakeholders | Warehouse operators (SH-05), programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Contextual Inquiry
Voice of Customer Reference | MR-1300; STN-003; IEC 60068.
User Journey | Transit; field handling.
Hazard Relationship | Impact damage (ISO 14971) → failure/loss of record.
Traceability | Back: MR-1300. Forward: PRD ruggedness, UN-ENV mechanical, verification.
Notes | Ruggedness supports reuse economics (fewer write-offs).
UN-PHY-007 — Be sealed against dust and moisture, including at the charge port
Need Statement | Users need the device sealed against dust and moisture — including the USB-C charge port — so ingress does not cause sensor drift, corrosion, or failure over its reusable life.
Primary Stakeholder | Service ops (SH-13), warehouse operators (SH-05).
Secondary Stakeholders | 3PLs (SH-04).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-026; FLD-011; IEC 60529.
User Journey | Cleaning; humid/dusty use.
Hazard Relationship | Ingress (ISO 14971) → sensor error → wrong verdict.
Traceability | Back: STN-026. Forward: PRD IP rating/port sealing, UN-ENV ingress, UN-SVC cleaning.
Notes | USB-C port sealing is a named concern; flag mechanical/materials engineers.
6.4 Cleanability & durability over reuse
UN-PHY-008 — Be cleanable and decontaminable without damage over many cycles
Need Statement | Service users need to clean and decontaminate the device repeatedly between trips without degrading its enclosure, seals, display, or labeling, so it stays safe and legible across its life.
Primary Stakeholder | Service ops (SH-13), technicians (SH-06).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-026; FLD-011; D1; IEC 60529.
User Journey | Refurbishment.
Hazard Relationship | Cleaning damage/ingress (ISO 14971) → sensor error.
Traceability | Back: STN-026. Forward: UN-SVC cleaning method, PRD materials compatibility.
Notes | Cleaning-chemical compatibility Evidence Pending; flag materials engineer.
6.5 Portability & labeling
UN-PHY-009 — Be light and portable enough not to burden handling or payload
Need Statement | Users need the device small and light enough that it neither burdens handlers nor meaningfully displaces vaccine payload, and that it carries durable, legible identification and safety labeling that survives its life.
Primary Stakeholder | 3PLs (SH-04), warehouse operators (SH-05).
Secondary Stakeholders | Service (SH-13), regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-012/024; WHO PQS E006; MR-1300.
User Journey | Handling; identification; audit.
Hazard Relationship | Illegible safety/ID label (ISO 14971) → mis-handling / mis-attribution.
Traceability | Back: STN-012. Forward: PRD weight/labeling, UN-REG labeling, UN-SEC identity.
Notes | Durable device ID underpins attribution (UN-FUN-016) and fleet management.
7. Validation Notes
Physical needs are validated by fit checks in representative shippers, mounting/retention and drop/vibration testing (IEC 60068), ingress testing (IEC 60529), cleaning-cycle durability, and gloved/cold usability studies. Fit, dexterity, and cleaning-compatibility thresholds are Evidence Pending (Stage 02 §6).
8. Traceability Placeholders
Upstream | MR-1300/1400; STN-024/026; concept subsystems; D1.
Downstream | UN-ENV, UN-UX, UN-SVC, UN-ECO, UN-REG; forward to PRD (ORCH-011), mechanical design. Matrix: ORCH-018 / EM-ENG-012.
UN-1200 — Service Needs (SVC)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on D1 (reusable multi-trip), STN-010/016/ 021/026, ISO/IEC 17025 concepts, and the service persona (MR-PER-011). Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what service, calibration, and reverse-logistics users need so the reusable model is operationally sustainable — fast, verifiable turnaround; calibration management; cleaning; recovery of strays; field diagnostics; remote support; warranty; and field upgrades. Because the device is reusable (D1), these are core needs, not aftersales afterthoughts.
2. Scope
In scope: device return/recovery, refurbishment turnaround, calibration verification/traceability, cleaning/decontamination workflow, diagnostics, remote support, warranty/repair, and field/OTA upgrades. Out of scope: the workflow steps (UN-FLOW), performance of charge readiness (UN-PERF), and the sealing/ cleanability physical needs (UN-PHY).
3. Applicable Standards & References
• ISO/IEC 17025 concepts — calibration/metrology traceability.
• EU GDP / 21 CFR Part 11 — calibration currency and record integrity across service.
• WHO PQS E006 — device-performance maintenance in EPI settings.
• UN 38.3 / IATA DGR — battery handling in reverse logistics.
• Calibration-interval and throughput figures are Evidence Pending (set in regulatory/service stage).
4. Background
Reuse only pays if devices come home, turn around fast, and stay trustworthy. The service depot is where the reusable loop closes: recover → clean → charge → recalibrate → update → re-provision → redeploy. Hidden drift or dead batteries destroy both the record's credibility and the reuse ROI, so service needs are tightly coupled to safety (UN-SAFE), performance (UN-PERF), and quality (UN-QUAL).
5. Need Categories
• 5.1 Recovery & return · 5.2 Turnaround & refurbishment · 5.3 Calibration management · 5.4 Cleaning & decontamination · 5.5 Diagnostics & remote support · 5.6 Warranty & field upgrades.
6. Detailed User Needs
6.1 Recovery & return
UN-SVC-001 — Recover and return reusable devices reliably and affordably
Need Statement | The reverse-logistics stakeholder needs a practical, affordable way to get devices back — including on hard export/last-mile lanes — so the reusable fleet is not steadily lost to attrition.
Primary Stakeholder | Service / reverse-logistics (SH-13).
Secondary Stakeholders | Distributors (SH-03), programs (SH-01), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | VoC / Market Analysis
Voice of Customer Reference | STN-016; VOC-006; FLD-007; LD-008; D1.
User Journey | Return / recovery.
Hazard Relationship | None direct; return rate governs reuse economics (conflict C6).
Traceability | Back: STN-016. Forward: UN-FLOW return, UN-BIZ return rate/TCO, UN-FUN location.
Notes | Return-rate target Evidence Pending (pilot); region/lane-specific strategy (conflict C6).
UN-SVC-002 — Locate strays to chase overdue devices
Need Statement | Service users need last-known-location for overdue/stray devices so they can be chased and recovered, within the energy budget.
Primary Stakeholder | Service ops (SH-13), IT/fleet (SH-09).
Secondary Stakeholders | Distributors (SH-03).
Importance | Desirable
Market Priority | Nice-to-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-015/016; FLD-005/012; conflict C3.
User Journey | Recovery.
Hazard Relationship | None direct; economic.
Traceability | Back: STN-016. Forward: UN-PERF location, UN-FUN locate.
Notes | Location energy-bound (conflict C3); Evidence Pending.
6.2 Turnaround & refurbishment
UN-SVC-003 — Turn a returned device around quickly and verifiably at scale
Need Statement | Service users need to refurbish returned devices quickly and at scale — with each step verifiable — so turnaround time does not throttle fleet utilization or reuse margin.
Primary Stakeholder | Service technicians (SH-06), service ops (SH-13).
Secondary Stakeholders | Procurement (SH-08), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / Stakeholder Interview
Voice of Customer Reference | STN-010; VOC-016/017; MR-PER-011; MR-JTBD-007.
User Journey | Refurbishment.
Hazard Relationship | Rushed/unverified refurbishment (ISO 14971) → bad device redeployed.
Traceability | Back: STN-010. Forward: UN-FUN readiness reporting, UN-PERF charge readiness, UN-BIZ TCO.
Notes | Throughput target Evidence Pending; device self-reports readiness (UN-FUN-017).
UN-SVC-004 — Know a returned device's state without hand-checking each one
Need Statement | Service users need each returned device to report its charge, calibration, and firmware state so triage does not require manual inspection of every unit.
Primary Stakeholder | Service technicians (SH-06).
Secondary Stakeholders | IT/fleet (SH-09), service ops (SH-13).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-010/021; CFB-012; VOC-016.
User Journey | Return triage.
Hazard Relationship | Undetected out-of-cal/low-charge (ISO 14971) → bad redeploy.
Traceability | Back: STN-010. Forward: UN-FUN readiness reporting.
Notes | Deploy blocked if out-of-cal (ties UN-SVC-005).
6.3 Calibration management
UN-SVC-005 — Keep every device's calibration current and traceable
Need Statement | Service and QA users need every circulating device to have current, traceable calibration, and to be blocked from deployment when calibration expires, so no out-of-cal device produces a disputed reading.
Primary Stakeholder | Service technicians (SH-06), regulators/QA (SH-07).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / Contextual Inquiry
Voice of Customer Reference | STN-021; CFB-012; ISO/IEC 17025 concepts.
User Journey | Recalibration; deployment gate.
Hazard Relationship | Out-of-cal reading (ISO 14971) → false verdict; severity anchor UN-VIS-001.
Traceability | Back: STN-021. Forward: UN-PERF accuracy, UN-FUN readiness, UN-REG calibration, UN-VAL.
Notes | Calibration interval Evidence Pending; traceable to a recognized standard.
UN-SVC-006 — Recalibrate efficiently, ideally without full teardown
Need Statement | Service users need a practical recalibration/verification method that keeps devices in circulation efficiently, so calibration currency does not become an economic bottleneck.
Primary Stakeholder | Service technicians (SH-06).
Secondary Stakeholders | Procurement (SH-08), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview
Voice of Customer Reference | STN-021; VOC-016; ISO/IEC 17025 concepts.
User Journey | Recalibration.
Hazard Relationship | Skipped recal (ISO 14971) → drift → false verdict.
Traceability | Back: STN-021. Forward: PRD calibration approach, UN-BIZ TCO.
Notes | Method/interval Evidence Pending; balances rigor vs throughput.
6.4 Cleaning & decontamination
UN-SVC-007 — Clean and decontaminate between trips with a defined, safe method
Need Statement | Service users need a defined cleaning/decontamination method that is effective and safe for the device, so devices are hygienically ready to redeploy without being damaged.
Primary Stakeholder | Service ops (SH-13), technicians (SH-06).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-026; FLD-011; D1.
User Journey | Refurbishment.
Hazard Relationship | Cleaning-induced ingress/damage (ISO 14971) → sensor error.
Traceability | Back: STN-026. Forward: UN-PHY cleanability/sealing, PRD cleaning method.
Notes | Chemical compatibility Evidence Pending; flag materials engineer.
6.5 Diagnostics & remote support
UN-SVC-008 — Diagnose device problems quickly, including remotely
Need Statement | Service and IT users need to diagnose device issues quickly — locally and remotely — so faults are identified without shipping every suspect unit back for inspection.
Primary Stakeholder | Service technicians (SH-06), IT/fleet (SH-09).
Secondary Stakeholders | Ecosystem partners (SH-14).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-010/013; CFB-012/016.
User Journey | Troubleshooting.
Hazard Relationship | Undiagnosed fault (ISO 14971) → bad device stays in service.
Traceability | Back: STN-010. Forward: UN-FUN device health, UN-ECO remote support.
Notes | Self-reported health/diagnostics (UN-FUN-017) enables remote triage.
UN-SVC-009 — Get responsive support and warranty across markets
Need Statement | Program, distributor, and partner users need responsive support and clear warranty terms in their region, so field problems are resolved without stranding a fleet.
Primary Stakeholder | Ecosystem partners (SH-14), programs (SH-01).
Secondary Stakeholders | Distributors (SH-03), procurement (SH-08).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-017; CFB-014; LD-005.
User Journey | Support/warranty.
Hazard Relationship | None direct; adoption/serviceability.
Traceability | Back: STN-017. Forward: UN-BIZ support/warranty, UN-ECO partner enablement.
Notes | Localized support enables channel scale (ties UN-STK-013).
6.6 Warranty & field upgrades
UN-SVC-010 — Keep devices current with safe field/OTA upgrades
Need Statement | IT/service users need to keep the fleet's firmware current via safe field/OTA upgrades with confirmation of success, so devices improve over life without depot recall or being bricked.
Primary Stakeholder | IT/fleet (SH-09), service (SH-06).
Secondary Stakeholders | Regulators (SH-07), ecosystem partners (SH-14).
Importance | Important
Market Priority | Must-have
Source | Stakeholder Interview / Regulatory
Voice of Customer Reference | STN-025; VOC-019; D2 (OTA).
User Journey | Firmware update over life.
Hazard Relationship | Failed/unsigned upgrade (ISO 14971 + security) → bricked/compromised device.
Traceability | Back: STN-025. Forward: UN-SEC signed/fail-safe OTA, UN-FLOW update, UN-SAFE-010.
Notes | Signed, resumable, fail-safe; per-device success confirmation.
7. Validation Notes
Service needs are validated by a service-depot time-and-motion study (turnaround throughput), calibration traceability audits, cleaning-cycle durability testing, remote-diagnostics trials, and OTA success-rate testing. Throughput, calibration interval, and return rate are Evidence Pending (pilot / regulatory stage; Stage 02 §6).
8. Traceability Placeholders
Upstream | D1; STN-010/016/021/025/026; MR-PER-011; ISO/IEC 17025 concepts.
Downstream | UN-PERF, UN-SAFE, UN-SEC, UN-PHY, UN-QUAL, UN-BIZ, UN-ECO; forward to PRD (ORCH-011). Matrix: ORCH-018 / EM-ENG-012.
UN-1300 — Regulatory Needs (REG)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on D4 (multi-market), MR-1100 (Regulatory Landscape), STN-011/012/022/025/026, and cited standards. Research-driven: identify → cite → confirm gaps. Forward to PRD (ORCH-011), regulatory stage, and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users and the business need from the SCM in regulatory/market-access terms — the compliances that gate whether the device can be sold, deployed, flown, and trusted across its target markets (D4). Each need cites the governing standard/regulation; where a designation or classification is not yet confirmed it is marked Evidence Pending, never invented.
2. Scope
In scope: WHO PQS eligibility, EU GDP fit, US FDA classification/data-integrity expectations, radio/EMC approvals, Li-battery transport, electrical/environmental safety, calibration traceability, data-integrity (Part 11), and regional privacy/environmental (RoHS/WEEE/battery) compliance. Out of scope: the technical means of compliance (design) and the verification tests themselves (verification stage).
3. Applicable Standards & References
• WHO PQS E006 — temperature-monitoring device performance/prequalification (exact category TBC — Evidence Pending, concept §7).
• EU GDP (2013/C 343/01) — good distribution practice; data-integrity expectations.
• US FDA — cold-chain/data-integrity expectations; device classification/pathway Evidence Pending (device is an information/accessory tool, not therapeutic — confirm).
• 21 CFR Part 11 / ALCOA+ — electronic records/signatures.
• UN 38.3 (Li-battery transport testing), IATA DGR (dangerous goods).
• FCC Part 15 / CE RED (radio: BLE/Wi-Fi/LTE-M/GPS); CISPR / IEC 61000 (EMC).
• IEC 62368-1 (electrical/thermal safety); IEC 60529 (IP), IEC 60068 (environmental).
• ISO/IEC 17025 concepts (calibration traceability); RoHS / WEEE / battery directives (EoL).
4. Background
Global vaccine distribution (D4) means the SCM must satisfy several regimes at once. Two are market-access gates: WHO PQS (the LMIC procurement gate) and radio/transport approvals (without which it cannot be sold or flown). Others (GDP, Part 11, IEC safety) make its records and use defensible. Two named gaps carry forward: the exact WHO PQS category and the US classification (concept §7).
5. Need Categories
• 5.1 Market-access eligibility · 5.2 Data-integrity & records compliance · 5.3 Radio, EMC & electrical safety · 5.4 Battery transport & dangerous goods · 5.5 Calibration & metrology · 5.6 Environmental, privacy & end-of-life.
6. Detailed User Needs
6.1 Market-access eligibility
UN-REG-001 — Be WHO PQS-eligible for the relevant device category
Need Statement | Programs and procurement need the device to be WHO PQS-listed in the relevant temperature-monitoring category, because PQS listing is a hard gate for donor-funded and LMIC procurement.
Primary Stakeholder | Procurement / GAVI / UNICEF (SH-08), EPI programs (SH-01).
Secondary Stakeholders | Executive (SH-15), regulators (SH-07).
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-011/022; VOC-012; MR-REG; WHO PQS E006.
User Journey | Tender qualification; procurement.
Hazard Relationship | None direct; gates whether the safety benefit reaches LMIC scale.
Traceability | Back: STN-011. Forward: PRD PQS conformance, UN-BIZ market access, UN-VAL.
Notes | Exact PQS category/code Evidence Pending (concept §7); plan listing early (long cycles).
UN-REG-002 — Have clear regulatory classification and market-access basis per market
Need Statement | Regulatory-affairs and program users need the device's status clear in each market — WHO PQS category, US FDA classification, EU GDP fit — so it can be placed, registered, and defended.
Primary Stakeholder | Regulatory affairs / QA (SH-07).
Secondary Stakeholders | Procurement (SH-08), executive (SH-15).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-022; LD-001; MR-REG; D4; concept §7.
User Journey | Registration; market entry.
Hazard Relationship | None direct; market access.
Traceability | Back: STN-022. Forward: PRD regulatory strategy, regulatory stage.
Notes | US classification and WHO PQS category Evidence Pending — named concept gaps.
6.2 Data-integrity & records compliance
UN-REG-003 — Produce records that satisfy EU GDP and 21 CFR Part 11 / ALCOA+
Need Statement | QA/regulatory users need trip records that satisfy EU GDP and 21 CFR Part 11 / ALCOA+ — attributable, tamper-evident, contemporaneous, complete — so records are admissible for release and audit.
Primary Stakeholder | QA/RA (SH-02), regulators/auditors (SH-07).
Secondary Stakeholders | IT (SH-09), distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-005/025; EU GDP; 21 CFR Part 11; ALCOA+.
User Journey | Release; audit.
Hazard Relationship | Non-compliant record (ISO 14971) → inadmissible → indefensible disposition.
Traceability | Back: STN-005. Forward: UN-FUN record, UN-SEC integrity, PRD Part 11 conformance, UN-VAL.
Notes | Standard basis; ties UN-VIS-003 and UN-SEC-001/009.
6.3 Radio, EMC & electrical safety
UN-REG-004 — Hold radio approvals for each market
Need Statement | Users need the device's radios (BLE/Wi-Fi/LTE-M/GPS) approved for the markets where it operates, so it can be legally sold and used across its target geographies.
Primary Stakeholder | Regulatory affairs (SH-07), IT (SH-09).
Secondary Stakeholders | Ecosystem partners (SH-14), executive (SH-15).
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | MR-1100; D4; FCC Part 15; CE RED.
User Journey | Market entry; use.
Hazard Relationship | None direct; legal-to-sell gate.
Traceability | Back: MR-1100. Forward: PRD radio approvals, UN-ENV RF, verification.
Notes | Per-market radio certification (FCC/CE and regional equivalents).
UN-REG-005 — Meet EMC and electrical/thermal safety standards
Need Statement | Users and operators need the device to meet EMC and electrical/thermal safety standards, so it is safe and accepted for aircraft, enterprise, and clinical environments.
Primary Stakeholder | 3PLs / air-cargo (SH-04), IT (SH-09).
Secondary Stakeholders | Regulators (SH-07), clinic staff (SH-05).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | MR-1100; D4; IEC 62368-1; CISPR / IEC 61000.
User Journey | Air transport; deployment.
Hazard Relationship | EMC/electrical fault (ISO 14971) → malfunction or injury.
Traceability | Back: MR-1100. Forward: UN-SAFE-008, UN-ENV-009, PRD safety/EMC, verification.
Notes | IEC 62368-1 safety; CISPR/61000 EMC.
6.4 Battery transport & dangerous goods
UN-REG-006 — Be UN 38.3-tested and IATA-compliant for Li-battery transport
Need Statement | Logistics users need the device's battery UN 38.3-tested and its transport IATA-compliant, with documentation provided, so it can be flown and trucked without a dangerous-goods hold.
Primary Stakeholder | 3PLs / air-cargo (SH-04).
Secondary Stakeholders | Distributors (SH-03), service (SH-13).
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-012; VOC-007; D3/D4; UN 38.3; IATA DGR.
User Journey | Air/ground transport.
Hazard Relationship | Non-compliant Li transport (ISO 14971) → transport-safety hazard / channel exclusion.
Traceability | Back: STN-012. Forward: UN-SAFE-009, PRD battery transport, verification.
Notes | UN 38.3 test summary + IATA guidance (documentation Evidence Pending).
6.5 Calibration & metrology
UN-REG-007 — Provide traceable calibration to a recognized standard
Need Statement | QA/regulatory users need device calibration traceable to a recognized metrology standard, so readings — and the dispositions based on them — are defensible.
Primary Stakeholder | Regulators/QA (SH-07), service (SH-06).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-021; ISO/IEC 17025 concepts.
User Journey | Calibration; audit.
Hazard Relationship | Untraceable calibration (ISO 14971) → disputed reading → wrong verdict.
Traceability | Back: STN-021. Forward: UN-SVC calibration, UN-PERF accuracy, UN-VAL.
Notes | Interval Evidence Pending; ties UN-SVC-005.
6.6 Environmental, privacy & end-of-life
UN-REG-008 — Comply with environmental/hazardous-substance and battery-EoL regulations
Need Statement | Program, service, and sustainability users need the device to comply with hazardous- substance and battery/electronics end-of-life regulations in its markets, so it can be sold and retired lawfully and responsibly.
Primary Stakeholder | Service / sustainability (SH-13), regulatory affairs (SH-07).
Secondary Stakeholders | Procurement (SH-08), 3PLs (SH-04).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-026; MR-1700; RoHS / WEEE / battery directives; UN 38.3.
User Journey | Market entry; end-of-life.
Hazard Relationship | Unsafe disposal (ISO 14971) → environmental/safety.
Traceability | Back: STN-026. Forward: UN-SVC decommission, UN-BIZ sustainability, PRD EoL.
Notes | RoHS/WEEE/battery per market; EoL path Evidence Pending (conflict C10).
UN-REG-009 — Meet applicable data-protection/privacy requirements per region
Need Statement | IT/regulatory users need any personal or location data handled in line with applicable regional data-protection law, so the platform is lawful to deploy in each market.
Primary Stakeholder | IT/security (SH-09), regulatory affairs (SH-07).
Secondary Stakeholders | Executive (SH-15).
Importance | Desirable
Market Priority | Nice-to-have
Source | Regulatory
Voice of Customer Reference | STN-013/015; MR-1100; D4.
User Journey | Platform deployment per region.
Hazard Relationship | None direct; legal/privacy.
Traceability | Back: STN-013. Forward: UN-SEC-008 privacy, PRD data handling.
Notes | Device monitors cargo not people; applicable-law mapping Evidence Pending.
UN-REG-010 — Provide labeling and documentation required by each regime
Need Statement | Users need the device to carry the labeling and be supported by the documentation each regime requires (safety marks, battery/transport marks, PQS/GDP documentation), so it passes acceptance and audit at every gate.
Primary Stakeholder | Regulatory affairs (SH-07), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01), procurement (SH-08).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-012/022; WHO PQS E006; IATA DGR; CE/FCC.
User Journey | Acceptance; audit; transport.
Hazard Relationship | Missing/incorrect labeling (ISO 14971) → mis-handling / non-acceptance.
Traceability | Back: STN-022. Forward: UN-PHY labeling, PRD labeling/docs, regulatory stage.
Notes | Durable labeling (UN-PHY-009); documentation set Evidence Pending.
7. Validation Notes
Regulatory needs are validated by achieving the listed certifications/listings and by regulatory review: WHO PQS prequalification, radio (FCC/CE + regional), EMC (CISPR/61000), safety (IEC 62368-1), transport (UN 38.3 / IATA), calibration traceability audit, and Part 11/GDP records validation. The WHO PQS category, US classification, calibration interval, and privacy-law mapping are Evidence Pending and are confirmed in the Regulatory stage (research-driven per this module's mandate).
8. Traceability Placeholders
Upstream | D4; MR-1100; STN-011/012/021/022/025/026; concept §7 gaps.
Downstream | UN-FUN, UN-SEC, UN-SAFE, UN-PERF, UN-SVC, UN-PHY, UN-BIZ; forward to PRD (ORCH-011), regulatory stage, verification/validation. Matrix: ORCH-018 / EM-ENG-012.
UN-1400 — Quality Needs (QUAL)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on D1 (reusable), STN-003/010/019/023, MR-EXEC, and the concept reliability drivers. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users need in terms of build quality, consistency, and reliability over the device's reusable life — the qualities that make the SCM trustworthy as an instrument of record and viable as a reusable asset. Needs are solution-neutral ("must perform consistently across units and over life").
2. Scope
In scope: unit-to-unit consistency, reliability/longevity over reuse cycles, consistent measurement across the fleet, professional build befitting an audit instrument, and manufacturing quality that supports supply assurance. Out of scope: specific performance targets (UN-PERF), calibration mechanics (UN-SVC), and manufacturing processes (later stages).
3. Applicable Standards & References
• ISO 13485 concepts / ISO 9001 — quality management for a regulated device.
• WHO PQS E006 — device performance/quality expectations.
• ISO/IEC 17025 concepts — measurement consistency/traceability.
• Reliability/cycle-life figures are Evidence Pending (set in PRD; part of MR-RISK-001).
4. Background
Because the SCM is an audit instrument and a reusable asset, quality is dual-purpose: readings must be consistent enough to trust across every device and every trip, and the hardware must last many cycles without degrading. Capacity fade over reuse and unit-to-unit measurement variation are the two quality risks that most directly threaten both record credibility and reuse economics.
5. Need Categories
• 5.1 Measurement consistency · 5.2 Reliability & longevity · 5.3 Unit-to-unit consistency · 5.4 Professional build & supply quality.
6. Detailed User Needs
6.1 Measurement consistency
UN-QUAL-001 — Every device reads consistently, so the fleet is interchangeable
Need Statement | Users need any device in the fleet to measure consistently with any other, so a trip's verdict does not depend on which unit happened to be used.
Primary Stakeholder | QA/RA (SH-02), regulators (SH-07).
Secondary Stakeholders | Programs (SH-01), service (SH-06).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-021/023; WHO PQS E006; ISO/IEC 17025 concepts.
User Journey | Whole fleet operation; audit.
Hazard Relationship | Unit-to-unit variation (ISO 14971) → inconsistent verdicts → wrong disposition.
Traceability | Back: STN-023. Forward: UN-PERF accuracy, UN-SVC calibration, PRD consistency.
Notes | Underpins UN-VIS-005 (one trustworthy approach); calibration keeps it in check.
UN-QUAL-002 — Readings stay consistent over the device's life
Need Statement | Users need each device's measurements to remain consistent over its reusable life (within calibration), so a trip in year three is as trustworthy as a trip in year one.
Primary Stakeholder | Regulators/QA (SH-07), service (SH-06).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-021; ISO/IEC 17025 concepts; D1.
User Journey | Whole reusable life.
Hazard Relationship | Drift over life (ISO 14971) → false verdict.
Traceability | Back: STN-021. Forward: UN-PERF-002, UN-SVC calibration.
Notes | Ties calibration currency (UN-SVC-005).
6.2 Reliability & longevity
UN-QUAL-003 — Last many reuse cycles without quality degradation
Need Statement | Procurement and service users need the device to survive its rated number of reuse cycles without meaningful degradation in sensing, battery, sealing, or display, so reuse economics hold.
Primary Stakeholder | Procurement (SH-08), service ops (SH-13).
Secondary Stakeholders | Finance (SH-10), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-010/019; D1/D3; MR-EXEC-008; MR-RISK-001.
User Journey | Whole reusable life.
Hazard Relationship | Capacity fade/wear (ISO 14971) → mid-life dead-in-cold or drift.
Traceability | Back: STN-019. Forward: UN-PERF-010 reliability, UN-BIZ TCO, PRD cycle life.
Notes | Cycle-life target Evidence Pending; part of MR-RISK-001.
UN-QUAL-004 — Low field-failure rate that protects trust and TCO
Need Statement | Users need a low rate of field failures, because each failed device is a lost record, a support cost, and a dent in trust that undermines both safety and the reuse business case.
Primary Stakeholder | Distributors (SH-03), service ops (SH-13).
Secondary Stakeholders | Programs (SH-01), finance (SH-10).
Importance | Important
Market Priority | Must-have
Source | Market Analysis
Voice of Customer Reference | STN-003/010; CFB-001/016; MR-EXEC-008.
User Journey | Whole life.
Hazard Relationship | Field failure (ISO 14971) → lost record / unmonitored trip.
Traceability | Back: STN-003. Forward: UN-PERF reliability, UN-SVC diagnostics, DFMEA (ERFKL).
Notes | Field-failure-rate target Evidence Pending; feeds reliability/DFMEA.
6.3 Unit-to-unit consistency
UN-QUAL-005 — Behave consistently unit-to-unit in use and interface
Need Statement | Users need every device to behave the same way in use — same interface, same verdict logic, same feedback — so training and trust transfer across the whole fleet without surprises.
Primary Stakeholder | Programs (SH-01), distributors (SH-03).
Secondary Stakeholders | Clinic staff (SH-05), service (SH-06).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Stakeholder Interview
Voice of Customer Reference | STN-004/023; MR-PER-006/007.
User Journey | Whole fleet operation.
Hazard Relationship | Inconsistent behaviour (IEC 62366; ISO 14971) → use-error.
Traceability | Back: STN-023. Forward: UN-UX consistency, PRD consistency.
Notes | Ties UN-UX-011 (verdict/record consistency).
6.4 Professional build & supply quality
UN-QUAL-006 — Present as a professional, audit-grade instrument
Need Statement | QA, regulatory, and executive users need the device to look and feel like a professional, audit-grade instrument, so it commands the trust its role as evidence-of-record requires.
Primary Stakeholder | QA/RA (SH-02), executive (SH-15).
Secondary Stakeholders | Regulators (SH-07), procurement (SH-08).
Importance | Desirable
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-005/023; MR-PER-001/009.
User Journey | Provisioning; audit; procurement.
Hazard Relationship | None direct; trust/adoption.
Traceability | Back: STN-023. Forward: UN-PHY build quality, UN-BIZ positioning.
Notes | Perceived quality supports premium positioning and audit confidence.
UN-QUAL-007 — Be available at scale with consistent supply quality
Need Statement | Procurement and programs need consistent quality across production volumes and supply assurance at scale, so a fleet-scale or tender commitment is not undermined by variable or interrupted supply.
Primary Stakeholder | Procurement (SH-08), executive (SH-15).
Secondary Stakeholders | Programs (SH-01), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Market Analysis
Voice of Customer Reference | STN-011/019; MR-EXEC-008; VOC-013.
User Journey | Procurement; scale-up.
Hazard Relationship | None direct; supply/quality risk.
Traceability | Back: STN-019. Forward: UN-BIZ supply assurance, manufacturing stages.
Notes | Supply assurance is a tender criterion (UN-STK-008); figures Evidence Pending.
7. Validation Notes
Quality needs are validated by reliability/cycle-life testing, unit-to-unit variation studies, field-failure- rate tracking, and manufacturing quality controls (later stages) — and cross-checked against the DFMEA (ERFKL, ORCH-016) and reliability analysis. Cycle-life, variation, and failure-rate targets are Evidence Pending.
8. Traceability Placeholders
Upstream | D1/D3; STN-003/010/019/021/023; MR-EXEC-008; MR-RISK-001.
Downstream | UN-PERF, UN-SVC, UN-UX, UN-PHY, UN-BIZ; forward to PRD (ORCH-011), reliability/DFMEA. Matrix: ORCH-018 / EM-ENG-012.
UN-1500 — Ecosystem Needs (ECO)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-1500 (Ecosystem Strategy), STN-013/ 014/017, and D1/D2. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users need from the SCM's wider ecosystem — the cloud platform/dashboard, mobile/BLE provisioning, notifications, integrations/APIs, and fleet management — expressed as needs, not architecture. The ecosystem is what turns a device into a fleet-scale program and what lets it live inside customers' existing logistics and quality systems.
2. Scope
In scope: cloud dashboard, fleet management, mobile/BLE provisioning app, notifications/alerting delivery, open/secured APIs and enterprise-system integration, multi-device management, and partner integration. Out of scope: on-device functions (UN-FUN), security mechanisms (UN-SEC), and commercial model (UN-BIZ).
3. Applicable Standards & References
• 21 CFR Part 11 / GAMP 5 / ALCOA+ — platform data-integrity and validation.
• EU GDP — distribution-record governance in the platform.
• Common enterprise-integration patterns (SSO, REST APIs, webhooks) — Judgment/Standard; specific integration scope is Evidence Pending (set in ecosystem/PRD stages).
4. Background
The concept includes a cloud dashboard, OTA, and mobile/BLE provisioning. Fleet visibility (STN-014) and enterprise integration (STN-013) turn per-device monitoring into network-level loss reduction and let the platform pass IT's bar. Partners (STN-017) need secured-open integration to embed and resell — reconciling open integration with enterprise security lockdown is conflict C9.
5. Need Categories
• 5.1 Fleet dashboard & visibility · 5.2 Provisioning & mobile · 5.3 Notifications & delivery · 5.4 Integration & APIs · 5.5 Multi-device & partner ecosystem.
6. Detailed User Needs
6.1 Fleet dashboard & visibility
UN-ECO-001 — See the whole fleet and network-level status in one place
Need Statement | Fleet and executive users need a network-level view of devices and shipments — status, alerts, location, utilization — so a program can be managed and loss reduction tracked across the network.
Primary Stakeholder | IT/fleet (SH-09), executive (SH-15).
Secondary Stakeholders | Distributors (SH-03), service ops (SH-13).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / JTBD
Voice of Customer Reference | STN-014/015; VOC-020; MR-1500; MR-JTBD-010.
User Journey | Fleet monitoring; reporting.
Hazard Relationship | None direct; supports loss reduction.
Traceability | Back: STN-014. Forward: UN-FUN fleet view, UN-BIZ ROI, PRD dashboard.
Notes | Liveness connectivity-bound (conflict C7); complements the offline edge (conflict C5).
UN-ECO-002 — Report network loss reduction, compliance, and cost-per-trip
Need Statement | Executive and finance users need the platform to report network-wide loss reduction, compliance exposure, utilization, and cost-per-trip, so they can justify and standardize the program.
Primary Stakeholder | Executive (SH-15), finance (SH-10).
Secondary Stakeholders | Procurement (SH-08).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-014/019; VOC-020; MR-1000.
User Journey | Executive reporting; business case.
Hazard Relationship | None direct; ROI/funding.
Traceability | Back: STN-014. Forward: UN-BIZ ROI/TCO, PRD analytics.
Notes | ROI figures Evidence Pending; validated by deployment metrics.
6.2 Provisioning & mobile
UN-ECO-003 — Provision and configure devices via mobile/BLE quickly
Need Statement | Dispatch and installer users need to provision, configure, and check devices via a mobile/BLE tool quickly, with an on-device fallback, so setup is fast where an app is available and still possible where it is not.
Primary Stakeholder | Dispatch (SH-02/SH-03), installers (SH-12).
Secondary Stakeholders | Service (SH-06).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008/017; VOC-002; FLD-001.
User Journey | Provisioning.
Hazard Relationship | App-only provisioning (ISO 14971) → cannot arm if app fails → void trip.
Traceability | Back: STN-008. Forward: UN-FUN arming, UN-FLOW provisioning, PRD mobile/BLE.
Notes | On-device fallback required (ties UN-FUN-012).
6.3 Notifications & delivery
UN-ECO-004 — Deliver alerts and notifications to the right people and channels
Need Statement | Users need alerts and status notifications delivered reliably to the right people through the channels they use, with escalation, so a warning reaches someone who can act.
Primary Stakeholder | Distributors / control tower (SH-03/SH-04).
Secondary Stakeholders | Programs (SH-01), IT (SH-09).
Importance | Important
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | STN-018; VOC-008; CFB-002.
User Journey | Excursion response.
Hazard Relationship | Undelivered alert (ISO 14971) → no intervention.
Traceability | Back: STN-018. Forward: UN-FUN routing/escalation, PRD notifications.
Notes | Configurable by role/shipment; acknowledgement tracked.
6.4 Integration & APIs
UN-ECO-005 — Integrate with existing logistics and quality systems via secured open APIs
Need Statement | IT and partner users need to integrate device data and events with their existing logistics/quality systems through secured, documented APIs, so the SCM lives inside their workflows rather than as a data silo.
Primary Stakeholder | IT/security (SH-09), ecosystem partners (SH-14).
Secondary Stakeholders | Distributors (SH-03), executive (SH-15).
Importance | Important
Market Priority | Must-have (enterprise)
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-013/017; VOC-018; CFB-010/014; MR-1500.
User Journey | Enterprise/partner integration.
Hazard Relationship | None direct; integrity via UN-SEC.
Traceability | Back: STN-013/017. Forward: UN-SEC access, UN-FUN export, PRD APIs.
Notes | Secured-open (not closed silo, not open-and-unsafe) — conflict C9 resolution direction.
UN-ECO-006 — Export records into quality/audit systems preserving integrity
Need Statement | QA/IT users need to export trip records into their quality/audit systems while preserving integrity and attribution, so evidence flows to where decisions and audits happen.
Primary Stakeholder | QA/RA (SH-02), IT (SH-09).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-005/013; 21 CFR Part 11; EU GDP.
User Journey | Audit; quality-system integration.
Hazard Relationship | Integrity lost in export (ISO 14971) → indefensible record.
Traceability | Back: STN-005/013. Forward: UN-FUN export, UN-SEC integrity, PRD export.
Notes | Ties UN-FUN-015; export preserves ALCOA+ integrity.
6.5 Multi-device & partner ecosystem
UN-ECO-007 — Manage many devices and programs across an organization
Need Statement | Fleet/IT users need to manage many devices, sites, and programs within one organization — grouping, roles, configuration at scale — so a large deployment stays governable.
Primary Stakeholder | IT/fleet (SH-09).
Secondary Stakeholders | Executive (SH-15), service ops (SH-13).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview
Voice of Customer Reference | STN-013/014; MR-1500.
User Journey | Fleet administration.
Hazard Relationship | None direct; governance.
Traceability | Back: STN-013. Forward: UN-SEC RBAC, UN-FLOW onboarding, PRD fleet management.
Notes | Bulk onboarding (UN-FLOW-001); role-based (UN-SEC-003).
UN-ECO-008 — Let partners embed, integrate, and support the product in their region
Need Statement | Ecosystem partners need provisioning tooling, secured integration, and localized enablement so they can embed, integrate, resell, and support the product in their market.
Primary Stakeholder | Ecosystem partners (SH-14).
Secondary Stakeholders | Installers (SH-12), IT (SH-09).
Importance | Desirable
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-017; CFB-013/014; MR-SEG-008/009.
User Journey | Partner integration and support.
Hazard Relationship | None direct; channel enablement (conflict C9).
Traceability | Back: STN-017. Forward: UN-SVC support, UN-PHY embeddable form, PRD partner APIs.
Notes | Localized docs/enablement; secured-open integration.
UN-ECO-009 — Localize the platform for the markets it serves
Need Statement | Program and partner users need the platform and any app localized for the languages and conventions of their markets, so it is usable and adoptable across global deployments.
Primary Stakeholder | Programs (SH-01), ecosystem partners (SH-14).
Secondary Stakeholders | Distributors (SH-03).
Importance | Desirable
Market Priority | Nice-to-have
Source | Market Analysis
Voice of Customer Reference | MR-1500; MR-SEG; D4.
User Journey | Deployment across markets.
Hazard Relationship | None direct (edge verdict is language-independent, UN-UX-003).
Traceability | Back: MR-1500. Forward: UN-BIZ localization, PRD i18n.
Notes | Edge verdict stays language-independent; platform localized.
7. Validation Notes
Ecosystem needs are validated by platform usability testing (fleet users), integration/API conformance and CSV review (GAMP 5), notification-delivery reliability testing, provisioning-app usability, and partner integration pilots. Integration scope and ROI reporting are Evidence Pending (ecosystem/PRD stages).
8. Traceability Placeholders
Upstream | MR-1500; STN-013/014/017/018; D1/D2/D4.
Downstream | UN-FUN, UN-SEC, UN-SVC, UN-BIZ, UN-UX; forward to PRD (ORCH-011), platform design. Matrix: ORCH-018 / EM-ENG-012.
UN-1600 — Business Needs (BIZ)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on D1 (reusable), MR-EXEC/800/1000/1400, STN-011/017/019, and conflicts C1/C6. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the commercial/business needs of the buying stakeholders — affordability, provable reuse TCO, scalability, subscription/service model, SKU/portfolio fit, global deployment, and localization — expressed as user needs, not a business plan. These needs gate whether the device is bought and whether the safety benefit is deployed at scale.
2. Scope
In scope: purchase affordability and price ceilings, cost-per-trip / TCO evidence, supply assurance, subscription/service model, SKU/portfolio fit (value vs professional), global deployment/localization, and future-expansion needs. Out of scope: internal pricing/margin strategy (business stage) and ROI reporting mechanics (UN-ECO-002).
3. Applicable Standards & References
• WHO PQS E006 (eligibility affecting procurement), UNICEF/GAVI procurement frameworks — pooled tender criteria (Cited/Standard).
• MR-EXEC-008 / MR-800 / MR-1000 / MR-1400 — reuse economics, commercial, metrics, portfolio.
• Price, cost, margin, TCO, and return-rate figures are Judgment/Assumption ("Evidence Pending") — no precise financials asserted as fact (per evidence posture).
4. Background
The reusable model (D1) only wins if lifetime cost-per-trip beats single-use, which requires high return rates (conflict C6) and a cold-capable build whose cost fits donor ceilings (conflict C1, the D3 tension / MR-RISK-001). A value/entry SKU may be needed to meet price ceilings while a professional SKU carries full endurance. These commercial needs are where the safety mission meets fixed budgets.
5. Need Categories
• 5.1 Affordability & price ceilings · 5.2 Reuse TCO & supply assurance · 5.3 Commercial model & subscription · 5.4 SKU/portfolio fit · 5.5 Global deployment & future expansion.
6. Detailed User Needs
6.1 Affordability & price ceilings
UN-BIZ-001 — Fit within donor/program budget ceilings
Need Statement | Procurement and program buyers need the device's lifetime cost to fit within fixed donor/program budget ceilings, because no capability justifies a price the budget cannot bear.
Primary Stakeholder | Procurement / GAVI / UNICEF (SH-08), EPI programs (SH-01).
Secondary Stakeholders | Finance (SH-10), executive (SH-15).
Importance | Essential
Market Priority | Must-have
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-011/019; VOC-012/013/064; LD-001/002; MR-800.
User Journey | Procurement decision.
Hazard Relationship | None direct; gates deployment of the safety benefit.
Traceability | Back: STN-011. Forward: UN-BIZ-004 SKU, UN-REG PQS, PRD SKU strategy.
Notes | Conflict C1 (cost vs −30 °C/90-day endurance — MR-RISK-001); price/TCO figures Evidence Pending.
6.2 Reuse TCO & supply assurance
UN-BIZ-002 — Have documented, provable cost-per-trip over the reusable life
Need Statement | Finance and procurement buyers need documented cost-per-trip and total cost of ownership across the device's reusable life, with return-rate assumptions, so the reuse business case can be justified against single-use.
Primary Stakeholder | Finance (SH-10), procurement (SH-08).
Secondary Stakeholders | Executive (SH-15), service ops (SH-13).
Importance | Important
Market Priority | Must-have
Source | Market Analysis
Voice of Customer Reference | STN-019; VOC-020/060; MR-EXEC-008; D1.
User Journey | Business-case approval.
Hazard Relationship | None direct; economic viability.
Traceability | Back: STN-019. Forward: UN-SVC return/turnaround, UN-ECO ROI reporting, UN-VAL pilot.
Notes | Cost-per-trip is the linchpin (MR-EXEC-008); return rate is the swing variable (conflict C6).
UN-BIZ-003 — Have supply assurance at the scale the program needs
Need Statement | Procurement and executive buyers need assurance the device can be supplied at program scale and over time, so a fleet-scale or tender commitment is not exposed to supply interruption.
Primary Stakeholder | Procurement (SH-08), executive (SH-15).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Market Analysis
Voice of Customer Reference | STN-011/019; MR-EXEC-008; VOC-013.
User Journey | Tender; scale-up.
Hazard Relationship | None direct; supply risk.
Traceability | Back: STN-011. Forward: UN-QUAL-007 supply quality, manufacturing stages.
Notes | Supply assurance is a tender criterion (UN-STK-008); figures Evidence Pending.
6.3 Commercial model & subscription
UN-BIZ-004 — Support a service/subscription model that fits how programs fund
Need Statement | Buyers need a commercial model (device + platform/service, e.g., subscription) aligned to how they fund and operate, so ongoing platform value is affordable and predictable, not a surprise cost.
Primary Stakeholder | Executive (SH-15), finance (SH-10).
Secondary Stakeholders | Procurement (SH-08), IT (SH-09).
Importance | Important
Market Priority | Differentiator
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-019; MR-800/1000; MR-1500.
User Journey | Procurement; ongoing operation.
Hazard Relationship | None direct; commercial fit.
Traceability | Back: STN-019. Forward: UN-ECO platform, PRD commercial model.
Notes | Recurring SaaS attach underpins the platform; terms Evidence Pending.
6.4 SKU/portfolio fit
UN-BIZ-005 — Offer SKUs that fit different lanes, budgets, and durability needs
Need Statement | Buyers need a portfolio that fits different needs — a value/entry option for tight budgets and manageable lanes, a professional option carrying full endurance — so each program can match device to budget and lane without over- or under-buying.
Primary Stakeholder | Procurement (SH-08), executive (SH-15).
Secondary Stakeholders | Distributors (SH-03), programs (SH-01).
Importance | Important
Market Priority | Differentiator
Source | Market Analysis
Voice of Customer Reference | STN-011; MR-1400; conflicts C1/C6.
User Journey | Procurement selection.
Hazard Relationship | None direct; commercial fit (ensure value SKU still meets safety floor).
Traceability | Back: STN-011. Forward: PRD SKU/portfolio, UN-REG PQS per SKU.
Notes | Value SKU must still meet the safety floor (UN-SAFE); ultra-cold −70 °C variant is roadmap (MR-1400), not a committed launch SKU.
6.5 Global deployment & future expansion
UN-BIZ-006 — Deploy across multiple markets from one product
Need Statement | Executive and program buyers need one product deployable across their global markets (multi-market compliant, localizable), so they can standardize rather than manage many regional products.
Primary Stakeholder | Executive (SH-15), manufacturers (SH-02).
Secondary Stakeholders | Programs (SH-01), ecosystem partners (SH-14).
Importance | Important
Market Priority | Differentiator
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-023; MR-SEG; D4.
User Journey | Global standardization.
Hazard Relationship | None direct; standardization supports consistent monitoring (UN-VIS-005).
Traceability | Back: STN-023. Forward: UN-REG multi-market, UN-ECO localization, PRD.
Notes | Ties UN-VIS-005 (one trustworthy approach across markets).
UN-BIZ-007 — Be upgradeable and future-proof over a multi-year deployment
Need Statement | Buyers need the device and platform to be upgradeable over a multi-year deployment (field/OTA improvements, new integrations) so a fleet investment does not become obsolete.
Primary Stakeholder | Executive (SH-15), IT (SH-09).
Secondary Stakeholders | Procurement (SH-08).
Importance | Desirable
Market Priority | Nice-to-have
Source | Market Analysis
Voice of Customer Reference | STN-025; MR-1600; D2.
User Journey | Multi-year ownership.
Hazard Relationship | None direct; investment protection.
Traceability | Back: STN-025. Forward: UN-SVC field upgrades, UN-SEC OTA, PRD roadmap.
Notes | Safe OTA (UN-SEC-005) enables this; roadmap per MR-1600.
UN-BIZ-008 — Justify itself through demonstrable loss reduction and compliance value
Need Statement | Executive buyers need the program to demonstrably reduce cold-chain loss and compliance exposure, so the investment is justified by outcomes, not just features.
Primary Stakeholder | Executive (SH-15).
Secondary Stakeholders | Finance (SH-10), procurement (SH-08).
Importance | Important
Market Priority | Differentiator
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-014; VOC-020; MR-1000; MR-EXEC-008.
User Journey | Sponsorship; renewal.
Hazard Relationship | None direct; funds the safety mission at scale.
Traceability | Back: STN-014. Forward: UN-ECO ROI reporting, UN-VAL market validation.
Notes | ROI/loss-reduction figures Evidence Pending; validated by deployment metrics.
7. Validation Notes
Business needs are validated by a return-rate/TCO pilot, a documented business-case workbook, tender- qualification outcomes, and deployment loss-reduction metrics (market validation, UN-VAL). All financials — price, cost-per-trip, TCO, return rate, ROI — are Evidence Pending and are set/confirmed in the business and validation stages (Stage 02 §6 quantitative gaps).
8. Traceability Placeholders
Upstream | D1/D3; MR-EXEC-008/800/1000/1400; STN-011/017/019/023; conflicts C1/C6.
Downstream | UN-SVC, UN-ECO, UN-REG, UN-QUAL, UN-VAL; forward to PRD (ORCH-011), business/validation stages. Matrix: ORCH-018 / EM-ENG-012.
UN-1700 — Validation Needs (VAL)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on the concept's closure-by-test posture (MR-VAL-006), IEC 62366 / ISO 14971, and the Stage 02 §6 VoC program. Forward to Validation (ORCH-015), Verification (ORCH-014), and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what stakeholders need in order to be confident the SCM actually meets their needs — the validation and acceptance needs. These are the "how we will know" counterparts to the needs in the other modules: customer/field validation, usability/human-factors validation, and the evidence each stakeholder requires before they will adopt, list, or fund the product.
2. Scope
In scope: acceptance criteria posture, customer/field validation, usability/human-factors validation, regulatory/audit validation, cold-endurance closure-by-test, reuse-economics pilot validation, and market validation. Out of scope: the detailed verification tests (ORCH-014) and validation protocols (ORCH-015), which this module seeds.
3. Applicable Standards & References
• IEC 62366-1 / FDA Human Factors guidance — formative + summative usability validation.
• ISO 14971 — validation of risk controls (safety-relevant needs).
• WHO PQS E006 — device-performance validation for listing.
• EU GDP / 21 CFR Part 11 — records validation (audit simulation).
• MR-VAL-006 — cold-endurance closure-by-test; Stage 02 §6 — recommended VoC program. Targets are Evidence Pending until set in PRD and closed by V&V.
4. Background
The SCM carries several needs whose satisfaction can only be confirmed by test or field study, not asserted: 90-day/−30 °C endurance (D3, the top risk), offline verdict usability by a novice nurse, record defensibility in an audit, alert latency vs real connectivity, and reuse return-rate/TCO. This module makes those validation needs explicit so they are planned, not discovered late.
5. Need Categories
• 5.1 Acceptance & customer validation · 5.2 Usability & human-factors validation · 5.3 Safety & regulatory validation · 5.4 Endurance & field validation · 5.5 Market & economic validation.
6. Detailed User Needs
6.1 Acceptance & customer validation
UN-VAL-001 — Every need has a clear, testable acceptance basis
Need Statement | Stakeholders need each user need to carry a clear, testable basis for judging it satisfied, so nothing is accepted on assertion and every need can be verified or validated.
Primary Stakeholder | QA/RA (SH-02), regulators (SH-07).
Secondary Stakeholders | All.
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | All STN success measures; concept evidence posture.
User Journey | Design controls; verification/validation.
Hazard Relationship | Untestable need (ISO 14971) → unverifiable safety control.
Traceability | Back: all modules. Forward: PRD acceptance criteria, PVL (ORCH-014), PVaL (ORCH-015).
Notes | Many targets are Evidence Pending; each carries the test that would confirm it.
UN-VAL-002 — Real users confirm the product meets their needs in context
Need Statement | Each stakeholder class needs to confirm, in their real context, that the product does the job for them, so adoption rests on demonstrated fit rather than a spec sheet.
Primary Stakeholder | All primary stakeholder classes (SH-01…SH-15).
Secondary Stakeholders | —
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Market Analysis
Voice of Customer Reference | Stage 02 §6 VoC program; STN success measures.
User Journey | Pilots; field validation.
Hazard Relationship | Unvalidated fit (ISO 14971) → use-error or wrong disposition in real use.
Traceability | Back: Stage 02 §6. Forward: PVaL (ORCH-015), pilot plans.
Notes | Ties the recommended VoC program (contextual inquiry, ride-alongs, audit sim, depot study).
6.2 Usability & human-factors validation
UN-VAL-003 — Confirm a novice, offline user can reach the right disposition
Need Statement | Programs and regulators need summative evidence that a minimally trained, offline last-mile user reads the verdict and reaches the correct accept/reject disposition, because this is the highest-consequence interaction.
Primary Stakeholder | EPI programs (SH-01), regulators (SH-07).
Secondary Stakeholders | Beneficiaries (SH-11), clinic staff (SH-05).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / Contextual Inquiry
Voice of Customer Reference | STN-004/020; MR-PER-007; IEC 62366.
User Journey | Summative usability study (receipt/point of care).
Hazard Relationship | Unvalidated use-error (IEC 62366; ISO 14971) → wrong disposition.
Traceability | Back: STN-004. Forward: Usability Engineering File, HA-### (ORCH-017), PVaL.
Notes | Critical-task focus; use-error rate target Evidence Pending.
UN-VAL-004 — Confirm gloved/cold and dispatch tasks succeed under real conditions
Need Statement | Operators and dispatch need validation that gloved/cold operation and fast, error-proof provisioning succeed under real conditions, so the cold-store and pack-out tasks are proven, not assumed.
Primary Stakeholder | Warehouse operators (SH-05), dispatch (SH-02/SH-03).
Secondary Stakeholders | Installers (SH-12).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008; FLD-003; conflict C4; MR-PER-005.
User Journey | Cold-store time-and-motion; provisioning study.
Hazard Relationship | Unvalidated provisioning (ISO 14971) → void trips.
Traceability | Back: STN-008. Forward: Usability Engineering File, PVaL.
Notes | Provisioning-time and cold-legibility targets Evidence Pending.
6.3 Safety & regulatory validation
UN-VAL-005 — Validate two-sided detection and correct disposition against the safety anchor
Need Statement | Programs, regulators, and beneficiaries (represented) need validated evidence that freeze and heat excursions are correctly detected and that the resulting disposition is correct — no false accept, no false reject — because this is the program's severity anchor.
Primary Stakeholder | Beneficiaries (SH-11, represented), programs (SH-01), regulators (SH-07).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / JTBD
Voice of Customer Reference | STN-002/020; MR-PROB-002/010; ISO 14971.
User Journey | Detection/verdict validation; hazard closure.
Hazard Relationship | The severity anchor (ISO 14971) — validated jointly with HA-### (ORCH-017).
Traceability | Back: STN-020. Forward: HA-###, Risk-Control Requirements, PVaL, PVL.
Notes | Accuracy/miss/false-accept targets Evidence Pending; safety-first Action Priority.
UN-VAL-006 — Validate audit defensibility of the record with real auditors
Need Statement | QA and regulatory stakeholders need the record's defensibility validated via audit simulation with real auditors, so the record is proven admissible before reliance in the field.
Primary Stakeholder | Regulators/auditors (SH-07), QA/RA (SH-02).
Secondary Stakeholders | IT (SH-09).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-005; VOC-001/019; 21 CFR Part 11; EU GDP.
User Journey | Audit-simulation study.
Hazard Relationship | Unvalidated record (ISO 14971) → indefensible disposition.
Traceability | Back: STN-005. Forward: PVaL audit sim, UN-REG-003, UN-SEC.
Notes | Stage 02 §6 audit-simulation study; ties UN-VIS-003.
6.4 Endurance & field validation
UN-VAL-007 — Close the cold-endurance risk by test (90 days at −30 °C)
Need Statement | Shippers, procurement, and engineering need the 90-day/−30 °C endurance and no-data-loss requirement closed by test, because it is the top program risk and cannot be asserted.
Primary Stakeholder | Distributors (SH-03), procurement (SH-08).
Secondary Stakeholders | 3PLs (SH-04), executive (SH-15).
Importance | Essential
Market Priority | Must-have
Source | Market Analysis / Regulatory
Voice of Customer Reference | STN-003; D3; MR-RISK-001; MR-VAL-006.
User Journey | Endurance/environmental test campaign.
Hazard Relationship | Dead-in-cold (ISO 14971) → lost record; the headline risk closed here.
Traceability | Back: STN-003, MR-VAL-006. Forward: PVL endurance test, power budget, HA-###.
Notes | The D3 tension closure-by-test; endurance/charge-safety/capacity-fade targets Evidence Pending.
UN-VAL-008 — Validate offline resilience and alert latency against real connectivity
Need Statement | Programs and distributors need offline operation, buffering, and alert latency validated against real last-mile connectivity, so store-and-forward behaviour is proven where losses concentrate.
Primary Stakeholder | EPI programs (SH-01), distributors (SH-03).
Secondary Stakeholders | 3PLs (SH-04), IT (SH-09).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Market Analysis
Voice of Customer Reference | STN-001/009; VOC-023; D2; conflict C7.
User Journey | Field connectivity trials; RF-in-container test.
Hazard Relationship | Unvalidated latency/offline (ISO 14971) → delayed/missed alert.
Traceability | Back: STN-009. Forward: PVL connectivity/latency, UN-ENV RF, PVaL.
Notes | RF-in-container performance is a named gap; latency/buffer targets Evidence Pending.
6.5 Market & economic validation
UN-VAL-009 — Validate reuse economics (return rate and cost-per-trip) by pilot
Need Statement | Finance, procurement, and service stakeholders need the reuse economics — return rate, turnaround, cost-per-trip — validated by pilot, because the reusable business case rests on assumptions that only field data can confirm.
Primary Stakeholder | Finance (SH-10), procurement (SH-08), service ops (SH-13).
Secondary Stakeholders | Executive (SH-15).
Importance | Important
Market Priority | Must-have
Source | Market Analysis
Voice of Customer Reference | STN-016/019; MR-EXEC-008; D1; conflict C6.
User Journey | Reuse/return pilot.
Hazard Relationship | None direct; economic viability that funds the safety mission at scale.
Traceability | Back: STN-019. Forward: UN-BIZ TCO, UN-SVC return, business-case workbook.
Notes | Return-rate/TCO Evidence Pending; region/lane-specific (conflict C6).
7. Validation Notes
This module is the validation backbone: it seeds the Validation Plan (ORCH-015) and cross-references the Verification Plan (ORCH-014), Usability Engineering File, and Hazard Analysis (ORCH-017). The recommended Stage 02 §6 VoC program (contextual inquiry, ride-alongs, audit simulation, service-depot study, importance–satisfaction survey) is the primary route to promote Evidence-Pending items to Cited. Safety- relevant validations use severity-first Action Priority (CLAUDE.md rule 3).
8. Traceability Placeholders
Upstream | all modules' success measures; MR-VAL-006; Stage 02 §6; ISO 14971 / IEC 62366; D1/D2/D3.
Downstream | PRD acceptance criteria (ORCH-011), Verification (ORCH-014), Validation (ORCH-015), Hazard Analysis (ORCH-017), Usability Engineering File. Matrix: ORCH-018 / EM-ENG-012.
UN-200 — Personas (PER)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-1200 personas, the stakeholder register, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module attaches, to each named persona, the capability the persona needs to succeed at their moment in the journey — expressed as a user need, not a feature. Where UN-STK captures the class-level anchor, UN-PER captures the persona-specific need shaped by that persona's literacy, environment, and accountability.
2. Scope
In scope: the eleven MR-1200 personas (MR-PER-001…011), including expressly the novice/low-training persona (last-mile nurse), the gloved/cold-environment persona (warehouse operator), and the occasional/expert split. Out of scope: persona demographics/biography (MR-1200) and topical needs.
3. Applicable Standards & References
• IEC 62366-1 (usability engineering / use-related risk) — frames the novice and gloved personas.
• WHO PQS E006, EU GDP, 21 CFR Part 11, IATA DGR / PCR, ISO/IEC 17025 — per persona.
• Cold-environment human-factors literature (gloved operation, OLED low-temperature legibility) — Judgment/Assumption; thresholds Evidence Pending.
4. Background
One device serves a supply-chain director and a rural nurse. Persona literacy, dexterity (gloved, cold), connectivity, and accountability differ enormously. Designing to the persona rather than a generic "operator" is how the SCM avoids the classic failure of a device that satisfies the auditor but defeats the edge user (or vice versa).
5. Need Categories
• 5.1 Provisioning personas · 5.2 Handling personas · 5.3 Point-of-use personas · 5.4 Governance personas · 5.5 Sustainment personas.
6. Detailed User Needs
6.1 Provisioning personas
UN-PER-001 — QA/cold-chain manager needs to release with confidence, fast
Need Statement | The pharma QA/cold-chain manager needs to provision, arm, and release a device for a trip in confidence that the record will be complete and defensible, in a time that does not bottleneck dispatch.
Primary Stakeholder | MR-PER-001 (SH-02).
Secondary Stakeholders | MR-PER-002 (SH-03), MR-PER-009 (SH-07).
Importance | Essential
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-005/008; VOC-001/002.
User Journey | Pack-out / release-to-ship.
Hazard Relationship | Rushed or forgotten arming → void trip / no record (ISO 14971).
Traceability | Back: STN-008. Forward: UN-FLOW provisioning, UN-FUN arming confirmation, UN-UX speed.
Notes | Expert, time-pressured user; the compliant path must be the fastest path (conflict C4).
UN-PER-002 — Distributor coordinator needs status at a glance across many shipments
Need Statement | The distributor logistics coordinator needs to see which of many concurrent shipments need attention, and be alerted only for genuine, action-worthy excursions.
Primary Stakeholder | MR-PER-002 (SH-03).
Secondary Stakeholders | MR-PER-004 (SH-04).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-007/018; VOC-005/008.
User Journey | Daily monitoring and excursion triage.
Hazard Relationship | Alarm fatigue → muted alerts → missed excursion (ISO 14971).
Traceability | Back: STN-007. Forward: UN-ECO fleet view, UN-FUN action-only alerts.
Notes | Not a data analyst; needs role-appropriate, action-only signaling.
UN-PER-003 — IT/data administrator needs to trust and govern the platform at fleet scale
Need Statement | The IT/data administrator needs to govern user access, data integrity, retention, and system validation across a device fleet in line with data-governance obligations.
Primary Stakeholder | MR-PER-003 (SH-09).
Secondary Stakeholders | MR-PER-009 (SH-07).
Importance | Important
Market Priority | Must-have (enterprise)
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-013/025; VOC-018.
User Journey | Onboarding, access management, validation, integration.
Hazard Relationship | Weak access control / integrity → corrupted disposition evidence (ISO 14971 + sec).
Traceability | Back: STN-013. Forward: UN-SEC, UN-ECO integration.
Notes | GAMP/CSV-literate; the gate for enterprise rollout.
6.2 Handling personas
UN-PER-004 — 3PL operations lead needs a device that never disrupts handling or transport
Need Statement | The 3PL operations lead needs the device to be invisible-until-relevant, non-disruptive to scanning/handling, and safe and compliant to fly and truck.
Primary Stakeholder | MR-PER-004 (SH-04).
Secondary Stakeholders | MR-PER-011 (SH-06/13).
Importance | Important
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-012; VOC-007.
User Journey | High-throughput handling; air/ground transport.
Hazard Relationship | Battery-transport non-compliance (UN 38.3/IATA) — transport safety.
Traceability | Back: STN-012. Forward: UN-REG transport, UN-PHY handling neutrality.
Notes | Never opens the box; endorsement multiplies channel adoption.
UN-PER-005 — Warehouse cold-store operator needs glove-friendly, cold-legible, one-action use
Need Statement | The warehouse/cold-store operator working gloved in +2/+8 °C rooms and −20 °C freezers needs to place, retrieve, arm, and read the device in a single action that works with cold hands and a cold display.
Primary Stakeholder | MR-PER-005 (SH-05 handling). (Gloved / cold-environment persona.)
Secondary Stakeholders | MR-PER-001 (SH-02), installers (SH-12).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008; FLD-003; MR-PER-005.
User Journey | Staging at trip start/end in the cold store.
Hazard Relationship | Fiddly cold interaction → mis-arm / mis-read → corrupt record or wrong disposition.
Traceability | Back: STN-008. Forward: UN-UX gloved/cold legibility, UN-PHY controls, UN-PERF cold-op.
Notes | OLED low-temperature legibility to −30 °C intersects display engineering; thresholds Evidence Pending.
6.3 Point-of-use personas
UN-PER-006 — EPI program manager needs a clear, PQS-aligned trip verdict
Need Statement | The EPI/immunization program manager needs an unambiguous, PQS-consistent trip verdict to make and stand behind consignment accept/reject decisions.
Primary Stakeholder | MR-PER-006 (SH-01).
Secondary Stakeholders | MR-PER-007 (SH-05), MR-PER-008 (SH-08).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | STN-004; VOC-009/014.
User Journey | Consignment arrival and disposition.
Hazard Relationship | Wrong verdict → spoiled dose accepted / safe dose rejected (ISO 14971 anchor).
Traceability | Back: STN-004. Forward: UN-FUN verdict, UN-REG PQS, UN-UX.
Notes | PQS/EVM-trained; decision-grade clarity over data volume.
UN-PER-007 — Last-mile nurse needs an instant, offline, self-evident status
Need Statement | The last-mile health worker/nurse, with minimal device training and often no connectivity, needs to read cold-chain status instantly on the device itself — no phone, network, or manual — to accept or reject stock at the point of care.
Primary Stakeholder | MR-PER-007 (SH-05). (Novice / low-training / vulnerable-context persona.)
Secondary Stakeholders | MR-PER-006 (SH-01), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004/027; VOC-014/015; FLD-006.
User Journey | Receipt at rural health post / mobile vaccination round.
Hazard Relationship | Use-error by an untrained user (IEC 62366; ISO 14971) → wrong disposition at point of care — highest-consequence edge decision.
Traceability | Back: STN-004. Forward: UN-UX low-literacy legibility, UN-FUN offline verdict, UN-VAL.
Notes | The design floor: if the nurse cannot use it offline and untrained, the product fails its purpose.
6.4 Governance personas
UN-PER-008 — Procurement officer needs PQS eligibility and reuse-TCO proof to qualify a tender
Need Statement | The pooled-procurement officer needs PQS prequalification and documented reuse total cost of ownership and supply assurance to qualify the device into a tender.
Primary Stakeholder | MR-PER-008 (SH-08).
Secondary Stakeholders | MR-PER-010 (SH-15), finance (SH-10).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-011/019; VOC-012/013.
User Journey | Tender qualification and award.
Hazard Relationship | None direct; gates LMIC-scale access to the safety benefit.
Traceability | Back: STN-011. Forward: UN-REG PQS, UN-BIZ TCO.
Notes | PQS listing cycles are long; plan early.
UN-PER-009 — Regulatory/quality auditor needs the record to be the easiest thing in the audit
Need Statement | The auditor/inspector needs tamper-evident, ALCOA+-aligned, traceable, exportable records with calibration currency, so verifying data integrity and excursion handling is fast and certain.
Primary Stakeholder | MR-PER-009 (SH-07).
Secondary Stakeholders | MR-PER-001 (SH-02), MR-PER-011 (SH-06).
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-005/021; VOC-019.
User Journey | Audit / inspection.
Hazard Relationship | Ambiguous records → inadmissible evidence → indefensible disposition (ISO 14971).
Traceability | Back: STN-005/021. Forward: UN-SEC integrity, UN-SVC calibration, UN-REG.
Notes | Inspector-level; export and audit-trail links to UN-ECO / UN-SEC.
UN-PER-010 — Executive/supply-chain director needs fleet ROI and risk visibility, not device detail
Need Statement | The executive/supply-chain director needs network-level visibility of loss reduction, compliance exposure, utilization, and cost-per-trip to authorize fleet-scale standardization.
Primary Stakeholder | MR-PER-010 (SH-15).
Secondary Stakeholders | MR-PER-003 (SH-09), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-014; VOC-020.
User Journey | Sponsorship and standardization decisions.
Hazard Relationship | None direct; funds scale.
Traceability | Back: STN-014. Forward: UN-ECO fleet dashboard, UN-BIZ.
Notes | Consumes fleet analytics; ROI figures Evidence Pending.
6.5 Sustainment personas
UN-PER-011 — Field-service/calibration technician needs verifiable, fast fleet turnaround
Need Statement | The field-service/calibration technician needs to clean, recharge, verify calibration, update firmware, and re-provision returned devices quickly and verifiably, at scale, so reuse is operationally sustainable.
Primary Stakeholder | MR-PER-011 (SH-06 / SH-13).
Secondary Stakeholders | MR-PER-009 (SH-07), procurement (SH-08).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / Stakeholder Interview
Voice of Customer Reference | STN-010/021/026; VOC-016/017.
User Journey | Return → refurbish → recalibrate → redeploy.
Hazard Relationship | Out-of-cal redeploy → disputed readings; under-charge → dead-in-cold next trip (ISO 14971).
Traceability | Back: STN-010. Forward: UN-SVC turnaround/calibration/cleaning, UN-PERF charge readiness.
Notes | Metrology/service-trained; turnaround throughput sets reuse margin (Evidence Pending).
7. Validation Notes
Persona needs are validated via IEC 62366-style formative and summative usability studies across the persona range — expressly the novice nurse (offline verdict), gloved cold-store operator, and expert QA manager (provisioning speed). Literacy/dexterity/legibility thresholds are Evidence Pending (Stage 02 §6 recommends ride-alongs and cold-store time-and-motion studies).
8. Traceability Placeholders
Upstream | MR-1200 (MR-PER-001…011); STN register; D1–D4.
Downstream | UN-UX, UN-FUN, UN-SVC, UN-PERF, UN-REG; forward to PRD (ORCH-011). Matrix: ORCH-018 / EM-ENG-012.
UN-300 — Jobs To Be Done (JTBD)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-400 (JTBD), VOC themes T1–T14, STN-001…027, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the jobs stakeholders hire the SCM to do — the progress they are trying to make in a given circumstance — expressed as solution-neutral needs. Each job is written as "when [situation], the [stakeholder] needs to [make progress], so that [outcome]," compressed into the standard need block.
2. Scope
In scope: primary jobs (monitor, warn, decide disposition), setup/configuration jobs, monitoring jobs, emergency/excursion-response jobs, maintenance/service jobs, troubleshooting jobs, and decommissioning/end-of-life jobs, across the reusable multi-trip lifecycle (D1). Out of scope: the detailed workflow steps (UN-FLOW) and the functions that satisfy each job (UN-FUN).
3. Applicable Standards & References
• WHO PQS E006 (excursion decision), EU GDP (custody/handling jobs), 21 CFR Part 11 (record jobs), UN 38.3 / IATA DGR (transport job), ISO/IEC 17025 concepts (calibration job), ISO 14971 (emergency/excursion-response framing). Quantitative targets Evidence Pending.
4. Background
The SCM's central job is not "log temperature" — it is "let the right person make the right disposition/rescue decision at the right moment, and defend it later." Because the device is reusable and crosses many custodians, the jobs multiply: provisioning and arming at pack-out, monitoring and warning in transit, deciding at receipt, servicing/recalibrating between trips, recovering strays, and retiring units.
5. Need Categories
• 5.1 Core jobs (monitor, warn, decide) · 5.2 Setup & configuration jobs · 5.3 Monitoring & reporting jobs · 5.4 Emergency / excursion-response jobs · 5.5 Maintenance & service jobs · 5.6 Troubleshooting & recovery jobs · 5.7 Decommissioning jobs.
6. Detailed User Needs
6.1 Core jobs
UN-JTBD-001 — Continuously know the product's thermal exposure through the whole trip
Need Statement | When a shipment is in the cold chain, the accountable stakeholder needs to have the product's temperature (and humidity/shock) exposure captured continuously for the full trip, so nothing material goes unrecorded.
Primary Stakeholder | Distributors (SH-03), manufacturers (SH-02).
Secondary Stakeholders | Programs (SH-01), 3PLs (SH-04), regulators (SH-07).
Importance | Essential
Market Priority | Must-have
Source | JTBD
Voice of Customer Reference | MR-JTBD-001; STN-001/005; VOC-010.
User Journey | Whole transit.
Hazard Relationship | A gap in capture (ISO 14971) hides an excursion → wrong disposition.
Traceability | Back: MR-JTBD-001. Forward: UN-FUN logging, UN-PERF endurance, UN-VAL.
Notes | "Continuous" means no unrecorded interval material to product stability; interval configurable.
UN-JTBD-002 — Decide accept / quarantine / reject at receipt, quickly and defensibly
Need Statement | When a consignment arrives, the receiver needs to decide accept, quarantine, or reject quickly and with a basis they can defend, so good stock is used and bad stock is caught.
Primary Stakeholder | Clinic/facility staff (SH-05), EPI programs (SH-01).
Secondary Stakeholders | Distributors (SH-03), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Contextual Inquiry
Voice of Customer Reference | MR-JTBD-002/004; STN-004/027; VOC-015.
User Journey | Receipt / point of use.
Hazard Relationship | Wrong disposition (ISO 14971 anchor) → spoiled dose used / safe dose wasted.
Traceability | Back: MR-JTBD-004. Forward: UN-FUN verdict, UN-UX offline legibility.
Notes | The edge decision; must work offline and untrained (UN-PER-007).
UN-JTBD-003 — Get warned of a break in time to rescue the product
Need Statement | When an excursion is developing, the accountable stakeholder needs to be warned while intervention is still possible, so the shipment can be rerouted, re-iced, or diverted before loss.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Essential
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | MR-JTBD-005; STN-001/018; VOC-008.
User Journey | Transit; hand-offs.
Hazard Relationship | Delayed/missing warning (ISO 14971) → rescuable loss becomes total loss.
Traceability | Back: MR-JTBD-005. Forward: UN-FUN alert/routing, UN-PERF latency.
Notes | D2 direction: immediate alert when connectivity available; latency connectivity-bound.
6.2 Setup & configuration jobs
UN-JTBD-004 — Provision, arm, and place a device for a trip in seconds, error-proof
Need Statement | At pack-out, the dispatch stakeholder needs to configure the trip profile, arm the device, and confirm correct placement in seconds, with no way to forget to start it or pick the wrong profile.
Primary Stakeholder | Manufacturer/distributor dispatch (SH-02/SH-03).
Secondary Stakeholders | Warehouse operators (SH-05), installers (SH-12).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | MR-JTBD-006; STN-008; VOC-002; FLD-001/003.
User Journey | Provisioning / arming.
Hazard Relationship | Forgot-to-arm or wrong profile (ISO 14971) → void trip / undetected excursion.
Traceability | Back: MR-JTBD-006. Forward: UN-FLOW setup, UN-FUN arming confirmation, UN-UX speed.
Notes | The compliant path must be the fastest path (conflict C4).
UN-JTBD-005 — Configure per-vaccine excursion rules
Need Statement | The QA/program stakeholder needs to set excursion thresholds and logic appropriate to the specific product's stability profile, so alerts and verdicts reflect what actually harms that product.
Primary Stakeholder | Manufacturer QA/RA (SH-02), EPI programs (SH-01).
Secondary Stakeholders | Distributors (SH-03), regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | STN-002/004/007; VOC-011.
User Journey | Profile setup before dispatch.
Hazard Relationship | Wrong thresholds (ISO 14971) → missed real excursion or false alarm.
Traceability | Back: STN-007. Forward: UN-FUN configurable thresholds, UN-SAFE two-sided detection.
Notes | Two-sided (freeze + heat), time-over-threshold / cumulative logic per product (conflict C8).
6.3 Monitoring & reporting jobs
UN-JTBD-006 — Watch a fleet of shipments and see only what needs action
Need Statement | The coordinator needs to monitor many concurrent shipments and be surfaced only the ones requiring action, so attention goes where it matters.
Primary Stakeholder | Distributors (SH-03), 3PL control tower (SH-04).
Secondary Stakeholders | Programs (SH-01), executive (SH-15).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | MR-JTBD-005/010; STN-007/018; VOC-005.
User Journey | Daily monitoring.
Hazard Relationship | Alarm fatigue → missed excursion (ISO 14971).
Traceability | Back: MR-JTBD-010. Forward: UN-ECO fleet view, UN-FUN action-only alerts.
Notes | Depends on connectivity for liveness (conflict C7); offline verdict remains the floor.
UN-JTBD-007 — Produce an audit-ready record and export it on demand
Need Statement | When challenged or audited, the QA/regulatory stakeholder needs to produce a complete, tamper-evident, exportable record of a trip on demand, so the disposition can be defended.
Primary Stakeholder | Manufacturer QA/RA (SH-02), auditors (SH-07).
Secondary Stakeholders | Distributors (SH-03), IT (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / JTBD
Voice of Customer Reference | MR-JTBD-003/008; STN-005; VOC-001/019.
User Journey | Audit / dispute / release.
Hazard Relationship | Incomplete export (ISO 14971) → indefensible disposition.
Traceability | Back: MR-JTBD-008. Forward: UN-SEC integrity, UN-ECO export, UN-REG Part 11/GDP.
Notes | ALCOA+; gaps flagged, never hidden.
UN-JTBD-008 — Attribute where and when a break happened
Need Statement | When product arrives out of range, the shipper needs objective, time- and location-stamped evidence of when and where the break occurred, so responsibility can be attributed fairly.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Differentiator
Source | VoC / JTBD
Voice of Customer Reference | MR-JTBD-003; STN-006; VOC-004; FLD-009.
User Journey | Dispute resolution at custody hand-offs.
Hazard Relationship | Not a direct patient hazard; commercial/attribution.
Traceability | Back: STN-006. Forward: UN-FUN timestamped events, UN-PERF location.
Notes | Depends on GPS-in-container performance (concept §7; conflict C3).
6.4 Emergency / excursion-response jobs
UN-JTBD-009 — Route an excursion alert to the person who can act, and escalate if unacknowledged
Need Statement | When an excursion alert fires, the operation needs it delivered to the right responder and escalated if not acknowledged, so a warning becomes an intervention.
Primary Stakeholder | Distributors / 3PL control tower (SH-03/SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | MR-JTBD-005; STN-018; VOC-008; CFB-002.
User Journey | Excursion response.
Hazard Relationship | Alert to the wrong/unavailable person (ISO 14971) → no intervention → loss.
Traceability | Back: STN-018. Forward: UN-FUN routing/escalation, UN-ECO notifications.
Notes | Cloud function; depends on connectivity (UN-JTBD path) and acknowledgement tracking.
6.5 Maintenance & service jobs
UN-JTBD-010 — Turn a returned device around: charge, calibrate, update, re-provision
Need Statement | When a device comes back, the service stakeholder needs to confirm and restore its readiness — charge, calibration currency, firmware — and re-provision it, without hand-checking each unit.
Primary Stakeholder | Service technicians (SH-06), service ops (SH-13).
Secondary Stakeholders | Regulators (SH-07), procurement (SH-08).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | MR-JTBD-007/009; STN-010/021; VOC-016/017.
User Journey | Reverse logistics / refurbishment.
Hazard Relationship | Out-of-cal or under-charged redeploy (ISO 14971) → disputed reading / dead-in-cold.
Traceability | Back: MR-JTBD-007. Forward: UN-SVC turnaround/calibration, UN-PERF charge readiness.
Notes | Charge-inhibit-below-0 °C (D3) affects readiness timing; cleaning between trips (UN-SVC).
6.6 Troubleshooting & recovery jobs
UN-JTBD-011 — Recover strays and get the reusable fleet back
Need Statement | The reverse-logistics stakeholder needs to locate and recover stray/overdue devices and drive a reliable return workflow, especially on export and last-mile lanes, so reuse economics hold.
Primary Stakeholder | Service / reverse-logistics (SH-13).
Secondary Stakeholders | Distributors (SH-03), programs (SH-01), finance (SH-10).
Importance | Important
Market Priority | Differentiator
Source | VoC / Market Analysis
Voice of Customer Reference | MR-JTBD-009/010; STN-016; VOC-006; FLD-007.
User Journey | Return / recovery.
Hazard Relationship | None direct; economic (conflict C6).
Traceability | Back: STN-016. Forward: UN-FUN last-known-location, UN-BIZ return rate.
Notes | Location (UN-PERF) supports recovery within the energy budget (conflict C3).
6.7 Decommissioning jobs
UN-JTBD-012 — Retire a device and dispose of its battery safely
Need Statement | At end of life, the service/sustainability stakeholder needs to retire a device and route it and its battery to a safe, compliant disposal/recycling path.
Primary Stakeholder | Service / sustainability ops (SH-13).
Secondary Stakeholders | 3PLs (SH-04), regulators (SH-07).
Importance | Desirable
Market Priority | Nice-to-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | MR-1700; STN-026; D1/D3; UN 38.3 / IATA.
User Journey | End-of-life.
Hazard Relationship | Unsafe Li-battery disposal (ISO 14971, transport) — safety/environmental.
Traceability | Back: STN-026. Forward: UN-REG battery EOL, UN-SVC decommission.
Notes | Cycle life / EOL path Evidence Pending; conflict C10.
7. Validation Notes
Jobs are validated by observing task completion in context (contextual inquiry, ride-alongs, service-depot time-and-motion — Stage 02 §6) and by summative usability testing. Completion-time and success-rate targets are Evidence Pending; the excursion-response job (UN-JTBD-003/009) is validated jointly with hazard analysis (HA-###, ORCH-017).
8. Traceability Placeholders
Upstream | MR-400 (JTBD); STN-001…027; D1–D4.
Downstream | UN-FLOW, UN-FUN, UN-PERF, UN-SVC; forward to PRD (ORCH-011). Matrix: ORCH-018 / EM-ENG-012.
UN-400 — Use Environments (ENV)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-1300 (Use Environments), D1–D4, and the concept operating range (−30 °C to +60 °C; 90-day endurance). Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users need the device to withstand and remain usable in, expressed as environment-of-use needs — the conditions the SCM must keep working through so that the users' jobs (UN-JTBD) succeed. It is solution-neutral (it says "must remain usable/accurate in condition X," not "use coating Y").
2. Scope
In scope: the thermal, humidity, mechanical (shock/vibration), ingress, RF/connectivity, and lighting environments across the whole journey — inside insulated shippers, in cold stores and freezers, on aircraft and trucks, at rural health posts, and at service depots. Out of scope: the qualification tests (UN-VAL, verification stage) and the enclosure implementation (UN-PHY / design outputs).
3. Applicable Standards & References
• IEC 60068 (environmental test conditions — thermal, humidity, vibration, shock).
• IEC 60529 (IP) (ingress protection ratings).
• CISPR / IEC 61000 (EMC — emissions/immunity); FCC Part 15 / CE RED (radio).
• IATA PCR/TCR, UN 38.3 (air/ground transport environment for Li batteries).
• WHO PQS E006 (cold-chain operating conditions). All cited Standard; exact test severities are set in PRD/verification and are Evidence Pending here.
4. Background
The concept places the device inside the insulated shipping container — a metallized, thermally isolated, RF-hostile enclosure — across a −30 °C to +60 °C range for up to 90 days (D3). It then must survive cold stores (−20 °C), freezer handling, aircraft holds (pressure/temperature), truck vibration, tropical humidity and condensation, dusty/rugged last-mile posts, and repeated cleaning at depots (D1). Each environment imposes a distinct usability or survival need.
5. Need Categories
• 5.1 Thermal environment · 5.2 Humidity & condensation · 5.3 Mechanical (shock/vibration) · 5.4 Ingress & cleaning environment · 5.5 RF / connectivity environment · 5.6 Lighting & point-of-use conditions · 5.7 EMC environment.
6. Detailed User Needs
6.1 Thermal environment
UN-ENV-001 — Remain fully functional across −30 °C to +60 °C
Need Statement | Users need the device to keep sensing, logging, alerting (when connected), and displaying accurately across the full −30 °C to +60 °C operating range, because the cold chain spans deep cold to tropical heat and the record must not degrade at the extremes.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01), service (SH-06).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Market Analysis
Voice of Customer Reference | STN-003; concept §2/§3 (D3); MR-1300.
User Journey | Whole transit; cold-store staging.
Hazard Relationship | Loss of function at temperature extreme (ISO 14971) → lost record exactly when needed; this is the dead-in-cold hazard.
Traceability | Back: STN-003, D3. Forward: UN-PERF endurance, UN-SAFE, IEC 60068 verification.
Notes | Top program risk (MR-RISK-001); closed by test (MR-VAL-006).
UN-ENV-002 — Display stays legible at low temperature
Need Statement | The warehouse operator and last-mile nurse need the device's on-device status to be legible when the device and display are cold (down to the operating extreme), so an offline verdict can still be read at the point of use.
Primary Stakeholder | Warehouse operators (SH-05), last-mile nurse (SH-05).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/008; MR-PER-005/007; FLD-003.
User Journey | Reading verdict at receipt / staging in the cold.
Hazard Relationship | Unreadable cold display (ISO 14971 use-error) → wrong disposition.
Traceability | Back: STN-004. Forward: UN-UX cold legibility, UN-PHY display.
Notes | OLED low-temp legibility thresholds Evidence Pending (Stage 02 §6).
6.2 Humidity & condensation
UN-ENV-003 — Tolerate high humidity and condensation cycles without measurement error
Need Statement | Users need the device to keep measuring accurately through tropical humidity and the condensation that forms when a cold device warms up, so humid lanes and freeze-thaw cycles don't corrupt readings.
Primary Stakeholder | Programs (SH-01), distributors (SH-03).
Secondary Stakeholders | Service (SH-06).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Regulatory
Voice of Customer Reference | MR-1300; concept §2; IEC 60068.
User Journey | Humid corridors; cold-to-warm transitions.
Hazard Relationship | Condensation-induced sensor error (ISO 14971) → false verdict.
Traceability | Back: MR-1300. Forward: UN-PERF accuracy, UN-PHY sealing, IEC 60068 verification.
Notes | Humidity sensing itself must remain valid; condensation on the temperature probe is a concern.
6.3 Mechanical (shock/vibration)
UN-ENV-004 — Survive transport shock and vibration without failure or false shock alerts
Need Statement | Users need the device to survive the drops, handling shocks, and sustained vibration of air and road transport without failing, and to distinguish real damaging shock from routine handling.
Primary Stakeholder | 3PLs (SH-04), distributors (SH-03).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Regulatory
Voice of Customer Reference | MR-1300; concept §2; IEC 60068.
User Journey | Air/road transport, handling.
Hazard Relationship | Device failure in transit (ISO 14971) → lost record; false shock alerts → fatigue.
Traceability | Back: MR-1300. Forward: UN-PHY durability, UN-FUN shock logic, IEC 60068 verification.
Notes | Shock detection must be product-relevant, not every bump (ties conflict C8).
6.4 Ingress & cleaning environment
UN-ENV-005 — Withstand dust, splash, and repeated cleaning between reuse cycles
Need Statement | Users need the device to keep out dust and moisture and to survive repeated cleaning and decontamination between trips without ingress, corrosion, or seal failure over its reusable life.
Primary Stakeholder | Service ops (SH-13), warehouse operators (SH-05).
Secondary Stakeholders | 3PLs (SH-04), programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-026; FLD-011; IEC 60529; D1.
User Journey | Depot cleaning; dusty last-mile posts.
Hazard Relationship | Ingress after cleaning (ISO 14971) → sensor drift/failure → false verdict.
Traceability | Back: STN-026. Forward: UN-PHY IP rating/sealing, UN-SVC cleaning method, IEC 60529.
Notes | USB-C port sealing and cleaning-chemical compatibility are named concerns (flag materials eng).
6.5 RF / connectivity environment
UN-ENV-006 — Get data out from inside a metallized, insulated container
Need Statement | Users need the device to still capture and eventually deliver its data even though it sits inside an RF-hostile, insulated/metallized shipper where wireless signals are attenuated.
Primary Stakeholder | Distributors (SH-03), IT/fleet (SH-09).
Secondary Stakeholders | Programs (SH-01), 3PLs (SH-04).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Market Analysis
Voice of Customer Reference | STN-009; FLD-004/005; concept §7 (RF-in-container); D2.
User Journey | Transit inside the shipper.
Hazard Relationship | No data egress (ISO 14971) → delayed/missed alert; store-and-forward mitigates.
Traceability | Back: STN-009. Forward: UN-PERF connectivity, UN-FUN store-and-forward, UN-PHY antenna.
Notes | Antenna/GPS performance inside metallized containers is a named concept gap; may drive external-antenna option (Evidence Pending; RF-in-container test).
UN-ENV-007 — Keep working through intermittent connectivity and power at the last mile
Need Statement | Last-mile users need the device to keep logging and give a verdict where connectivity and mains power are unreliable or absent, uploading opportunistically when a link appears.
Primary Stakeholder | EPI programs (SH-01), clinic staff (SH-05).
Secondary Stakeholders | Distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / Market Analysis
Voice of Customer Reference | STN-009; VOC-023; FLD-004; D2.
User Journey | Remote health-post receipt and storage.
Hazard Relationship | Dependence on live link (ISO 14971) → no verdict where losses concentrate.
Traceability | Back: STN-009. Forward: UN-FUN offline verdict, UN-PERF buffering.
Notes | Offline-first (D2); live data is upside, never a dependency (conflict C7).
6.6 Lighting & point-of-use conditions
UN-ENV-008 — Be readable in the lighting of a real health post or dock
Need Statement | Users need the on-device status to be readable across the lighting conditions of real use — dim rural posts, bright loading docks, direct sun — so the verdict is never lost to glare or gloom.
Primary Stakeholder | Clinic/last-mile staff (SH-05).
Secondary Stakeholders | Warehouse operators (SH-05), 3PLs (SH-04).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | MR-1300; STN-004/027.
User Journey | Receipt/handling in varied lighting.
Hazard Relationship | Unreadable in glare/dark (ISO 14971 use-error) → wrong disposition.
Traceability | Back: MR-1300. Forward: UN-UX readability, UN-PHY display.
Notes | Complements cold legibility (UN-ENV-002).
6.7 EMC environment
UN-ENV-009 — Coexist electromagnetically on aircraft and in facilities
Need Statement | Users and operators need the device to neither emit interference that disrupts aircraft/facility systems nor malfunction due to external interference, so it is accepted for air transport and enterprise use.
Primary Stakeholder | 3PLs / air-cargo (SH-04), IT (SH-09).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | MR-1100; D4; CISPR/IEC 61000; FCC Part 15 / CE RED.
User Journey | Air transport; enterprise deployment.
Hazard Relationship | EMC malfunction (ISO 14971) → corrupted data or unsafe transport.
Traceability | Back: MR-1100. Forward: UN-REG EMC/radio, verification (EMC tests).
Notes | Radio approvals per market (D4); air-transport RF acceptance.
7. Validation Notes
Environmental needs are validated by IEC 60068 (thermal/humidity/vibration/shock), IEC 60529 (ingress), CISPR/IEC 61000 + FCC/CE (EMC/radio), and a dedicated RF-in-container and cold-endurance test (MR-VAL-006). Exact severities/durations are set in PRD/verification and are Evidence Pending here.
8. Traceability Placeholders
Upstream | MR-1300 (ENV); MR-1100 (regulatory); concept §2/§3/§7; D3/D4.
Downstream | UN-PERF, UN-PHY, UN-SAFE, UN-REG; forward to PRD (ORCH-011) and verification. Matrix: ORCH-018 / EM-ENG-012.
UN-500 — User Workflow (FLOW)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-400 (JTBD), MR-1200/1300, STN records, and D1–D4 (reusable multi-trip lifecycle). Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the needs users have at each step of the workflow as a reusable device moves through its lifecycle — from acquisition and setup, through the trip, to return, refurbishment, and eventual retirement. Needs are solution-neutral ("the user needs the step to be X"), and each step is written so it can be observed and validated.
2. Scope
In scope: discovery/onboarding, unboxing/first-use, provisioning/arming, in-trip handling, receipt, excursion response, return/reverse-logistics, cleaning, charging, recalibration/re-provisioning, firmware update, troubleshooting, and disposal — across the D1 reusable loop. Out of scope: the functions used at each step (UN-FUN) and the service-depot capabilities (UN-SVC).
3. Applicable Standards & References
• WHO PQS E006 (arming/receipt workflow expectations), EU GDP (custody handling), 21 CFR Part 11 (record continuity across steps), ISO/IEC 17025 concepts (recalibration step), IEC 62366-1 (workflow use-error). Step-time targets Evidence Pending.
4. Background
Unlike a single-use logger, the SCM's workflow is a loop, not a line: recover → clean → charge → recalibrate → re-provision → redeploy, many times over. The most failure-prone steps are provisioning (forgot-to-arm), receipt (wrong disposition), and return (device never comes home). This module surfaces the user need at each transition so the loop stays fast, correct, and closed.
5. Need Categories
• 5.1 Acquire & onboard · 5.2 Provision & arm · 5.3 In-trip & hand-off · 5.4 Receive & decide · 5.5 Return & refurbish · 5.6 Charge, update & troubleshoot · 5.7 Retire.
6. Detailed User Needs
6.1 Acquire & onboard
UN-FLOW-001 — Onboard a new device fleet without specialist effort
Need Statement | The IT/fleet stakeholder needs to register and onboard new devices into the fleet simply and in bulk, so a program can stand up monitoring without specialist integration effort per unit.
Primary Stakeholder | IT/fleet admins (SH-09).
Secondary Stakeholders | Service ops (SH-13), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Stakeholder Interview
Voice of Customer Reference | STN-013; MR-1500 ECO.
User Journey | First deployment / fleet expansion.
Hazard Relationship | Mis-registered device (ISO 14971) → orphaned record / wrong attribution.
Traceability | Back: STN-013. Forward: UN-ECO fleet onboarding, UN-SEC identity.
Notes | Bulk onboarding supports fleet scale (SH-15).
6.2 Provision & arm
UN-FLOW-002 — Provision and arm with a single confirmed action
Need Statement | The dispatch user needs to select the trip profile and arm the device in a single, confirmed action at pack-out, so arming is never forgotten and the wrong profile is never silently applied.
Primary Stakeholder | Dispatch (SH-02/SH-03), warehouse operators (SH-05).
Secondary Stakeholders | Installers (SH-12).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-008; VOC-002; FLD-001/003; MR-JTBD-006.
User Journey | Provisioning / arming.
Hazard Relationship | Forgot-to-arm / wrong profile (ISO 14971) → void trip / missed excursion.
Traceability | Back: STN-008. Forward: UN-FUN arming confirmation, UN-UX one-action, UN-VAL provisioning.
Notes | On-device confirmation of "armed/placed"; on-device fallback if BLE/app fails (conflict C4).
UN-FLOW-003 — Confirm correct placement before the box is sealed
Need Statement | The user needs clear confirmation that the device and its probe are placed correctly before sealing the shipper, so readings reflect the product zone from the first minute.
Primary Stakeholder | Warehouse operators (SH-05), installers (SH-12).
Secondary Stakeholders | Dispatch (SH-02/SH-03).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-024; FLD-002; MR-JTBD-006.
User Journey | Placement at pack-out.
Hazard Relationship | Misplaced probe (ISO 14971) → biased freeze/heat detection → wrong verdict.
Traceability | Back: STN-024. Forward: UN-FUN placement guidance, UN-PHY probe/mount.
Notes | Guidance for internal vs external probe placement.
6.3 In-trip & hand-off
UN-FLOW-004 — Pass through handling and hand-offs without any user action
Need Statement | Handling users (never opening the box) need the device to require no action from them in transit, so it neither slows the dock nor depends on handlers doing anything.
Primary Stakeholder | 3PLs / handlers (SH-04).
Secondary Stakeholders | Distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-012; VOC-007; MR-PER-004.
User Journey | Transit; cross-dock / tarmac hand-offs.
Hazard Relationship | None direct; handling-neutrality avoids workarounds that defeat monitoring.
Traceability | Back: STN-012. Forward: UN-PHY handling neutrality, UN-FUN autonomous logging.
Notes | "Invisible-until-relevant" (MR-PER-004).
6.4 Receive & decide
UN-FLOW-005 — Read the trip verdict at receipt without tools or training
Need Statement | The receiver needs to read a clear accept/quarantine/reject verdict directly from the device at receipt, without a phone, network, login, or manual.
Primary Stakeholder | Clinic/last-mile staff (SH-05), EPI programs (SH-01).
Secondary Stakeholders | Distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004/027; VOC-014/015; FLD-006.
User Journey | Receipt / point of use.
Hazard Relationship | Misread verdict (ISO 14971 use-error) → wrong disposition at point of care.
Traceability | Back: STN-004. Forward: UN-FUN offline verdict, UN-UX legibility.
Notes | The archetypal edge step; design for the least-trained user (UN-PER-007).
UN-FLOW-006 — Respond to an in-trip excursion alert and act
Need Statement | When alerted to an excursion in transit, the responder needs enough context to act (what, where, how severe) and a way to acknowledge and resolve it, so the warning becomes an intervention.
Primary Stakeholder | Distributors / control tower (SH-03/SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | STN-001/018; VOC-008; CFB-002.
User Journey | Excursion response.
Hazard Relationship | No actionable context (ISO 14971) → no intervention → loss.
Traceability | Back: STN-018. Forward: UN-FUN alert content/ack, UN-ECO notifications.
Notes | Acknowledgement/escalation tracked; depends on connectivity.
6.5 Return & refurbish
UN-FLOW-007 — Return the device home through a simple, defined workflow
Need Statement | The user handling returns needs a simple, defined return path (labelling, drop-off, or pooled return) so reusable devices reliably come back, especially on hard export/last-mile lanes.
Primary Stakeholder | Service / reverse-logistics (SH-13).
Secondary Stakeholders | Programs (SH-01), distributors (SH-03).
Importance | Important
Market Priority | Differentiator
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-016; VOC-006; FLD-007; LD-008; D1.
User Journey | Return / reverse logistics.
Hazard Relationship | None direct; return rate governs reuse economics (conflict C6).
Traceability | Back: STN-016. Forward: UN-SVC return workflow, UN-BIZ return rate.
Notes | Last-known-location (UN-PERF) helps chase strays.
UN-FLOW-008 — Clean and decontaminate the device between trips
Need Statement | The service user needs to clean and decontaminate a returned device between trips without damaging it or compromising its seal, so it is safe and ready to redeploy.
Primary Stakeholder | Service ops (SH-13), technicians (SH-06).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-026; FLD-011; IEC 60529; D1.
User Journey | Refurbishment.
Hazard Relationship | Cleaning-induced ingress (ISO 14971) → sensor error next trip.
Traceability | Back: STN-026. Forward: UN-SVC cleaning method, UN-PHY sealing.
Notes | Cleaning-chemical compatibility and port sealing — flag materials engineer.
6.6 Charge, update & troubleshoot
UN-FLOW-009 — Recharge a returned device and know when it is trip-ready
Need Statement | The service user needs to recharge a returned device and get a clear indication of when it is fully ready for the next trip, including when a cold device must warm before it will charge.
Primary Stakeholder | Service technicians (SH-06).
Secondary Stakeholders | Service ops (SH-13).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-010; CFB-016; D3.
User Journey | Charging at depot.
Hazard Relationship | Charging a cold battery (ISO 14971, battery safety) — charge-inhibit-below-0 °C; under-charge → dead-in-cold next trip.
Traceability | Back: STN-010, D3. Forward: UN-PERF charge readiness, UN-SAFE charge inhibit.
Notes | Clear "warming / will charge" state (D3 charge-inhibit-below-0 °C; self-heat under evaluation).
UN-FLOW-010 — Update firmware safely and know it succeeded
Need Statement | The IT/service user needs to update device firmware over the fleet safely and confirm each update succeeded, so devices stay current without being bricked or left in an unknown state.
Primary Stakeholder | IT/fleet (SH-09), service (SH-06).
Secondary Stakeholders | Regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Stakeholder Interview / Regulatory
Voice of Customer Reference | STN-025; VOC-019; D2 (OTA).
User Journey | Firmware update (OTA or at depot).
Hazard Relationship | Failed/unsigned update (ISO 14971 + security) → bricked or compromised device.
Traceability | Back: STN-025. Forward: UN-SEC signed/fail-safe OTA, UN-SVC field upgrade.
Notes | Signed, resumable, fail-safe (UN-SEC); confirm success per device.
6.7 Retire
UN-FLOW-011 — Retire and dispose of a device at end of life
Need Statement | The service/sustainability user needs a clear way to identify a device at end of life and route it and its battery to compliant disposal/recycling.
Primary Stakeholder | Service / sustainability (SH-13).
Secondary Stakeholders | Regulators (SH-07), 3PLs (SH-04).
Importance | Desirable
Market Priority | Nice-to-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-026; MR-1700; UN 38.3 / IATA; D3.
User Journey | End-of-life.
Hazard Relationship | Unsafe battery disposal (ISO 14971, transport) — safety/environmental.
Traceability | Back: STN-026. Forward: UN-REG battery EOL, UN-SVC decommission.
Notes | Health/cycle-life indication supports retire timing (Evidence Pending).
7. Validation Notes
Workflow needs are validated by end-to-end walkthroughs and time-and-motion studies at dispatch, receipt, and service depots (Stage 02 §6), and by summative usability testing of the most error-prone steps (provisioning, receipt). Step-time and error-rate targets are Evidence Pending.
8. Traceability Placeholders
Upstream | MR-400 (JTBD); STN-004/008/010/016/024/025/026; D1–D3.
Downstream | UN-FUN, UN-UX, UN-SVC, UN-SEC, UN-PERF; forward to PRD (ORCH-011). Matrix: ORCH-018 / EM-ENG-012.
UN-600 — Functional Needs (FUN)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-300/400, VOC themes T1–T14, STN-001…027, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states, in solution-neutral terms, what the SCM must do for its users — the customer-driven functions. Each need says what capability the user needs ("the user needs to be warned…"), never how it is implemented ("the system shall send MQTT…"). Functions here are the backbone from which product/system requirements (PRD/SYS) are derived.
2. Scope
In scope: sensing/logging, excursion detection (two-sided), alerting and routing, on-device verdict, store-and-forward, configuration, record/export, attribution (time/location), fleet visibility, provisioning and readiness reporting, and firmware update — as user needs. Out of scope: performance targets (UN-PERF), safety-specific behaviours (UN-SAFE), security mechanisms (UN-SEC), and implementation.
3. Applicable Standards & References
• WHO PQS E006 (monitoring/verdict function), EU GDP + 21 CFR Part 11 + ALCOA+ (record function), ISO 14971 (excursion/alert as risk controls), ISO/IEC 17025 concepts (calibration reporting). Functional thresholds/latencies are set in PRD and are Evidence Pending here.
4. Background
The SCM's functions exist to serve one chain of value: capture → detect → warn → decide → prove. D2 shapes several of them: local logging + scheduled cadence upload + immediate excursion alert when connected, with offline buffering. Two-sided detection (freeze as well as heat) is central because freezing damage is common, silent, and under-detected. The on-device verdict exists because the last-mile user has no phone, network, or training.
5. Need Categories
• 5.1 Sense & log · 5.2 Detect & alert · 5.3 Decide (on-device verdict) · 5.4 Store & forward · 5.5 Configure · 5.6 Record, prove & export · 5.7 Locate & attribute · 5.8 Fleet & manage.
6. Detailed User Needs
6.1 Sense & log
UN-FUN-001 — Continuously measure and log temperature over the whole trip
Need Statement | Users need the device to measure and record temperature continuously at a configurable interval for the whole trip, so the thermal history is complete and defensible.
Primary Stakeholder | Manufacturers (SH-02), distributors (SH-03).
Secondary Stakeholders | Programs (SH-01), regulators (SH-07).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | MR-JTBD-001; STN-001/005; VOC-010.
User Journey | Whole transit.
Hazard Relationship | Logging gap (ISO 14971) → hidden excursion → wrong disposition.
Traceability | Back: STN-005. Forward: PRD logging/interval, UN-PERF endurance, UN-VAL.
Notes | Interval default 5–10 min, configurable (D2); no unrecorded material interval.
UN-FUN-002 — Measure humidity and shock alongside temperature
Need Statement | Users need the device to also record humidity and shock/impact events, so exposures beyond temperature that affect product or indicate mishandling are captured.
Primary Stakeholder | Distributors (SH-03), manufacturers (SH-02).
Secondary Stakeholders | 3PLs (SH-04).
Importance | Important
Market Priority | Must-have
Source | JTBD / Market Analysis
Voice of Customer Reference | MR-JTBD-001; concept §1; STN-006.
User Journey | Whole transit.
Hazard Relationship | Missed humidity/shock exposure (ISO 14971) → undetected product risk.
Traceability | Back: concept §1. Forward: PRD humidity/accelerometer, UN-ENV-003/004.
Notes | Shock detection must be product-relevant, not every bump (conflict C8).
UN-FUN-003 — Support measuring the product zone, not just the air near the device
Need Statement | Users need to measure the temperature that represents the product's actual zone (via an appropriately placed probe), so the record reflects what the vaccine experienced, not just ambient near the electronics.
Primary Stakeholder | Manufacturers (SH-02), installers (SH-12).
Secondary Stakeholders | Programs (SH-01), regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-002/024; FLD-002; MR-JTBD-001.
User Journey | Placement and transit.
Hazard Relationship | Unrepresentative measurement (ISO 14971) → false accept/reject.
Traceability | Back: STN-024. Forward: PRD probe support, UN-FUN-013 placement guidance.
Notes | Internal + external-probe temperature supports freeze detection (STN-002).
6.2 Detect & alert
UN-FUN-004 — Detect both freezing and overheating (two-sided)
Need Statement | Users need the device to detect and flag inadvertent freezing as reliably as overheating, because freeze-sensitive vaccines are silently ruined by over-cold exposure that heat-biased indicators miss.
Primary Stakeholder | EPI programs (SH-01), manufacturers (SH-02).
Secondary Stakeholders | Distributors (SH-03), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-002; VOC-011; FLD-002; MR-PROB-002.
User Journey | Transit; receipt verdict.
Hazard Relationship | Missed freeze (ISO 14971) → spoiled freeze-sensitive dose administered — a top loss cause; severity anchor UN-VIS-001.
Traceability | Back: STN-002. Forward: PRD two-sided thresholds, UN-SAFE, UN-VAL.
Notes | Configurable low/high thresholds; Cited to WHO/UNICEF freezing reporting.
UN-FUN-005 — Judge excursions by product-relevant logic, not raw transients
Need Statement | Users need excursion detection based on product-stability-relevant exposure (e.g., time-over/under-threshold or cumulative logic), not every momentary transient, so alerts and verdicts reflect actual product harm.
Primary Stakeholder | Distributors (SH-03), manufacturers (SH-02).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | VoC / JTBD
Voice of Customer Reference | STN-007; VOC-005/022; FLD-010; MR-PROB-009.
User Journey | Transit; receipt.
Hazard Relationship | Over-sensitivity → alarm fatigue → muted real alert (ISO 14971); under-sensitivity → missed excursion. The C8 balance.
Traceability | Back: STN-007. Forward: PRD excursion logic, UN-SAFE, UN-UX alert quality.
Notes | Configurable per vaccine stability profile (ties UN-FUN-011).
UN-FUN-006 — Push an immediate alert on a detected excursion when connectivity allows
Need Statement | Users need an immediate excursion alert when connectivity is available, so they learn of a break in time to intervene rather than at delivery.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Essential
Market Priority | Must-have
Source | VoC / JTBD
Voice of Customer Reference | STN-001; VOC-008; MR-JTBD-005; D2.
User Journey | Transit.
Hazard Relationship | Delayed/missing alert (ISO 14971) → rescuable loss becomes total.
Traceability | Back: STN-001. Forward: PRD alerting/latency, UN-PERF latency, HA-###.
Notes | Latency is connectivity-bound; event captured locally and surfaced on reconnection (D2).
UN-FUN-007 — Route and escalate alerts to the responder who can act
Need Statement | Users need excursion alerts routed to the right responder by role/shipment and escalated if unacknowledged, so a warning reliably reaches someone who can intervene.
Primary Stakeholder | Distributors / control tower (SH-03/SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | STN-018; VOC-008; CFB-002; MR-JTBD-005.
User Journey | Excursion response.
Hazard Relationship | Alert to wrong/unavailable person (ISO 14971) → no intervention.
Traceability | Back: STN-018. Forward: PRD routing/escalation, UN-ECO notifications.
Notes | Acknowledgement tracked; configurable by role; depends on connectivity.
6.3 Decide (on-device verdict)
UN-FUN-008 — Give a clear accept/reject verdict on the device itself, offline
Need Statement | Users need an unambiguous accept/reject (and quarantine) verdict readable on the device itself with no phone, network, login, or manual, so a non-technical last-mile user can decide in seconds.
Primary Stakeholder | Clinic/last-mile staff (SH-05), EPI programs (SH-01).
Secondary Stakeholders | Distributors (SH-03), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004/027; VOC-009/014/015; FLD-006.
User Journey | Receipt / point of use.
Hazard Relationship | Wrong/misread verdict (ISO 14971) → spoiled dose accepted / safe dose rejected — highest-consequence edge decision.
Traceability | Back: STN-004. Forward: PRD on-device verdict, UN-UX legibility, UN-VAL persona test.
Notes | Verdict must be PQS-consistent (category TBC); layered vs full record (conflict C5).
6.4 Store & forward
UN-FUN-009 — Keep working and buffer data when offline
Need Statement | Users need the device to keep logging and give a verdict with no live connection, and to buffer data and upload it opportunistically when a link appears, so the last mile is not a black box.
Primary Stakeholder | EPI programs (SH-01), distributors (SH-03).
Secondary Stakeholders | 3PLs (SH-04), IT (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / Market Analysis
Voice of Customer Reference | STN-009; VOC-023; FLD-004; D2.
User Journey | Transit; remote receipt.
Hazard Relationship | Dependence on live link (ISO 14971) → no verdict/record where losses concentrate.
Traceability | Back: STN-009. Forward: PRD store-and-forward/buffer, UN-PERF buffering.
Notes | Buffered data uploaded in intact time order; no data loss (target Evidence Pending).
UN-FUN-010 — Upload on a schedule and offload over whatever link is available
Need Statement | Users need the device to upload its data on a configurable cadence and to use whatever connectivity is available (cellular, Wi-Fi, or short-range), so data reaches the cloud without draining the 90-day budget.
Primary Stakeholder | IT/fleet (SH-09), distributors (SH-03).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | JTBD / Market Analysis
Voice of Customer Reference | STN-009/014; D2; concept §1.
User Journey | Transit; opportunistic offload.
Hazard Relationship | Aggressive live upload (ISO 14971) → battery exhaustion → dead-in-cold (conflict C2).
Traceability | Back: D2. Forward: PRD cadence/offload, UN-PERF energy budget.
Notes | Cadence default hourly, configurable; energy budget is the constraint.
6.5 Configure
UN-FUN-011 — Configure trip profiles and thresholds per product
Need Statement | QA/program users need to define and select trip profiles — thresholds, excursion logic, logging interval — appropriate to each product, so monitoring matches what harms that specific vaccine.
Primary Stakeholder | Manufacturer QA/RA (SH-02), EPI programs (SH-01).
Secondary Stakeholders | Distributors (SH-03), regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | STN-004/007; VOC-011; MR-JTBD-006.
User Journey | Profile setup before dispatch.
Hazard Relationship | Wrong profile (ISO 14971) → missed excursion / false alarm.
Traceability | Back: STN-007. Forward: PRD profiles/config, UN-FLOW provisioning.
Notes | Error-proof selection; sane defaults; wrong profile must not be silently applied (conflict C4).
UN-FUN-012 — Arm the device and confirm it is armed and placed
Need Statement | The dispatch user needs to arm the device and get positive on-device confirmation that it is armed and correctly placed, so a trip never starts unmonitored.
Primary Stakeholder | Dispatch (SH-02/SH-03), warehouse operators (SH-05).
Secondary Stakeholders | Installers (SH-12).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-008; VOC-002; FLD-001/003.
User Journey | Arming.
Hazard Relationship | Forgot-to-arm (ISO 14971) → void trip → undetected excursion.
Traceability | Back: STN-008. Forward: PRD arming confirmation, UN-UX one-action.
Notes | On-device fallback if BLE/app unavailable.
UN-FUN-013 — Guide correct probe/device placement
Need Statement | The installer/operator needs clear guidance on where to place the device and probe so readings represent the product zone and freeze/heat detection is unbiased.
Primary Stakeholder | Installers (SH-12), warehouse operators (SH-05).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-024; FLD-002; MR-JTBD-006.
User Journey | Placement at pack-out / install.
Hazard Relationship | Wrong placement (ISO 14971) → biased detection → wrong verdict.
Traceability | Back: STN-024. Forward: PRD placement guidance, UN-PHY probe/mount.
Notes | Internal vs external probe guidance; ties UN-FUN-003.
6.6 Record, prove & export
UN-FUN-014 — Produce a complete, tamper-evident, attributable record
Need Statement | Users need a complete, tamper-evident, attributable record of each trip — with any gap explicitly flagged, never hidden — so a disposition can be defended to an inspector.
Primary Stakeholder | Manufacturer QA/RA (SH-02), regulators/auditors (SH-07).
Secondary Stakeholders | Distributors (SH-03), IT (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / VoC
Voice of Customer Reference | STN-005; VOC-001/019; MR-JTBD-008.
User Journey | Release; audit; dispute.
Hazard Relationship | Editable/gapped record (ISO 14971) → indefensible or wrong disposition.
Traceability | Back: STN-005. Forward: PRD record/audit-trail, UN-SEC integrity, UN-REG Part 11/GDP.
Notes | ALCOA+ aligned; gaps flagged. Overlaps UN-SEC-004 (integrity mechanism side).
UN-FUN-015 — Export a trip record on demand in a usable, shareable form
Need Statement | Users need to export a trip's record on demand in a form they can share with partners, auditors, and quality systems, so evidence can move where the decision or dispute is.
Primary Stakeholder | QA/RA (SH-02), IT (SH-09).
Secondary Stakeholders | Regulators (SH-07), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-005/013; MR-JTBD-003/008.
User Journey | Audit / dispute / handover.
Hazard Relationship | Incomplete export (ISO 14971) → indefensible disposition.
Traceability | Back: STN-005/013. Forward: PRD export, UN-ECO integration/export.
Notes | Export must preserve integrity/attribution; ties UN-ECO API.
6.7 Locate & attribute
UN-FUN-016 — Time- and location-stamp events to attribute where a break occurred
Need Statement | Users need excursion and custody events time- and location-stamped, so it can be shown when and where a break happened and responsibility attributed at a hand-off.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Differentiator
Source | VoC / JTBD
Voice of Customer Reference | STN-006/015; VOC-004; FLD-009; MR-PROB-006.
User Journey | Dispute at custody hand-off.
Hazard Relationship | None direct; commercial attribution.
Traceability | Back: STN-006. Forward: PRD timestamped/geostamped events, UN-PERF location.
Notes | Location within the energy budget (conflict C3); GPS-in-container performance Evidence Pending.
6.8 Fleet & manage
UN-FUN-017 — Report device readiness (charge, calibration, firmware) to the fleet
Need Statement | Service and fleet users need each device to report its charge state, calibration status, and firmware version, so returned devices can be triaged and out-of-cal or under-charged units are blocked from redeployment without hand-checking.
Primary Stakeholder | Service technicians (SH-06), IT/fleet (SH-09).
Secondary Stakeholders | Regulators (SH-07), service ops (SH-13).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-010/021; VOC-016/017; CFB-012.
User Journey | Return / refurbishment triage.
Hazard Relationship | Out-of-cal redeploy (ISO 14971) → disputed readings; under-charge → dead-in-cold.
Traceability | Back: STN-010/021. Forward: PRD readiness reporting, UN-SVC turnaround/calibration.
Notes | Deploy blocked if out-of-cal; ties UN-PERF charge readiness.
UN-FUN-018 — See the fleet: which devices, where, and which need action
Need Statement | Fleet and program users need a network-level view of devices and shipments — status, location, and which need action — so a fleet can be managed and strays chased.
Primary Stakeholder | IT/fleet (SH-09), executive (SH-15).
Secondary Stakeholders | Service ops (SH-13), distributors (SH-03).
Importance | Important
Market Priority | Differentiator
Source | JTBD / Stakeholder Interview
Voice of Customer Reference | STN-014/015/016; VOC-020; MR-JTBD-010.
User Journey | Fleet monitoring and recovery.
Hazard Relationship | None direct; supports loss reduction and recovery.
Traceability | Back: STN-014. Forward: UN-ECO fleet dashboard, UN-BIZ ROI.
Notes | Liveness connectivity-bound (conflict C7); location energy-bound (conflict C3).
7. Validation Notes
Functional needs are validated by verification of the derived PRD/SYS requirements and by validation scenarios (UN-VAL): two-sided detection accuracy, excursion-logic false-alarm/miss rates, offline verdict correctness, store-and-forward integrity, alert latency vs connectivity, and record completeness/export in audit simulation. Thresholds and rates are Evidence Pending; safety-relevant functions (UN-FUN-004/005/ 006/008) are cross-validated with hazard analysis (HA-###, ORCH-017).
8. Traceability Placeholders
Upstream | MR-300/400; STN-001…027; D2 (store-and-forward), D1/D3.
Downstream | UN-PERF, UN-SAFE, UN-SEC, UN-UX, UN-ECO, UN-SVC; forward to PRD (ORCH-011) and hazard analysis. Matrix: ORCH-018 / EM-ENG-012.
UN-700 — User Experience Needs (UX)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on MR-1200 personas, contextual-inquiry themes, IEC 62366, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states what users need the SCM to feel like to use — legible, fast, error-resistant, trustworthy — across the enormous literacy range from a supply-chain director to a rural nurse. UX needs here are the human-factors layer of the same functions defined in UN-FUN; they are the primary defence against use-error, which is a patient-safety concern (IEC 62366 / ISO 14971).
2. Scope
In scope: offline legibility, low-literacy/low-training operability, gloved/cold interaction, minimal training, fast setup, error prevention, consistency, feedback, guidance/help, and trust/confidence. Out of scope: the functions themselves (UN-FUN), display hardware (UN-PHY), and validation methods (UN-VAL).
3. Applicable Standards & References
• IEC 62366-1 — usability engineering for medical devices; use-related risk.
• FDA Human Factors guidance — use-error and critical-task focus.
• WHO PQS E006 — operator expectations in EPI settings.
• ISO 9241 concepts — ergonomics of human-system interaction.
• Cold-environment / low-literacy HMI literature — Judgment/Assumption; thresholds Evidence Pending.
4. Background
The central UX tension (MR-1200, conflict C5) is one device, many literacies: the OLED-plus-button interface must be self-evident to the least-trained user while the data model satisfies the most exacting auditor. The highest-consequence interaction is the last-mile nurse reading an offline accept/reject verdict with cold, gloved hands, in poor lighting, having had minimal training. Every UX need is anchored to preventing a wrong disposition through use-error.
5. Need Categories
• 5.1 Legibility & readability · 5.2 Low-literacy / minimal-training operability · 5.3 Speed & simplicity · 5.4 Error prevention · 5.5 Feedback, guidance & help · 5.6 Consistency & trust.
6. Detailed User Needs
6.1 Legibility & readability
UN-UX-001 — Read the verdict at a glance, offline, in any real lighting
Need Statement | The receiver needs to read the accept/reject verdict at a glance from the device, offline, in dim, bright, or sunlit conditions, so the disposition decision is never lost to glare or gloom.
Primary Stakeholder | Clinic/last-mile staff (SH-05), programs (SH-01).
Secondary Stakeholders | Warehouse operators (SH-05).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/027; VOC-014/015; MR-PER-007.
User Journey | Receipt / point of use.
Hazard Relationship | Unreadable verdict (IEC 62366 use-error; ISO 14971) → wrong disposition.
Traceability | Back: STN-004. Forward: UN-PHY display, UN-ENV lighting/cold, UN-VAL.
Notes | Complements UN-ENV-002/008; layered with full record for QA (conflict C5).
UN-UX-002 — Stay legible with cold hands and a cold display
Need Statement | The cold-store operator and nurse need the display and controls to remain legible and operable when device and hands are cold, so the verdict can be read and the device used in the freezer/field.
Primary Stakeholder | Warehouse operators (SH-05), nurse (SH-05).
Secondary Stakeholders | Service (SH-06).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008; FLD-003; MR-PER-005.
User Journey | Cold-store staging; cold-field receipt.
Hazard Relationship | Cold-illegible display (ISO 14971) → misread → wrong disposition.
Traceability | Back: STN-008. Forward: UN-ENV-002, UN-PHY display.
Notes | OLED low-temp legibility thresholds Evidence Pending.
6.2 Low-literacy / minimal-training operability
UN-UX-003 — Understand the status without language, training, or a manual
Need Statement | The least-trained user needs the device's status to be understandable through clear, language-independent indication, so someone with minimal training and no manual can act correctly.
Primary Stakeholder | Last-mile nurse (SH-05). (Novice / low-literacy user.)
Secondary Stakeholders | Programs (SH-01), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / JTBD
Voice of Customer Reference | STN-004; VOC-014; MR-PER-007.
User Journey | Receipt at rural post.
Hazard Relationship | Misunderstood status (IEC 62366) → wrong disposition — highest-consequence.
Traceability | Back: STN-004. Forward: UN-FUN verdict, UN-VAL low-literacy usability.
Notes | Icon/colour + minimal text; avoid ambiguity between "cold-OK" and "frozen-bad".
UN-UX-004 — Operate the core tasks with a single, obvious control
Need Statement | Users need the core tasks (read verdict, arm) to be doable with a single, obvious control, so there is little to learn and little to get wrong.
Primary Stakeholder | Clinic staff (SH-05), warehouse operators (SH-05).
Secondary Stakeholders | Dispatch (SH-02/SH-03).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/008; FLD-003.
User Journey | Arming and receipt.
Hazard Relationship | Complex control (IEC 62366) → mis-operation → void trip or wrong disposition.
Traceability | Back: STN-008. Forward: UN-PHY controls, UN-FUN one-action arming.
Notes | Glove-friendly single-action interaction (MR-PER-005).
6.3 Speed & simplicity
UN-UX-005 — Provision and arm in seconds, with the compliant path the fastest path
Need Statement | The dispatch user needs provisioning and arming to take seconds, with the correct, complete, compliant path being the fastest path, so rigor never tempts a skipped or void trip.
Primary Stakeholder | Dispatch (SH-02/SH-03).
Secondary Stakeholders | Warehouse operators (SH-05), installers (SH-12).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-008; VOC-002; FLD-001; conflict C4.
User Journey | Provisioning.
Hazard Relationship | Slow rigor → skipped arming (ISO 14971) → void trip.
Traceability | Back: STN-008. Forward: UN-FLOW provisioning, UN-FUN arming.
Notes | Provisioning-time target Evidence Pending; sane defaults reduce steps.
UN-UX-006 — Get answers fast, without hunting through screens or data
Need Statement | Coordinators and receivers need the information they need immediately (status, verdict, what needs action) without digging through raw data, so time-pressured users are not slowed.
Primary Stakeholder | Distributors (SH-03), clinic staff (SH-05).
Secondary Stakeholders | Programs (SH-01).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/007; VOC-005/015.
User Journey | Monitoring; receipt.
Hazard Relationship | Buried information (ISO 14971) → delayed/wrong action.
Traceability | Back: STN-007. Forward: UN-ECO dashboard, UN-FUN verdict.
Notes | Layered information — verdict first, detail on demand (conflict C5).
6.4 Error prevention
UN-UX-007 — Be prevented from selecting the wrong trip profile
Need Statement | The dispatch user needs the interface to make selecting the wrong trip profile hard and to confirm the choice, so a mis-set profile cannot silently invalidate a trip.
Primary Stakeholder | Dispatch (SH-02/SH-03).
Secondary Stakeholders | QA/RA (SH-02).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-008; VOC-002; conflict C4.
User Journey | Configuration/arming.
Hazard Relationship | Wrong profile (IEC 62366; ISO 14971) → missed excursion / false alarm.
Traceability | Back: STN-008. Forward: UN-FUN configurable profiles, UN-SAFE.
Notes | Error-proofing (poka-yoke) at the interface level.
UN-UX-008 — Not confuse a "safe cold" reading with a damaging "frozen" one
Need Statement | Users need the interface to make the difference between acceptable cold and damaging freezing unmistakable, so a freeze excursion is never read as normal cold.
Primary Stakeholder | Clinic/last-mile staff (SH-05), programs (SH-01).
Secondary Stakeholders | Distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-002/004; VOC-011.
User Journey | Receipt verdict.
Hazard Relationship | Freeze read as OK (IEC 62366; ISO 14971) → spoiled freeze-sensitive dose given.
Traceability | Back: STN-002. Forward: UN-FUN two-sided detection, UN-VAL.
Notes | Directly serves the two-sided detection need (UN-FUN-004).
6.5 Feedback, guidance & help
UN-UX-009 — Get clear feedback that an action succeeded
Need Statement | Users need immediate, unambiguous feedback that an action (armed, placed, charging, updated) succeeded or failed, so they are never left unsure whether monitoring is active.
Primary Stakeholder | Dispatch (SH-02/SH-03), service (SH-06).
Secondary Stakeholders | Warehouse operators (SH-05).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-008/010; FLD-003.
User Journey | Arming; charging; update.
Hazard Relationship | No confirmation (ISO 14971) → assumed-armed-but-not → void trip.
Traceability | Back: STN-008. Forward: UN-FUN arming confirmation, UN-FLOW charge readiness.
Notes | Feedback must survive offline (on-device), not depend on app/cloud.
UN-UX-010 — Be guided to correct placement and use in context
Need Statement | Installers and operators need in-context guidance on correct placement and use, so correct behaviour is prompted at the moment of action rather than buried in a manual.
Primary Stakeholder | Installers (SH-12), warehouse operators (SH-05).
Secondary Stakeholders | Dispatch (SH-02/SH-03).
Importance | Desirable
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-024; FLD-002.
User Journey | Placement/install.
Hazard Relationship | No guidance → wrong placement (ISO 14971) → biased detection.
Traceability | Back: STN-024. Forward: UN-FUN placement guidance, UN-PER installer.
Notes | Complements UN-FUN-013; can be on-device, app, or labeling.
6.6 Consistency & trust
UN-UX-011 — Trust that the verdict and record agree and are consistent
Need Statement | Users across roles need the on-device verdict and the cloud/exported record to reflect the same underlying truth, so no one is undermined by a device that "says one thing and the report says another."
Primary Stakeholder | QA/RA (SH-02), programs (SH-01).
Secondary Stakeholders | Regulators (SH-07), distributors (SH-03).
Importance | Important
Market Priority | Must-have
Source | Stakeholder Interview / Contextual Inquiry
Voice of Customer Reference | STN-004/005; conflict C5.
User Journey | Receipt (edge) and audit (back office).
Hazard Relationship | Inconsistent verdict vs record (ISO 14971) → disputed/wrong disposition, lost trust.
Traceability | Back: STN-005. Forward: UN-FUN verdict + record, UN-VAL cross-persona.
Notes | "Same truth, role-appropriate presentation" (conflict C5 resolution direction).
UN-UX-012 — Have confidence the device is working without checking
Need Statement | Users need visible, passive reassurance that the device is armed and healthy, so they can trust that monitoring is active without having to interrogate it.
Primary Stakeholder | Distributors (SH-03), programs (SH-01).
Secondary Stakeholders | 3PLs (SH-04), warehouse operators (SH-05).
Importance | Desirable
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-008; VOC-002.
User Journey | Whole trip.
Hazard Relationship | False sense of monitoring (ISO 14971) → unmonitored trip believed monitored.
Traceability | Back: STN-008. Forward: UN-FUN status indication, UN-PHY indicator.
Notes | Balance reassurance against energy budget (must not waste the 90-day budget on indicators).
7. Validation Notes
UX needs are validated by IEC 62366 formative and summative usability studies across the persona range, with critical-task focus on the offline verdict (novice nurse), profile selection (dispatch), and cold/gloved operation (cold-store). Use-error rates and legibility thresholds are Evidence Pending (Stage 02 §6: ride-alongs, cold-store time-and-motion, receipt-workflow observation). Use-error findings feed the Usability Engineering File and hazard analysis (HA-###, ORCH-017).
8. Traceability Placeholders
Upstream | MR-1200 personas; STN-004/005/008/024; conflict C5; IEC 62366.
Downstream | UN-FUN, UN-PHY, UN-ENV, UN-SAFE, UN-VAL; forward to PRD (ORCH-011), Usability Engineering File, hazard analysis. Matrix: ORCH-018 / EM-ENG-012.
UN-800 — Performance Needs (PERF)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on the concept performance drivers (accuracy, 90-day endurance at −30 °C, alert latency), STN records, and D1–D4. Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the levels of performance users need — accuracy, endurance, latency, capacity, availability, reliability — in solution-neutral terms ("the user needs accuracy sufficient to X"). Specific numeric targets are set in Product Requirements; here they are expressed as the user-facing performance the target must satisfy, with quantitative values marked Evidence Pending where not yet evidence-backed.
2. Scope
In scope: measurement accuracy/resolution, 90-day cold endurance, alert latency, log capacity/buffering, connectivity performance, location performance within the energy budget, charge readiness, and reliability over reuse cycles. Out of scope: the functions themselves (UN-FUN), environmental survival (UN-ENV), and the tests that verify them (verification stage).
3. Applicable Standards & References
• WHO PQS E006 — temperature-monitoring device accuracy/performance expectations (exact class TBC — Evidence Pending).
• IEC 60068 (environmental performance), ISO/IEC 17025 concepts (measurement traceability).
• Concept §4 performance drivers; MR-RISK-001 (cold endurance). Numeric targets Evidence Pending; the D3 tension is closed by test (MR-VAL-006).
4. Background
Three performance needs dominate: accuracy good enough that an accept/reject verdict is trustworthy; endurance to survive 90 days at −30 °C without dying or losing data (the top program risk, D3); and alert latency low enough to enable rescue (bounded by connectivity, D2). Location and liveness are subordinate to the energy budget (conflicts C2/C3).
5. Need Categories
• 5.1 Measurement accuracy · 5.2 Endurance & energy · 5.3 Latency & responsiveness · 5.4 Capacity & buffering · 5.5 Connectivity & location · 5.6 Reliability & availability.
6. Detailed User Needs
6.1 Measurement accuracy
UN-PERF-001 — Measure temperature accurately enough to trust the verdict
Need Statement | Users need temperature measured accurately and repeatably enough across the operating range that an accept/reject decision based on it can be trusted and defended.
Primary Stakeholder | Manufacturers (SH-02), regulators (SH-07).
Secondary Stakeholders | Programs (SH-01), distributors (SH-03).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / JTBD
Voice of Customer Reference | STN-005/021; WHO PQS E006; MR-JTBD-001.
User Journey | Whole trip; verdict; audit.
Hazard Relationship | Inaccurate reading (ISO 14971) → false accept (spoiled given) or false reject (safe wasted) — severity anchor UN-VIS-001.
Traceability | Back: STN-005. Forward: PRD accuracy/resolution, UN-SVC calibration, verification.
Notes | Accuracy target per WHO PQS class (TBC — Evidence Pending); traceable calibration underpins it.
UN-PERF-002 — Hold accuracy across the whole range and over the calibration interval
Need Statement | Users need measurement accuracy to hold across −30 °C to +60 °C and to remain within spec over the defined calibration interval, so readings stay trustworthy at the extremes and over time.
Primary Stakeholder | Regulators/QA (SH-07), service technicians (SH-06).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-021; ISO/IEC 17025 concepts; concept §4.
User Journey | Whole life; recalibration cycle.
Hazard Relationship | Drift over interval (ISO 14971) → disputed/wrong verdict.
Traceability | Back: STN-021. Forward: PRD range accuracy, UN-SVC calibration interval, verification.
Notes | Calibration interval Evidence Pending (set in regulatory/service stage).
6.2 Endurance & energy
UN-PERF-003 — Survive a 90-day trip through deep cold without dying or losing data
Need Statement | Users need the device to keep operating and retain all data for the full 90-day trip across −30 °C to +60 °C, so a device never quits in the cold and loses the record exactly when it matters.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01), service (SH-06).
Importance | Essential
Market Priority | Must-have
Source | VoC / Market Analysis
Voice of Customer Reference | STN-003; VOC-010/024; CFB-001; D3; MR-RISK-001.
User Journey | Whole cold trip.
Hazard Relationship | Dead-in-cold (ISO 14971) → lost record → indefensible/unmonitored trip — the top program risk.
Traceability | Back: STN-003, D3. Forward: PRD endurance/energy budget, UN-ENV thermal, MR-VAL-006.
Notes | 90-day/−30 °C combination is the headline tension (conflict C1/C2/C3); closed by test.
UN-PERF-004 — Be ready to charge and reach trip-readiness after a cold return
Need Statement | Service users need a returned device to reach full trip-readiness after charging, including a clear state when a cold device must warm before it will safely charge, so turnaround is predictable and no under-charged unit is redeployed.
Primary Stakeholder | Service technicians (SH-06).
Secondary Stakeholders | Service ops (SH-13).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry
Voice of Customer Reference | STN-010; CFB-016; D3.
User Journey | Depot charging.
Hazard Relationship | Charging cold battery (ISO 14971, battery safety); under-charge → dead-in-cold.
Traceability | Back: STN-010, D3. Forward: PRD charge behaviour, UN-SAFE charge inhibit, UN-FLOW charge.
Notes | Charge-inhibit-below-0 °C (D3); self-heat-before-charge under evaluation.
6.3 Latency & responsiveness
UN-PERF-005 — Deliver an excursion alert fast enough to enable rescue (when connected)
Need Statement | Users need an excursion alert delivered quickly enough, when connectivity exists, that they can still act, so real-time monitoring translates into prevented loss.
Primary Stakeholder | Distributors (SH-03), 3PLs (SH-04).
Secondary Stakeholders | Programs (SH-01).
Importance | Essential
Market Priority | Must-have
Source | VoC / JTBD
Voice of Customer Reference | STN-001; VOC-008; MR-JTBD-005; D2.
User Journey | Transit.
Hazard Relationship | Excessive latency (ISO 14971) → rescuable loss becomes total.
Traceability | Back: STN-001. Forward: PRD alert-latency target, UN-FUN alerting, HA-###.
Notes | Latency connectivity-bound; target Evidence Pending; offline capture on reconnection (D2).
UN-PERF-006 — Respond promptly to on-device interaction
Need Statement | Users need the device to respond promptly when read or operated (verdict shown, arming confirmed) so a time-pressured user is not left waiting or unsure.
Primary Stakeholder | Clinic staff (SH-05), dispatch (SH-02/SH-03).
Secondary Stakeholders | Warehouse operators (SH-05).
Importance | Important
Market Priority | Differentiator
Source | Contextual Inquiry
Voice of Customer Reference | STN-004/008.
User Journey | Receipt; arming.
Hazard Relationship | Slow/uncertain response (ISO 14971) → mis-operation or abandonment.
Traceability | Back: STN-004. Forward: PRD responsiveness, UN-UX feedback.
Notes | Must remain responsive at low temperature and low battery.
6.4 Capacity & buffering
UN-PERF-007 — Store the full trip's data, even offline for the whole trip
Need Statement | Users need the device to retain the complete log for the full trip duration even with no connectivity, so an entirely offline trip still yields a complete record.
Primary Stakeholder | Programs (SH-01), distributors (SH-03).
Secondary Stakeholders | Regulators (SH-07), IT (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-009; VOC-023; FLD-004; D2.
User Journey | Offline transit; late upload.
Hazard Relationship | Buffer overflow/data loss (ISO 14971) → incomplete record → indefensible.
Traceability | Back: STN-009. Forward: PRD log capacity/buffer, UN-FUN store-and-forward.
Notes | Buffer duration target Evidence Pending; time order preserved on upload.
6.5 Connectivity & location
UN-PERF-008 — Establish connectivity opportunistically across markets
Need Statement | Users need the device to obtain connectivity wherever it reasonably can across global lanes, so buffered data and alerts get through without manual intervention.
Primary Stakeholder | IT/fleet (SH-09), distributors (SH-03).
Secondary Stakeholders | Programs (SH-01), 3PLs (SH-04).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / JTBD
Voice of Customer Reference | STN-009/014; D2/D4; concept §4.
User Journey | Transit; opportunistic offload.
Hazard Relationship | No egress (ISO 14971) → delayed alert; store-and-forward mitigates.
Traceability | Back: STN-009. Forward: PRD connectivity, UN-ENV RF-in-container, UN-FUN offload.
Notes | Multi-bearer opportunistic offload; RF-in-container performance Evidence Pending.
UN-PERF-009 — Provide useful location without breaking the energy budget
Need Statement | Fleet users need periodic location adequate for fleet management and stray recovery, delivered within the 90-day energy budget rather than continuous fixes that would exhaust it.
Primary Stakeholder | IT/fleet (SH-09), executive (SH-15).
Secondary Stakeholders | Service ops (SH-13).
Importance | Desirable
Market Priority | Nice-to-have
Source | Stakeholder Interview / Market Analysis
Voice of Customer Reference | STN-015; FLD-005/012; conflict C3; D3.
User Journey | Fleet management; recovery.
Hazard Relationship | None direct; energy trade-off vs endurance (conflict C3).
Traceability | Back: STN-015. Forward: PRD location cadence, UN-FUN locate/attribute, power budget.
Notes | Explicit tradeoff; may need external-antenna/event-triggered fixes; Evidence Pending.
6.6 Reliability & availability
UN-PERF-010 — Be reliable across many reuse cycles
Need Statement | Users need the device to remain reliable — sensing, logging, charging, connecting — across its full reusable life of many trips, so reuse economics and record credibility hold over time.
Primary Stakeholder | Service ops (SH-13), procurement (SH-08).
Secondary Stakeholders | Distributors (SH-03), finance (SH-10).
Importance | Important
Market Priority | Must-have
Source | Market Analysis / Stakeholder Interview
Voice of Customer Reference | STN-010/019; D1/D3; MR-EXEC-008.
User Journey | Whole reusable life.
Hazard Relationship | Capacity fade / component wear (ISO 14971) → mid-life dead-in-cold or drift.
Traceability | Back: STN-010/019. Forward: PRD reliability/cycle life, UN-QUAL, UN-BIZ TCO.
Notes | Capacity fade over reuse cycles is part of MR-RISK-001; cycle-life target Evidence Pending.
UN-PERF-011 — Be available and monitoring whenever a trip is underway
Need Statement | Users need confidence that an armed device is actually monitoring for the entire trip with no silent outages, so "armed" reliably means "monitoring."
Primary Stakeholder | Distributors (SH-03), programs (SH-01).
Secondary Stakeholders | Manufacturers (SH-02).
Importance | Important
Market Priority | Must-have
Source | JTBD / VoC
Voice of Customer Reference | STN-005/008; VOC-002.
User Journey | Whole trip.
Hazard Relationship | Silent outage (ISO 14971) → believed-monitored but unmonitored trip.
Traceability | Back: STN-005. Forward: PRD availability/self-check, UN-UX confidence, UN-SAFE.
Notes | Ties UN-UX-012 (passive reassurance) and self-check.
7. Validation Notes
Performance needs are verified by test in the verification stage: accuracy against a traceable reference across range; 90-day/−30 °C endurance and no-data-loss (MR-VAL-006); alert-latency vs connectivity; buffer-capacity full-trip offline; charge behaviour at low temperature; and reliability/cycle-life. All numeric targets are Evidence Pending until set in PRD and closed by test.
8. Traceability Placeholders
Upstream | concept §4; MR-RISK-001; STN-001/003/005/010/015/021; D2/D3/D4.
Downstream | UN-FUN, UN-ENV, UN-SAFE, UN-SVC, UN-QUAL; forward to PRD (ORCH-011), power budget, verification, MR-VAL-006. Matrix: ORCH-018 / EM-ENG-012.
UN-900 — Safety Needs (SAFE)
Master User Needs Library — SCM · Stage 04 (ORCH-010). Based on ISO 14971, the STN severity anchor (STN-020), the hazard-relevant STNs, and D1–D4. These needs cross-check against, and seed, the Preliminary Hazard Analysis (ORCH-017, HA-###). Forward to PRD (ORCH-011) and ORCH-018 / EM-ENG-012.
1. Purpose
This module states the safety-relevant user needs — what users need so that the SCM neither causes harm directly (e.g., battery hazard) nor contributes to harm indirectly by enabling a wrong disposition (a spoiled dose administered, or a safe dose wasted). Per CLAUDE.md rule 3, priority is severity-first, with the patient/beneficiary outcome (STN-020 / UN-VIS-001) as the tie-breaker.
2. Scope
In scope: the information-safety needs (correct disposition support, no missed excursion, no false reassurance) and the physical-safety needs of a reusable Li-ion, cellular device (battery, charge, transport, thermal). Out of scope: the risk controls themselves (derived in HA-###, ORCH-017) and security threats (UN-SEC) except where they are safety-relevant.
3. Applicable Standards & References
• ISO 14971 — application of risk management (severity-first Action Priority).
• IEC 62366-1 — use-related risk (misread verdict, mis-arm).
• IEC 62368-1 — electrical/thermal safety of the device.
• UN 38.3 / IATA DGR — Li-battery transport safety.
• WHO PQS E006 — disposition-decision performance. The SCM is an information device (not therapeutic); its dominant safety contribution is correct/incorrect disposition support.
4. Background
The SCM's most consequential safety role is informational: it supports the accept/reject decision that determines whether a potentially spoiled dose is given or a potent dose is discarded. Both error directions harm the beneficiary (STN-020). Physical hazards are those of any reusable Li-ion cellular device — charge safety at low temperature (D3), transport safety (D4), and thermal/electrical safety — which must be controlled so the device is never itself a source of injury.
5. Need Categories
• 5.1 Correct-disposition safety · 5.2 No missed excursion / no false reassurance · 5.3 Use-error safety · 5.4 Battery & charge safety · 5.5 Transport & thermal safety · 5.6 Safe recovery & fail-safe behaviour.
6. Detailed User Needs
6.1 Correct-disposition safety
UN-SAFE-001 — Never mislead a user into giving a compromised dose or wasting a safe one
Need Statement | Users need the device to never present information that leads to administering a temperature-compromised dose or to discarding a potent one, because either error harms the beneficiary.
Primary Stakeholder | Beneficiaries/patients (SH-11, represented), clinic staff (SH-05).
Secondary Stakeholders | Programs (SH-01), regulators (SH-07).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview (represented)
Voice of Customer Reference | STN-020; VOC-011/021; MR-PROB-010.
User Journey | Disposition at receipt / point of care.
Hazard Relationship | The severity anchor (ISO 14971): false accept → patient harm / ineffective immunization; false reject → stock-out → unvaccinated patient.
Traceability | Back: STN-020. Forward: UN-FUN verdict/detection, HA-### (missed excursion), UN-VAL.
Notes | The tie-breaker across all conflicts; seeds the top hazard in HA-### (ORCH-017).
6.2 No missed excursion / no false reassurance
UN-SAFE-002 — Catch a freezing excursion as reliably as an overheating one
Need Statement | Users need freezing excursions detected and flagged as reliably as heat excursions, so silent freeze damage to freeze-sensitive vaccines is not missed.
Primary Stakeholder | Programs (SH-01), manufacturers (SH-02).
Secondary Stakeholders | Beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | VoC / Regulatory
Voice of Customer Reference | STN-002; VOC-011; MR-PROB-002.
User Journey | Transit; receipt verdict.
Hazard Relationship | Missed freeze (ISO 14971) → spoiled freeze-sensitive dose administered.
Traceability | Back: STN-002. Forward: UN-FUN two-sided detection, HA-###, UN-VAL.
Notes | Two-sided detection is a safety requirement, not just a feature.
UN-SAFE-003 — Not create false reassurance when monitoring has failed or lapsed
Need Statement | Users need the device to make any failure, gap, or lapse in monitoring visible rather than appear healthy, so no one trusts an unmonitored or incomplete trip as if it were fully monitored.
Primary Stakeholder | Distributors (SH-03), QA/RA (SH-02).
Secondary Stakeholders | Programs (SH-01), regulators (SH-07).
Importance | Essential
Market Priority | Must-have
Source | JTBD / Regulatory
Voice of Customer Reference | STN-005/008; VOC-002; MR-JTBD-008.
User Journey | Whole trip; receipt; audit.
Hazard Relationship | False reassurance (ISO 14971) → wrong disposition on an incomplete record.
Traceability | Back: STN-005. Forward: UN-FUN gap-flagging, UN-UX confidence, HA-###.
Notes | Gaps flagged, never hidden (ties UN-FUN-014); self-check availability (UN-PERF-011).
UN-SAFE-004 — Not drown the real alert in false ones
Need Statement | Users need alerting tuned so genuine excursions are not lost among false alarms, so alarm fatigue does not cause a real, harmful excursion to be ignored.
Primary Stakeholder | Distributors (SH-03), clinic staff (SH-05).
Secondary Stakeholders | 3PLs (SH-04), programs (SH-01).
Importance | Important
Market Priority | Must-have
Source | VoC / Contextual Inquiry
Voice of Customer Reference | STN-007; VOC-005/022; MR-PROB-009; conflict C8.
User Journey | Monitoring; excursion response.
Hazard Relationship | Alarm fatigue (ISO 14971) → muted alerts → missed excursion.
Traceability | Back: STN-007. Forward: UN-FUN excursion logic, HA-###, alarm-management validation.
Notes | The C8 balance is safety-relevant in both directions.
6.3 Use-error safety
UN-SAFE-005 — Protect against foreseeable use errors by low-training users
Need Statement | Users need the device designed so that foreseeable use errors by minimally trained, gloved, offline users do not lead to a wrong disposition, because the edge user is the least-trained and most safety-critical operator.
Primary Stakeholder | Last-mile nurse (SH-05). (Vulnerable/low-training user.)
Secondary Stakeholders | Programs (SH-01), beneficiaries (SH-11).
Importance | Essential
Market Priority | Must-have
Source | Contextual Inquiry / Regulatory
Voice of Customer Reference | STN-004; MR-PER-007; IEC 62366.
User Journey | Receipt / point of care.
Hazard Relationship | Use-error (IEC 62366; ISO 14971) → wrong disposition — highest-consequence.
Traceability | Back: STN-004. Forward: UN-UX low-literacy/error-prevention, Usability Eng File, HA-###.
Notes | Drives the summative usability study critical tasks.
UN-SAFE-006 — Prevent a trip from starting or being accepted unmonitored
Need Statement | Users need it to be effectively impossible to dispatch a trip believing it is monitored when it is not (forgot-to-arm, wrong profile), so unmonitored shipments are not silently trusted.
Primary Stakeholder | Dispatch (SH-02/SH-03).
Secondary Stakeholders | Programs (SH-01), QA (SH-02).
Importance | Important
Market Priority | Must-have
Source | Contextual Inquiry / VoC
Voice of Customer Reference | STN-008; VOC-002; FLD-001.
User Journey | Provisioning.
Hazard Relationship | Void trip believed monitored (ISO 14971) → undetected excursion.
Traceability | Back: STN-008. Forward: UN-FUN arming confirmation, UN-UX error prevention, HA-###.
Notes | Ties UN-SAFE-003 (no false reassurance).
6.4 Battery & charge safety
UN-SAFE-007 — Never charge the battery in an unsafe (cold) condition
Need Statement | Users need the device to refuse to charge when it is too cold to do so safely and to indicate that clearly, so charging never damages the battery or creates a safety hazard.
Primary Stakeholder | Service technicians (SH-06).
Secondary Stakeholders | Service ops (SH-13), IT/fleet (SH-09).
Importance | Essential
Market Priority | Must-have
Source | Regulatory / Market Analysis
Voice of Customer Reference | STN-010; D3; IEC 62368-1.
User Journey | Depot charging after cold return.
Hazard Relationship | Charging below 0 °C (ISO 14971, battery safety) → lithium plating / thermal event.
Traceability | Back: D3. Forward: UN-PERF charge readiness, PRD charge-inhibit, HA-### (battery).
Notes | Charge-inhibit-below-0 °C (D3); self-heat-before-charge under evaluation.
UN-SAFE-008 — Not become a fire/thermal hazard in any use or fault condition
Need Statement | All users near the device need it to not present a fire, burn, or thermal-runaway hazard in normal use, foreseeable misuse, or single-fault conditions, so it is safe in aircraft, cold stores, clinics, and homes.
Primary Stakeholder | 3PLs (SH-04), clinic staff (SH-05).
Secondary Stakeholders | All handlers.
Importance | Essential
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | MR-1100; D3/D4; IEC 62368-1; UN 38.3.
User Journey | Whole life, including transport and storage.
Hazard Relationship | Thermal runaway / electrical fault (ISO 14971) → fire/injury.
Traceability | Back: MR-1100. Forward: UN-REG safety standards, PRD BMS/protection, HA-### (battery).
Notes | BMS/protection derived in design; flag battery + electrical engineers.
6.5 Transport & thermal safety
UN-SAFE-009 — Be safe and compliant to transport by air and ground
Need Statement | Logistics users need the device (and its battery) to be safe and compliant to fly and truck, so it never endangers a shipment or crew and never causes a transport hold.
Primary Stakeholder | 3PLs / air-cargo (SH-04).
Secondary Stakeholders | Distributors (SH-03), service (SH-13).
Importance | Important
Market Priority | Must-have
Source | Regulatory
Voice of Customer Reference | STN-012; D3/D4; UN 38.3; IATA DGR.
User Journey | Air/ground transport.
Hazard Relationship | Non-compliant Li-battery transport (ISO 14971) → transport-safety hazard.
Traceability | Back: STN-012. Forward: UN-REG transport, UN-SVC decommission, HA-###.
Notes | UN 38.3 test summary + IATA guidance provided (documentation Evidence Pending).
6.6 Safe recovery & fail-safe behaviour
UN-SAFE-010 — Fail safe and recover without data loss or unsafe state
Need Statement | Users need the device to fail safe — preserving logged data and returning to a known, safe state — after a power, connectivity, firmware, or reset event, so a fault never silently corrupts the record or leaves the device unsafe.
Primary Stakeholder | IT/fleet (SH-09), service (SH-06).
Secondary Stakeholders | QA/RA (SH-02), regulators (SH-07).
Importance | Important
Market Priority | Must-have
Source | Regulatory / Stakeholder Interview
Voice of Customer Reference | STN-025; VOC-019; D2 (OTA).
User Journey | Fault/recovery; firmware update.
Hazard Relationship | Unsafe/data-losing failure (ISO 14971) → corrupted record or bricked device.
Traceability | Back: STN-025. Forward: UN-SEC fail-safe OTA, UN-FUN gap-flagging, HA-###.
Notes | Signed, resumable, fail-safe OTA (UN-SEC); data preserved across resets.
7. Validation Notes
Safety needs are validated jointly with the Preliminary Hazard Analysis (HA-###, ORCH-017) and the Usability Engineering File: correct-disposition validation, two-sided detection accuracy, false-reassurance/gap handling, summative use-error study, and battery/charge/transport safety testing (IEC 62368-1, UN 38.3). Severity is fixed by outcome; probability/detection targets are Evidence Pending. Action Priority is severity-first (CLAUDE.md rule 3).
8. Traceability Placeholders
Upstream | STN-020 (anchor), STN-002/004/007/008/010/012/025; ISO 14971; IEC 62366/62368-1; D3/D4.
Downstream | UN-FUN, UN-UX, UN-PERF, UN-SEC, UN-REG; forward to HA-### (ORCH-017), PRD (ORCH-011), Risk-Control Requirements, Validation. Matrix: ORCH-018 / EM-ENG-012.

## 5 Product Requirements
Product Requirements
The single source of truth for what the product must do, with a plain description and how the parts fit together.
↩ Contents
In this section:
• Master Product Requirements Library — Smart Cold Chain Monitor (SCM)
• PRD-VIS (000) — Product Vision
• PRD-STK (100) — Stakeholders
• PRD-MOB (1000) — Mobile Application
• PRD-SEC (1100) — Cybersecurity
• PRD-REL (1200) — Reliability
• PRD-SAFE (1300) — Safety
• PRD-REG (1400) — Regulatory
• PRD-MFG (1500) — Manufacturing
• PRD-SVC (1600) — Serviceability
• PRD-PERF (1700) — Performance
• PRD-QUAL (1800) — Quality
• PRD-VER (1900) — Verification
• PRD-ENV (200) — Use Environment
• PRD-ARCH (300) — Product Architecture
• PRD-FUN (400) — Functional Requirements
• PRD-UX (500) — User Experience
• PRD-MECH (600) — Mechanical Requirements
• PRD-ELEC (700) — Electrical Requirements
• PRD-CONN (800) — Connectivity Requirements
• PRD-CLOUD (900) — Cloud Ecosystem
• Device / Design Architecture — Smart Cold Chain Monitor (SCM)
• Product Description — Smart Cold Chain Monitor (SCM)
• Requirement Rationale & Traceability — Smart Cold Chain Monitor (SCM)
• System Block Diagram — Smart Cold Chain Monitor (SCM)
Master Product Requirements Library — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 05 — Product Requirements (ORCH-011) · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Category owners: Principal Engineers per discipline; Quality/Regulatory (EM-ENG-006) owns PRD-REG · Status: Authored — instantiated, product-specific. Single source of truth for the SCM's detailed design inputs.
This library converts the Master User Needs Library (Stage 04) into concrete, testable product requirements — the authoritative system-level baseline from which the System/Subsystem Requirements (Stage 06), Software Requirements (ORCH-012), Design Outputs (ORCH-013), Hazard Analysis (ORCH-017), ERFKL (ORCH-016), Verification (ORCH-014), and Validation (ORCH-015) are derived. It follows 00-governance/product-requirements-standard.md and the 20-module map at 14-tools/product-requirements-copilot/product-requirements-library/references/module-map.md, and is based on 00-concept/concept-and-decisions.md (foundational decisions D1–D4).
Every requirement uses the semantic ID scheme PRD-<CODE>-NNN and carries the mandated 11-field block: ID · Title · Requirement Statement (single "shall", atomic/measurable/verifiable, no implementation detail) · Rationale · Priority (Must/Should/Could) · Source · Verification Method (Test/Analysis/Inspection/ Demonstration) · Parent Requirement (trace to UN-…/MR-…) · Traceability (forward to system req/design output/verification) · Risk Category (Safety/Security/Reliability/Performance/Compliance/Usability/None) · Notes.
---
Product in one line
The SCM is a reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shipping containers to continuously log temperature/humidity/shock, detect two-sided excursions, alert stakeholders in time to act, give an offline accept/reject verdict at the door, and provide an auditable, tamper-evident cold-chain record. It is a monitoring/information device — it does not control temperature and is not therapeutic. Operating range −30 °C to +60 °C; endurance target 90 days (see D1–D4).
Supporting documents: [product-description.md](product-description.md) (§5) · [system-block-diagram.md](system-block-diagram.md) (§5.1, Mermaid) · [design-architecture.md](design-architecture.md) (§7) · [requirement-rationale-traceability.md](requirement-rationale-traceability.md) (§8).
---
Module table
Module | Code | File | Reqs | Applicability
PRD-000 Product Vision | VIS | [PRD-000-product-vision.md](PRD-000-product-vision.md) | 7 | Applicable
PRD-100 Stakeholders | STK | [PRD-100-stakeholders.md](PRD-100-stakeholders.md) | 6 | Applicable
PRD-200 Use Environment | ENV | [PRD-200-use-environment.md](PRD-200-use-environment.md) | 12 | Applicable
PRD-300 Product Architecture | ARCH | [PRD-300-product-architecture.md](PRD-300-product-architecture.md) | 10 | Applicable
PRD-400 Functional | FUN | [PRD-400-functional.md](PRD-400-functional.md) | 22 | Applicable
PRD-500 User Experience | UX | [PRD-500-user-experience.md](PRD-500-user-experience.md) | 12 | Applicable
PRD-600 Mechanical | MECH | [PRD-600-mechanical.md](PRD-600-mechanical.md) | 12 | Applicable
PRD-700 Electrical | ELEC | [PRD-700-electrical.md](PRD-700-electrical.md) | 16 | Applicable
PRD-800 Connectivity | CONN | [PRD-800-connectivity.md](PRD-800-connectivity.md) | 12 | Applicable
PRD-900 Cloud Ecosystem | CLOUD | [PRD-900-cloud.md](PRD-900-cloud.md) | 12 | Applicable
PRD-1000 Mobile Application | MOB | [PRD-1000-mobile.md](PRD-1000-mobile.md) | 8 | Applicable
PRD-1100 Cybersecurity | SEC | [PRD-1100-cybersecurity.md](PRD-1100-cybersecurity.md) | 12 | Applicable
PRD-1200 Reliability | REL | [PRD-1200-reliability.md](PRD-1200-reliability.md) | 12 | Applicable
PRD-1300 Safety | SAFE | [PRD-1300-safety.md](PRD-1300-safety.md) | 12 | Applicable (preliminary — reconcile w/ HAL)
PRD-1400 Regulatory | REG | [PRD-1400-regulatory.md](PRD-1400-regulatory.md) | 12 | Applicable
PRD-1500 Manufacturing | MFG | [PRD-1500-manufacturing.md](PRD-1500-manufacturing.md) | 8 | Applicable
PRD-1600 Serviceability | SVC | [PRD-1600-serviceability.md](PRD-1600-serviceability.md) | 8 | Applicable
PRD-1700 Performance | PERF | [PRD-1700-performance.md](PRD-1700-performance.md) | 12 | Applicable
PRD-1800 Quality | QUAL | [PRD-1800-quality.md](PRD-1800-quality.md) | 7 | Applicable
PRD-1900 Verification | VER | [PRD-1900-verification.md](PRD-1900-verification.md) | 8 | Applicable
Total |  | 20 modules | 220 | 20 / 20 applicable
---
Applicability Register
Per the standard, scope is never silently dropped: every module is emitted; any exclusion or borderline topic is recorded with a rationale. For the SCM, all 20 modules are Applicable — it is a globally distributed, regulated, reusable, connected, safety-significant IoT device with device + cloud + mobile tiers, so every module carries real requirements.
Module (code) | Status | Rationale
All 20 (VIS … VER) | Applicable | Multi-market (WHO PQS/EU GDP/US FDA/air-transport), reusable, connected, audit-instrument device serving ~15 stakeholder classes across a multi-trip lifecycle; every module maps to real product requirements.
Borderline topics flagged inside modules (confirm, not excluded):
• PRD-SAFE — the entire module is PRELIMINARY; safety/risk-control requirements originate in the Hazard Analysis Library (HAL, Stage 10, ORCH-017). Each PRD-SAFE requirement is derived from Stage-04 safety needs and the concept, and will be traced to a HA-<CODE>-NNN hazard and RCR-### risk-control requirement, refined by the Preliminary/Detailed Hazard Analyses, and re-baselined here.
• PRD-REG — exact WHO PQS device category/code and US FDA classification/pathway are Evidence Pending (concept §7); confirmed in the Regulatory stage. Standards cited by designation; no clause numbers asserted as fact; revisions to confirm. ISO 13485 applicability depends on classification.
• PRD-PERF / PRD-CONN / PRD-ENV — the RF-in-container link budget/antenna strategy, GNSS-in-container performance, OLED cold-legibility thresholds, and several numeric targets (±0.5 °C, 5-min latency, 3-s responsiveness, ±5% RH, IP67, storage sizing) are Judgment/Evidence Pending and closed by V&V.
• PRD-ELEC / PRD-REL — the D3 −30 °C/90-day battery tension (endurance, low-temperature charge safety, capacity fade over reuse) is carried as an explicit Reliability/Safety requirement set (PRD-REL-001/002/003, PRD-ELEC-002, PRD-SAFE-007/012, PRD-PERF-005) with acceptance criteria and "validate by test" (MR-VAL-006). Final chemistry/datasheet Evidence Pending.
• PRD-VIS / PRD-QUAL / PRD-MFG — commercial/yield/cosmetic targets are Judgment/Evidence Pending; no financials or field rates asserted as fact.
---
Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, every requirement is based and labelled Standard, Cited, or Judgment; unresolved quantities are Evidence Pending, never invented. Standards (WHO PQS E006, EU GDP, 21 CFR Part 11 / ALCOA+, GAMP 5, ISO 14971, IEC 62366-1 / 62368-1 / 60068 / 60529, CISPR/IEC 61000, FCC Part 15 / CE RED, UN 38.3 / IATA DGR, ISO/IEC 17025, IEC 62443 concepts, RoHS/WEEE) are cited by designation with revision to confirm — no clause number is asserted as fact. Numeric targets marked Judgment are confirmed in Verification (ORCH-014) / Validation (ORCH-015). The D3 battery tension is the top program risk and is closed by test.
---
Priority & risk-category summary (orientation)
Priority | the safety-, function-, and compliance-critical requirements are Must; enabling and differentiator requirements are Should; upside/optional are Could.
Risk-first | Action Priority is severity-first (CLAUDE.md rule 3). Safety-risk-category requirements (PRD-SAFE-###, plus safety-tagged FUN/ELEC/REL/PERF/UX/CONN items) dominate trade-offs; the beneficiary-outcome anchor is PRD-VIS-002 / PRD-SAFE-001 (STN-020).
---
Traceability note (chain position)
This library is the authoritative detailed design-input baseline for the SCM. Backward, each requirement traces to a User Need (UN-<CODE>-NNN, Stage 04) and through it to STN/VOC/MR and D1–D4. Forward, each seeds one or more System Requirements (SYS-###, Stage 06), Software Requirements (SRS-<CODE>-NNN, ORCH-012), Design Outputs (DO-<CODE>-NNN, ORCH-013), and Verification/Validation (PV-…/VAL-…). Safety requirements (PRD-SAFE) seed and reconcile with the Hazard Analysis (HA-###/RCR-###, ORCH-017, Stage 10).
Flow: MR (01) → VOC (02) → STN (03) → UN (04) → PRD (05, this library) → System/Subsystem Req (06) → Architecture → Design Outputs (ORCH-013) → Hazard Analysis (ORCH-017) / ERFKL (ORCH-016) → Risk-Control Requirements → Verification (ORCH-014) / Validation (ORCH-015).
The end-to-end trace matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18); each requirement's Traceability field carries the forward placeholders (SYS-…, DO-…, PV-…, VAL-…, HA-…) the matrix will resolve. See [requirement-rationale-traceability.md](requirement-rationale-traceability.md).
Convening: ORCH-011 (Product Requirements). Integrator: EM-ENG-001. Category owners: Principal Engineers per discipline; EM-ENG-006 owns PRD-REG (research-driven). Feeds ORCH-012 (Software Requirements) and ORCH-018 / EM-ENG-012 (lifecycle traceability, Stage 18).
PRD-VIS (000) — Product Vision
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Chief Systems Engineer (EM-ENG-001). Requirements use PRD-VIS-NNN, single "shall", atomic/measurable/verifiable, with the 11-field block. Based on UN-000 (Product Vision), UN-VIS/BIZ, MR, and D1–D4. Forward to Stage 06 and ORCH-018 / EM-ENG-012.
1. Purpose
Fix the product-level vision, success metrics, lifecycle, and variant constraints that bound every other module — so downstream requirements optimize the right outcome (protecting dose potency and providing defensible evidence), not a local sub-optimum.
2. Scope
Product identity and intended use, top-level success metrics, lifecycle/reuse model (D1), end-of-life, and variant/SKU constraints. Excludes detailed functions (PRD-FUN) and commercial terms (business case).
3. Applicable Standards
WHO PQS E006 (device role — revision to confirm); ISO 14971 (severity-first intent); EU GDP / 21 CFR Part 11 (evidence intent). Cited by designation; revisions to confirm.
4. Background
The SCM's reason to exist is the beneficiary outcome (STN-020): no compromised dose administered, no safe dose wasted. It is an information device, reusable and multi-market (D1/D4), whose credibility rests on accuracy, auditability, and 90-day cold survival (D3).
5. Requirement Categories
Identity & intended use · Success metrics · Lifecycle & reuse · End-of-life · Variants & configurability.
6. Detailed Requirements
PRD-VIS-001 — Intended-use conformance
Requirement Statement | The SCM shall function as a reusable, in-container environmental monitoring and information device that records temperature, humidity, and shock and supports (never automates or replaces) the human product-disposition decision.
Rationale | Locks the product class so no downstream requirement makes it a therapeutic or a temperature-control device.
Priority | Must
Source | User Need / Business
Verification Method | Inspection
Parent Requirement | UN-VIS-001, UN-STK-014 (intended use); concept §2.
Traceability | → SYS intended-use spec; → PVL design review; → labeling (PRD-REG-010).
Risk Category | Safety
Notes | Evidence: Standard/Judgment. Severity anchor STN-020.
PRD-VIS-002 — Beneficiary-outcome primacy (severity-first)
Requirement Statement | The SCM shall be designed so that, wherever requirements conflict, the option that best prevents administering a temperature-compromised dose or wasting a potent dose is selected.
Rationale | Encodes CLAUDE.md rule 3 and STN-020 as a binding product-level tie-breaker.
Priority | Must
Source | Safety / User Need
Verification Method | Analysis
Parent Requirement | UN-VIS-001, UN-SAFE-001 (STN-020).
Traceability | → design-decision records; → HAL (Stage 10) severity basis; → PVaL.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10). Evidence: Standard (ISO 14971).
PRD-VIS-003 — Two-sided protection scope
Requirement Statement | The SCM shall detect and record both freezing and overheating excursions as first-class, equally reliable capabilities.
Rationale | Freeze damage is silent, common, and under-detected; heat-biased monitoring misses it.
Priority | Must
Source | User Need / Market
Verification Method | Test
Parent Requirement | UN-VIS-002, UN-FUN-004, UN-SAFE-002.
Traceability | → PRD-FUN-006/007; → HAL missed-freeze hazard; → PVL detection test.
Risk Category | Safety
Notes | Cited to WHO/UNICEF freezing reporting.
PRD-VIS-004 — Program success metric: no undetected material excursion
Requirement Statement | The SCM shall meet a validated detection performance target such that no material temperature excursion (per the active trip profile) goes undetected in the trip record.
Rationale | The product's headline promise; makes "protect the dose" measurable.
Priority | Must
Source | Business / User Need
Verification Method | Test
Parent Requirement | UN-VIS-002, UN-VAL-###; MR-1000 success metrics.
Traceability | → PRD-FUN-006/007, PRD-PERF-001; → PVaL detection validation.
Risk Category | Safety
Notes | Numeric miss-rate target Evidence Pending; set in validation. Evidence: Judgment.
PRD-VIS-005 — Reusable multi-trip lifecycle (D1)
Requirement Statement | The SCM shall support a reusable multi-trip lifecycle including return, cleaning, recharge, readiness triage, and re-provisioning without loss of calibration or record integrity.
Rationale | D1 direction; reuse economics and record credibility depend on a defined loop.
Priority | Must
Source | Business (D1)
Verification Method | Demonstration
Parent Requirement | UN-VIS-006, UN-SVC-001; D1.
Traceability | → PRD-SVC, PRD-REL cycle-life; → PVaL refurbishment scenario.
Risk Category | Reliability
Notes | Cycle-life target Evidence Pending (PRD-REL). Evidence: Standard (D1).
PRD-VIS-006 — Environmentally-responsible end-of-life
Requirement Statement | The SCM shall support a defined end-of-life path for the device and battery compliant with applicable RoHS/WEEE/battery regulations in its target markets.
Rationale | Reusable Li-ion device must be retired lawfully and responsibly (C10).
Priority | Should
Source | Regulatory / Sustainability
Verification Method | Inspection
Parent Requirement | UN-REG-008; MR-1700; conflict C10.
Traceability | → PRD-REG-008, PRD-SVC decommission; → EoL documentation.
Risk Category | Compliance
Notes | EoL path Evidence Pending. Evidence: Standard (designations, rev. to confirm).
PRD-VIS-007 — Configurable across products without hardware change
Requirement Statement | The SCM shall be configurable to different vaccine/product thermal profiles through trip-profile configuration alone, without hardware variants.
Rationale | One reusable fleet must serve many products/markets; reduces SKU proliferation.
Priority | Should
Source | Business / User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-011, UN-VIS-007.
Traceability | → PRD-FUN-012/013 profiles; → PRD-CLOUD config management.
Risk Category | Usability
Notes | Evidence: Judgment.
7. Verification Notes
Vision requirements are verified by design review (PRD-VIS-001/002/007), and by the safety/detection validations that close PRD-VIS-003/004 (with HAL, Stage 10) and the refurbishment scenario (PRD-VIS-005).
8. Traceability Placeholders
Upstream | UN-000/VIS/BIZ; MR-000/1000/1700; STN-020; D1/D4.
Downstream | all modules (VIS bounds them); HAL (Stage 10); ORCH-018 / EM-ENG-012.
PRD-STK (100) — Stakeholders
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001. PRD-STK-NNN, single "shall", 11-field block. Based on UN-100 (Stakeholders), UN-200 (Personas), and D1–D4. Forward to Stage 06 and ORCH-018 / EM-ENG-012.
1. Purpose
State the product-level requirements that exist because a specific stakeholder class must be served — the requirements that would be missed if the product were designed only for the "primary user."
2. Scope
The stakeholder-class obligations of the product (minimally-trained edge user, dispatch, service/fleet, QA/regulatory, IT/security, OEM partner, 3PL/air-cargo). Excludes detailed workflow (PRD-UX) and functions.
3. Applicable Standards
IEC 62366-1 (use by defined user profiles); 21 CFR Part 11 (QA/IT roles); WHO PQS E006 (program user). Cited by designation; revisions to confirm.
4. Background
The SCM serves ~15 stakeholder classes across a multi-trip lifecycle; the least-trained (last-mile nurse) is the most safety-critical operator, and enterprise deployment is gated by IT/QA (STN-013).
5. Requirement Categories
Edge/low-training user · Dispatch/operations · Service/fleet · QA/regulatory · IT/security · OEM/channel.
6. Detailed Requirements
PRD-STK-001 — Serve the minimally-trained edge user without training or a phone
Requirement Statement | The SCM shall enable a minimally-trained user to obtain a correct accept/reject/quarantine verdict using only the device, with no phone, network, login, or manual.
Rationale | The last-mile nurse is the least-trained, most safety-critical operator.
Priority | Must
Source | User Need / Safety
Verification Method | Test (summative usability)
Parent Requirement | UN-STK-005, UN-FUN-008, UN-SAFE-005.
Traceability | → PRD-UX-001/002; → PRD-FUN-009; → PVaL usability study.
Risk Category | Usability
Notes | Drives summative critical tasks. Evidence: Standard (IEC 62366-1, rev. to confirm).
PRD-STK-002 — Serve dispatch: unmissable arming
Requirement Statement | The SCM shall provide the dispatch user positive confirmation that the device is armed, configured with the correct profile, and placed before a trip begins.
Rationale | Forgot-to-arm / wrong-profile voids a trip believed monitored.
Priority | Must
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-STK-002, UN-FUN-013/014, UN-SAFE-006.
Traceability | → PRD-FUN-014/015; → PRD-UX-004; → HAL void-trip hazard.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-STK-003 — Serve service/fleet: machine-readable readiness
Requirement Statement | The SCM shall report its charge state, calibration status, and firmware version in a machine-readable form to the fleet/service tooling.
Rationale | Enables triage and blocks out-of-cal/under-charged units from redeployment without hand-checking.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-STK-006, UN-FUN-017.
Traceability | → PRD-FUN-019, PRD-SVC-004, PRD-CLOUD-006; → PVL readiness test.
Risk Category | Reliability
Notes | Ties PRD-ELEC fuel gauge and PRD-SVC calibration.
PRD-STK-004 — Serve QA/regulatory: defensible records
Requirement Statement | The SCM shall provide QA/regulatory stakeholders trip records that are complete, attributable, and tamper-evident to support release and audit.
Rationale | QA/RA and auditors decide disposition and defend it; the record is their instrument.
Priority | Must
Source | Regulatory / User Need
Verification Method | Analysis
Parent Requirement | UN-STK-007, UN-FUN-016, UN-REG-003.
Traceability | → PRD-FUN-016/017, PRD-SEC-001/002; → PVaL audit simulation.
Risk Category | Compliance
Notes | ALCOA+ / Part 11 (rev. to confirm).
PRD-STK-005 — Serve IT/security: enterprise governance
Requirement Statement | The SCM cloud platform shall provide IT/security stakeholders role-based access control and enterprise single-sign-on for all configuration, data-access, and administrative functions.
Rationale | Enterprise deployment is gated by IT security and CSV (STN-013).
Priority | Must
Source | User Need / Security
Verification Method | Test
Parent Requirement | UN-STK-009, UN-SEC-003.
Traceability | → PRD-SEC-005/006, PRD-CLOUD-009; → PVL access-control test.
Risk Category | Security
Notes | Must not gate the login-free edge read (trust-boundary split, AD-5).
PRD-STK-006 — Serve OEM/channel partners: embeddable module
Requirement Statement | The SCM shall be available as a standard embeddable physical/logical module that OEM partners can integrate into their shippers/containers.
Rationale | Designed-in distribution channel (STN-017/024).
Priority | Should
Source | Business / User Need
Verification Method | Inspection
Parent Requirement | UN-STK-012, UN-PHY-003.
Traceability | → PRD-MECH embeddable form, PRD-CLOUD partner API.
Risk Category | None
Notes | Evidence: Judgment.
7. Verification Notes
Stakeholder requirements are verified through the module they drive (usability study for STK-001, access-control test for STK-005, audit simulation for STK-004) and by design review of stakeholder coverage.
8. Traceability Placeholders
Upstream | UN-100/200; MR-1200 personas; STN register; D1/D4.
Downstream | PRD-UX, PRD-FUN, PRD-SVC, PRD-SEC, PRD-CLOUD, PRD-MECH; HAL; ORCH-018 / EM-ENG-012.
PRD-MOB (1000) — Mobile Application
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Mobile/App engineer with Human-Factors and EM-ENG-001. PRD-MOB-NNN, single "shall", 11-field block. Based on UN-500 (workflow), UN-700 (UX), UN-1000 (security), and D2. Forward to Stage 06, app design outputs, and ORCH-018.
1. Purpose
Specify the mobile/BLE application requirements — provisioning, configuration, local read, notifications — without making the app a dependency for safety-critical edge use.
2. Scope
BLE pairing, trip-profile configuration, arming aid, local device read, notifications, and sync. Excludes device functions (PRD-FUN), the login-free on-device verdict (PRD-UX/FUN), and cloud functions (PRD-CLOUD).
3. Applicable Standards
IEC 62366-1 (usability); platform accessibility guidelines (iOS/Android); TLS 1.2+/1.3 (transport, PRD-SEC). Cited by designation; revisions to confirm.
4. Background
The app is a provisioning/configuration and notification aid for dispatch, service, and fleet users. The frictionless offline edge read must remain on-device only (no app dependency) — the app augments, never gates, the safety-critical verdict (AD-5/C9).
5. Requirement Categories
Pairing · Configuration · Arming aid · Local read · Notifications · Sync/offline · Permissions/accessibility.
6. Detailed Requirements
PRD-MOB-001 — Secure BLE pairing
Requirement Statement | The mobile app shall pair with a device over authenticated BLE such that only authorized users can configure or read controlled data.
Rationale | Provisioning must be authorized without breaking edge simplicity (C9).
Priority | Must
Source | Security / User Need
Verification Method | Test
Parent Requirement | UN-SEC-003, PRD-CONN-003/006.
Traceability | → PRD-SEC-005/006; → PVL pairing test.
Risk Category | Security
Notes | Edge read remains login-free (AD-5).
PRD-MOB-002 — Trip-profile configuration
Requirement Statement | The mobile app shall let an authorized user select/define and apply a trip profile to a device, presenting a valid default and preventing invalid configurations.
Rationale | Configuration must match the product and be error-proof (C4).
Priority | Must
Source | User Need / Safety
Verification Method | Test
Parent Requirement | UN-FUN-011, UN-SAFE-006; STN-007.
Traceability | → PRD-FUN-014, PRD-UX-007; → PVaL config-error scenario.
Risk Category | Safety
Notes | Poka-yoke selection; to be reconciled with HAL.
PRD-MOB-003 — Arming aid with confirmation
Requirement Statement | The mobile app shall assist arming and display confirmation of armed + profile-valid + placement state, without being required for on-device arming fallback.
Rationale | The app speeds arming; the device must still arm without it (offline resilience).
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-012, UN-SAFE-006.
Traceability | → PRD-FUN-015, PRD-UX-004; → PVL arming test.
Risk Category | Safety
Notes | On-device fallback mandatory (PRD-FUN-015).
PRD-MOB-004 — Local device read
Requirement Statement | The mobile app shall read a device's current status, recent history, and diagnostics over BLE for field triage.
Rationale | Field/service users need detail beyond the on-device summary.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-017/018; STN-010.
Traceability | → PRD-CONN-003, PRD-FUN-020; → PVL read test.
Risk Category | None
Notes | Complements on-device verdict; not a substitute.
PRD-MOB-005 — Alert notifications
Requirement Statement | The mobile app shall present excursion and device-fault notifications routed to the user via the cloud, with acknowledgement.
Rationale | Mobile is a primary responder channel for excursion alerts.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-007; STN-018.
Traceability | → PRD-CLOUD-003; → PVL notification test.
Risk Category | Safety
Notes | Channel among several (PRD-CLOUD-003).
PRD-MOB-006 — Offline app operation and sync
Requirement Statement | The mobile app shall function for local (BLE) provisioning/read when the phone has no network and shall sync with the cloud when connectivity returns.
Rationale | Field users often lack phone connectivity at the point of work.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-ENV-007; D2.
Traceability | → PRD-CONN-003, PRD-CLOUD-010; → PVL offline-sync test.
Risk Category | Reliability
Notes | Sync conflicts resolved with audit.
PRD-MOB-007 — Role-based permissions
Requirement Statement | The mobile app shall enforce the user's role-based permissions (via SSO/RBAC) for configuration and data actions.
Rationale | App actions must obey enterprise governance (Part 11 access).
Priority | Must
Source | Security
Verification Method | Test
Parent Requirement | UN-SEC-003, PRD-CLOUD-009.
Traceability | → PRD-SEC-005; → PVL app-permission test.
Risk Category | Security
Notes | Consistent with cloud RBAC.
PRD-MOB-008 — Accessibility and localization
Requirement Statement | The mobile app shall meet the target-platform accessibility guidelines and support localization for the target markets.
Rationale | Diverse users across markets (D4) and inclusive design.
Priority | Could
Source | User Need
Verification Method | Inspection
Parent Requirement | UN-UX-###; D4.
Traceability | → PRD-UX-010; → accessibility review.
Risk Category | Usability
Notes | Target language set Evidence Pending.
7. Verification Notes
Mobile requirements are verified by pairing/security tests (with PRD-SEC), configuration/arming demonstrations (with PRD-FUN/UX), notification and offline-sync tests (with PRD-CLOUD), and accessibility/localization review.
8. Traceability Placeholders
Upstream | UN-500/700/1000; D2/D4.
Downstream | PRD-FUN, PRD-CONN, PRD-CLOUD, PRD-SEC, PRD-UX; app design outputs; HAL; ORCH-018.
PRD-SEC (1100) — Cybersecurity
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Cybersecurity engineer with EM-ENG-001. PRD-SEC-NNN, single "shall", 11-field block. Based on UN-1000 (Security & Privacy), 21 CFR Part 11 / ALCOA+, GAMP 5, IEC 62443 concepts, and D2. Forward to Stage 06, cybersecurity design outputs, and ORCH-018. A formal threat model is Evidence Pending and set with the cybersecurity persona.
1. Purpose
Specify the security requirements that keep the SCM's data trustworthy and its connectivity/updates controlled — integrity, identity, access, secure transport, trustworthy update, and auditability.
2. Scope
Data integrity/tamper-evidence, device identity/secure boot, access control, secure connectivity, signed fail-safe update, key management, data retention/privacy, and audit. Excludes safety framing (PRD-SAFE-010) and cloud endpoint functions (PRD-CLOUD).
3. Applicable Standards
21 CFR Part 11 / ALCOA+ (records/signatures/audit); GAMP 5 (CSV); IEC 62443 concepts (device/system security); TLS 1.2+/1.3; NIST-style crypto good practice. Cited by designation; revisions to confirm.
4. Background
A cellular, OTA-updatable device on shipments is both an attack surface and a data-integrity dependency for disposition decisions. Security must be strong without defeating the frictionless offline edge read (C9); enterprise deployment is gated by IT security and CSV (STN-013). Threat model Evidence Pending.
5. Requirement Categories
Data integrity · Identity & secure boot · Access control · Secure connectivity · Trustworthy update · Key management · Retention/privacy · Audit.
6. Detailed Requirements
PRD-SEC-001 — Tamper-evident records
Requirement Statement | The SCM shall protect logged records end-to-end (device, transport, cloud) such that any alteration of a stored record is detectable.
Rationale | A disposition/audit rests on the record; silent edits must be impossible to hide.
Priority | Must
Source | Regulatory / Security
Verification Method | Test
Parent Requirement | UN-SEC-001, UN-FUN-014; 21 CFR Part 11 / ALCOA+.
Traceability | → PRD-FUN-017, PRD-CLOUD-002; → PVaL tamper-evidence test.
Risk Category | Security
Notes | Cryptographic integrity (e.g., signing/hash chain); mechanism in design.
PRD-SEC-002 — Attribution of records and actions
Requirement Statement | The SCM shall bind each record and configuration action to an authenticated source (device identity, user, action) with a trusted timestamp.
Rationale | ALCOA+ attributable/original; evidence must be assignable to a source.
Priority | Must
Source | Regulatory / Security
Verification Method | Test
Parent Requirement | UN-SEC-002, UN-FUN-016; ALCOA+.
Traceability | → PRD-ARCH-007 time base, PRD-CLOUD-001 identity; → PVaL attribution test.
Risk Category | Compliance
Notes | Ties device identity (PRD-CLOUD-001) and durable ID (PRD-MECH-010).
PRD-SEC-003 — Device identity and secure boot
Requirement Statement | The SCM device shall hold a unique protected identity/credential and shall verify the integrity/authenticity of its firmware at boot, refusing to run unverified firmware.
Rationale | Prevents impersonation and running of tampered firmware.
Priority | Must
Source | Security
Verification Method | Test
Parent Requirement | UN-SEC-005; IEC 62443 concepts.
Traceability | → PRD-SEC-008 signed OTA, PRD-ARCH-006; → PVL secure-boot test.
Risk Category | Security
Notes | Secure element / protected key store; threat model Evidence Pending.
PRD-SEC-004 — Audit trail (device + cloud)
Requirement Statement | The SCM shall maintain a secure, non-editable, attributable audit trail of configuration changes, access, and data actions across device and cloud, retained per policy.
Rationale | The system's governance must itself be inspectable (Part 11).
Priority | Must
Source | Regulatory / Security
Verification Method | Test
Parent Requirement | UN-SEC-009; 21 CFR Part 11.
Traceability | → PRD-CLOUD-007; → PVaL audit-trail review.
Risk Category | Compliance
Notes | Standard (21 CFR Part 11, rev. to confirm).
PRD-SEC-005 — Role-based access control
Requirement Statement | The SCM platform and app shall enforce least-privilege role-based access control for all configuration, data-access, and administrative functions.
Rationale | Only authorized roles may configure devices or read/modify data.
Priority | Must
Source | Security / Regulatory
Verification Method | Test
Parent Requirement | UN-SEC-003, PRD-STK-005; 21 CFR Part 11.
Traceability | → PRD-CLOUD-009, PRD-MOB-007; → PVL access-control test.
Risk Category | Security
Notes | Must not gate login-free edge read (AD-5).
PRD-SEC-006 — Secure connectivity (transit)
Requirement Statement | The SCM shall protect all device-to-cloud and app-to-cloud communication with authenticated, encrypted transport (TLS 1.2 or later / equivalent) resistant to interception and spoofing.
Rationale | Records/alerts must not be intercepted or spoofed in transit.
Priority | Must
Source | Security
Verification Method | Test
Parent Requirement | UN-SEC-006; TLS.
Traceability | → PRD-CONN-006, PRD-CLOUD-001; → PVL transport-security test.
Risk Category | Security
Notes | Mutual authentication where feasible; cipher suite in design.
PRD-SEC-007 — Data protection at rest
Requirement Statement | The SCM shall protect controlled data at rest on the device and in the cloud against unauthorized access and exposure.
Rationale | Stored records must not be exposed if a device or store is compromised.
Priority | Must
Source | Security
Verification Method | Test
Parent Requirement | UN-SEC-006/007.
Traceability | → PRD-CLOUD-011 retention; → PVL at-rest-encryption test.
Risk Category | Security
Notes | Encryption/access controls; key management (PRD-SEC-009).
PRD-SEC-008 — Signed, resumable, fail-safe firmware update
Requirement Statement | The SCM shall apply firmware updates only after cryptographic integrity/authenticity verification, shall resume interrupted updates, and shall never enter an unrecoverable or insecure state due to a failed update.
Rationale | An update must not be malicious, corrupt a device, or brick the fleet.
Priority | Must
Source | Security / Safety
Verification Method | Test
Parent Requirement | UN-SEC-005, UN-SAFE-010; D2.
Traceability | → PRD-FUN-022, PRD-ARCH-008, PRD-CLOUD-008; → HAL; → PVL OTA-security test.
Risk Category | Security
Notes | Ties safety fail-safe (PRD-SAFE-010). To be reconciled with HAL.
PRD-SEC-009 — Key and credential management
Requirement Statement | The SCM shall generate, store, rotate, and revoke cryptographic keys and credentials through a defined, protected lifecycle across device and platform.
Rationale | Integrity/identity/transport security all depend on sound key management.
Priority | Must
Source | Security
Verification Method | Analysis + Test
Parent Requirement | UN-SEC-005/006; IEC 62443 concepts.
Traceability | → PRD-SEC-003/006/008; → PVL key-lifecycle review.
Risk Category | Security
Notes | Protected key store; process Evidence Pending (design/CSV).
PRD-SEC-010 — Privacy of location/personal data
Requirement Statement | The SCM shall handle any location or personal data proportionately, limiting collection/retention to operational purpose and applicable privacy law.
Rationale | Asset/location tracking must not create disproportionate privacy exposure.
Priority | Should
Source | Regulatory / Security
Verification Method | Inspection
Parent Requirement | UN-SEC-008, UN-REG-009.
Traceability | → PRD-CLOUD-011, PRD-REG-009; → privacy review.
Risk Category | Compliance
Notes | Device monitors cargo, not people; regional law mapping Evidence Pending.
PRD-SEC-011 — Security monitoring and incident response
Requirement Statement | The SCM platform shall detect, log, and alert on security-relevant events and support a defined incident-response process.
Rationale | A connected fleet needs detection/response, not just prevention.
Priority | Should
Source | Security
Verification Method | Demonstration
Parent Requirement | UN-SEC-009; IEC 62443 concepts.
Traceability | → PRD-CLOUD-007/012; → security-operations review.
Risk Category | Security
Notes | IR runbook Evidence Pending.
PRD-SEC-012 — Security-by-design threat modeling
Requirement Statement | The SCM design shall be developed against a documented threat model whose identified risks are traced to security requirements and controls.
Rationale | Requirements must be driven by an explicit, reviewed threat model.
Priority | Must
Source | Security
Verification Method | Analysis
Parent Requirement | UN-SEC (all); IEC 62443 concepts.
Traceability | → all PRD-SEC; → HAL security-relevant hazards; → security assessment.
Risk Category | Security
Notes | Threat model Evidence Pending; to be reconciled with HAL for safety-relevant security.
7. Verification Notes
Security requirements are verified by security assessment against the threat model, tamper-evidence and attribution tests, secure-boot/OTA-security tests, access-control and transport/at-rest encryption tests, key-lifecycle review, and CSV/validation (GAMP 5). Safety-relevant security (SEC-008) reconciles with HAL.
8. Traceability Placeholders
Upstream | UN-1000; 21 CFR Part 11 / ALCOA+ / GAMP 5 / IEC 62443 concepts; D2; conflict C9.
Downstream | PRD-FUN, PRD-CONN, PRD-CLOUD, PRD-MOB, PRD-SAFE, PRD-REG; cybersecurity design outputs; HAL; ORCH-018 / EM-ENG-012.
PRD-REL (1200) — Reliability
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Reliability/Materials engineer with EM-ENG-001. PRD-REL-NNN, single "shall", 11-field block. Based on UN-800 (endurance/availability/ reliability), UN-900 (fail-safe), and D1/D3. Forward to Stage 06, reliability program, DFMEA (ERFKL), and ORCH-018. Carries the D3 −30 °C/90-day tension with explicit acceptance criteria and "validate by test".
1. Purpose
Specify the reliability, endurance, availability, storage-integrity, and graceful-degradation requirements — including the top program risk (survive 90 days at −30 °C without dying or losing data).
2. Scope
Endurance, cycle life, availability/self-check, watchdog/recovery, storage integrity, graceful degradation, robustness over reuse, and platform reliability. Excludes battery electrical design (PRD-ELEC), environmental survival (PRD-ENV), and safety controls (PRD-SAFE).
3. Applicable Standards
IEC 60068 (environmental reliability); JEDEC/handbook methods (reliability prediction, method to confirm); IEC 62368-1 (single-fault); ISO 14971 (fail-safe). Cited by designation; revisions/method to confirm.
4. Background
The 90-day/−30 °C combination is in genuine tension (D3): endurance at deep cold, charge safety, and capacity fade over reuse are the top program risks (MR-RISK-001). This module states them as measurable requirements closed by test, not asserted as met.
5. Requirement Categories
Endurance (D3) · Cycle life & fade · Availability & self-check · Recovery & watchdog · Storage integrity · Graceful degradation · Platform reliability.
6. Detailed Requirements
PRD-REL-001 — 90-day cold endurance (D3, validate by test)
Requirement Statement | The SCM shall operate continuously — sensing, logging, on-device verdict, and scheduled connectivity attempts — for at least 90 days from full charge under the defined worst-case cold duty profile spanning −30 °C, with zero loss of logged data.
Rationale | The headline endurance promise and the top program risk (D3, MR-RISK-001); a dead-in-cold device loses the record exactly when it matters.
Priority | Must
Source | Reliability (D3)
Verification Method | Test (90-day cold-endurance, MR-VAL-006)
Parent Requirement | UN-PERF-003, UN-ENV-001; D3.
Traceability | → PRD-ELEC-002/008, PRD-ARCH-004; → PVL endurance test (MR-VAL-006).
Risk Category | Reliability
Notes | Acceptance: ≥90 days at worst-case cold profile + zero data loss + verdict/record intact. Validated by test. Cross-flag PRD-SAFE (dead-in-cold as safety-relevant informational failure).
PRD-REL-002 — Cycle life over reuse (D1/D3, validate by test)
Requirement Statement | The SCM shall meet PRD-REL-001 over its specified number of reuse/charge cycles, or shall reliably flag when battery health can no longer meet it and block redeployment.
Rationale | Capacity fade over reuse can silently cause mid-life dead-in-cold (part of MR-RISK-001).
Priority | Must
Source | Reliability (D1/D3)
Verification Method | Test (cycle-life)
Parent Requirement | UN-PERF-010; D1/D3.
Traceability | → PRD-ELEC-003, PRD-FUN-020 readiness; → PVL cycle-life test.
Risk Category | Reliability
Notes | Acceptance: end-of-cycle-life unit still meets PRD-REL-001 or self-flags. Cycle count Evidence Pending. Validated by test.
PRD-REL-003 — Low-temperature charge safety reliability (D3, validate by test)
Requirement Statement | The SCM charge-inhibit and (if provided) self-heat functions shall operate reliably across the operating temperature range and over reuse life so charging never occurs below the safe cell limit.
Rationale | The charge-safety leg of the D3 tension must be reliable, not just present.
Priority | Must
Source | Safety / Reliability (D3)
Verification Method | Test
Parent Requirement | UN-SAFE-007, UN-PERF-004; D3.
Traceability | → PRD-ELEC-005/006, PRD-SAFE-007; → HAL battery hazard; → PVL charge-inhibit reliability.
Risk Category | Safety
Notes | Acceptance: no charge below limit across life and temperature. To be reconciled with HAL.
PRD-REL-004 — Fail-safe on fault
Requirement Statement | The SCM shall, on any power, connectivity, firmware, or reset fault, enter a known safe state that preserves logged data and honest monitoring status and recovers to normal operation.
Rationale | A fault must never corrupt the record or leave the device unsafe or falsely reassuring.
Priority | Must
Source | Reliability / Safety
Verification Method | Test (fault injection)
Parent Requirement | UN-SAFE-010; UN-FUN-014.
Traceability | → PRD-ARCH-008, PRD-ELEC-011, PRD-SAFE-010; → HAL; → PVL fault-injection.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-REL-005 — Storage integrity (power-fail-safe writes)
Requirement Statement | The SCM non-volatile log shall preserve committed records across power loss, reset, brown-out, and firmware update, with no corruption of committed data.
Rationale | The record's completeness/integrity is the product's core value.
Priority | Must
Source | Reliability
Verification Method | Test
Parent Requirement | UN-FUN-009/014; UN-SAFE-010.
Traceability | → PRD-ELEC-011 brown-out, PRD-FUN-017; → PVL power-cycle-integrity test.
Risk Category | Reliability
Notes | Journaled/atomic writes; capacity per PRD-PERF-006.
PRD-REL-006 — Watchdog and autonomous recovery
Requirement Statement | The SCM firmware shall be supervised by a watchdog that recovers a hung device to a monitoring state without user intervention and logs the recovery event.
Rationale | A silent hang would create an unmonitored trip believed monitored.
Priority | Must
Source | Reliability / Safety
Verification Method | Test
Parent Requirement | UN-PERF-011, UN-SAFE-003.
Traceability | → PRD-ARCH-008, PRD-SAFE-003; → PVL watchdog test.
Risk Category | Reliability
Notes | Recovery event flagged (no false reassurance).
PRD-REL-007 — Availability / silent-outage prevention
Requirement Statement | The SCM shall self-check its sensing/logging chain periodically and flag any outage such that an armed device is either monitoring or visibly flagged as not.
Rationale | "Armed" must reliably mean "monitoring"; silent outages are a false-reassurance hazard.
Priority | Must
Source | Reliability / Safety
Verification Method | Test
Parent Requirement | UN-PERF-011, UN-SAFE-003.
Traceability | → PRD-SAFE-003, PRD-UX-012; → PVL self-check test.
Risk Category | Safety
Notes | To be reconciled with HAL.
PRD-REL-008 — Graceful degradation under low energy
Requirement Statement | The SCM shall degrade gracefully as energy depletes, shedding discretionary functions (uploads, location) before core logging so the record is preserved to the last.
Rationale | A depleting device must protect the record over convenience.
Priority | Must
Source | Reliability
Verification Method | Test
Parent Requirement | UN-PERF-003; UN-SAFE-010.
Traceability | → PRD-ELEC-010; → PVL depletion test.
Risk Category | Reliability
Notes | Degradation order Evidence Pending.
PRD-REL-009 — Robustness over reuse (mechanical/sensor)
Requirement Statement | The SCM sensing accuracy and mechanical integrity (seals, probe, controls) shall remain within spec over the reuse life, cleaning cycles, and transport stresses.
Rationale | Reuse economics and record credibility require durability, not just first-trip performance.
Priority | Must
Source | Reliability
Verification Method | Test (life/aging)
Parent Requirement | UN-PERF-010, UN-PHY-006/008.
Traceability | → PRD-MECH-005..009, PRD-PERF-002; → PVL life/aging test.
Risk Category | Reliability
Notes | Ties DFMEA (ERFKL) wear mechanisms.
PRD-REL-010 — Reliability program and prediction
Requirement Statement | The SCM shall have a documented reliability program (DFMEA, reliability prediction/allocation, life testing) demonstrating the endurance/cycle-life/availability targets.
Rationale | Reliability must be engineered and demonstrated, not assumed.
Priority | Should
Source | Reliability / Quality
Verification Method | Analysis
Parent Requirement | UN-QUAL-###, UN-PERF-010.
Traceability | → ERFKL DFMEA (Stage 11), PRD-QUAL; → reliability report.
Risk Category | Reliability
Notes | Prediction method Evidence Pending.
PRD-REL-011 — Communication robustness
Requirement Statement | The SCM shall tolerate lossy/intermittent links, retrying and buffering such that transient network conditions never cause data loss or duplicate records.
Rationale | Global lanes have unreliable connectivity; robustness protects the record.
Priority | Must
Source | Reliability
Verification Method | Test
Parent Requirement | UN-ENV-007, UN-FUN-009; D2.
Traceability | → PRD-CONN-008, PRD-CLOUD-002; → PVL lossy-link test.
Risk Category | Reliability
Notes | Idempotent, ordered upload (PRD-CONN-008).
PRD-REL-012 — Platform reliability and durability
Requirement Statement | The cloud platform shall meet its availability target and shall not lose or corrupt ingested data under single-component failure (durable storage, backup/DR).
Rationale | Alerts and records depend on a durable, available platform.
Priority | Must
Source | Reliability
Verification Method | Test/Analysis
Parent Requirement | UN-FUN-006/014; STN-013.
Traceability | → PRD-CLOUD-012; → PVL availability/DR test.
Risk Category | Reliability
Notes | Availability/SLA target Evidence Pending.
7. Verification Notes
Reliability requirements are closed by the 90-day cold-endurance and cycle-life tests (MR-VAL-006), charge-inhibit reliability, fault-injection, power-cycle storage-integrity, watchdog/self-check, depletion, life/aging, lossy-link, and platform DR tests, plus DFMEA/reliability analysis (ERFKL, Stage 11). The D3 tension (REL-001/002/003) is explicitly validated by test.
8. Traceability Placeholders
Upstream | UN-800/900; MR-RISK-001; D1/D3; power budget (Stage 00).
Downstream | PRD-ELEC, PRD-ARCH, PRD-SAFE, PRD-CONN, PRD-CLOUD, PRD-QUAL, PRD-VER; ERFKL DFMEA (Stage 11); HAL (Stage 10); ORCH-018 / EM-ENG-012.
PRD-SAFE (1300) — Safety
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001 with Safety/Risk and Battery engineers. PRD-SAFE-NNN, single "shall", 11-field block. Based on UN-900 (Safety), ISO 14971, IEC 62366-1/62368-1, UN 38.3, and D3/D4.
STATUS — PRELIMINARY (to be reconciled with HAL, Stage 10). Per the Product Requirements Standard,
Safety/risk-control requirements originate in the Hazard Analysis Library (HAL, ORCH-017). These are
preliminary safety/risk-control requirements derived from the Stage-04 safety needs and the concept.
Each will be traced to a HAL HA-<CODE>-NNN hazard record and a Risk-Control Requirement (RCR-###),
refined by the Preliminary and Detailed Hazard Analyses, and re-baselined here. Action Priority is
severity-first (CLAUDE.md rule 3); severity is fixed by outcome, probability/detection are Evidence
Pending until HAL.
1. Purpose
State the preliminary safety requirements so that the SCM neither harms directly (battery/thermal/transport) nor contributes to harm indirectly by enabling a wrong disposition (spoiled dose given / safe dose wasted).
2. Scope
Correct-disposition safety, no-missed-excursion / no-false-reassurance, use-error safety, battery/charge safety, transport/thermal safety, and fail-safe behaviour. Excludes the risk-control mechanisms' detailed design (design outputs) and security-only threats (PRD-SEC) except where safety-relevant.
3. Applicable Standards
ISO 14971 (risk management, severity-first); IEC 62366-1 (use-related risk); IEC 62368-1 (electrical/thermal safety); UN 38.3 / IATA DGR (Li-battery transport); WHO PQS E006 (disposition performance). Cited by designation; revisions to confirm.
4. Background
The SCM's most consequential safety role is informational — supporting the accept/reject decision. Both error directions harm the beneficiary (severity anchor STN-020 / UN-VIS-001). Physical hazards are those of a reusable Li-ion cellular device (charge at low temperature D3, transport D4, thermal/electrical).
5. Requirement Categories
Correct-disposition · No missed excursion / no false reassurance · Use-error · Battery/charge · Transport/ thermal · Fail-safe/recovery.
6. Detailed Requirements
PRD-SAFE-001 — Do not mislead disposition (severity anchor)
Requirement Statement | The SCM shall not present information that would lead a user to administer a temperature-compromised dose or to discard a potent one, within the validated performance of its detection and verdict functions.
Rationale | Either error harms the beneficiary — the dominant safety contribution (STN-020).
Priority | Must
Source | Safety (ISO 14971)
Verification Method | Test (validation)
Parent Requirement | UN-SAFE-001; STN-020.
Traceability | → PRD-FUN-006/010/011, PRD-VIS-002; → HA-### (top hazard), RCR-###; → PVaL.
Risk Category | Safety
Notes | PRELIMINARY — seeds the top HAL hazard. Severity-first tie-breaker.
PRD-SAFE-002 — Reliable two-sided (freeze) detection as a risk control
Requirement Statement | The SCM shall detect and flag freezing excursions with reliability equal to overheating excursions, per the validated detection performance target.
Rationale | Missed freeze → spoiled freeze-sensitive dose administered (a top loss cause).
Priority | Must
Source | Safety
Verification Method | Test
Parent Requirement | UN-SAFE-002, UN-FUN-004; STN-002.
Traceability | → PRD-FUN-006/007; → HA-### missed-freeze, RCR-###; → PVaL detection validation.
Risk Category | Safety
Notes | PRELIMINARY. Cited to WHO/UNICEF freezing reporting.
PRD-SAFE-003 — No false reassurance (honest gaps/outages)
Requirement Statement | The SCM shall make any monitoring gap, outage, or lapse visible in the record and on-device status rather than appear healthy.
Rationale | False reassurance → wrong disposition on an incomplete record.
Priority | Must
Source | Safety
Verification Method | Test
Parent Requirement | UN-SAFE-003, UN-FUN-014; STN-005/008.
Traceability | → PRD-FUN-017, PRD-REL-006/007, PRD-UX-008; → HA-###, RCR-###.
Risk Category | Safety
Notes | PRELIMINARY. Ties self-check (PRD-REL-007).
PRD-SAFE-004 — Alarm balance (no fatigue-induced miss)
Requirement Statement | The SCM shall tune excursion alerting so that false-alarm and missed-excursion rates meet the validated targets that avoid alarm fatigue muting real, harmful excursions.
Rationale | Over-sensitivity → fatigue → muted real alert; under-sensitivity → missed excursion (C8).
Priority | Must
Source | Safety
Verification Method | Test (validation)
Parent Requirement | UN-SAFE-004, UN-FUN-005/007; STN-007.
Traceability | → PRD-FUN-007, PRD-CLOUD-003; → HA-###, RCR-###; → PVaL false-alarm/miss.
Risk Category | Safety
Notes | PRELIMINARY. Balance is safety-relevant both directions.
PRD-SAFE-005 — Use-error protection for low-training users
Requirement Statement | The SCM shall be designed so foreseeable use errors by minimally-trained, gloved, offline users do not lead to a wrong disposition, per the summative usability acceptance criteria.
Rationale | The edge user is the least-trained, most safety-critical operator.
Priority | Must
Source | Safety (IEC 62366-1)
Verification Method | Test (summative usability)
Parent Requirement | UN-SAFE-005, UN-FUN-008; STN-004.
Traceability | → PRD-UX-001/007, PRD-FUN-010; → HA-### use-error, RCR-###; → PVaL usability.
Risk Category | Safety
Notes | PRELIMINARY. Drives summative critical tasks. Usability engineering file.
PRD-SAFE-006 — Prevent unmonitored trip start (mis-arm)
Requirement Statement | The SCM shall prevent a trip from being credited as monitored unless it is armed with a valid profile, making an un-armed/void state unmistakable.
Rationale | A void trip believed monitored → undetected excursion.
Priority | Must
Source | Safety
Verification Method | Test
Parent Requirement | UN-SAFE-006, UN-FUN-012; STN-008.
Traceability | → PRD-FUN-015, PRD-UX-004; → HA-### void-trip, RCR-###.
Risk Category | Safety
Notes | PRELIMINARY. Ties no-false-reassurance (PRD-SAFE-003).
PRD-SAFE-007 — No charge in unsafe (cold) condition
Requirement Statement | The SCM shall inhibit charging below the safe cell temperature limit (≥0 °C) and clearly indicate the inhibited state.
Rationale | Charging a cold Li-ion cell risks lithium plating / thermal event (D3).
Priority | Must
Source | Safety (D3)
Verification Method | Test
Parent Requirement | UN-SAFE-007; D3; IEC 62368-1.
Traceability | → PRD-ELEC-005/006, PRD-REL-003; → HA-### battery, RCR-###; → PVL charge-inhibit.
Risk Category | Safety
Notes | PRELIMINARY. To be reconciled with HAL battery hazard.
PRD-SAFE-008 — No fire/thermal hazard in any condition
Requirement Statement | The SCM shall not present a fire, burn, or thermal-runaway hazard in normal use, foreseeable misuse, or single-fault conditions.
Rationale | It must be safe in aircraft, cold stores, clinics, and homes (all handlers).
Priority | Must
Source | Safety
Verification Method | Test (IEC 62368-1, abuse)
Parent Requirement | UN-SAFE-008; IEC 62368-1; UN 38.3.
Traceability | → PRD-ELEC-007/014; → HA-### thermal, RCR-###; → PVL safety/abuse.
Risk Category | Safety
Notes | PRELIMINARY. BMS/protection in PRD-ELEC; flag battery + electrical engineers.
PRD-SAFE-009 — Safe/compliant Li-battery transport
Requirement Statement | The SCM (and battery) shall be safe and compliant to transport by air and ground per UN 38.3 and IATA DGR, with documentation provided.
Rationale | Non-compliant Li transport → transport-safety hazard / channel exclusion.
Priority | Must
Source | Safety / Regulatory
Verification Method | Test (UN 38.3)
Parent Requirement | UN-SAFE-009, UN-REG-006; UN 38.3 / IATA DGR.
Traceability | → PRD-ELEC-015, PRD-REG-006; → HA-### transport, RCR-###.
Risk Category | Safety
Notes | PRELIMINARY. Test summary Evidence Pending.
PRD-SAFE-010 — Fail-safe and recover without data loss
Requirement Statement | The SCM shall fail safe — preserving logged data and returning to a known safe state — after any power, connectivity, firmware, or reset event, including OTA.
Rationale | A fault must never corrupt the record or leave the device unsafe.
Priority | Must
Source | Safety
Verification Method | Test (fault injection)
Parent Requirement | UN-SAFE-010; STN-025; D2.
Traceability | → PRD-ARCH-008, PRD-REL-004/005, PRD-SEC-008; → HA-###, RCR-###.
Risk Category | Safety
Notes | PRELIMINARY. Signed, resumable, fail-safe OTA (PRD-SEC-008).
PRD-SAFE-011 — Residual-risk acceptability
Requirement Statement | The SCM shall demonstrate that residual risks, after risk controls, are acceptable per the ISO 14971 risk-management plan and benefit-risk analysis.
Rationale | ISO 14971 requires overall residual-risk evaluation, not just per-hazard control.
Priority | Must
Source | Safety (ISO 14971)
Verification Method | Analysis
Parent Requirement | UN-SAFE (all); ISO 14971.
Traceability | → HAL risk-management file (Stage 10), RCR-###; → PVaL.
Risk Category | Safety
Notes | PRELIMINARY. Closed in HAL/risk-management file.
PRD-SAFE-012 — Dead-in-cold as a safety-relevant informational failure
Requirement Statement | The SCM shall meet the 90-day cold-endurance and data-retention requirements (PRD-REL-001) so that loss of the record in deep cold — which can cause a wrong disposition — is prevented.
Rationale | A dead-in-cold device loses the record and can force an unsafe disposition (the D3 risk framed as safety).
Priority | Must
Source | Safety / Reliability (D3)
Verification Method | Test (90-day cold-endurance, MR-VAL-006)
Parent Requirement | UN-PERF-003, UN-SAFE-001; D3; MR-RISK-001.
Traceability | → PRD-REL-001, PRD-ELEC-002; → HA-### lost-record, RCR-###; → PVL endurance.
Risk Category | Safety
Notes | PRELIMINARY. Links the D3 tension to the disposition-safety anchor. Validated by test.
7. Verification Notes
Safety requirements are validated jointly with the Hazard Analysis (HAL, Stage 10) and the Usability Engineering File: correct-disposition and detection validation, false-reassurance/gap handling, summative use-error study, battery/charge/transport safety (IEC 62368-1, UN 38.3), and fault-injection. Severity is fixed by outcome; probability/detection targets are Evidence Pending until HAL. Action Priority is severity-first (CLAUDE.md rule 3).
8. Traceability Placeholders
Upstream | UN-900; ISO 14971; IEC 62366-1/62368-1; UN 38.3; STN-020; D3/D4.
Downstream | HAL (HA-###, RCR-###, Stage 10); PRD-FUN, PRD-UX, PRD-ELEC, PRD-REL, PRD-SEC, PRD-REG; Validation (PVaL); ORCH-018 / EM-ENG-012.
PRD-REG (1400) — Regulatory
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Quality/Regulatory engineer (EM-ENG-006) with EM-ENG-001. PRD-REG-NNN, single "shall", 11-field block. Research-driven per the standard (identify → search KB → ingest missing → cite). Based on UN-1300 (Regulatory), MR-1100, and D4. Forward to Stage 06, regulatory stage, and ORCH-018.
1. Purpose
State the regulatory/market-access requirements that gate whether the SCM can be sold, deployed, flown, and trusted across its target markets (D4), each cited to a governing standard/regulation by designation.
2. Scope
WHO PQS eligibility, EU GDP fit, US FDA classification/data-integrity, radio/EMC/electrical safety, Li-battery transport, calibration traceability, data-integrity (Part 11), environmental/EoL, and privacy. Excludes the technical means of compliance (design) and verification tests (verification stage).
3. Applicable Standards
WHO PQS E006 (category TBC — Evidence Pending); EU GDP (2013/C 343/01); US FDA cold-chain/data-integrity (classification Evidence Pending); 21 CFR Part 11 / ALCOA+; UN 38.3 / IATA DGR; FCC Part 15 / CE RED; CISPR / IEC 61000; IEC 62368-1; IEC 60529 / IEC 60068; ISO/IEC 17025 concepts; RoHS / WEEE / battery directives. Cited by designation; revisions to confirm; standard clause numbers are not asserted as fact.
4. Background
Global vaccine distribution (D4) means several regimes at once. WHO PQS and radio/transport approvals are market-access gates; GDP, Part 11, and IEC safety make records and use defensible. Two named gaps carry forward: exact WHO PQS category and US classification (concept §7).
5. Requirement Categories
Market-access · Data-integrity/records · Radio/EMC/electrical safety · Battery transport · Calibration/ metrology · Environmental/privacy/EoL · Labeling/documentation.
6. Detailed Requirements
PRD-REG-001 — WHO PQS eligibility
Requirement Statement | The SCM shall conform to the requirements of the applicable WHO PQS temperature-monitoring device category sufficient to achieve PQS listing.
Rationale | PQS listing is a hard gate for donor-funded/LMIC procurement.
Priority | Must
Source | Regulatory
Verification Method | Test + Inspection (PQS conformance)
Parent Requirement | UN-REG-001; WHO PQS E006.
Traceability | → PRD-PERF-001 accuracy, PRD-FUN detection/verdict; → regulatory stage.
Risk Category | Compliance
Notes | Exact PQS category/code Evidence Pending (concept §7); plan listing early (long cycles).
PRD-REG-002 — Per-market regulatory classification basis
Requirement Statement | The SCM shall have a documented regulatory status/classification per target market (WHO PQS category, US FDA classification, EU GDP fit) supporting placement and registration.
Rationale | The device must be placeable, registrable, and defensible per market (D4).
Priority | Must
Source | Regulatory
Verification Method | Analysis
Parent Requirement | UN-REG-002; D4; concept §7.
Traceability | → regulatory strategy (regulatory stage); → PRD-REG-001.
Risk Category | Compliance
Notes | US classification and WHO PQS category Evidence Pending (named concept gaps).
PRD-REG-003 — EU GDP + 21 CFR Part 11 / ALCOA+ records
Requirement Statement | The SCM records shall satisfy EU GDP and 21 CFR Part 11 / ALCOA+ — attributable, legible, contemporaneous, original, accurate, complete, tamper-evident — to be admissible for release/audit.
Rationale | Records must be admissible evidence for disposition and audit.
Priority | Must
Source | Regulatory
Verification Method | Analysis + Test (CSV)
Parent Requirement | UN-REG-003, UN-FUN-014, UN-SEC-001; EU GDP; 21 CFR Part 11 / ALCOA+.
Traceability | → PRD-FUN-017, PRD-SEC-001/002/004, PRD-CLOUD-007; → PVaL audit simulation.
Risk Category | Compliance
Notes | Standard (rev. to confirm); CSV per GAMP 5.
PRD-REG-004 — Radio approvals per market
Requirement Statement | The SCM radios (BLE/Wi-Fi/LTE-M/GNSS) shall hold the radio approvals required in each target market (e.g., FCC Part 15, CE RED, and regional equivalents).
Rationale | Radio approval is a legal-to-sell/use gate per geography (D4).
Priority | Must
Source | Regulatory
Verification Method | Test (radio conformance)
Parent Requirement | UN-REG-004; D4; FCC Part 15 / CE RED.
Traceability | → PRD-CONN-012, PRD-ENV-009; → PVL radio-conformance test.
Risk Category | Compliance
Notes | May drive per-market hardware variants (PRD-ARCH-010). Standard (rev. to confirm).
PRD-REG-005 — EMC and electrical/thermal safety
Requirement Statement | The SCM shall meet EMC (emissions and immunity, CISPR/IEC 61000) and electrical/thermal safety (IEC 62368-1) requirements for its use and transport environments.
Rationale | Safe and accepted for aircraft, enterprise, and clinical environments.
Priority | Must
Source | Regulatory / Safety
Verification Method | Test (EMC/safety)
Parent Requirement | UN-REG-005, UN-SAFE-008, UN-ENV-009; IEC 62368-1; CISPR/IEC 61000.
Traceability | → PRD-ELEC-013/014, PRD-SAFE-008; → PVL EMC/safety test.
Risk Category | Compliance
Notes | Standard (rev. to confirm).
PRD-REG-006 — Li-battery transport (UN 38.3 / IATA)
Requirement Statement | The SCM battery shall be UN 38.3-tested and the device transport IATA-compliant, with the required documentation provided.
Rationale | Legal, safe air/ground transport of the Li battery (D4); avoids DG holds.
Priority | Must
Source | Regulatory
Verification Method | Test (UN 38.3)
Parent Requirement | UN-REG-006, UN-SAFE-009; UN 38.3 / IATA DGR.
Traceability | → PRD-ELEC-015, PRD-SAFE-009; → PVL UN 38.3 report.
Risk Category | Compliance
Notes | Test summary + IATA guidance Evidence Pending. Standard (rev. to confirm).
PRD-REG-007 — Traceable calibration
Requirement Statement | The SCM shall have calibration traceable to a recognized metrology standard, with calibration status and interval recorded.
Rationale | Readings — and dispositions based on them — must be defensible.
Priority | Must
Source | Regulatory
Verification Method | Inspection (calibration cert)
Parent Requirement | UN-REG-007, UN-PERF-002; ISO/IEC 17025 concepts.
Traceability | → PRD-PERF-002, PRD-SVC-005, PRD-FUN-020; → calibration records.
Risk Category | Compliance
Notes | Interval Evidence Pending (regulatory/service stage).
PRD-REG-008 — Environmental / hazardous-substance / battery EoL
Requirement Statement | The SCM shall comply with applicable hazardous-substance (RoHS/REACH) and battery/electronics end-of-life (WEEE/battery directives) regulations in its target markets.
Rationale | It must be sold and retired lawfully and responsibly (C10).
Priority | Must
Source | Regulatory / Sustainability
Verification Method | Inspection
Parent Requirement | UN-REG-008, PRD-VIS-006; RoHS/WEEE/battery directives.
Traceability | → PRD-MECH-009 materials, PRD-SVC decommission; → compliance declarations.
Risk Category | Compliance
Notes | Per-market; EoL path Evidence Pending. Standard (rev. to confirm).
PRD-REG-009 — Regional data-protection/privacy
Requirement Statement | The SCM shall handle any personal/location data in line with applicable regional data-protection law for each market of deployment.
Rationale | The platform must be lawful to deploy per region (D4).
Priority | Should
Source | Regulatory
Verification Method | Inspection
Parent Requirement | UN-REG-009, UN-SEC-008.
Traceability | → PRD-SEC-010, PRD-CLOUD-011; → privacy review.
Risk Category | Compliance
Notes | Device monitors cargo not people; applicable-law mapping Evidence Pending.
PRD-REG-010 — Labeling and documentation per regime
Requirement Statement | The SCM shall carry the labeling and be supported by the documentation each regime requires (safety marks, battery/transport marks, PQS/GDP documentation, device ID).
Rationale | It must pass acceptance and audit at every gate.
Priority | Must
Source | Regulatory
Verification Method | Inspection
Parent Requirement | UN-REG-010, UN-PHY-009; WHO PQS E006; IATA DGR; CE/FCC.
Traceability | → PRD-MECH-010 durable labeling; → regulatory stage documentation set.
Risk Category | Compliance
Notes | Documentation set Evidence Pending. Standard (rev. to confirm).
PRD-REG-011 — Quality-system basis for a regulated device
Requirement Statement | The SCM shall be developed under a quality system appropriate to its confirmed regulatory classification, with design controls and a risk-management file (ISO 14971).
Rationale | Market access and defensibility require an appropriate QMS and design-control record.
Priority | Must
Source | Regulatory / Quality
Verification Method | Analysis (QMS audit)
Parent Requirement | UN-REG-002; ISO 14971; (ISO 13485 if classified as a device — confirm).
Traceability | → PRD-QUAL, HAL risk-management file; → QMS records.
Risk Category | Compliance
Notes | Applicable QMS depends on classification (Evidence Pending); ISO 13485 applicability to confirm.
PRD-REG-012 — Computerized-system validation (CSV)
Requirement Statement | The SCM cloud/software shall be validated per GAMP 5 / applicable CSV expectations for the regulated data it manages.
Rationale | Records/decisions on a computerized system must be validated for regulated customers.
Priority | Should
Source | Regulatory
Verification Method | Analysis (CSV)
Parent Requirement | UN-SEC-003, UN-REG-003; GAMP 5.
Traceability | → PRD-CLOUD-007/009, PRD-SEC; → CSV documentation.
Risk Category | Compliance
Notes | Standard (GAMP 5, rev. to confirm).
7. Verification Notes
Regulatory requirements are closed by achieving the listed certifications/listings and by regulatory review: WHO PQS prequalification, radio (FCC/CE + regional), EMC (CISPR/61000), safety (IEC 62368-1), transport (UN 38.3/IATA), calibration traceability, Part 11/GDP records validation, and RoHS/WEEE declarations. The WHO PQS category, US classification, calibration interval, and privacy-law mapping are Evidence Pending and confirmed in the Regulatory stage (research-driven).
8. Traceability Placeholders
Upstream | UN-1300; MR-1100; D4; concept §7 gaps.
Downstream | PRD-PERF, PRD-FUN, PRD-ELEC, PRD-CONN, PRD-SEC, PRD-CLOUD, PRD-SAFE, PRD-SVC, PRD-QUAL; regulatory stage; HAL; ORCH-018 / EM-ENG-012.
PRD-MFG (1500) — Manufacturing
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Manufacturing/Test engineer with Quality and EM-ENG-001. PRD-MFG-NNN, single "shall", 11-field block. Based on UN-1400 (Quality), UN-1300 (traceability), and D1. Forward to Stage 06, manufacturing transfer, PFMEA (ERFKL), and ORCH-018.
1. Purpose
Specify the manufacturing, calibration, programming, serialization, testability, and packaging requirements that make the SCM producible, traceable, and reliably conformant at volume.
2. Scope
Assembly/testability, calibration at manufacture, secure programming/provisioning, serialization/traceability, end-of-line test, packaging/labeling for shipment, and yield. Excludes field service (PRD-SVC) and product design detail.
3. Applicable Standards
IPC assembly standards (rev. to confirm); ISO/IEC 17025 concepts (calibration); UN 38.3 / IATA (battery shipment); ISO 13485 concepts (if classified as a device — confirm). Cited by designation; revisions to confirm.
4. Background
A reusable, calibrated, secure-provisioned, serialized device (D1) must be built with production calibration, secure key/firmware injection, and full unit traceability so field records and fleet management are credible.
5. Requirement Categories
Assembly/testability · Production calibration · Secure programming · Serialization/traceability · End-of-line test · Packaging/labeling · Yield.
6. Detailed Requirements
PRD-MFG-001 — Design for manufacture and assembly
Requirement Statement | The SCM shall be designed for reliable assembly with defined test-access points enabling automated end-of-line test of all safety- and function-critical parameters.
Rationale | Testability and DFM/DFA drive yield and conformance at volume.
Priority | Should
Source | Manufacturing
Verification Method | Analysis (DFM review)
Parent Requirement | UN-QUAL-###; D1.
Traceability | → PRD-MFG-005, PFMEA (ERFKL Stage 11); → DFM review.
Risk Category | Reliability
Notes | Test-access points defined with test engineering.
PRD-MFG-002 — Production calibration with traceability
Requirement Statement | Each SCM unit shall be calibrated at manufacture against a reference traceable to a recognized metrology standard, and the calibration data/date shall be recorded to the unit.
Rationale | Accuracy and PQS/GDP defensibility require per-unit traceable calibration.
Priority | Must
Source | Regulatory / Quality
Verification Method | Inspection (calibration records)
Parent Requirement | UN-REG-007, UN-PERF-001; ISO/IEC 17025 concepts.
Traceability | → PRD-REG-007, PRD-PERF-001, PRD-SVC-005; → calibration records.
Risk Category | Compliance
Notes | Ties field readiness (PRD-FUN-020).
PRD-MFG-003 — Secure firmware and credential provisioning
Requirement Statement | Each SCM unit shall be provisioned in manufacture with verified firmware and its unique protected identity/credentials via a controlled, auditable process.
Rationale | Device identity, secure boot, and integrity depend on trustworthy factory provisioning.
Priority | Must
Source | Security / Manufacturing
Verification Method | Inspection + Test
Parent Requirement | UN-SEC-003/005; IEC 62443 concepts.
Traceability | → PRD-SEC-003/009, PRD-CLOUD-001; → provisioning-process audit.
Risk Category | Security
Notes | Key-injection process protected; ties CSV.
PRD-MFG-004 — Serialization and unit traceability
Requirement Statement | Each SCM unit shall carry a unique serial/identifier linkable to its build/calibration/component records for lifecycle traceability.
Rationale | Fleet management, recalls, and attribution require per-unit traceability.
Priority | Must
Source | Quality / Regulatory
Verification Method | Inspection
Parent Requirement | UN-FUN-017/018; STN-010; ISO 13485 concepts.
Traceability | → PRD-MECH-010 durable ID, PRD-CLOUD-001; → traceability records.
Risk Category | Compliance
Notes | Durable marking (PRD-MECH-010).
PRD-MFG-005 — End-of-line functional and safety test
Requirement Statement | Each SCM unit shall pass an end-of-line test verifying sensing accuracy, connectivity, charge/charge-inhibit, display/verdict, and safety-critical protections before release.
Rationale | No unit ships without confirmed safety-/function-critical conformance.
Priority | Must
Source | Manufacturing / Quality
Verification Method | Test (EOL)
Parent Requirement | UN-QUAL-###; PRD-SAFE (protections).
Traceability | → PRD-ELEC-005/007, PRD-FUN, PRD-PERF; → EOL test records.
Risk Category | Safety
Notes | Test spec derived from this library (single source of truth).
PRD-MFG-006 — Packaging protects device and battery in shipment
Requirement Statement | The SCM packaging shall protect the device and battery through distribution and comply with Li-battery shipment marking/documentation (UN 38.3 / IATA).
Rationale | The product must reach customers safely and compliantly (battery shipment).
Priority | Must
Source | Manufacturing / Regulatory
Verification Method | Inspection + Test
Parent Requirement | UN-REG-006; UN 38.3 / IATA DGR.
Traceability | → PRD-ELEC-015, PRD-REG-006; → packaging/shipping test.
Risk Category | Compliance
Notes | Distinct from in-use sealing (PRD-MECH).
PRD-MFG-007 — Production process control and PFMEA
Requirement Statement | The SCM production shall be governed by controlled processes with a PFMEA and control plan for safety- and function-critical characteristics.
Rationale | Volume conformance and safety require process control, not inspection alone.
Priority | Should
Source | Quality / Manufacturing
Verification Method | Analysis (PFMEA/control plan)
Parent Requirement | UN-QUAL-###.
Traceability | → ERFKL PFMEA (Stage 11), PRD-QUAL; → control plan.
Risk Category | Reliability
Notes | Severity-first Action Priority (CLAUDE.md rule 3).
PRD-MFG-008 — Production yield and first-pass target
Requirement Statement | The SCM production shall meet a first-pass-yield target set in the manufacturing plan, with rework/scrap dispositioned under quality control.
Rationale | Yield governs cost and reuse economics; escapes threaten field credibility.
Priority | Could
Source | Business / Quality
Verification Method | Analysis
Parent Requirement | UN-BIZ-###; MR-900.
Traceability | → PRD-QUAL, PRD-MFG-005; → yield reports.
Risk Category | None
Notes | Yield target Evidence Pending.
7. Verification Notes
Manufacturing requirements are verified by DFM review, calibration/provisioning-process audit, EOL test records, packaging/shipping tests, and PFMEA/control-plan review (ERFKL, Stage 11). The EOL test spec is derived directly from this library.
8. Traceability Placeholders
Upstream | UN-1400/1300; D1; MR-900.
Downstream | PRD-ELEC, PRD-FUN, PRD-PERF, PRD-SEC, PRD-QUAL, PRD-REG, PRD-SVC; ERFKL PFMEA; manufacturing transfer; ORCH-018 / EM-ENG-012.
PRD-SVC (1600) — Serviceability
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Service/Reliability engineer with EM-ENG-001. PRD-SVC-NNN, single "shall", 11-field block. Based on UN-1200 (Service), UN-800 (reliability), and D1 (reusable). Forward to Stage 06, service documentation, and ORCH-018.
1. Purpose
Specify the serviceability requirements for the reusable multi-trip lifecycle (D1): return, cleaning, recharge, readiness triage, recalibration, firmware maintenance, spares, and decommission.
2. Scope
Reverse-logistics readiness, cleaning/decontamination, recharge turnaround, readiness triage, recalibration, diagnostics/logs, spares/repair, upgrade paths, and end-of-life decommission. Excludes manufacture (PRD-MFG) and field-user operation (PRD-FUN/UX).
3. Applicable Standards
ISO/IEC 17025 concepts (recalibration traceability); IEC 60529 (cleaning vs ingress); WEEE/battery directives (decommission); IEC 62368-1 (safe service). Cited by designation; revisions to confirm.
4. Background
A reusable device (D1) lives or dies on its refurbishment loop: a returned unit must be cleaned, recharged (charge-inhibited if cold), triaged for calibration/charge/firmware, and re-provisioned, with unfit units blocked from redeployment.
5. Requirement Categories
Return/refurbishment · Cleaning · Recharge turnaround · Readiness triage · Recalibration · Diagnostics/logs · Spares/repair · Decommission/EoL.
6. Detailed Requirements
PRD-SVC-001 — Refurbishment turnaround loop
Requirement Statement | The SCM shall support a defined refurbishment loop (receive → clean → recharge → triage → re-provision) whose steps a trained service operator can complete within the turnaround target.
Rationale | Reuse economics depend on a predictable, fast turnaround (D1, C6).
Priority | Must
Source | User Need (D1)
Verification Method | Demonstration
Parent Requirement | UN-SVC-001, UN-VIS-005; D1.
Traceability | → PRD-SVC-002..005; → PVaL refurbishment scenario.
Risk Category | Reliability
Notes | Turnaround target Evidence Pending.
PRD-SVC-002 — Cleaning/decontamination without damage
Requirement Statement | The SCM shall be cleanable/decontaminable with the specified agents and method between trips without ingress, corrosion, seal, display, or label degradation.
Rationale | Repeated cleaning (D1) must not create ingress or degrade the device.
Priority | Must
Source | User Need
Verification Method | Test (cleaning-cycle)
Parent Requirement | UN-SVC-###, UN-PHY-008; STN-026.
Traceability | → PRD-ENV-008, PRD-MECH-009; → PVL cleaning-durability.
Risk Category | Reliability
Notes | Agent list/method Evidence Pending; flag materials engineer.
PRD-SVC-003 — Predictable recharge and trip-readiness
Requirement Statement | The SCM shall reach full trip-readiness after depot recharge within the charge- time target, and shall clearly indicate when a cold-returned device must warm before it will safely charge.
Rationale | Predictable turnaround; no under-charged or cold-charged unit redeployed (D3).
Priority | Must
Source | User Need (D3)
Verification Method | Test
Parent Requirement | UN-PERF-004, UN-SVC-###; D3.
Traceability | → PRD-ELEC-004/005/006, PRD-SAFE-007; → PVL charge-turnaround test.
Risk Category | Safety
Notes | Charge-inhibit-below-0 °C indicated (PRD-ELEC-005). To be reconciled with HAL.
PRD-SVC-004 — Readiness triage with deploy-block
Requirement Statement | The SCM shall present, and report to fleet tooling, its charge/calibration/ firmware readiness such that an out-of-calibration or under-charged unit is blocked from redeployment.
Rationale | Unfit units must not silently redeploy (out-of-cal disputes; dead-in-cold).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-017, UN-SVC-###; STN-010/021.
Traceability | → PRD-FUN-020, PRD-CLOUD-006, PRD-ELEC-009; → PVL readiness/triage test.
Risk Category | Reliability
Notes | Deploy-block on out-of-cal.
PRD-SVC-005 — Recalibration to traceable standard
Requirement Statement | The SCM shall support recalibration to a recognized traceable standard at a defined interval, updating the recorded calibration status/date.
Rationale | Accuracy over life and PQS/GDP defensibility require periodic traceable recalibration.
Priority | Must
Source | Regulatory / User Need
Verification Method | Demonstration
Parent Requirement | UN-REG-007, UN-PERF-002; ISO/IEC 17025 concepts.
Traceability | → PRD-REG-007, PRD-MFG-002, PRD-PERF-002; → calibration records.
Risk Category | Compliance
Notes | Interval Evidence Pending.
PRD-SVC-006 — Service diagnostics and logs
Requirement Statement | The SCM shall provide service-accessible diagnostics and event/error logs sufficient to triage returned devices without disassembly.
Rationale | Efficient triage and root-cause need non-destructive diagnostics.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-SVC-###, UN-FUN-017.
Traceability | → PRD-MOB-004, PRD-CONN-010; → service-tool review.
Risk Category | None
Notes | Logs preserve integrity/attribution.
PRD-SVC-007 — Firmware maintenance and upgrade path
Requirement Statement | The SCM shall support firmware maintenance/upgrade in the field (OTA) and at depot (USB-C) via the signed, fail-safe update process, without data loss.
Rationale | A reusable fleet must stay current and secure over years.
Priority | Must
Source | User Need / Security
Verification Method | Test
Parent Requirement | UN-SEC-005, UN-SAFE-010; D2.
Traceability | → PRD-FUN-022, PRD-SEC-008, PRD-CLOUD-008; → PVL OTA test.
Risk Category | Security
Notes | Same process as PRD-SEC-008.
PRD-SVC-008 — Safe decommission and end-of-life
Requirement Statement | The SCM shall support safe decommissioning including data sanitization and compliant battery/electronics disposal/recycling per applicable regulations.
Rationale | Reusable Li-ion device must be retired safely, lawfully, and with data protected (C10).
Priority | Should
Source | Regulatory / Sustainability
Verification Method | Demonstration
Parent Requirement | UN-REG-008, UN-SVC-###; RoHS/WEEE/battery directives.
Traceability | → PRD-REG-008, PRD-VIS-006, PRD-SEC-007; → EoL procedure.
Risk Category | Compliance
Notes | Data sanitization ties PRD-SEC; EoL path Evidence Pending.
7. Verification Notes
Serviceability requirements are verified by the refurbishment-loop demonstration, cleaning-durability and charge-turnaround tests, readiness/triage and recalibration demonstrations, service-tool review, OTA test, and the decommission procedure review.
8. Traceability Placeholders
Upstream | UN-1200/800; D1/D3; STN-010/021/026.
Downstream | PRD-ELEC, PRD-MECH, PRD-FUN, PRD-CLOUD, PRD-SEC, PRD-REG, PRD-PERF; service documentation; HAL; ORCH-018 / EM-ENG-012.
PRD-PERF (1700) — Performance
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001 with Reliability, RF, and Electronics engineers. PRD-PERF-NNN, single "shall", 11-field block. Based on UN-800 (Performance), WHO PQS E006, IEC 60068, and D2/D3. Forward to Stage 06, verification, and ORCH-018.
1. Purpose
Set the measurable performance targets — accuracy, endurance, latency, capacity, responsiveness, location, and availability — that make the SCM's verdict trustworthy and its record complete.
2. Scope
Measurement accuracy/resolution, 90-day energy performance, alert latency, log capacity, responsiveness, location cadence, connectivity performance, and availability. Excludes the functions (PRD-FUN), environmental survival (PRD-ENV), and reliability program (PRD-REL).
3. Applicable Standards
WHO PQS E006 (accuracy/performance class TBC); IEC 60068 (performance vs environment); ISO/IEC 17025 concepts (measurement traceability). Cited by designation; class/revision to confirm (Evidence Pending).
4. Background
Three targets dominate: accuracy good enough to trust the verdict; endurance to survive 90 days at −30 °C without dying or losing data (top risk, D3); and alert latency low enough to enable rescue (connectivity- bound, D2). Numeric targets marked Judgment are confirmed in verification/validation.
5. Requirement Categories
Accuracy · Endurance/energy · Latency · Capacity · Responsiveness · Location · Connectivity · Availability.
6. Detailed Requirements
PRD-PERF-001 — Temperature measurement accuracy
Requirement Statement | The SCM shall measure temperature with an accuracy of ±0.5 °C or better across the −30 °C to +60 °C operating range (against a traceable reference).
Rationale | Accuracy must be good enough that an accept/reject decision can be trusted and defended.
Priority | Must
Source | Regulatory / Performance
Verification Method | Test (traceable reference)
Parent Requirement | UN-PERF-001; WHO PQS E006; STN-005.
Traceability | → PRD-REG-001/007, PRD-MECH-011, PRD-FUN-002; → PVL accuracy test.
Risk Category | Safety
Notes | ±0.5 °C is a Judgment target — confirm against WHO PQS E006 class (rev. to confirm).
PRD-PERF-002 — Accuracy stability over range and calibration interval
Requirement Statement | The SCM shall hold PRD-PERF-001 accuracy across the full range and remain within spec over the defined calibration interval and reuse life.
Rationale | Readings must stay trustworthy at extremes and over time (drift risk).
Priority | Must
Source | Regulatory / Reliability
Verification Method | Test (aging/drift)
Parent Requirement | UN-PERF-002; ISO/IEC 17025 concepts.
Traceability | → PRD-REG-007, PRD-SVC-005, PRD-REL-009; → PVL drift test.
Risk Category | Compliance
Notes | Calibration interval Evidence Pending.
PRD-PERF-003 — Temperature resolution
Requirement Statement | The SCM shall record temperature with a resolution of 0.1 °C or finer.
Rationale | Resolution must support the accuracy target and meaningful excursion logic.
Priority | Should
Source | Performance
Verification Method | Test
Parent Requirement | UN-PERF-001.
Traceability | → PRD-FUN-001/007; → PVL resolution test.
Risk Category | Performance
Notes | 0.1 °C is Judgment; confirm vs PQS class.
PRD-PERF-004 — Excursion-alert latency (when connected)
Requirement Statement | The SCM shall deliver an excursion alert to the cloud within 5 minutes of detection when connectivity is available at detection time.
Rationale | Fast alerting converts monitoring into prevented loss; connectivity-bound (D2).
Priority | Must
Source | User Need / Safety
Verification Method | Test
Parent Requirement | UN-PERF-005, UN-FUN-006; STN-001; D2.
Traceability | → PRD-FUN-008, PRD-CONN-009, PRD-CLOUD-003; → HAL delayed-alert; → PVL latency test.
Risk Category | Safety
Notes | 5 min is a Judgment target — confirm in validation. Offline-captured otherwise.
PRD-PERF-005 — 90-day energy performance
Requirement Statement | The SCM shall achieve ≥90 days of operation from full charge under the defined worst-case cold-duty profile with a documented energy margin.
Rationale | The headline endurance target (D3); the energy budget is the constraint.
Priority | Must
Source | User Need (D3)
Verification Method | Test (endurance, MR-VAL-006)
Parent Requirement | UN-PERF-003; D3; power budget (Stage 00).
Traceability | → PRD-REL-001, PRD-ELEC-002/008, PRD-ARCH-004; → PVL 90-day endurance.
Risk Category | Reliability
Notes | Margin % Evidence Pending. Validated by test (top risk).
PRD-PERF-006 — Log capacity for a full offline trip
Requirement Statement | The SCM shall store the complete multi-channel log for a full 90-day trip at the minimum logging interval with at least 50% spare capacity, even with no connectivity.
Rationale | An entirely offline trip must still yield a complete record with headroom.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-PERF-007, UN-FUN-009; D2.
Traceability | → PRD-FUN-012, PRD-REL-005; → PVL capacity test.
Risk Category | Reliability
Notes | 90 d at 1-min interval ≈ 1.3×10^5 samples/channel; exact sizing Evidence Pending. 50% spare is a Judgment target.
PRD-PERF-007 — Interaction responsiveness
Requirement Statement | The SCM shall present the requested status/verdict within 3 seconds of a user action, including at low temperature and low battery.
Rationale | A time-pressured user must not be left waiting or unsure.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-PERF-006; STN-004/008.
Traceability | → PRD-UX-006; → PVL response-time test.
Risk Category | Usability
Notes | 3 s is a Judgment target; confirm in usability study.
PRD-PERF-008 — Humidity measurement accuracy
Requirement Statement | The SCM shall measure relative humidity with an accuracy of ±5% RH or better over the operating range.
Rationale | Humidity data must be meaningful for products/conditions that depend on it.
Priority | Should
Source | Performance
Verification Method | Test
Parent Requirement | UN-FUN-002.
Traceability | → PRD-FUN-003, PRD-ENV-004; → PVL humidity-accuracy test.
Risk Category | Performance
Notes | ±5% RH is a Judgment target.
PRD-PERF-009 — Availability (no silent outage)
Requirement Statement | The SCM shall be available and actively monitoring for at least a defined fraction of the armed trip, with any outage flagged rather than silent.
Rationale | "Armed" must reliably mean "monitoring"; silent outages are a false-reassurance hazard.
Priority | Must
Source | Reliability / Safety
Verification Method | Test
Parent Requirement | UN-PERF-011, UN-SAFE-003.
Traceability | → PRD-REL-007, PRD-SAFE-003; → PVL availability/self-check test.
Risk Category | Safety
Notes | Availability fraction Evidence Pending. To be reconciled with HAL.
PRD-PERF-010 — Location cadence within energy budget
Requirement Statement | The SCM shall provide periodic/event-triggered location adequate for fleet management and stray recovery within the energy allocation for location, without breaking PRD-PERF-005.
Rationale | Location is useful but must not exhaust the 90-day budget (C3).
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-PERF-009, UN-FUN-016; conflict C3.
Traceability | → PRD-CONN-005, PRD-FUN-019, PRD-ELEC-008; → PVL location-cadence test.
Risk Category | Performance
Notes | GNSS-in-container performance Evidence Pending; may need external antenna.
PRD-PERF-011 — Data-delivery success over a trip
Requirement Statement | The SCM shall deliver at least a defined fraction of buffered data to the cloud over a representative trip when any usable connectivity occurs, without data loss.
Rationale | Store-and-forward must actually get the data out over global lanes.
Priority | Must
Source | Performance / Reliability
Verification Method | Test (RF-in-container/field)
Parent Requirement | UN-PERF-008, UN-ENV-006; D2.
Traceability | → PRD-CONN-007/008, PRD-CLOUD-002; → PVL delivery-success test.
Risk Category | Performance
Notes | Success fraction Evidence Pending; principal open risk (RF-in-container).
PRD-PERF-012 — Charge time to trip-readiness
Requirement Statement | The SCM shall recharge from empty to full trip-readiness within a defined charge-time target at permitted charge temperature.
Rationale | Depot turnaround (reuse economics) depends on predictable charge time.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-PERF-004; D1.
Traceability | → PRD-ELEC-004, PRD-SVC-003; → PVL charge-time test.
Risk Category | Reliability
Notes | Charge-time target Evidence Pending.
7. Verification Notes
Performance requirements are closed by test: accuracy against a traceable reference across range and over drift; 90-day/−30 °C endurance and capacity (MR-VAL-006); alert latency vs connectivity; responsiveness; humidity accuracy; availability/self-check; location cadence; data-delivery success (RF-in-container); and charge time. All Judgment targets (±0.5 °C, 5 min, 3 s, ±5% RH, 50% spare) are confirmed in V&V.
8. Traceability Placeholders
Upstream | UN-800; WHO PQS E006; IEC 60068; MR-RISK-001; D2/D3; power budget (Stage 00).
Downstream | PRD-FUN, PRD-ELEC, PRD-CONN, PRD-REL, PRD-REG, PRD-SVC, PRD-VER; verification/validation; HAL; ORCH-018 / EM-ENG-012.
PRD-QUAL (1800) — Quality
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Quality engineer (EM-ENG-006) with EM-ENG-001. PRD-QUAL-NNN, single "shall", 11-field block. Based on UN-1400 (Quality), and D1. Forward to Stage 06, quality plan, and ORCH-018.
1. Purpose
State the product-quality, durability, appearance, and acceptance requirements that make the SCM dependable and credible over its reusable life and across a fleet.
2. Scope
Product quality metrics, durability/fit-and-finish, appearance, inspection/acceptance criteria, defect targets, and quality records. Excludes manufacturing process control (PRD-MFG) and reliability program (PRD-REL).
3. Applicable Standards
ISO 9001 concepts (quality management); ISO 13485 concepts (if classified as a device — confirm); AQL sampling standards (rev. to confirm). Cited by designation; revisions to confirm.
4. Background
A reusable, audit-instrument device (D1) must look and function credibly for years; visible defects, delamination of labels, or drift erode trust in both the device and its records.
5. Requirement Categories
Quality metrics · Durability/fit-and-finish · Appearance · Inspection/acceptance · Defect targets · Quality records.
6. Detailed Requirements
PRD-QUAL-001 — Conformance to acceptance criteria
Requirement Statement | Each SCM unit shall conform to the acceptance criteria derived from this requirements library before release, with nonconformances dispositioned under quality control.
Rationale | The library is the single source of truth; acceptance derives from it.
Priority | Must
Source | Quality
Verification Method | Inspection + Test
Parent Requirement | UN-QUAL-###; PRD-VER.
Traceability | → PRD-MFG-005 EOL, PRD-VER; → acceptance records.
Risk Category | Compliance
Notes | Ties verification matrix (PRD-VER).
PRD-QUAL-002 — Durability over reuse life
Requirement Statement | The SCM shall retain functional performance, sealing, and legibility of display/ labels over its specified reuse life and cleaning cycles.
Rationale | Durability sustains reuse economics and record credibility (D1).
Priority | Must
Source | Quality / Reliability
Verification Method | Test (life/aging)
Parent Requirement | UN-QUAL-###, UN-PERF-010; D1.
Traceability | → PRD-REL-009, PRD-MECH-009/010; → PVL life test.
Risk Category | Reliability
Notes | Overlaps PRD-REL-009 (durability engineering side).
PRD-QUAL-003 — Fit and finish
Requirement Statement | The SCM shall meet defined cosmetic/fit-and-finish criteria (no cracks, burrs, gaps, or defects that impair sealing, legibility, or credibility).
Rationale | Appearance affects trust and can indicate functional defects (seal/ingress).
Priority | Should
Source | Quality
Verification Method | Inspection
Parent Requirement | UN-QUAL-###.
Traceability | → PRD-MECH; → cosmetic acceptance spec.
Risk Category | None
Notes | Cosmetic spec Evidence Pending.
PRD-QUAL-004 — Label/marking legibility over life
Requirement Statement | The SCM identification and safety/regulatory markings shall remain legible and adhered over the reuse life and cleaning cycles.
Rationale | Attribution, safety, and regulatory acceptance depend on durable markings.
Priority | Must
Source | Quality / Regulatory
Verification Method | Test (label durability)
Parent Requirement | UN-PHY-009, UN-REG-010.
Traceability | → PRD-MECH-010, PRD-REG-010; → PVL label-durability test.
Risk Category | Compliance
Notes | Overlaps PRD-MECH-010.
PRD-QUAL-005 — Field defect / escape target
Requirement Statement | The SCM shall meet a field-defect/escape-rate target set in the quality plan, with field failures analyzed for root cause and corrective action.
Rationale | Field escapes threaten record credibility and reuse economics; must be bounded and learned from.
Priority | Should
Source | Quality / Business
Verification Method | Analysis
Parent Requirement | UN-QUAL-###; MR-1000.
Traceability | → CAPA/field-feedback (Stage 17), ERFKL; → quality metrics.
Risk Category | Reliability
Notes | Target Evidence Pending.
PRD-QUAL-006 — Quality records and traceability
Requirement Statement | The SCM shall have quality records (calibration, EOL test, nonconformance, CAPA) linkable per unit/lot for traceability and audit.
Rationale | Regulated/enterprise customers and recalls require traceable quality records.
Priority | Must
Source | Quality / Regulatory
Verification Method | Inspection
Parent Requirement | UN-QUAL-###, UN-REG-003; ISO 13485 concepts.
Traceability | → PRD-MFG-002/004, PRD-REG-011; → quality records.
Risk Category | Compliance
Notes | Ties serialization (PRD-MFG-004).
PRD-QUAL-007 — Incoming/component quality control
Requirement Statement | The SCM shall have incoming quality control for safety- and function-critical components (cells, sensors, connectors) with supplier qualification.
Rationale | Battery/sensor quality directly drives safety and accuracy; must be controlled upstream.
Priority | Should
Source | Quality
Verification Method | Analysis (supplier qual)
Parent Requirement | UN-QUAL-###; PRD-SAFE-008.
Traceability | → PRD-MFG-007, ERFKL; → supplier qualification records.
Risk Category | Reliability
Notes | Critical-component list from this library.
7. Verification Notes
Quality requirements are verified by inspection/acceptance against the library-derived criteria, life/aging and label-durability tests, cosmetic inspection, quality-records audit, supplier qualification, and field defect-rate analysis (with CAPA, Stage 17).
8. Traceability Placeholders
Upstream | UN-1400; D1; MR-1000.
Downstream | PRD-MFG, PRD-REL, PRD-MECH, PRD-REG, PRD-VER; CAPA/field feedback (Stage 17); ORCH-018.
PRD-VER (1900) — Verification
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001 with Verification lead (ORCH-014). PRD-VER-NNN, single "shall", 11-field block. Based on UN-1700 (Validation), the whole PRD library, and D1–D4. Forward to the Verification Plan (PVL, ORCH-014), Validation (PVaL, ORCH-015), and ORCH-018 / EM-ENG-012.
1. Purpose
State the verification-strategy requirements: how the SCM's requirements are traced, verified (Test/Analysis/ Inspection/Demonstration), and accepted — including the headline D3 validate-by-test and the safety/HAL reconciliation.
2. Scope
Verification strategy, requirement traceability, method allocation, acceptance criteria, key qualification tests, and validation linkage. Excludes the detailed test procedures (Verification Plan, ORCH-014) and the validation scenarios (Validation, ORCH-015).
3. Applicable Standards
ISO/IEC/IEEE 29148 (requirements verification); IEC 60068 / 60529 / 62368-1 / UN 38.3 / CISPR-IEC 61000 (qualification); ISO 14971 (risk-control verification); GAMP 5 (CSV). Cited by designation; revisions to confirm.
4. Background
This library is the single source of truth from which the Verification Plan, System Test Plan, and Verification/Traceability Matrices are derived. Every requirement carries a Verification Method; safety requirements (PRD-SAFE) are verified jointly with the Hazard Analysis (HAL, Stage 10); the D3 tension is closed by dedicated tests (MR-VAL-006).
5. Requirement Categories
Strategy · Traceability · Method allocation · Acceptance · Qualification tests · Validation linkage.
6. Detailed Requirements
PRD-VER-001 — Every requirement is verifiable and has a method
Requirement Statement | Every SCM product requirement shall carry a verification method (Test/Analysis/Inspection/Demonstration) and objective acceptance criteria.
Rationale | An unverifiable requirement cannot be accepted or defended.
Priority | Must
Source | Verification
Verification Method | Inspection
Parent Requirement | UN-VAL-###; ISO/IEC/IEEE 29148.
Traceability | → all PRD modules; → Verification Matrix (ORCH-014).
Risk Category | Compliance
Notes | Enforced by this library's 11-field block.
PRD-VER-002 — Bidirectional requirement traceability
Requirement Statement | The SCM verification program shall maintain bidirectional traceability from each requirement back to a user need/MR and forward to design output(s) and verification result(s).
Rationale | Traceability proves coverage and enables impact analysis and audit.
Priority | Must
Source | Verification / Regulatory
Verification Method | Inspection
Parent Requirement | UN-VAL-###; ISO 14971 / 29148.
Traceability | → requirement-rationale-traceability.md; → Traceability Matrix (ORCH-018 / EM-ENG-012).
Risk Category | Compliance
Notes | Matrix assembled by ORCH-018 / EM-ENG-012.
PRD-VER-003 — Method allocation appropriate to risk
Requirement Statement | The SCM verification program shall allocate the most rigorous appropriate method to safety- and function-critical requirements (Test preferred for safety-critical), justifying any use of Analysis/Inspection in their place.
Rationale | Severity-first (CLAUDE.md rule 3): critical requirements warrant the strongest evidence.
Priority | Must
Source | Verification / Safety
Verification Method | Analysis
Parent Requirement | PRD-SAFE; ISO 14971.
Traceability | → PRD-SAFE, PRD-REL; → Verification Plan (ORCH-014).
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-VER-004 — D3 endurance/charge validated by dedicated test
Requirement Statement | The SCM verification program shall include dedicated 90-day cold-endurance, cold-capacity, cycle-life, and charge-inhibit/charge-safety tests closing the D3 tension (MR-VAL-006).
Rationale | The top program risk (D3) must be closed by test, not analysis alone.
Priority | Must
Source | Verification (D3)
Verification Method | Test
Parent Requirement | PRD-REL-001/002/003, PRD-ELEC-002; D3; MR-VAL-006.
Traceability | → PRD-REL, PRD-ELEC, PRD-PERF-005; → Verification Plan / Validation.
Risk Category | Reliability
Notes | Headline validate-by-test.
PRD-VER-005 — RF-in-container qualification test
Requirement Statement | The SCM verification program shall include an RF-in-container test qualifying data-egress/alert delivery from inside representative metallized/insulated shippers.
Rationale | The principal open technical risk (RF attenuation; concept §7) must be qualified.
Priority | Must
Source | Verification
Verification Method | Test
Parent Requirement | PRD-CONN-007, PRD-ENV-009, PRD-PERF-011.
Traceability | → PRD-CONN, PRD-ENV; → Verification Plan.
Risk Category | Performance
Notes | Antenna strategy Evidence Pending; test informs it.
PRD-VER-006 — Environmental and regulatory qualification suite
Requirement Statement | The SCM verification program shall include the environmental (IEC 60068), ingress (IEC 60529), EMC/radio (CISPR/IEC 61000, FCC/CE), safety (IEC 62368-1), and transport (UN 38.3) qualification tests required for its markets.
Rationale | Market access and safety depend on the full qualification suite (D4).
Priority | Must
Source | Verification / Regulatory
Verification Method | Test
Parent Requirement | PRD-ENV, PRD-ELEC, PRD-REG; D4.
Traceability | → PRD-ENV/ELEC/REG; → Verification Plan; → certifications.
Risk Category | Compliance
Notes | Test houses Evidence Pending (concept §7).
PRD-VER-007 — Safety verification reconciled with HAL
Requirement Statement | The SCM verification of safety requirements (PRD-SAFE) shall be reconciled with the Hazard Analysis, each safety verification tracing to a HA-###/RCR-### record and demonstrating the risk control is effective.
Rationale | Safety requirements originate in and are closed against the HAL (Stage 10).
Priority | Must
Source | Verification / Safety
Verification Method | Analysis + Test
Parent Requirement | PRD-SAFE (all); ISO 14971.
Traceability | → HAL (HA-###/RCR-###, Stage 10); → Validation (PVaL).
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-VER-008 — Validation of intended use and needs
Requirement Statement | The SCM shall be validated against the intended use and user needs — including the offline-verdict usability study, detection/false-alarm validation, and audit-record simulation — to confirm it meets user needs in representative use.
Rationale | Verification proves requirements met; validation proves the right product was built.
Priority | Must
Source | Validation
Verification Method | Test (validation)
Parent Requirement | UN-VAL-###; UN-FUN/UX/SAFE.
Traceability | → Validation (PVaL, ORCH-015); → UN library; → ORCH-018.
Risk Category | Safety
Notes | Usability study is a critical validation (IEC 62366-1).
7. Verification Notes
This module governs the Verification Plan (PVL, ORCH-014), System Test Plan, and Verification/Traceability Matrices derived from this library. Safety verification reconciles with HAL (Stage 10); validation is executed under PVaL (ORCH-015). The D3 tests and RF-in-container test are the headline qualifications.
8. Traceability Placeholders
Upstream | UN-1700; whole PRD library; MR-VAL-006; D1–D4.
Downstream | Verification Plan (ORCH-014); Validation (ORCH-015); HAL (Stage 10); Traceability Matrix (ORCH-018 / EM-ENG-012).
PRD-ENV (200) — Use Environment
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001 (Reliability/Materials contributor). PRD-ENV-NNN, single "shall", 11-field block. Based on UN-400 (Use Environments), UN-1100 (Physical), MR-1300, and D3/D4. Forward to Stage 06, verification, and ORCH-018 / EM-ENG-012.
1. Purpose
Specify the environmental conditions the SCM shall remain functional, accurate, and safe within — the survival and usability envelope from which mechanical, electrical, and reliability requirements derive.
2. Scope
Thermal, humidity/condensation, mechanical (shock/vibration), ingress/cleaning, RF/connectivity, lighting, altitude/pressure, and EMC environments across the whole journey. Excludes qualification test severities (PRD-VER / verification stage) and enclosure implementation (design outputs).
3. Applicable Standards
IEC 60068 (environmental test conditions); IEC 60529 (IP ingress); CISPR / IEC 61000 (EMC); FCC Part 15 / CE RED (radio); IATA PCR/TCR, UN 38.3 (transport); WHO PQS E006 (cold-chain conditions). Cited by designation; revisions and exact severities to confirm (Evidence Pending) and set in verification.
4. Background
The device lives inside a metallized, thermally-isolated, RF-hostile shipper across −30 °C to +60 °C for up to 90 days (D3), then survives cold stores, aircraft holds, truck vibration, tropical humidity, dusty last-mile posts, and repeated depot cleaning (D1).
5. Requirement Categories
Thermal · Humidity/condensation · Shock/vibration · Ingress/cleaning · RF environment · Lighting · Altitude/pressure · EMC.
6. Detailed Requirements
PRD-ENV-001 — Operating temperature range
Requirement Statement | The SCM shall remain fully functional (sensing, logging, on-device display, and alerting when connected) across an operating temperature range of −30 °C to +60 °C.
Rationale | The cold chain spans deep cold to tropical heat; the record must not degrade at extremes.
Priority | Must
Source | User Need (D3) / Market
Verification Method | Test (IEC 60068-2-1/-2)
Parent Requirement | UN-ENV-001, UN-PERF-003.
Traceability | → PRD-REL-001, PRD-PERF-005; → PVL thermal chamber test.
Risk Category | Reliability
Notes | Top program risk (MR-RISK-001); Standard (IEC 60068, rev. to confirm).
PRD-ENV-002 — Storage and transport temperature range
Requirement Statement | The SCM shall survive, without permanent degradation, a non-operating storage/transport temperature range of at least −40 °C to +70 °C.
Rationale | Depot, warehouse, and transport extremes exceed operating conditions.
Priority | Must
Source | Market / Regulatory
Verification Method | Test (IEC 60068-2-1/-2)
Parent Requirement | UN-ENV-001; concept §4; MR-1300.
Traceability | → PRD-MECH, PRD-ELEC battery storage; → PVL storage test.
Risk Category | Reliability
Notes | Range is Judgment; confirm against battery datasheet (Evidence Pending).
PRD-ENV-003 — Cold display legibility
Requirement Statement | The SCM display shall remain legible (verdict readable) at temperatures down to the −30 °C operating extreme.
Rationale | An offline verdict must be readable at cold-store staging and receipt.
Priority | Must
Source | User Need
Verification Method | Test (cold-chamber legibility)
Parent Requirement | UN-ENV-002, UN-PHY-004.
Traceability | → PRD-UX-005, PRD-MECH display; → PVL cold legibility test.
Risk Category | Usability
Notes | OLED low-temp legibility thresholds Evidence Pending; flag display engineer.
PRD-ENV-004 — Humidity operating range
Requirement Statement | The SCM shall remain functional and within measurement accuracy across relative humidity from 10% to 95% non-condensing, and shall not sustain damage from occasional condensation during cold-to-warm transitions.
Rationale | Tropical lanes and freeze-thaw cycles must not corrupt readings or the device.
Priority | Must
Source | Market / Regulatory
Verification Method | Test (IEC 60068-2-30/-78)
Parent Requirement | UN-ENV-003.
Traceability | → PRD-MECH sealing, PRD-PERF accuracy; → PVL humidity/condensation test.
Risk Category | Reliability
Notes | Condensation on the temperature probe is a named concern. Standard (IEC 60068).
PRD-ENV-005 — Transport vibration survival
Requirement Statement | The SCM shall survive the random and sinusoidal vibration of air and road transport, per an IEC 60068-2-6/-64 profile to be set in verification, without failure or measurement error.
Rationale | Sustained transport vibration must not fail the device or shift the probe.
Priority | Must
Source | Market / Regulatory
Verification Method | Test (IEC 60068-2-6/-64)
Parent Requirement | UN-ENV-004, UN-PHY-006.
Traceability | → PRD-MECH ruggedness; → PVL vibration test.
Risk Category | Reliability
Notes | Profile severity Evidence Pending. Standard (IEC 60068).
PRD-ENV-006 — Mechanical shock / drop survival
Requirement Statement | The SCM shall survive handling drops and mechanical shock, per an IEC 60068-2-27/-31 profile to be set in verification, without loss of function or logged data.
Rationale | Drops and rough handling are routine in logistics and field use.
Priority | Must
Source | Market / Contextual Inquiry
Verification Method | Test (IEC 60068-2-27/-31)
Parent Requirement | UN-ENV-004, UN-PHY-006.
Traceability | → PRD-MECH-004, PRD-REL storage integrity; → PVL drop test.
Risk Category | Reliability
Notes | Drop height Evidence Pending. Standard (IEC 60068).
PRD-ENV-007 — Ingress protection rating
Requirement Statement | The SCM enclosure, including the charge port, shall meet an ingress protection rating of at least IP67 per IEC 60529.
Rationale | Dust and moisture (including immersion during cleaning) must not cause sensor drift, corrosion, or failure over the reusable life.
Priority | Must
Source | User Need / Regulatory
Verification Method | Test (IEC 60529)
Parent Requirement | UN-ENV-005, UN-PHY-007.
Traceability | → PRD-MECH-005/006 sealing; → PVL IP test.
Risk Category | Reliability
Notes | IP67 is a Judgment target; confirm against cleaning method (Evidence Pending). USB-C port seal is a named concern.
PRD-ENV-008 — Cleaning/decontamination chemical compatibility
Requirement Statement | The SCM external surfaces shall withstand repeated cleaning and decontamination with the agents specified in the service procedure over the device's full reuse life without material, seal, display, or label degradation.
Rationale | Reusable device is cleaned between every trip (D1); cleaning must not create ingress paths.
Priority | Must
Source | User Need / Regulatory
Verification Method | Test (cleaning-cycle aging)
Parent Requirement | UN-ENV-005, UN-PHY-008.
Traceability | → PRD-MECH materials, PRD-SVC-002 cleaning method; → PVL cleaning-durability test.
Risk Category | Reliability
Notes | Agent list Evidence Pending; flag materials engineer.
PRD-ENV-009 — RF operation inside the shipper
Requirement Statement | The SCM shall capture all data locally and deliver buffered data and alerts when a usable link is available even while operating inside a representative metallized, insulated shipping container.
Rationale | The device sits inside an RF-hostile enclosure where signals are attenuated; data must not be lost to poor RF.
Priority | Must
Source | User Need (D2) / Market
Verification Method | Test (RF-in-container)
Parent Requirement | UN-ENV-006, UN-PERF-008.
Traceability | → PRD-CONN-001/007, PRD-FUN-010 store-and-forward; → PVL RF-in-container test.
Risk Category | Performance
Notes | Antenna strategy / external-antenna option Evidence Pending (concept §7). Principal open risk.
PRD-ENV-010 — Intermittent connectivity/power tolerance at the last mile
Requirement Statement | The SCM shall continue logging and provide an on-device verdict with no live network connection and no external power for the full trip duration.
Rationale | The last mile has unreliable or absent connectivity and mains power (D2).
Priority | Must
Source | User Need (D2)
Verification Method | Test
Parent Requirement | UN-ENV-007, UN-FUN-011.
Traceability | → PRD-FUN-009/011, PRD-PERF-006 capacity; → PVL offline-trip test.
Risk Category | Reliability
Notes | Offline-first (D2); live data is upside, never dependency.
PRD-ENV-011 — Ambient-lighting legibility
Requirement Statement | The SCM on-device status shall be readable across ambient lighting from dim (≈10 lux) to bright/direct-sun (≥50,000 lux) conditions.
Rationale | Verdict must not be lost to glare or gloom at real posts and docks.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-ENV-008.
Traceability | → PRD-UX-005, PRD-MECH display; → PVL legibility test.
Risk Category | Usability
Notes | Lux thresholds Judgment; confirm in usability study.
PRD-ENV-012 — Altitude / cargo-hold pressure operation
Requirement Statement | The SCM shall operate and remain safe through the reduced-pressure conditions of an aircraft cargo hold (equivalent to at least 4,570 m / 15,000 ft cabin-altitude), including any pressure cycling.
Rationale | Air transport exposes the device and battery to reduced/cycling pressure.
Priority | Must
Source | Regulatory / Market
Verification Method | Test (altitude/UN 38.3 T.1)
Parent Requirement | UN-ENV-004; D4; UN 38.3.
Traceability | → PRD-ELEC battery, PRD-SAFE transport, PRD-REG-006; → PVL altitude test.
Risk Category | Safety
Notes | Standard (UN 38.3, rev. to confirm); altitude value per DGR.
7. Verification Notes
Environmental requirements are closed by IEC 60068 (thermal/humidity/vibration/shock), IEC 60529 (ingress), CISPR/IEC 61000 + FCC/CE (EMC/radio, see PRD-REG), cleaning-durability aging, and the dedicated RF-in-container and 90-day cold-endurance tests (MR-VAL-006). Severities are Evidence Pending until set in the Verification Plan.
8. Traceability Placeholders
Upstream | UN-400/1100; MR-1300/1100; concept §2/§3/§7; D3/D4.
Downstream | PRD-MECH, PRD-ELEC, PRD-CONN, PRD-REL, PRD-SAFE, PRD-REG, PRD-VER; HAL; ORCH-018.
PRD-ARCH (300) — Product Architecture
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001. PRD-ARCH-NNN, single "shall", 11-field block. Based on the concept subsystems, the Stage-04 needs, and D1–D4. Forward to Architecture (Stage 06), design outputs, and ORCH-018 / EM-ENG-012. See design-architecture.md.
1. Purpose
State the product-level architectural requirements — the subsystem structure, partitioning, and interface obligations that must hold for the design to realise the needs (offline-first, duty-cycled, trust-split).
2. Scope
System composition (device + cloud + mobile), subsystem existence and partitioning, trust boundaries, power-state governance, and interface conformance. Excludes detailed functions (PRD-FUN) and detailed electrical/mechanical design (PRD-ELEC/MECH).
3. Applicable Standards
ISO/IEC/IEEE 29148 (requirements/architecture); ISO 14971 (safe-state architecture); IEC 62443 concepts (security zoning). Cited by designation; revisions to confirm.
4. Background
The architecture is organized around capture→detect→warn→decide→prove under the 90-day/−30 °C constraint (D3), as an offline-first edge node with opportunistic connectivity (D2) and a split trust boundary (edge read vs platform governance, C9).
5. Requirement Categories
System composition · Subsystem partitioning · Power-state governance · Trust boundaries · Interface conformance · Extensibility.
6. Detailed Requirements
PRD-ARCH-001 — Three-tier system composition
Requirement Statement | The SCM system shall comprise a device tier, a cloud platform tier, and a mobile/BLE application tier, each able to perform its safety-relevant functions within its defined connectivity assumptions.
Rationale | Establishes the top-level structure and its independence assumptions.
Priority | Must
Source | Architecture
Verification Method | Inspection
Parent Requirement | UN-VIS; concept §4; D2.
Traceability | → SYS architecture; → design outputs; → block diagram.
Risk Category | None
Notes | See system-block-diagram.md. Evidence: Judgment.
PRD-ARCH-002 — Device subsystem set
Requirement Statement | The SCM device shall provide enclosure/mechanical, power, compute/firmware, sensing, connectivity, HMI, and non-volatile storage subsystems as defined in the block diagram.
Rationale | Ensures every concept subsystem is realised and allocated.
Priority | Must
Source | Architecture
Verification Method | Inspection
Parent Requirement | concept §4 subsystems.
Traceability | → PRD-MECH/ELEC/FUN/CONN/UX; → design outputs.
Risk Category | None
Notes | Evidence: Standard (concept).
PRD-ARCH-003 — Edge-autonomous core function
Requirement Statement | The SCM device shall perform sensing, logging, excursion detection, and on-device verdict entirely on the device without dependence on cloud or mobile connectivity.
Rationale | Offline-first (D2); the last mile must never be a black box.
Priority | Must
Source | Architecture / User Need (D2)
Verification Method | Test
Parent Requirement | UN-FUN-009/011, UN-ENV-007; D2.
Traceability | → PRD-FUN-001..009; → PVL offline test.
Risk Category | Safety
Notes | AD-1. Evidence: Standard (D2).
PRD-ARCH-004 — Duty-cycled power-state governor
Requirement Statement | The SCM device shall implement a power-state architecture that keeps the device in low-power states between sampling and upload activities to meet the 90-day energy budget.
Rationale | The energy budget, not chemistry alone, delivers 90-day endurance at −30 °C (D3).
Priority | Must
Source | Architecture / Reliability (D3)
Verification Method | Analysis + Test
Parent Requirement | UN-PERF-003; D3; power budget (Stage 00).
Traceability | → PRD-ELEC-008 power budget, PRD-PERF-005; → PVL endurance test.
Risk Category | Reliability
Notes | AD-2. Validated by test (top risk).
PRD-ARCH-005 — Multi-bearer connectivity, LTE-M primary
Requirement Statement | The SCM device shall support LTE-M as the primary bearer and Wi-Fi and BLE as additional bearers, selecting opportunistically to move buffered data and alerts within the energy budget.
Rationale | Global lanes and RF-hostile shippers demand opportunistic multi-bearer offload (D2/D4).
Priority | Must
Source | Architecture / User Need
Verification Method | Test
Parent Requirement | UN-PERF-008, UN-FUN-010; D2/D4.
Traceability | → PRD-CONN-001..004; → PVL connectivity test.
Risk Category | Performance
Notes | AD-1. Bearer priority configurable.
PRD-ARCH-006 — Split trust boundary (edge read vs platform governance)
Requirement Statement | The SCM architecture shall separate a read-only, login-free on-device verdict trust boundary from a controlled configuration/data trust boundary enforced by the cloud platform.
Rationale | Frictionless offline edge use must coexist with enterprise/Part 11 governance (C9).
Priority | Must
Source | Architecture / Security
Verification Method | Analysis + Test
Parent Requirement | UN-SEC-003/004; conflict C9.
Traceability | → PRD-SEC-005..007, PRD-UX-001; → PVL security review.
Risk Category | Security
Notes | AD-5. To be reconciled with cybersecurity threat model (Evidence Pending).
PRD-ARCH-007 — Trusted time base
Requirement Statement | The SCM device shall maintain a trusted, monotonic time base for timestamping that persists across sleep, power loss, and reset, and shall flag any period of time-base uncertainty.
Rationale | Every record and verdict depends on trustworthy time; ALCOA+ "contemporaneous".
Priority | Must
Source | Architecture / Regulatory
Verification Method | Test
Parent Requirement | UN-FUN-016, UN-SEC-002, UN-REG-003.
Traceability | → PRD-FUN-016, PRD-SEC-002; → PVL clock test.
Risk Category | Compliance
Notes | RTC + time-sync on connect; drift budget Evidence Pending.
PRD-ARCH-008 — Safe-state architecture
Requirement Statement | The SCM device shall define and enter a known safe state (data preserved, charge inhibited if unsafe, monitoring status honest) on any power, firmware, connectivity, or reset fault.
Rationale | A fault must never corrupt the record or leave the device unsafe or falsely reassuring.
Priority | Must
Source | Architecture / Safety
Verification Method | Test (fault injection)
Parent Requirement | UN-SAFE-010, UN-SAFE-003; ISO 14971.
Traceability | → PRD-SAFE-010, PRD-REL-004; → HAL; → PVL fault-injection.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10). AD-7.
PRD-ARCH-009 — Interface conformance to the block diagram
Requirement Statement | Each subsystem interface shall conform to the typed interface inventory (power/data/RF/thermal/mechanical) defined in the system block diagram.
Rationale | Prevents interface drift and integration defects across disciplines.
Priority | Must
Source | Architecture
Verification Method | Inspection
Parent Requirement | block diagram I-01..I-12.
Traceability | → interface control docs (Stage 06); → design outputs.
Risk Category | None
Notes | RF-in-container interface (I-08) is the principal open risk.
PRD-ARCH-010 — Configuration-driven variant strategy
Requirement Statement | The SCM shall realise product/market variants through configuration (trip profiles, regional radio/regulatory profiles) rather than hardware forks wherever feasible.
Rationale | One reusable fleet serves many products/markets; limits SKU proliferation (PRD-VIS-007).
Priority | Should
Source | Architecture / Business
Verification Method | Demonstration
Parent Requirement | PRD-VIS-007, UN-FUN-011.
Traceability | → PRD-FUN-012, PRD-CLOUD config; → PVaL variant scenario.
Risk Category | None
Notes | Radio variants may still require per-market hardware/certification (PRD-REG-004).
7. Verification Notes
Architecture requirements are verified by design review/inspection against the block diagram, and by the endurance (ARCH-004), offline (ARCH-003), security (ARCH-006), clock (ARCH-007), and fault-injection (ARCH-008) tests that close the safety/reliability-relevant decisions.
8. Traceability Placeholders
Upstream | concept §4; UN-FUN/PERF/SEC/SAFE; D1–D4; power budget (Stage 00).
Downstream | all device/cloud modules; HAL (Stage 10); ORCH-018 / EM-ENG-012.
PRD-FUN (400) — Functional Requirements
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: EM-ENG-001 with Embedded Software engineer. PRD-FUN-NNN, single "shall", atomic/measurable/verifiable, 11-field block. Based on UN-600 (Functional), UN-500 (Workflow), and D1–D4. Forward to SyRS/SRS (ORCH-012), Stage 06, and ORCH-018.
1. Purpose
Specify what the SCM shall do — the testable functions realising capture → detect → warn → decide → prove.
2. Scope
Sensing/logging, two-sided excursion detection, alerting/routing, on-device verdict, store-and-forward, configuration/provisioning/arming, record/audit/export, time/location attribution, fleet readiness, and firmware update behaviour. Excludes performance targets (PRD-PERF), safety controls (PRD-SAFE), security mechanisms (PRD-SEC), UX presentation (PRD-UX).
3. Applicable Standards
WHO PQS E006 (monitor/verdict); EU GDP + 21 CFR Part 11 + ALCOA+ (record); ISO 14971 (detect/alert as risk controls); ISO/IEC 17025 concepts (calibration reporting). Cited by designation; revisions to confirm.
4. Background
D2 shapes the store-and-forward functions; two-sided detection is central (freeze is silent); the on-device verdict exists because the last-mile user has no phone, network, or training.
5. Requirement Categories
Sense & log · Detect & alert · Decide (verdict) · Store & forward · Configure/provision/arm · Record & export · Locate & attribute · Fleet, readiness & OTA.
6. Detailed Requirements
6.1 Sense & log
PRD-FUN-001 — Continuous interval temperature logging
Requirement Statement | The SCM shall measure and record temperature at a configurable interval (default 5 min, range 1–60 min) continuously for the armed trip duration with no unrecorded interval.
Rationale | A complete, defensible thermal history requires gap-free interval logging (D2).
Priority | Must
Source | User Need / Regulatory
Verification Method | Test
Parent Requirement | UN-FUN-001; STN-005.
Traceability | → PRD-PERF-006 capacity, PRD-REL storage; → PVL logging test.
Risk Category | Compliance
Notes | Default 5–10 min (D2); any gap is flagged, never hidden (PRD-FUN-016).
PRD-FUN-002 — Dual-channel temperature (internal + external probe)
Requirement Statement | The SCM shall measure temperature on both an internal sensor and an external probe channel and record each channel distinctly.
Rationale | The probe represents the product zone; internal supports device/ambient context.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-003; STN-002/024.
Traceability | → PRD-MECH probe, PRD-PERF-001 accuracy; → PVL dual-channel test.
Risk Category | Safety
Notes | External probe underpins freeze detection.
PRD-FUN-003 — Humidity logging
Requirement Statement | The SCM shall measure and record relative humidity at the configured logging interval across 10–95% RH.
Rationale | Humidity exposure affects some products and indicates condition.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-002.
Traceability | → PRD-ENV-004, PRD-PERF accuracy; → PVL humidity test.
Risk Category | Performance
Notes | Accuracy target set in PRD-PERF.
PRD-FUN-004 — Shock/impact event capture
Requirement Statement | The SCM shall detect and record shock/impact events exceeding a configurable acceleration threshold, with timestamp and peak magnitude, without logging routine handling.
Rationale | Product-relevant mishandling must be captured without alarm-fatigue noise (C8).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-002; STN-006.
Traceability | → PRD-ENV-005/006, PRD-FUN-016; → PVL shock test.
Risk Category | Performance
Notes | Threshold configurable per profile; ties C8 balance.
6.2 Detect & alert
PRD-FUN-005 — Configurable two-sided thresholds
Requirement Statement | The SCM shall support configurable independent low (freeze) and high (heat) temperature thresholds per trip profile.
Rationale | Freeze- and heat-sensitive products need distinct, product-specific limits.
Priority | Must
Source | User Need / Safety
Verification Method | Test
Parent Requirement | UN-FUN-004/011; STN-002.
Traceability | → PRD-FUN-012 profiles, PRD-SAFE-002; → PVL threshold test.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-FUN-006 — Two-sided excursion detection (equal reliability)
Requirement Statement | The SCM shall detect freezing excursions with reliability equal to overheating excursions and flag each in the trip record.
Rationale | Freeze damage is silent and under-detected; must not be a lesser capability.
Priority | Must
Source | Safety / User Need
Verification Method | Test
Parent Requirement | UN-FUN-004, UN-SAFE-002; STN-002.
Traceability | → PRD-VIS-003, PRD-SAFE-002; → HAL missed-freeze hazard; → PVaL detection validation.
Risk Category | Safety
Notes | Cited to WHO/UNICEF freezing reporting. To be reconciled with HAL.
PRD-FUN-007 — Product-relevant excursion logic
Requirement Statement | The SCM shall determine an excursion using configurable product-stability logic (threshold with time-over/under-threshold and/or cumulative exposure) rather than instantaneous transients alone.
Rationale | Alerts and verdicts must reflect actual product harm, not momentary noise (C8).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-005; STN-007.
Traceability | → PRD-SAFE-004 alarm balance, PRD-FUN-012; → PVaL false-alarm/miss validation.
Risk Category | Safety
Notes | Balances over- vs under-sensitivity (C8).
PRD-FUN-008 — Immediate excursion alert when connected
Requirement Statement | The SCM shall issue an excursion alert to the cloud within the PRD-PERF-004 latency target when connectivity is available at the time of detection, and shall queue the alert for delivery on next connection otherwise.
Rationale | Learning of a break in time to intervene converts monitoring into prevented loss (D2).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-006, UN-PERF-005; STN-001; D2.
Traceability | → PRD-PERF-004 latency, PRD-CLOUD-003 routing; → HAL delayed-alert hazard.
Risk Category | Safety
Notes | Latency connectivity-bound; offline-captured. To be reconciled with HAL.
PRD-FUN-009 — Alert routing and escalation
Requirement Statement | The SCM cloud platform shall route each excursion alert to the responder(s) configured for the shipment/role and escalate it if unacknowledged within a configurable time.
Rationale | A warning must reach someone who can act; unacknowledged alerts must escalate.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-007; STN-018.
Traceability | → PRD-CLOUD-003/004; → PVL routing/escalation test.
Risk Category | Safety
Notes | Acknowledgement tracked; configurable by role.
6.3 Decide (on-device verdict)
PRD-FUN-010 — Offline on-device accept/reject/quarantine verdict
Requirement Statement | The SCM shall compute and display on the device an unambiguous accept / reject / quarantine verdict for the trip, readable with no phone, network, login, or manual.
Rationale | The highest-consequence edge decision is made by a non-technical, offline user.
Priority | Must
Source | User Need / Safety
Verification Method | Test (summative usability)
Parent Requirement | UN-FUN-008, UN-SAFE-005; STN-004/027.
Traceability | → PRD-UX-001/002, PRD-ENV-003 legibility; → HAL misread-verdict hazard; → PVaL.
Risk Category | Safety
Notes | PQS-consistent (category TBC); layered vs full record (C5). To be reconciled with HAL.
PRD-FUN-011 — Verdict integrity vs record
Requirement Statement | The SCM shall derive the on-device verdict from the same recorded data and active trip profile as the cloud record, such that device verdict and record disposition never disagree.
Rationale | A device saying "accept" while the record says "reject" would be catastrophic.
Priority | Must
Source | Safety / User Need
Verification Method | Test
Parent Requirement | UN-FUN-008/014; UN-SAFE-001.
Traceability | → PRD-FUN-010/016, PRD-CLOUD-002; → PVaL cross-check.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
6.4 Store & forward
PRD-FUN-012 — Offline buffering for full trip
Requirement Statement | The SCM shall buffer all logged data locally for the full 90-day trip when offline and upload it in intact chronological order when a link becomes available, with no data loss.
Rationale | An entirely offline trip must still yield a complete record (D2).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-009, UN-PERF-006; STN-009; D2.
Traceability | → PRD-PERF-006 capacity, PRD-REL storage; → PVL offline-then-upload test.
Risk Category | Reliability
Notes | No-data-loss target closed by test (MR-VAL-006).
PRD-FUN-013 — Scheduled cadence upload with multi-bearer offload
Requirement Statement | The SCM shall upload buffered data on a configurable cadence (default hourly, range 15 min–24 h) using the available bearer (LTE-M/Wi-Fi/BLE) selected to conserve the energy budget.
Rationale | Data must reach the cloud without draining the 90-day budget (D2, C2).
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-010, UN-PERF-008; D2.
Traceability | → PRD-CONN-001..004, PRD-ELEC-008 budget; → PVL cadence test.
Risk Category | Performance
Notes | Cadence default hourly; energy budget is the constraint.
6.5 Configure, provision & arm
PRD-FUN-014 — Trip-profile configuration
Requirement Statement | The SCM shall allow an authorized user to define and select trip profiles (thresholds, excursion logic, logging interval, shock threshold) and shall record which profile is active.
Rationale | Monitoring must match what harms the specific product; the active profile is part of the record.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-011; STN-004/007.
Traceability | → PRD-FUN-005/007, PRD-MOB-002, PRD-SEC-005; → PVL config test.
Risk Category | Safety
Notes | Error-proof selection; sane defaults; wrong profile must not be silently applied (C4).
PRD-FUN-015 — Arm with positive placement confirmation
Requirement Statement | The SCM shall require an explicit arm action and shall give positive on-device confirmation that it is armed, has a valid profile, and (where detectable) is correctly placed before monitoring is credited as started.
Rationale | A trip must never silently start unmonitored (forgot-to-arm).
Priority | Must
Source | User Need / Safety
Verification Method | Demonstration
Parent Requirement | UN-FUN-012, UN-SAFE-006; STN-008.
Traceability | → PRD-UX-004, PRD-STK-002; → HAL void-trip hazard; → PVL arming test.
Risk Category | Safety
Notes | On-device fallback if BLE/app unavailable. To be reconciled with HAL.
PRD-FUN-016 — Placement guidance
Requirement Statement | The SCM shall provide guidance (on-device and/or app) for correct device and probe placement appropriate to the selected profile/shipper.
Rationale | Wrong placement biases detection and the verdict.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-013; STN-024.
Traceability | → PRD-MECH mounting, PRD-UX; → PVaL placement scenario.
Risk Category | Usability
Notes | Internal vs external probe guidance.
6.6 Record, prove & export
PRD-FUN-017 — Complete, tamper-evident, attributable record with gap flagging
Requirement Statement | The SCM shall produce, for each trip, a complete record that is tamper-evident, attributable to device and configuration, and that explicitly flags any monitoring gap rather than concealing it.
Rationale | A disposition must be defensible to an inspector; hidden gaps are indefensible.
Priority | Must
Source | Regulatory / User Need
Verification Method | Analysis + Test
Parent Requirement | UN-FUN-014, UN-SEC-001; STN-005; 21 CFR Part 11 / ALCOA+.
Traceability | → PRD-SEC-001/002, PRD-REG-003, PRD-ARCH-007; → PVaL audit simulation.
Risk Category | Compliance
Notes | ALCOA+ aligned. Overlaps PRD-SEC-001 (integrity mechanism side).
PRD-FUN-018 — On-demand record export
Requirement Statement | The SCM cloud platform shall export a trip's record on demand in a shareable, integrity-preserving format (human-readable and machine-readable) suitable for partners, auditors, and quality systems.
Rationale | Evidence must move where the decision or dispute is, without losing integrity.
Priority | Must
Source | Regulatory / User Need
Verification Method | Test
Parent Requirement | UN-FUN-015; STN-005/013.
Traceability | → PRD-CLOUD-005 export/API, PRD-SEC-002; → PVL export test.
Risk Category | Compliance
Notes | Export preserves attribution and tamper-evidence.
6.7 Locate & attribute
PRD-FUN-019 — Time- and location-stamped events
Requirement Statement | The SCM shall time-stamp every sample and event against the trusted time base and shall associate excursion and custody events with a location fix where one is available within the energy budget.
Rationale | Attribution of when/where a break occurred supports dispute resolution at hand-offs.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-016, UN-PERF-009; STN-006.
Traceability | → PRD-ARCH-007 time base, PRD-PERF-007 location; → PVL geo-stamp test.
Risk Category | Performance
Notes | Location within energy budget (C3); GPS-in-container performance Evidence Pending.
6.8 Fleet, readiness & OTA
PRD-FUN-020 — Device readiness reporting
Requirement Statement | The SCM shall report charge state, calibration status/date, and firmware version to the fleet, and shall indicate when it is not fit to deploy (out-of-calibration or under-charged).
Rationale | Returned devices must be triaged; unfit units must not be redeployed unnoticed.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-017; STN-010/021.
Traceability | → PRD-STK-003, PRD-SVC-004, PRD-CLOUD-006; → PVL readiness test.
Risk Category | Reliability
Notes | Deploy blocked if out-of-cal.
PRD-FUN-021 — Fleet visibility
Requirement Statement | The SCM cloud platform shall present, for a fleet, each device/shipment status, last-known location, and required action.
Rationale | A fleet must be managed and strays chased.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-018; STN-014/015.
Traceability | → PRD-CLOUD-001/002 dashboard; → PVaL fleet scenario.
Risk Category | None
Notes | Liveness connectivity-bound (C7); location energy-bound (C3).
PRD-FUN-022 — Firmware update behaviour
Requirement Statement | The SCM shall accept firmware updates over the air and via USB-C, applying an update only after integrity/authenticity verification, resuming an interrupted update, and preserving logged data and monitoring state across the update.
Rationale | Updates must be genuine, resumable, and non-destructive to the fleet and the record.
Priority | Must
Source | User Need / Security
Verification Method | Test
Parent Requirement | UN-SEC-005, UN-SAFE-010; D2 (OTA).
Traceability | → PRD-SEC-008 signed OTA, PRD-ARCH-008 safe-state, PRD-CLOUD-008; → PVL OTA test.
Risk Category | Security
Notes | Signing/keys in PRD-SEC; fail-safe in PRD-SAFE-010.
7. Verification Notes
Functional requirements are verified by test (logging, detection, alerting, buffering, arming, OTA), by validation scenarios (offline verdict, false-alarm/miss, audit export), and — for the safety-relevant functions (FUN-006/007/008/010/011/015) — jointly with the Hazard Analysis (HAL, Stage 10).
8. Traceability Placeholders
Upstream | UN-600/500; STN-001..027; D1/D2/D3.
Downstream | PRD-PERF, PRD-SAFE, PRD-SEC, PRD-UX, PRD-CONN, PRD-CLOUD, PRD-SVC; SRS (ORCH-012); HAL; ORCH-018 / EM-ENG-012.
PRD-UX (500) — User Experience
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Human-Factors/Usability engineer with EM-ENG-001. PRD-UX-NNN, single "shall", 11-field block. Based on UN-700 (UX), UN-500 (Workflow), UN-200 (Personas), IEC 62366-1, and D1/D2. Forward to Stage 06, usability engineering file, and ORCH-018.
1. Purpose
Specify the on-device and app user-experience the SCM shall deliver so that a minimally-trained, gloved, offline user reaches a correct disposition, and so that use errors do not lead to harm.
2. Scope
On-device verdict presentation, indicators/controls, feedback/response, onboarding/arming UX, error messages, accessibility, localization, and low-literacy comprehension. Excludes functions (PRD-FUN), display hardware (PRD-MECH), and mobile-app functions (PRD-MOB).
3. Applicable Standards
IEC 62366-1 (usability engineering / use-related risk); ISO 9241 concepts (interaction); WHO PQS E006 (user comprehension). Cited by designation; revisions to confirm.
4. Background
The edge user is the least-trained, most safety-critical operator (UN-SAFE-005): no phone, no login, gloves, poor light, cold. UX is therefore a safety instrument (misread verdict → wrong disposition).
5. Requirement Categories
Verdict comprehension · Controls & feedback · Onboarding/arming · Error prevention & messaging · Accessibility & localization · Confidence/reassurance.
6. Detailed Requirements
PRD-UX-001 — Unambiguous verdict comprehension without training
Requirement Statement | The SCM shall present the trip verdict using distinct, non-color-dependent symbols/text such that a first-time, untrained user correctly interprets accept vs reject vs quarantine in a summative usability study meeting the acceptance criterion set in the usability plan.
Rationale | Misreading the verdict is the highest-consequence use error.
Priority | Must
Source | User Need / Safety
Verification Method | Test (summative usability)
Parent Requirement | UN-UX-###, UN-FUN-008, UN-SAFE-005; STN-004/027.
Traceability | → PRD-FUN-010, PRD-STK-001; → HAL misread hazard; → PVaL usability study.
Risk Category | Usability
Notes | Non-color-dependent for accessibility. Pass criterion Evidence Pending. To reconcile with HAL.
PRD-UX-002 — No-login, no-manual offline read
Requirement Statement | The SCM shall make the verdict readable on-device with no login, account, network, companion device, or reference to a manual.
Rationale | The edge use context has none of these available.
Priority | Must
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-UX-###, UN-SEC-004; conflict C9.
Traceability | → PRD-ARCH-006 trust split, PRD-FUN-010; → PVaL edge scenario.
Risk Category | Usability
Notes | Trust-boundary split (AD-5).
PRD-UX-003 — Single-action, glove-operable control
Requirement Statement | The SCM shall enable its primary interactions (wake/read, arm) with a single physical action operable with gloved, cold hands.
Rationale | Fiddly controls cause mis-operation in the cold/field.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-UX-###, UN-PHY-005; STN-008.
Traceability | → PRD-MECH-007 controls; → PVaL gloved-usability test.
Risk Category | Usability
Notes | Dexterity thresholds Evidence Pending.
PRD-UX-004 — Positive arming feedback
Requirement Statement | The SCM shall give unambiguous positive feedback (visual and, where present, indicator) confirming armed + profile-valid state, and shall make an un-armed/void state equally obvious.
Rationale | The user must never believe a trip is monitored when it is not.
Priority | Must
Source | User Need / Safety
Verification Method | Demonstration
Parent Requirement | UN-UX-###, UN-FUN-012, UN-SAFE-006; STN-008.
Traceability | → PRD-FUN-015; → HAL void-trip hazard; → PVL arming test.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-UX-005 — Legibility in cold, glare, and low light
Requirement Statement | The SCM display shall present the verdict legibly across the operating temperature range and the ambient-lighting range of PRD-ENV-011.
Rationale | Verdict must be readable in freezer, dim post, and direct sun.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-ENV-002/008, UN-PHY-004.
Traceability | → PRD-ENV-003/011, PRD-MECH display; → PVL legibility test.
Risk Category | Usability
Notes | OLED cold-legibility thresholds Evidence Pending.
PRD-UX-006 — Prompt interaction response
Requirement Statement | The SCM shall present the requested status/verdict within the PRD-PERF-006 responsiveness target after a user action, including at low temperature and low battery.
Rationale | A time-pressured user must not be left waiting or unsure.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-UX-###, UN-PERF-006; STN-004/008.
Traceability | → PRD-PERF-006; → PVL response-time test.
Risk Category | Usability
Notes | Response target set in PRD-PERF.
PRD-UX-007 — Error-preventing configuration
Requirement Statement | The SCM configuration UX shall present a valid default profile and prevent arming with an invalid or missing profile, requiring explicit confirmation for non-default choices.
Rationale | Wrong-profile selection causes missed excursions or false alarms (C4).
Priority | Must
Source | User Need / Safety
Verification Method | Test
Parent Requirement | UN-UX-###, UN-FUN-014, UN-SAFE-006.
Traceability | → PRD-FUN-014, PRD-MOB-002; → PVaL config-error scenario.
Risk Category | Safety
Notes | Poka-yoke selection. To be reconciled with HAL.
PRD-UX-008 — Actionable error/status messaging
Requirement Statement | The SCM shall present device faults, low-battery, out-of-calibration, and monitoring-gap conditions as clear, actionable status rather than silent failure or generic codes.
Rationale | Honest, actionable status prevents false reassurance and mis-operation.
Priority | Must
Source | User Need / Safety
Verification Method | Demonstration
Parent Requirement | UN-UX-###, UN-SAFE-003, UN-FUN-020.
Traceability | → PRD-SAFE-003, PRD-ELEC low-battery; → PVL status-message review.
Risk Category | Safety
Notes | Ties no-false-reassurance (PRD-SAFE-003).
PRD-UX-009 — Low-literacy / language-independent edge cues
Requirement Statement | The SCM on-device verdict shall be comprehensible using language-independent cues (symbols/iconography) so a low-literacy user in any market reads it correctly.
Rationale | The global edge user may not read the configured language.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-UX-###, UN-SAFE-005; MR-PER-007.
Traceability | → PRD-UX-001; → PVaL cross-market usability.
Risk Category | Usability
Notes | Complements localization (UX-010).
PRD-UX-010 — App/dashboard localization
Requirement Statement | The SCM mobile app and cloud dashboard shall support localization of user- facing text and units for the target markets.
Rationale | Multi-market deployment (D4) requires localized enterprise/config UX.
Priority | Should
Source | Business / User Need
Verification Method | Inspection
Parent Requirement | UN-UX-###; D4.
Traceability | → PRD-MOB, PRD-CLOUD dashboard; → localization review.
Risk Category | None
Notes | Target language set Evidence Pending.
PRD-UX-011 — Layered information (verdict first, detail on demand)
Requirement Statement | The SCM shall present the accept/reject/quarantine verdict as the primary, immediate information and make detailed history/record available only on demand (device summary and cloud).
Rationale | The edge decision needs a simple verdict; analysts need the full record (C5).
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-VIS-007, UN-UX-###; conflict C5.
Traceability | → PRD-FUN-010/017, PRD-CLOUD dashboard; → PVaL.
Risk Category | Usability
Notes | AD-6 layered information.
PRD-UX-012 — Passive monitoring reassurance
Requirement Statement | The SCM shall provide a passive, glanceable indication that it is armed and actively monitoring, without requiring user action.
Rationale | Users need confidence that "armed" means "monitoring" throughout the trip.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-UX-###, UN-PERF-011.
Traceability | → PRD-PERF-009 availability, PRD-ELEC indicator; → PVaL.
Risk Category | Usability
Notes | Energy-cheap indication (must not burden budget).
7. Verification Notes
UX requirements are verified primarily by summative usability study (IEC 62366-1) with the critical tasks (read verdict, arm, configure), plus legibility/response tests and localization inspection. Safety-relevant UX (UX-001/004/007/008) is cross-validated with the Hazard Analysis (HAL, Stage 10).
8. Traceability Placeholders
Upstream | UN-700/500/200; IEC 62366-1; STN-004/008/027; D2.
Downstream | PRD-FUN, PRD-MECH, PRD-MOB, PRD-CLOUD, PRD-SAFE; usability engineering file; HAL; ORCH-018.
PRD-MECH (600) — Mechanical Requirements
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Mechanical/Structural engineer with EM-ENG-001. PRD-MECH-NNN, single "shall", 11-field block. Based on UN-1100 (Physical), UN-400 (Environments), IEC 60529/60068, and D1. Forward to Stage 06, mechanical design outputs, and ORCH-018.
1. Purpose
Specify the physical, structural, sealing, thermal-coupling, and human-fit mechanical requirements of the SCM device and probe.
2. Scope
Form factor, mass, mounting/retention, external probe, display/control housing, sealing/ingress, ruggedness, materials/cleanability, connectors, labeling, and packaging interface. Excludes environmental conditions (PRD-ENV), electrical design (PRD-ELEC), and detailed tolerances (design outputs).
3. Applicable Standards
IEC 60529 (IP); IEC 60068 (mechanical robustness); WHO PQS E006 (physical/label expectations); RoHS/REACH (materials). Cited by designation; revisions to confirm.
4. Background
The device lives inside insulated shippers, is embeddable by OEMs, handled gloved in the cold, cleaned repeatedly, and reused for years (D1). Ruggedness, sealing, cleanability, and durable labeling are requirements, not niceties.
5. Requirement Categories
Form/fit/mount · Probe & thermal coupling · Display/control housing · Sealing/ingress · Ruggedness · Materials/cleanability · Connectors · Labeling/packaging.
6. Detailed Requirements
PRD-MECH-001 — Fit inside standard insulated shippers
Requirement Statement | The SCM device shall fit within the internal volume of the target insulated shippers/cold boxes without displacing payload beyond an allowance set in the product spec.
Rationale | It must be placed with the product across many shipper sizes.
Priority | Must
Source | User Need
Verification Method | Inspection (fit check)
Parent Requirement | UN-PHY-001; STN-024.
Traceability | → PRD-STK-006 embeddable, PRD-VIS-007; → PVL fit check.
Risk Category | Usability
Notes | Envelope dimensions Evidence Pending (multiple shipper sizes).
PRD-MECH-002 — Mass limit
Requirement Statement | The SCM device (with battery, excluding external packaging) shall not exceed a mass limit set in the product spec, targeted so it does not meaningfully displace payload or burden handlers.
Rationale | Weight burdens handling and payload economics.
Priority | Should
Source | User Need
Verification Method | Inspection
Parent Requirement | UN-PHY-009; STN-024.
Traceability | → PRD-ELEC battery size; → PVL mass check.
Risk Category | None
Notes | Target mass Evidence Pending (Judgment).
PRD-MECH-003 — Mounting and retention
Requirement Statement | The SCM shall provide a defined means to mount and retain the device and probe in position within the shipper such that they do not shift under the transport vibration/shock of PRD-ENV.
Rationale | A shifted device/probe gives unrepresentative readings and can be lost in the box.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-PHY-002; STN-024.
Traceability | → PRD-ENV-005/006, PRD-FUN-016; → PVL retention test.
Risk Category | Safety
Notes | Ties placement guidance (PRD-FUN-016).
PRD-MECH-004 — External temperature probe
Requirement Statement | The SCM shall provide an external temperature probe with a lead length and tip form suitable to place in the product zone, sealed and strain-relieved against handling and cleaning.
Rationale | The probe must represent the vaccine's actual thermal zone reliably over reuse.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-003, UN-PHY-002.
Traceability | → PRD-FUN-002 dual-channel, PRD-PERF-001; → PVL probe test.
Risk Category | Safety
Notes | Lead length Evidence Pending; strain-relief and seal named concerns.
PRD-MECH-005 — Ingress-protected enclosure
Requirement Statement | The SCM enclosure shall be sealed to at least IP67 (IEC 60529), including all seams and the display window.
Rationale | Dust/moisture must not reach the electronics or sensors over the reusable life.
Priority | Must
Source | User Need / Regulatory
Verification Method | Test (IEC 60529)
Parent Requirement | UN-PHY-007, UN-ENV-005.
Traceability | → PRD-ENV-007; → PVL IP test.
Risk Category | Reliability
Notes | IP67 Judgment target; confirm vs cleaning immersion.
PRD-MECH-006 — Sealed charge/service port
Requirement Statement | The SCM USB-C port shall be sealed (integral cover/gasket or sealed connector) to maintain the enclosure IP rating when closed.
Rationale | The charge port is the most likely ingress path over repeated use/cleaning.
Priority | Must
Source | User Need
Verification Method | Test (IEC 60529)
Parent Requirement | UN-PHY-007; STN-026.
Traceability | → PRD-ELEC-005 USB-C, PRD-ENV-007; → PVL port-seal test.
Risk Category | Reliability
Notes | Named concern; flag materials engineer.
PRD-MECH-007 — Rugged housing
Requirement Statement | The SCM housing shall protect the electronics, display, probe interface, and battery through the drop/shock/vibration of PRD-ENV-005/006 without functional damage or measurement error.
Rationale | Ruggedness sustains reuse economics (fewer write-offs) and record integrity.
Priority | Must
Source | User Need
Verification Method | Test (IEC 60068)
Parent Requirement | UN-PHY-006, UN-ENV-004.
Traceability | → PRD-ENV-005/006, PRD-REL; → PVL drop/vibe test.
Risk Category | Reliability
Notes | Standard (IEC 60068).
PRD-MECH-008 — Glove-operable control and readable display
Requirement Statement | The SCM shall provide a physical control operable with gloved, cold hands and a display sized/positioned to be read at arm's length at the point of use.
Rationale | Field/cold-store ergonomics for the edge user.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-PHY-004/005; STN-008.
Traceability | → PRD-UX-003/005, PRD-ENV-003; → PVaL gloved usability.
Risk Category | Usability
Notes | Button size/force and display size Evidence Pending.
PRD-MECH-009 — Cleanable, compatible materials
Requirement Statement | The SCM external materials shall be compatible with the specified cleaning/ decontamination agents and shall not degrade, craze, or lose sealing over the reuse-life cleaning cycles.
Rationale | Repeated cleaning (D1) must not create ingress or degrade the device.
Priority | Must
Source | User Need / Regulatory
Verification Method | Test (cleaning aging)
Parent Requirement | UN-PHY-008, UN-ENV-005.
Traceability | → PRD-ENV-008, PRD-SVC-002; → PVL cleaning-durability test.
Risk Category | Reliability
Notes | Agent list Evidence Pending; RoHS/REACH compliant materials (PRD-REG-008).
PRD-MECH-010 — Durable identification and safety labeling
Requirement Statement | The SCM shall carry durable, legible device identification (serial/UDI-style) and required safety/transport markings that remain legible over the reuse life and cleaning cycles.
Rationale | Attribution, fleet management, and regulatory acceptance depend on durable labeling.
Priority | Must
Source | User Need / Regulatory
Verification Method | Test (label durability)
Parent Requirement | UN-PHY-009, UN-REG-010; STN-012.
Traceability | → PRD-REG-010 labeling, PRD-SEC-002 identity, PRD-MFG-004 serialization; → PVL.
Risk Category | Compliance
Notes | Marking set Evidence Pending (regulatory stage).
PRD-MECH-011 — Probe/thermal-coupling accuracy support
Requirement Statement | The mechanical design shall couple the internal and external temperature sensors to their measured zones (device ambient / product zone) with a thermal lag within the budget allocated to meet PRD-PERF-001 accuracy.
Rationale | Poor thermal coupling degrades measured accuracy regardless of sensor accuracy.
Priority | Must
Source | Performance / User Need
Verification Method | Test/Analysis
Parent Requirement | UN-PERF-001, UN-FUN-003.
Traceability | → PRD-PERF-001/002; → PVL thermal-response test.
Risk Category | Performance
Notes | Thermal-lag budget Evidence Pending.
PRD-MECH-012 — Embeddable module interface
Requirement Statement | The SCM shall provide a defined mechanical mounting interface that allows OEM partners to embed it into their shippers/containers as a standard module.
Rationale | Designed-in channel enablement (STN-017/024).
Priority | Should
Source | Business / User Need
Verification Method | Inspection
Parent Requirement | UN-PHY-003, PRD-STK-006.
Traceability | → PRD-ARCH-010, PRD-CLOUD partner API; → integration guide.
Risk Category | None
Notes | Interface spec Evidence Pending.
7. Verification Notes
Mechanical requirements are verified by IP testing (IEC 60529), drop/vibration/shock (IEC 60068), fit checks in representative shippers, cleaning-durability aging, label-durability testing, thermal-response tests, and gloved-usability studies (with PRD-UX).
8. Traceability Placeholders
Upstream | UN-1100/400; IEC 60529/60068; STN-024/026; D1.
Downstream | PRD-ENV, PRD-ELEC, PRD-PERF, PRD-UX, PRD-SVC, PRD-REG, PRD-MFG; mechanical design outputs; ORCH-018 / EM-ENG-012.
PRD-ELEC (700) — Electrical Requirements
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: PCB/Electronics engineer with Battery/Power engineer and EM-ENG-001. PRD-ELEC-NNN, single "shall", 11-field block. Based on UN-800 (Performance/energy), UN-900 (battery safety), IEC 62368-1 / UN 38.3, and D3. Forward to Stage 06, PCB/EE design outputs, and ORCH-018. Battery-safety items are preliminary and to be reconciled with HAL (Stage 10).
1. Purpose
Specify the power, battery, charging, protection, and electrical-safety requirements — the subsystem that carries the top program risk (90-day/−30 °C endurance and safe low-temperature charging, D3).
2. Scope
Battery/energy source, charge path (USB-C), charge-inhibit control, BMS/protection, power budget, power rails, low-battery behaviour, indicators, evidence/isolation, and electrical safety. Excludes EMC/radio (PRD-CONN/REG), enclosure (PRD-MECH), and firmware power-state logic (PRD-ARCH-004/PRD-FUN).
3. Applicable Standards
IEC 62368-1 (electrical/thermal safety); UN 38.3 (Li-battery transport); IEC 62133 concepts (cell safety); USB-C / USB-PD (charge); CISPR/IEC 61000 (EMC, see PRD-REG). Cited by designation; revisions to confirm.
4. Background
The 90-day/−30 °C combination is in genuine tension (D3): standard Li-ion loses usable capacity in deep cold and cannot be safely charged below 0 °C. Direction: low-temperature-tolerant chemistry, aggressive duty-cycling (PRD-ARCH-004), charge-inhibit below 0 °C, self-heat-before-charge under evaluation — all validated by test.
5. Requirement Categories
Battery/energy · Charging & charge-inhibit · BMS/protection · Power budget & rails · Low-battery behaviour · Indicators · Evidence/isolation · Electrical safety.
6. Detailed Requirements
6.1 Battery & energy
PRD-ELEC-001 — Low-temperature-capable rechargeable energy source
Requirement Statement | The SCM shall use a rechargeable battery chemistry specified to deliver its rated capacity for device operation across the −30 °C to +60 °C operating range.
Rationale | Deep-cold operation and reuse (D1/D3) require a low-temperature-tolerant rechargeable cell.
Priority | Must
Source | User Need (D3) / Reliability
Verification Method | Test
Parent Requirement | UN-PERF-003, UN-PERF-010; D3.
Traceability | → PRD-REL-001 endurance, PRD-ELEC-008 budget; → PVL cold-capacity test.
Risk Category | Reliability
Notes | Final chemistry & low-temp datasheet Evidence Pending (concept §7). Validated by test.
PRD-ELEC-002 — 90-day endurance capacity
Requirement Statement | The SCM battery, with the duty-cycled power architecture, shall provide at least 90 days of operation from a full charge under a defined worst-case cold-duty profile with margin.
Rationale | The headline endurance target (D3); the record must not die in the cold.
Priority | Must
Source | User Need (D3)
Verification Method | Test
Parent Requirement | UN-PERF-003; D3; MR-RISK-001.
Traceability | → PRD-ARCH-004, PRD-ELEC-008, PRD-REL-001; → PVL 90-day endurance (MR-VAL-006).
Risk Category | Reliability
Notes | Margin % Evidence Pending. Top risk — validated by test. Cross-flag PRD-SAFE.
PRD-ELEC-003 — Capacity retention over reuse cycles
Requirement Statement | The SCM battery shall retain sufficient capacity to meet PRD-ELEC-002 over the device's specified number of reuse/charge cycles, or the device shall flag when it can no longer meet it.
Rationale | Capacity fade over reuse can silently cause mid-life dead-in-cold (part of MR-RISK-001).
Priority | Must
Source | Reliability (D1/D3)
Verification Method | Test (cycle-life)
Parent Requirement | UN-PERF-010; D1/D3.
Traceability | → PRD-REL-002 cycle life, PRD-FUN-020 readiness; → PVL cycle-life test.
Risk Category | Reliability
Notes | Cycle count Evidence Pending. Ties readiness reporting (charge/health).
6.2 Charging & charge-inhibit
PRD-ELEC-004 — USB-C charging
Requirement Statement | The SCM shall recharge through a USB-C interface from a standard compliant source and reach full trip-readiness within the charge-time target set in the product spec.
Rationale | Reusability (D1) requires convenient, standard recharge at the depot.
Priority | Must
Source | User Need (D1)
Verification Method | Test
Parent Requirement | UN-PERF-004; concept §1; D1.
Traceability | → PRD-MECH-006 sealed port, PRD-SVC-003 turnaround; → PVL charge test.
Risk Category | Reliability
Notes | Charge time Evidence Pending. USB-C per USB spec (rev. to confirm).
PRD-ELEC-005 — Charge inhibit below 0 °C
Requirement Statement | The SCM shall inhibit battery charging whenever the cell temperature is below 0 °C (or the cell-defined safe lower charge limit) and shall clearly indicate the inhibited state.
Rationale | Charging a cold Li-ion cell risks lithium plating and a thermal event (D3).
Priority | Must
Source | Safety (D3)
Verification Method | Test
Parent Requirement | UN-SAFE-007, UN-PERF-004; D3; IEC 62368-1.
Traceability | → PRD-SAFE-007, PRD-UX-008 status; → HAL battery hazard; → PVL charge-inhibit test.
Risk Category | Safety
Notes | Threshold per cell datasheet. To be reconciled with HAL (Stage 10).
PRD-ELEC-006 — Self-heat-before-charge (evaluation)
Requirement Statement | The SCM shall, where the selected design provides it, safely warm the cell to a permitted charge temperature before charging a cold-returned device, without exceeding safe limits.
Rationale | Improves depot turnaround for cold returns without unsafe cold charging (D3).
Priority | Could
Source | Reliability (D3)
Verification Method | Test
Parent Requirement | UN-PERF-004; D3.
Traceability | → PRD-ELEC-005, PRD-SVC-003; → PVL self-heat test.
Risk Category | Safety
Notes | Under evaluation (D3); feasibility Evidence Pending. To be reconciled with HAL.
6.3 BMS & protection
PRD-ELEC-007 — Battery protection (OV/UV/OC/OT/short)
Requirement Statement | The SCM battery management shall protect against overvoltage, undervoltage, overcurrent, over-temperature, and short-circuit in normal use, foreseeable misuse, and single-fault.
Rationale | Prevents the device from becoming a fire/thermal hazard (UN-SAFE-008).
Priority | Must
Source | Safety
Verification Method | Test (fault injection)
Parent Requirement | UN-SAFE-008; IEC 62368-1 / IEC 62133 concepts.
Traceability | → PRD-SAFE-008, PRD-REL-004; → HAL battery hazard; → PVL protection test.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10). Cross-flag battery engineer.
PRD-ELEC-008 — Power budget conformance
Requirement Statement | The SCM electrical design shall meet a documented power budget that allocates energy across sleep, sensing, logging, display, and each radio bearer such that PRD-ELEC-002 is achieved.
Rationale | Endurance is an energy-budget outcome; radios are the dominant load and must be bounded.
Priority | Must
Source | Reliability / Performance
Verification Method | Analysis + Test
Parent Requirement | UN-PERF-003; power budget (Stage 00); D3.
Traceability | → PRD-ARCH-004, PRD-CONN duty, PRD-PERF-005; → PVL current-profile test.
Risk Category | Reliability
Notes | Budget document (Stage 00) is the baseline; radio energy Evidence Pending pending RF test.
PRD-ELEC-009 — Fuel gauge / state-of-charge accuracy
Requirement Statement | The SCM shall measure and report battery state of charge and state of health accurately enough across temperature to support readiness decisions and a meaningful low-battery warning.
Rationale | Readiness triage and low-battery warning depend on trustworthy SoC/SoH.
Priority | Must
Source | Reliability / User Need
Verification Method | Test
Parent Requirement | UN-FUN-017, UN-PERF-010.
Traceability | → PRD-FUN-020, PRD-ELEC-010; → PVL gauge-accuracy test.
Risk Category | Reliability
Notes | Accuracy target Evidence Pending (temperature-dependent).
6.4 Low-battery behaviour & rails
PRD-ELEC-010 — Low-battery graceful degradation and warning
Requirement Statement | The SCM shall warn of low battery and, as energy depletes, prioritize logging and data preservation over discretionary functions (uploads, location), preserving the record to the last.
Rationale | A depleting device must protect the record, not lose it silently.
Priority | Must
Source | Reliability / Safety
Verification Method | Test
Parent Requirement | UN-PERF-003, UN-SAFE-010.
Traceability | → PRD-REL-004, PRD-FUN-012 buffering, PRD-UX-008; → PVL depletion test.
Risk Category | Reliability
Notes | Graceful-degradation order Evidence Pending.
PRD-ELEC-011 — Power rails and brown-out protection
Requirement Statement | The SCM shall provide regulated power rails and brown-out protection so a falling supply causes a clean safe-state entry with no corruption of logged data or configuration.
Rationale | Brown-out during a write must never corrupt the record.
Priority | Must
Source | Reliability
Verification Method | Test
Parent Requirement | UN-SAFE-010; UN-FUN-014.
Traceability | → PRD-ARCH-008 safe-state, PRD-REL-005 storage integrity; → PVL brown-out test.
Risk Category | Reliability
Notes | Ties power-fail-safe writes (PRD-REL-005).
6.5 Indicators, isolation & safety
PRD-ELEC-012 — Status indicator
Requirement Statement | The SCM shall provide a low-energy status/charge indicator distinguishing at least armed/monitoring, fault, charging, and charge-inhibited states.
Rationale | Passive reassurance and service triage without draining the budget.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-PERF-011, UN-SAFE-007.
Traceability | → PRD-UX-012, PRD-ELEC-005; → PVL indicator test.
Risk Category | Usability
Notes | Energy-cheap indication.
PRD-ELEC-013 — Evidence, isolation & ESD
Requirement Statement | The SCM electrical design shall provide evidence, isolation, and ESD protection sufficient to meet the immunity levels of PRD-REG EMC requirements without malfunction.
Rationale | ESD and coupling must not corrupt data or damage the device.
Priority | Must
Source | Regulatory / Reliability
Verification Method | Test (IEC 61000-4-2)
Parent Requirement | UN-ENV-009, UN-REG-005; CISPR/IEC 61000.
Traceability | → PRD-REG-005 EMC; → PVL ESD/EMC test.
Risk Category | Reliability
Notes | Immunity levels set in PRD-REG (rev. to confirm).
PRD-ELEC-014 — Electrical/thermal safety compliance
Requirement Statement | The SCM shall meet IEC 62368-1 electrical and thermal safety requirements in normal, abnormal, and single-fault conditions across its use environments.
Rationale | It must be safe on aircraft, in cold stores, clinics, and homes (UN-SAFE-008).
Priority | Must
Source | Regulatory / Safety
Verification Method | Test (IEC 62368-1)
Parent Requirement | UN-SAFE-008, UN-REG-005; IEC 62368-1.
Traceability | → PRD-SAFE-008, PRD-REG-005; → HAL; → PVL safety test.
Risk Category | Safety
Notes | To be reconciled with HAL. Standard (IEC 62368-1, rev. to confirm).
PRD-ELEC-015 — UN 38.3 battery transport testing
Requirement Statement | The SCM battery/pack shall pass UN 38.3 transport testing and the device shall carry the resulting documentation.
Rationale | Legal, safe air/ground transport of the Li battery (D4).
Priority | Must
Source | Regulatory / Safety
Verification Method | Test (UN 38.3)
Parent Requirement | UN-REG-006, UN-SAFE-009; UN 38.3 / IATA DGR.
Traceability | → PRD-REG-006, PRD-SAFE-009; → PVL UN 38.3 report.
Risk Category | Compliance
Notes | Test summary Evidence Pending. Standard (UN 38.3, rev. to confirm).
PRD-ELEC-016 — Reverse-polarity / fault-tolerant charge input
Requirement Statement | The SCM charge input shall tolerate reverse polarity, over-voltage, and non-compliant sources without damage or hazard.
Rationale | Field/depot chargers vary; a bad source must not damage or endanger the device.
Priority | Should
Source | Reliability / Safety
Verification Method | Test
Parent Requirement | UN-SAFE-008, UN-PERF-004.
Traceability | → PRD-ELEC-004/007; → PVL abuse-input test.
Risk Category | Safety
Notes | To be reconciled with HAL.
7. Verification Notes
Electrical requirements are verified by cold-capacity and 90-day endurance testing (MR-VAL-006), charge/charge-inhibit and abuse testing, BMS fault-injection, current-profile measurement against the power budget, UN 38.3, IEC 62368-1 safety, and EMC/ESD (with PRD-REG). Battery-safety items reconcile with HAL.
8. Traceability Placeholders
Upstream | UN-800/900; IEC 62368-1 / UN 38.3 / IEC 62133 concepts; D3/D4; power budget (Stage 00).
Downstream | PRD-ARCH, PRD-REL, PRD-SAFE, PRD-CONN, PRD-REG, PRD-SVC; PCB/EE design outputs; HAL; ORCH-018 / EM-ENG-012.
PRD-CONN (800) — Connectivity Requirements
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: RF/Connectivity engineer with Embedded Software and EM-ENG-001. PRD-CONN-NNN, single "shall", 11-field block. Based on UN-800 (connectivity/location), UN-400 (RF environment), and D2/D4. Forward to Stage 06, RF design outputs, and ORCH-018.
1. Purpose
Specify how the SCM moves data and alerts out and is provisioned — bearers, opportunistic offload, provisioning/pairing, offline behaviour, and RF performance inside a shipper.
2. Scope
LTE-M, Wi-Fi, BLE, GNSS, bearer selection, provisioning/pairing, offline/recovery behaviour, and RF-in- container performance. Excludes radio certification (PRD-REG-004), cloud endpoints (PRD-CLOUD), and mobile app functions (PRD-MOB).
3. Applicable Standards
3GPP LTE-M (Cat-M1); IEEE 802.11 (Wi-Fi); Bluetooth Core (BLE); FCC Part 15 / CE RED (radio, see PRD-REG); TLS 1.2+/1.3 (secure transport, see PRD-SEC). Cited by designation; revisions to confirm.
4. Background
The device sits inside a metallized, insulated shipper (RF-hostile) and must offload opportunistically without draining the 90-day budget (D2). LTE-M is primary (low-power WAN, global lanes); Wi-Fi/BLE offload where available. RF-in-container performance is the principal open technical risk (concept §7).
5. Requirement Categories
Bearers · Opportunistic offload · Provisioning/pairing · Offline & recovery · RF performance · Diagnostics.
6. Detailed Requirements
PRD-CONN-001 — LTE-M primary bearer
Requirement Statement | The SCM shall use LTE-M (Cat-M1) as its primary wide-area bearer for store-and-forward upload and excursion alerting.
Rationale | Low-power WAN suited to duty-cycled global lanes (D2/D4).
Priority | Must
Source | Architecture / User Need
Verification Method | Test
Parent Requirement | UN-PERF-008, UN-FUN-010; D2/D4.
Traceability | → PRD-ARCH-005, PRD-ELEC-008 budget; → PVL cellular test.
Risk Category | Performance
Notes | Roaming/coverage per market Evidence Pending (SIM/connectivity strategy).
PRD-CONN-002 — Wi-Fi offload
Requirement Statement | The SCM shall support Wi-Fi for opportunistic bulk data offload where a provisioned network is available.
Rationale | Cheaper/higher-throughput offload at depots and facilities preserves the energy budget.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-010, UN-PERF-008.
Traceability | → PRD-CONN-004 selection; → PVL Wi-Fi test.
Risk Category | Performance
Notes | Enterprise Wi-Fi auth Evidence Pending.
PRD-CONN-003 — BLE provisioning and short-range offload
Requirement Statement | The SCM shall support BLE for local provisioning/configuration, local read, and short-range data offload to a paired app/gateway.
Rationale | Frictionless local setup and offload where WAN is absent.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-011, UN-PERF-008; D2.
Traceability | → PRD-MOB-001 pairing, PRD-CONN-006; → PVL BLE test.
Risk Category | Performance
Notes | BLE pairing security in PRD-SEC.
PRD-CONN-004 — Opportunistic multi-bearer selection
Requirement Statement | The SCM shall select among available bearers to deliver buffered data and alerts, preferring the lowest-energy path that satisfies the cadence/latency targets.
Rationale | Energy-aware bearer choice keeps data flowing without exhausting the budget (C2).
Priority | Must
Source | Architecture / Reliability
Verification Method | Test
Parent Requirement | UN-FUN-010, UN-PERF-008; D2.
Traceability | → PRD-ARCH-005, PRD-ELEC-008, PRD-FUN-013; → PVL bearer-selection test.
Risk Category | Performance
Notes | Selection policy configurable.
PRD-CONN-005 — GNSS location acquisition
Requirement Statement | The SCM shall acquire a GNSS location fix on a configurable/event-triggered cadence within the energy allocation for location, and shall record fix quality.
Rationale | Fleet management and attribution need location without breaking the budget (C3).
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-PERF-009, UN-FUN-016; conflict C3.
Traceability | → PRD-PERF-007 location, PRD-FUN-019; → PVL GNSS test.
Risk Category | Performance
Notes | GNSS-in-container performance Evidence Pending; may need external antenna.
PRD-CONN-006 — Secure provisioning and pairing
Requirement Statement | The SCM shall support authenticated provisioning and pairing such that only authorized users/gateways can configure or read controlled data over any bearer.
Rationale | Provisioning is an attack surface; only authorized parties may configure the device.
Priority | Must
Source | Security / User Need
Verification Method | Test
Parent Requirement | UN-SEC-003/005; conflict C9.
Traceability | → PRD-SEC-005/006, PRD-ARCH-006; → PVL pairing-security test.
Risk Category | Security
Notes | Edge read stays login-free (AD-5); config is controlled.
PRD-CONN-007 — RF performance inside representative shipper
Requirement Statement | The SCM shall achieve a link budget sufficient to deliver buffered data and alerts from inside a representative metallized/insulated shipper within a defined delivery-success target.
Rationale | RF attenuation inside the shipper must not prevent data egress over a trip.
Priority | Must
Source | User Need / Reliability
Verification Method | Test (RF-in-container)
Parent Requirement | UN-ENV-006, UN-PERF-008; concept §7.
Traceability | → PRD-ENV-009, PRD-CONN-001; → PVL RF-in-container test.
Risk Category | Performance
Notes | Principal open risk; antenna strategy / external-antenna option Evidence Pending.
PRD-CONN-008 — Store-and-forward on connectivity loss
Requirement Statement | The SCM shall detect connectivity loss, continue buffering, and resume upload in intact chronological order on reconnection without duplication or data loss.
Rationale | Intermittent links must not corrupt or lose the record (D2).
Priority | Must
Source | Reliability / User Need
Verification Method | Test
Parent Requirement | UN-FUN-009, UN-ENV-007; D2.
Traceability | → PRD-FUN-012, PRD-CLOUD-002 ingest; → PVL reconnection test.
Risk Category | Reliability
Notes | Idempotent upload; ordering preserved.
PRD-CONN-009 — Alert-path prioritization
Requirement Statement | The SCM shall prioritize excursion-alert delivery over routine data upload when connectivity is constrained, subject to the energy budget.
Rationale | A rescue-enabling alert must not queue behind bulk data.
Priority | Must
Source | Safety / User Need
Verification Method | Test
Parent Requirement | UN-FUN-006, UN-PERF-005.
Traceability | → PRD-FUN-008, PRD-PERF-004 latency; → HAL delayed-alert hazard; → PVL priority test.
Risk Category | Safety
Notes | To be reconciled with HAL (Stage 10).
PRD-CONN-010 — Connectivity diagnostics
Requirement Statement | The SCM shall record connectivity diagnostics (bearer used, signal quality, attempts/failures, backlog depth) available to fleet tooling.
Rationale | Fleet operators must diagnose silent-comms devices and coverage gaps.
Priority | Should
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-018; STN-014.
Traceability | → PRD-CLOUD-006 fleet health; → PVL diagnostics test.
Risk Category | None
Notes | Energy-cheap diagnostics.
PRD-CONN-011 — Firmware-update transport
Requirement Statement | The SCM shall support firmware-update download over LTE-M/Wi-Fi (OTA) and USB-C, resuming interrupted transfers, and shall only apply verified updates (per PRD-SEC-008).
Rationale | The fleet must be updatable in the field and at depot without bricking (D2).
Priority | Must
Source | User Need / Security
Verification Method | Test
Parent Requirement | UN-SEC-005, UN-SAFE-010; D2.
Traceability | → PRD-FUN-022, PRD-SEC-008, PRD-CLOUD-008; → PVL OTA-transport test.
Risk Category | Security
Notes | Verification/signing in PRD-SEC-008.
PRD-CONN-012 — Regional radio configurability
Requirement Statement | The SCM shall support configuration of radio operation to the bands/regulatory profiles of its target markets.
Rationale | Multi-market operation (D4) requires per-region radio conformance.
Priority | Must
Source | Regulatory
Verification Method | Inspection + Test
Parent Requirement | UN-REG-004; D4; FCC Part 15 / CE RED.
Traceability | → PRD-REG-004 radio approvals, PRD-ARCH-010; → PVL regional radio test.
Risk Category | Compliance
Notes | Some regions may need per-market hardware variants (PRD-REG-004).
7. Verification Notes
Connectivity requirements are verified by bearer tests (LTE-M/Wi-Fi/BLE/GNSS), the RF-in-container test, reconnection/store-and-forward integrity tests, alert-priority and latency tests (with PRD-PERF), pairing- security tests (with PRD-SEC), and radio conformance (with PRD-REG).
8. Traceability Placeholders
Upstream | UN-800/400; concept §7; D2/D4.
Downstream | PRD-ARCH, PRD-ELEC (budget), PRD-CLOUD, PRD-MOB, PRD-SEC, PRD-REG, PRD-PERF; RF design outputs; HAL; ORCH-018 / EM-ENG-012.
PRD-CLOUD (900) — Cloud Ecosystem
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Cloud/Platform engineer with Cybersecurity engineer and EM-ENG-001. PRD-CLOUD-NNN, single "shall", 11-field block. Based on UN-600 (record/fleet/alert), UN-1000 (security), UN-1300 (Part 11/GDP), and D2. Forward to Stage 06, cloud design outputs, and ORCH-018.
1. Purpose
Specify the cloud platform requirements — device registry, store-and-forward ingest, alert routing, record assembly with audit trail, fleet dashboard/reporting/export, OTA service, and enterprise integration.
2. Scope
Registry, ingest, rules/alerting, record/audit, dashboard/reporting/export, OTA orchestration, APIs/ integration, and platform monitoring. Excludes device-side functions (PRD-FUN), security mechanisms detail (PRD-SEC), and mobile app (PRD-MOB).
3. Applicable Standards
21 CFR Part 11 / ALCOA+ (records/audit); EU GDP (data governance); GAMP 5 (CSV); ISO/IEC 27001 concepts (platform security). Cited by designation; revisions to confirm.
4. Background
The cloud is where the record becomes defensible evidence, alerts reach responders, and the fleet is managed. It ingests store-and-forward data (D2), must preserve integrity/attribution end to end, and must be validatable (CSV/GAMP 5) for regulated customers.
5. Requirement Categories
Registry · Ingest/store-and-forward · Alerting/routing · Record/audit · Dashboard/reporting/export · OTA · APIs/integration · Platform monitoring/availability.
6. Detailed Requirements
PRD-CLOUD-001 — Device registry and identity
Requirement Statement | The cloud platform shall maintain a registry of every device with unique identity, configuration, calibration status, and firmware version, and shall only accept data from registered, authenticated devices.
Rationale | Attribution and fleet control require an authoritative device identity.
Priority | Must
Source | Architecture / Security
Verification Method | Test
Parent Requirement | UN-FUN-018, UN-SEC-002/003.
Traceability | → PRD-SEC-006, PRD-FUN-020; → PVL registry test.
Risk Category | Security
Notes | Device identity ties PRD-MECH-010 durable ID.
PRD-CLOUD-002 — Store-and-forward ingest with integrity preservation
Requirement Statement | The cloud platform shall ingest buffered device data in intact chronological order, de-duplicate idempotently, preserve tamper-evidence and attribution, and flag any gaps.
Rationale | The uploaded record must remain complete and defensible after opportunistic offload (D2).
Priority | Must
Source | Regulatory / Reliability
Verification Method | Test
Parent Requirement | UN-FUN-009/014, UN-SEC-001; D2.
Traceability | → PRD-CONN-008, PRD-FUN-017, PRD-SEC-001; → PVaL audit simulation.
Risk Category | Compliance
Notes | Gaps flagged, never hidden.
PRD-CLOUD-003 — Alert routing and escalation
Requirement Statement | The cloud platform shall route excursion and device-fault alerts to the responders configured per shipment/role via the configured channels and escalate unacknowledged alerts.
Rationale | A warning must reliably reach someone who can act.
Priority | Must
Source | User Need
Verification Method | Test
Parent Requirement | UN-FUN-007; STN-018.
Traceability | → PRD-FUN-009, PRD-PERF-004; → PVL routing/escalation test.
Risk Category | Safety
Notes | Acknowledgement tracked; to be reconciled with HAL.
PRD-CLOUD-004 — Configurable alert rules
Requirement Statement | The cloud platform shall allow authorized users to configure alert rules (thresholds, recipients, channels, escalation timing) per product/shipment within governance controls.
Rationale | Alerts must match product stability and organizational response processes.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-007/011; STN-018.
Traceability | → PRD-FUN-014, PRD-SEC-005; → PVL rules test.
Risk Category | Usability
Notes | Change is audit-trailed (PRD-CLOUD-006).
PRD-CLOUD-005 — Trip record assembly and export
Requirement Statement | The cloud platform shall assemble a complete, attributable, tamper-evident trip record and export it on demand in human- and machine-readable, integrity-preserving formats.
Rationale | Evidence must be assembled and shareable for release, audit, and dispute.
Priority | Must
Source | Regulatory / User Need
Verification Method | Test
Parent Requirement | UN-FUN-014/015, UN-REG-003.
Traceability | → PRD-FUN-017/018, PRD-SEC-002; → PVaL audit export.
Risk Category | Compliance
Notes | Formats Evidence Pending (e.g., PDF + CSV/JSON); integrity preserved.
PRD-CLOUD-006 — Fleet dashboard and health
Requirement Statement | The cloud platform shall present fleet status (device/shipment status, location, readiness, connectivity health, required action) to authorized users.
Rationale | A fleet must be managed, strays chased, and unfit devices flagged.
Priority | Should
Source | User Need
Verification Method | Demonstration
Parent Requirement | UN-FUN-018/017; STN-014/015.
Traceability | → PRD-FUN-020/021, PRD-CONN-010; → PVaL fleet scenario.
Risk Category | None
Notes | Liveness connectivity-bound (C7).
PRD-CLOUD-007 — Audit trail (Part 11-style)
Requirement Statement | The cloud platform shall maintain a secure, time-stamped, attributable, non-editable audit trail of configuration changes, access, and data actions, retained per policy.
Rationale | The governance of the record must itself be inspectable and defensible (Part 11).
Priority | Must
Source | Regulatory / Security
Verification Method | Test
Parent Requirement | UN-SEC-009, UN-REG-003; 21 CFR Part 11 / ALCOA+.
Traceability | → PRD-SEC-004 audit, PRD-CLOUD-004; → PVaL audit-trail review.
Risk Category | Compliance
Notes | Standard (21 CFR Part 11, rev. to confirm).
PRD-CLOUD-008 — OTA orchestration
Requirement Statement | The cloud platform shall orchestrate signed firmware updates to the fleet with staged rollout, per-device targeting, status tracking, and rollback of a failed campaign.
Rationale | Fleet updates must be controlled, observable, and reversible to avoid mass bricking.
Priority | Must
Source | User Need / Security
Verification Method | Test
Parent Requirement | UN-SEC-005, UN-SAFE-010; D2.
Traceability | → PRD-FUN-022, PRD-SEC-008, PRD-CONN-011; → PVL OTA-campaign test.
Risk Category | Security
Notes | Signing/keys in PRD-SEC-008.
PRD-CLOUD-009 — Enterprise access control (RBAC/SSO)
Requirement Statement | The cloud platform shall enforce role-based access control and support enterprise single-sign-on for all configuration, data-access, and administrative functions.
Rationale | Enterprise/regulated deployment is gated by IT security and CSV.
Priority | Must
Source | Security / User Need
Verification Method | Test
Parent Requirement | UN-SEC-003, PRD-STK-005.
Traceability | → PRD-SEC-005/006, PRD-ARCH-006; → PVL access-control test.
Risk Category | Security
Notes | Does not gate the login-free edge read (AD-5).
PRD-CLOUD-010 — Integration APIs
Requirement Statement | The cloud platform shall expose secured APIs enabling authorized enterprise/OEM systems to retrieve records, alerts, and fleet status.
Rationale | Records/alerts must flow into ERP/QMS/partner systems where decisions live.
Priority | Should
Source | User Need / Business
Verification Method | Test
Parent Requirement | UN-FUN-015; UN-ECO-###; STN-013.
Traceability | → PRD-FUN-018, PRD-SEC-006; → PVL API test.
Risk Category | Security
Notes | API auth per PRD-SEC.
PRD-CLOUD-011 — Data retention and ownership controls
Requirement Statement | The cloud platform shall let the customer own its data and configure retention and deletion per policy, applying deletion only through audited, authorized action.
Rationale | Regulated customers must control retention/ownership (governance).
Priority | Must
Source | Regulatory / User Need
Verification Method | Demonstration
Parent Requirement | UN-SEC-007, UN-REG-009.
Traceability | → PRD-SEC-007, PRD-CLOUD-007; → PVaL retention scenario.
Risk Category | Compliance
Notes | Retention periods Evidence Pending (per customer/market).
PRD-CLOUD-012 — Platform availability and integrity
Requirement Statement | The cloud platform shall meet a defined availability target for ingest and alerting and shall not lose or silently corrupt ingested data under component failure.
Rationale | Alerts and records depend on a reliable, durable platform.
Priority | Must
Source | Reliability
Verification Method | Test/Analysis
Parent Requirement | UN-FUN-006/014; STN-013.
Traceability | → PRD-REL-006, PRD-CLOUD-002; → PVL availability/DR test.
Risk Category | Reliability
Notes | Availability target Evidence Pending (SLA); durable storage.
7. Verification Notes
Cloud requirements are verified by ingest-integrity, routing/escalation, audit-trail, OTA-campaign, access-control, export, and availability/DR testing, and by CSV/validation review (GAMP 5) for regulated customers. Safety-relevant routing (CLOUD-003) reconciles with the Hazard Analysis (HAL, Stage 10).
8. Traceability Placeholders
Upstream | UN-600/1000/1300; 21 CFR Part 11 / GDP / GAMP 5; D2.
Downstream | PRD-SEC, PRD-FUN, PRD-CONN, PRD-MOB, PRD-REL, PRD-REG; cloud design outputs; HAL; ORCH-018.
Device / Design Architecture — Smart Cold Chain Monitor (SCM)
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · §7 of the Product Requirements Standard. Owner: Chief Systems Engineer (EM-ENG-001). Describes how the product requirements map onto the architecture — subsystem allocation, key design decisions, and interfaces. Forward to System/Subsystem Requirements and Architecture (Stage 06) and ORCH-018 / EM-ENG-012.
---
1. Architectural intent
The SCM architecture is organized around one value chain — capture → detect → warn → decide → prove — under a single overriding constraint: survive 90 days at −30 °C without dying or losing data (D3, the top program risk). Every architectural decision below is a resolution of that constraint against the other user needs (accuracy, offline verdict, auditability, security, reusability). The device is deliberately an offline-first edge node with opportunistic connectivity (D2), not a live-streaming telemetry endpoint.
2. Subsystem allocation (requirements → architecture)
Subsystem | Owns (requirement modules) | Principal design decisions
S1 Enclosure / mechanical | PRD-MECH, PRD-PHY-derived, part of PRD-ENV | Sealed IP-rated housing sized to fit standard shippers; external probe to represent the product zone; sealed USB-C port; glove-friendly single-button HMI cutout; durable labeling.
S2 Power | PRD-ELEC, PRD-REL, part of PRD-SAFE | Low-temperature-capable rechargeable Li-ion + BMS; USB-C charge path with charge-inhibit below 0 °C and optional self-heat-before-charge; fuel gauge; protection (OV/UV/OC/OT). The 90-day/−30 °C budget is met by duty-cycling, not chemistry alone.
S3 Compute / MCU + FW | PRD-FUN, PRD-SEC (boot/keys), PRD-PERF, PRD-REL | Low-power MCU with deep-sleep between samples; RTC for trusted timestamps; secure element for keys/secure boot; watchdog; power-state machine as the energy governor.
S4 Sensing | PRD-FUN, PRD-PERF, part of PRD-MECH | Internal + external (probe) temperature for two-sided detection and product-zone representation; humidity; 3-axis accelerometer for product-relevant shock; GNSS duty-cycled/event-triggered within the energy budget.
S5 Connectivity | PRD-CONN, part of PRD-ENV (RF) | LTE-M primary (low-power WAN), Wi-Fi and BLE offload; multi-bearer opportunistic upload; antenna strategy to fight shipper attenuation (external-antenna option Evidence Pending).
S6 HMI | PRD-UX, part of PRD-MECH | OLED verdict + single button + indicator; cold-legible, glare-legible, glove-operable, offline, no-login.
S7 Storage | PRD-FUN, PRD-SEC, PRD-REL | Non-volatile tamper-evident log sized for a full 90-day trip at minimum interval with margin; append-only audit; power-fail-safe writes.
S8 Cloud + dashboard | PRD-CLOUD, PRD-SEC, part of PRD-REG | Device registry, store-and-forward ingest, rules/alert routing + escalation, fleet dashboard, reporting/export, signed OTA, RBAC/SSO, 21 CFR Part 11-style audit trail.
S9 Mobile / BLE app | PRD-MOB, part of PRD-UX/PRD-CONN | Provisioning, trip-profile configuration, arming aid, local read, notifications; edge read stays login-free while platform enforces governance (trust-boundary split).
3. Key architectural decisions and the tensions they resolve
• AD-1 — Offline-first, store-and-forward (D2). Local logging + scheduled cadence upload + immediate excursion alert when connected. Resolves C2 (live data vs battery) and C7 (real-time vs no-connectivity): live telemetry is upside, never a dependency. Evidence: D2; power budget (Stage 00).
• AD-2 — Duty-cycled energy governor (D3). The MCU power-state machine sleeps between samples and batches uploads; radios are the dominant load and are gated hard. Resolves C1/C2/C3 (cost/battery/GPS vs endurance). The 90-day/−30 °C budget is the headline tension and is validated by test (PRD-REL / PRD-PERF acceptance criteria). Evidence: D3; MR-RISK-001; MR-VAL-006.
• AD-3 — Charge-inhibit below 0 °C + self-heat evaluation (D3). Charging a cold Li-ion cell risks lithium plating / thermal event; the BMS refuses charge below 0 °C and indicates the state. Resolves the charge-safety leg of the D3 tension. Evidence: D3; IEC 62368-1 (rev. to confirm).
• AD-4 — Two-sided detection with internal + external sensing. Freeze detection is a safety requirement, not a feature (freeze damage is silent and under-detected). The external probe represents the product zone. Resolves the missed-freeze hazard (UN-SAFE-002). Evidence: STN-002; WHO/UNICEF freezing reporting (Cited).
• AD-5 — Trust-boundary split (edge read vs platform governance). The on-device verdict is readable with no login (frictionless edge); configuration, data, and access are controlled on the platform (RBAC/SSO, Part 11 audit). Resolves C9 (security lockdown vs open edge). Evidence: UN-SEC-003/004; 21 CFR Part 11 (rev. to confirm).
• AD-6 — Layered information (verdict → detail). The edge shows an unambiguous accept/reject/quarantine verdict; the full record lives in the cloud/export. Resolves C5 (rich data vs simple verdict). Evidence: UN-VIS-007; UN-UX; STN-004/027.
• AD-7 — Signed, resumable, fail-safe OTA. Firmware is verifiable-genuine, resumable, and cannot brick the fleet; data is preserved across updates/resets. Resolves the OTA-integrity threat and the fail-safe safety need. Evidence: UN-SEC-005; UN-SAFE-010; D2.
• AD-8 — Reusable-fleet lifecycle (D1). Return → clean → recharge → readiness triage (charge/cal/firmware) → re-provision, with deploy blocked when out-of-calibration. Resolves C6 (reuse TCO vs reverse-logistics cost). Evidence: D1; UN-SVC; UN-FUN-017.
4. Functional decomposition (top level)
• F1 Measure & timestamp (S4+S3+S7) → PRD-FUN sense/log, PRD-PERF accuracy.
• F2 Detect excursions (two-sided, product-relevant) (S3) → PRD-FUN detect, PRD-SAFE.
• F3 Alert & escalate (S3+S5+S8) → PRD-FUN alert, PRD-CLOUD routing, PRD-PERF latency.
• F4 Decide (offline verdict) (S3+S6) → PRD-FUN verdict, PRD-UX.
• F5 Store & forward (S7+S5) → PRD-FUN buffer, PRD-CONN offload, PRD-PERF capacity.
• F6 Configure & provision & arm (S9+S6+S3) → PRD-FUN config/arm, PRD-UX, PRD-CONN provisioning.
• F7 Record, prove & export (S7+S8) → PRD-FUN record, PRD-SEC integrity, PRD-CLOUD export, PRD-REG.
• F8 Locate & attribute (S4+S3+S8) → PRD-FUN locate, PRD-PERF location.
• F9 Fleet & manage & OTA (S8+S5) → PRD-CLOUD fleet/OTA, PRD-SEC OTA, PRD-FUN readiness.
• F10 Power, charge & self-protect (S2+S3) → PRD-ELEC, PRD-SAFE, PRD-REL.
5. Interface architecture
Typed interfaces are enumerated in [system-block-diagram.md](system-block-diagram.md) §3 (I-01…I-12): THRM (payload/shipper → probe/enclosure → temp sensors), MECH (shipper fit/retention), PWR (USB-C → BMS → cell → rails, charge-inhibit gated), DATA (sensors → MCU → storage/HMI; device ↔ cloud/app), RF (antennas ↔ cellular/Wi-Fi/GNSS, attenuated by the shipper). The RF-in-container interface (I-08) is the principal open architectural risk (antenna strategy Evidence Pending; UN-ENV-006, concept §7).
6. Neighbouring-discipline flags (CLAUDE.md rule 2)
• Battery / power electronics engineer — cell chemistry selection, BMS, charge-inhibit, self-heat, and the 90-day/−30 °C budget (AD-2/AD-3; top risk).
• RF/antenna engineer — RF-in-container performance and antenna strategy (AD-1/I-08).
• Mechanical / materials engineer — sealing (incl. USB-C port), cleaning-chemical compatibility, probe form and retention (S1).
• Cybersecurity engineer — threat model, key management, OTA signing, RBAC/SSO (AD-5/AD-7); threat model Evidence Pending.
• Quality / Regulatory engineer (EM-ENG-006) — WHO PQS category, US FDA classification, Part 11/GDP, transport/radio/EMC certification (PRD-REG); named gaps Evidence Pending.
• Hazard analysis (ORCH-017) — all PRD-SAFE requirements here are preliminary and to be reconciled with the Hazard Analysis Library (HAL, Stage 10); each will trace to a HA-###/RCR-### record.
Evidence posture: architectural decisions cite D1–D4, the Stage-04 needs, and standards by designation (revision to confirm). Quantitative budgets (energy, RF link, storage sizing) are set as requirements in the modules and closed by verification/validation; unresolved values are Evidence Pending, never invented.
Product Description — Smart Cold Chain Monitor (SCM)
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · Owner: Chief Systems Engineer (EM-ENG-001). This is the §5 "Product Description" of the Product Requirements Standard (00-governance/product-requirements-standard.md). Based on 00-concept/concept-and-decisions.md (D1–D4) and the Master User Needs Library (Stage 04). Forward to System Requirements (Stage 06) and ORCH-018 / EM-ENG-012.
---
1. What the SCM is
The Smart Cold Chain Monitor (SCM) is a reusable, battery-operated, cellular-connected environmental data logger placed inside insulated vaccine shipping containers, cartons, and cold boxes. Across the global vaccine distribution chain — from manufacturer pack-out, through distributors, 3PLs, and air/ground transport, to national immunization programs and the last-mile point of use — it continuously measures and records temperature (with humidity and shock), detects both freezing and overheating excursions against configurable product-specific thresholds, alerts accountable stakeholders when connectivity permits, presents an unambiguous offline accept/reject verdict on the device itself, and provides an auditable, tamper-evident cold-chain record for release, acceptance, and disposition decisions.
The SCM is an information / monitoring device that supports human and organizational decisions. It does not control temperature, is not the primary means of preserving the vaccine, and is not therapeutic. Its dominant safety contribution is informational — supporting the accept/reject decision that determines whether a potentially spoiled dose is administered or a potent dose is discarded (both error directions harm the beneficiary; severity anchor UN-VIS-001 / UN-SAFE-001 / STN-020).
Foundational decisions carried from concept (D1–D4):
D1 — Reusable / rechargeable (multi-trip) | USB-C charging, rechargeable battery, and OTA imply a multi-trip device; cleaning, reverse-logistics return, and re-provisioning are lifecycle requirements.
D2 — Local logging + scheduled cadence upload + immediate excursion alerts | store-and-forward operation — log at a configurable interval (default 5–10 min), upload on a configurable cadence (default hourly), push an immediate alert on a detected excursion when connectivity is available, buffer when offline and upload opportunistically (LTE-M primary; Wi-Fi/BLE offload where available).
D3 — Rechargeable, low-temperature-capable Li-ion, duty-cycled | the −30 °C + 90-day combination is in genuine tension (top program risk); direction is a low-temperature-tolerant chemistry, aggressive duty-cycling, charge inhibited below 0 °C, and self-heat-before-charge under evaluation — validated by test (carried through PRD-REL / PRD-SAFE / PRD-PERF).
D4 — Multi-market | WHO PQS + EU GDP + US FDA data-integrity + air/ground transport (UN 38.3 / IATA), radio (FCC Part 15 / CE RED), EMC/safety (IEC 62368-1, CISPR/IEC 61000), environmental (IEC 60529 IP, IEC 60068).
Operating range −30 °C to +60 °C; endurance target 90 days per charge.
---
2. Operating concept (capture → detect → warn → decide → prove)
1. Provision & arm. QA/dispatch selects a product-specific trip profile (thresholds, excursion logic, logging interval) over BLE/mobile or on-device, places the device (and external probe) in the product zone, and arms it; the device gives positive on-device confirmation that it is armed and placed. A trip cannot silently start unmonitored.
2. Sense & log. The device measures internal and external (probe) temperature, humidity, and shock at the configured interval, timestamps each sample against a trusted clock, and writes it to tamper-evident non-volatile storage. Between samples it sleeps to conserve energy (D2/D3).
3. Detect & alert. Two-sided excursion logic (freeze and heat) evaluates product-relevant exposure (time-over/under-threshold or cumulative). On a detected excursion the device captures the event locally and, when connectivity exists, pushes an immediate alert routed and escalated to the right responder (D2).
4. Store & forward. The device uploads on a configurable cadence over whatever bearer is available (LTE-M primary, Wi-Fi, BLE offload), buffering when offline and uploading in intact time order when a link appears — the last mile is never a black box.
5. Decide (offline verdict). At receipt the last-mile user reads an unambiguous accept / reject / quarantine verdict on the device's own display, with no phone, network, login, or manual.
6. Prove & manage. The cloud platform assembles a complete, tamper-evident, attributable trip record (ALCOA+ / 21 CFR Part 11-style audit trail), exports it on demand, and presents fleet status, device readiness (charge/calibration/firmware), and location for management and stray recovery. Signed, fail-safe OTA keeps the fleet current.
7. Return & refurbish (D1). After a trip the device is returned, cleaned, recharged (charge inhibited below 0 °C), triaged for calibration/charge/firmware readiness, and re-provisioned for the next trip.
---
3. Major subsystems
# | Subsystem | Role | Key elements
S1 | Enclosure / mechanical | Houses and protects the electronics; fits inside insulated shippers; sealed and cleanable | Sealed IP-rated housing, external temperature probe + lead, mount/retention, USB-C port seal, labeling
S2 | Power | Stores and manages energy for 90-day cold endurance; safe charge | Low-temperature-capable rechargeable Li-ion cell, Battery Management System (BMS) with protection, USB-C charge path, charge-inhibit-below-0 °C control, fuel gauge
S3 | Compute / MCU + firmware | Runs the capture→detect→decide logic; manages power states, storage, connectivity, OTA | Low-power MCU, real-time clock, secure element/secure boot, firmware application
S4 | Sensing | Measures the monitored environment | Internal temperature sensor, external temperature probe, humidity sensor, 3-axis accelerometer (shock), GNSS/GPS receiver
S5 | Connectivity | Moves data and alerts out; provisioning | LTE-M cellular modem (primary), Wi-Fi, BLE; antennas
S6 | HMI (human-machine interface) | On-device status/verdict and single-action control | OLED display, single glove-friendly button, status indicator
S7 | Storage | Retains the full trip log tamper-evidently | Non-volatile memory (log + config + audit)
S8 | Cloud platform + dashboard | Fleet registry, ingest, alert routing, record assembly, audit trail, OTA, APIs | Device registry, ingest/store-and-forward endpoint, rules/alerting, fleet dashboard, reporting/export, OTA service, RBAC/SSO, audit log
S9 | Mobile / BLE app | Provisioning, configuration, local read, notifications | BLE pairing, trip-profile config, arming aid, alert notifications
External interfaces: the cold-chain product/payload (thermal), the insulated shipper (mechanical fit / RF attenuation), the charger/USB-C source (power), cellular / Wi-Fi networks (RF/data), GNSS constellation (RF), and the enterprise/quality systems the cloud integrates with (data).
---
4. Key performance drivers (set as requirements in this library)
• Measurement accuracy & calibration — temperature accuracy target ±0.5 °C across range (Judgment — confirm against WHO PQS E006 class, revision to confirm); traceable calibration.
• 90-day energy budget at −30 °C — the headline D3 tension; duty-cycled operation; validated by test.
• Excursion-detection logic — two-sided, product-relevant, configurable; false-alarm vs miss balance.
• Alert latency — target ≤ 5 min from detection to cloud notification when connected (Judgment — confirm in validation); connectivity-bound, offline-captured.
• Data integrity, security & auditability — tamper-evident, attributable, signed FW, encrypted transport, Part 11-style audit trail.
• Cold-chain auditability — complete, gap-flagged, exportable record admissible to EU GDP / 21 CFR Part 11 / ALCOA+.
---
5. Cross-references
• System Block Diagram: [system-block-diagram.md](system-block-diagram.md) (§5.1 of the standard).
• Design / product architecture: [design-architecture.md](design-architecture.md) (§7).
• Requirement rationale & traceability: [requirement-rationale-traceability.md](requirement-rationale-traceability.md) (§8).
• Module index & Applicability Register: [index.md](index.md).
Evidence posture (honesty): every quantitative target above is labelled Standard, Cited, or Judgment; unresolved quantities are Evidence Pending and closed by verification/validation. No standard clause number is asserted as fact — standards are cited by designation with revision to confirm.
Requirement Rationale & Traceability — Smart Cold Chain Monitor (SCM)
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · §8 of the Product Requirements Standard. Owner: EM-ENG-001. Two views: (A) a per-requirement rationale table (why each requirement exists), and (B) an upstream→PRD traceability rollup (UN/MR/D → PRD) with the forward placeholders resolved by ORCH-018 / EM-ENG-012 (Stage 18). Each module file carries the full 11-field block including per-requirement forward Traceability; this document is the consolidated map.
---
A. Requirement rationale table
Evidence key: S = Standard · C = Cited · J = Judgment · EP = contains Evidence-Pending value. Full rationale text is in each requirement's block; this is the consolidated why + parent.
PRD-VIS (Product Vision)
ID | Why it exists | Parent | Ground
PRD-VIS-001 | Lock the product class (monitoring/information, not therapeutic/control) | UN-VIS-001, concept §2 | S/J
PRD-VIS-002 | Binding severity-first tie-breaker (beneficiary outcome) | UN-SAFE-001 / STN-020 | S
PRD-VIS-003 | Two-sided (freeze+heat) protection as first-class scope | UN-FUN-004, UN-SAFE-002 | C
PRD-VIS-004 | Make "protect the dose" a measurable detection target | UN-VIS-002, MR-1000 | J/EP
PRD-VIS-005 | Reusable multi-trip lifecycle (D1) | UN-VIS-006, D1 | S
PRD-VIS-006 | Lawful/responsible end-of-life | UN-REG-008, MR-1700 | S/EP
PRD-VIS-007 | Configurable across products without HW forks | UN-FUN-011 | J
PRD-STK (Stakeholders)
ID | Why it exists | Parent | Ground
PRD-STK-001 | Serve least-trained, most safety-critical edge user | UN-STK-005, UN-SAFE-005 | S
PRD-STK-002 | Unmissable arming for dispatch | UN-FUN-013, UN-SAFE-006 | S
PRD-STK-003 | Machine-readable readiness for service/fleet | UN-FUN-017 | S
PRD-STK-004 | Defensible records for QA/regulatory | UN-REG-003 | S
PRD-STK-005 | Enterprise governance for IT/security | UN-SEC-003 | S
PRD-STK-006 | Embeddable module for OEM channel | UN-PHY-003 | J
PRD-ENV (Use Environment)
ID | Why it exists | Parent | Ground
PRD-ENV-001 | Full function −30..+60 °C | UN-ENV-001, UN-PERF-003 | S
PRD-ENV-002 | Survive storage/transport extremes | UN-ENV-001, MR-1300 | J/EP
PRD-ENV-003 | Cold display legibility | UN-ENV-002, UN-PHY-004 | S/EP
PRD-ENV-004 | Humidity/condensation tolerance | UN-ENV-003 | S
PRD-ENV-005 | Transport vibration survival | UN-ENV-004, UN-PHY-006 | S/EP
PRD-ENV-006 | Drop/shock survival | UN-ENV-004, UN-PHY-006 | S/EP
PRD-ENV-007 | IP67 ingress incl. port | UN-ENV-005, UN-PHY-007 | S/J
PRD-ENV-008 | Cleaning-chemical compatibility | UN-ENV-005, UN-PHY-008 | S/EP
PRD-ENV-009 | RF egress inside metallized shipper | UN-ENV-006, UN-PERF-008 | J/EP
PRD-ENV-010 | Offline logging/verdict full trip | UN-ENV-007, UN-FUN-011 | S
PRD-ENV-011 | Ambient-lighting legibility | UN-ENV-008 | J
PRD-ENV-012 | Altitude/pressure operation (air) | UN-ENV-004, D4 | S
PRD-ARCH (Product Architecture)
ID | Why it exists | Parent | Ground
PRD-ARCH-001 | Three-tier composition | concept §4, D2 | J
PRD-ARCH-002 | Realise all device subsystems | concept §4 | S
PRD-ARCH-003 | Edge-autonomous core (offline-first) | UN-FUN-009/011, D2 | S
PRD-ARCH-004 | Duty-cycled energy governor (90-day) | UN-PERF-003, D3 | S
PRD-ARCH-005 | Multi-bearer, LTE-M primary | UN-PERF-008, D2/D4 | S
PRD-ARCH-006 | Split trust boundary (edge vs platform) | UN-SEC-003/004, C9 | S
PRD-ARCH-007 | Trusted time base | UN-FUN-016, UN-REG-003 | S
PRD-ARCH-008 | Safe-state architecture | UN-SAFE-010 | S
PRD-ARCH-009 | Interface conformance to block diagram | block diagram | J
PRD-ARCH-010 | Config-driven variant strategy | PRD-VIS-007 | J
PRD-FUN (Functional)
ID | Why it exists | Parent | Ground
PRD-FUN-001 | Gap-free interval temperature logging | UN-FUN-001 | S
PRD-FUN-002 | Dual-channel temp (product zone) | UN-FUN-003 | S
PRD-FUN-003 | Humidity logging | UN-FUN-002 | S
PRD-FUN-004 | Product-relevant shock capture | UN-FUN-002 | S
PRD-FUN-005 | Configurable two-sided thresholds | UN-FUN-004/011 | S
PRD-FUN-006 | Equal-reliability two-sided detection | UN-FUN-004, UN-SAFE-002 | C
PRD-FUN-007 | Product-relevant excursion logic | UN-FUN-005 | S
PRD-FUN-008 | Immediate alert when connected | UN-FUN-006, D2 | S
PRD-FUN-009 | Alert routing/escalation | UN-FUN-007 | S
PRD-FUN-010 | Offline on-device verdict | UN-FUN-008, UN-SAFE-005 | S
PRD-FUN-011 | Verdict/record never disagree | UN-FUN-008/014, UN-SAFE-001 | S
PRD-FUN-012 | Full-trip offline buffering | UN-FUN-009, D2 | S
PRD-FUN-013 | Cadence upload + multi-bearer offload | UN-FUN-010, D2 | S
PRD-FUN-014 | Trip-profile configuration | UN-FUN-011 | S
PRD-FUN-015 | Arm with placement confirmation | UN-FUN-012, UN-SAFE-006 | S
PRD-FUN-016 | Placement guidance | UN-FUN-013 | S
PRD-FUN-017 | Complete tamper-evident record + gap flag | UN-FUN-014, UN-SEC-001 | S
PRD-FUN-018 | On-demand export | UN-FUN-015 | S
PRD-FUN-019 | Time/location-stamped events | UN-FUN-016, UN-PERF-009 | S/EP
PRD-FUN-020 | Readiness reporting + deploy-block | UN-FUN-017 | S
PRD-FUN-021 | Fleet visibility | UN-FUN-018 | J
PRD-FUN-022 | Verified, resumable, non-destructive FW update | UN-SEC-005, UN-SAFE-010, D2 | S
PRD-UX (User Experience)
ID | Why it exists | Parent | Ground
PRD-UX-001 | Verdict comprehension w/o training | UN-FUN-008, UN-SAFE-005 | S/EP
PRD-UX-002 | No-login, no-manual offline read | UN-SEC-004, C9 | S
PRD-UX-003 | Single-action glove-operable control | UN-PHY-005 | S/EP
PRD-UX-004 | Positive arming feedback | UN-FUN-012, UN-SAFE-006 | S
PRD-UX-005 | Cold/glare/low-light legibility | UN-ENV-002/008, UN-PHY-004 | S/EP
PRD-UX-006 | Prompt interaction response | UN-PERF-006 | J
PRD-UX-007 | Error-preventing configuration | UN-FUN-014, UN-SAFE-006 | S
PRD-UX-008 | Actionable status/error messaging | UN-SAFE-003, UN-FUN-020 | S
PRD-UX-009 | Low-literacy/language-independent cues | UN-SAFE-005 | J
PRD-UX-010 | App/dashboard localization | D4 | J/EP
PRD-UX-011 | Layered info (verdict → detail) | UN-VIS-007, C5 | J
PRD-UX-012 | Passive monitoring reassurance | UN-PERF-011 | J
PRD-MECH (Mechanical)
ID | Why it exists | Parent | Ground
PRD-MECH-001 | Fit inside standard shippers | UN-PHY-001 | S/EP
PRD-MECH-002 | Mass limit | UN-PHY-009 | J/EP
PRD-MECH-003 | Mounting/retention | UN-PHY-002 | S
PRD-MECH-004 | External probe (product zone) | UN-FUN-003, UN-PHY-002 | S/EP
PRD-MECH-005 | IP67 enclosure | UN-PHY-007, UN-ENV-005 | S/J
PRD-MECH-006 | Sealed USB-C port | UN-PHY-007 | S
PRD-MECH-007 | Rugged housing | UN-PHY-006, UN-ENV-004 | S
PRD-MECH-008 | Glove-operable control + readable display | UN-PHY-004/005 | S/EP
PRD-MECH-009 | Cleanable compatible materials | UN-PHY-008, UN-ENV-005 | S/EP
PRD-MECH-010 | Durable ID + safety labeling | UN-PHY-009, UN-REG-010 | S/EP
PRD-MECH-011 | Thermal-coupling supports accuracy | UN-PERF-001, UN-FUN-003 | J/EP
PRD-MECH-012 | Embeddable module interface | UN-PHY-003 | J/EP
PRD-ELEC (Electrical)
ID | Why it exists | Parent | Ground
PRD-ELEC-001 | Low-temp rechargeable chemistry | UN-PERF-003, D3 | S/EP
PRD-ELEC-002 | 90-day endurance capacity | UN-PERF-003, D3 | S/EP
PRD-ELEC-003 | Capacity retention over reuse | UN-PERF-010, D1/D3 | J/EP
PRD-ELEC-004 | USB-C charging | UN-PERF-004, D1 | S/EP
PRD-ELEC-005 | Charge inhibit below 0 °C | UN-SAFE-007, D3 | S
PRD-ELEC-006 | Self-heat-before-charge (eval) | UN-PERF-004, D3 | J/EP
PRD-ELEC-007 | Battery protection (OV/UV/OC/OT/short) | UN-SAFE-008 | S
PRD-ELEC-008 | Power-budget conformance | UN-PERF-003, D3 | S/EP
PRD-ELEC-009 | Fuel-gauge SoC/SoH accuracy | UN-FUN-017, UN-PERF-010 | J/EP
PRD-ELEC-010 | Low-battery graceful degradation | UN-PERF-003, UN-SAFE-010 | S/EP
PRD-ELEC-011 | Rails + brown-out protection | UN-SAFE-010, UN-FUN-014 | S
PRD-ELEC-012 | Status indicator | UN-PERF-011, UN-SAFE-007 | J
PRD-ELEC-013 | Evidence/isolation/ESD | UN-ENV-009, UN-REG-005 | S
PRD-ELEC-014 | Electrical/thermal safety (IEC 62368-1) | UN-SAFE-008, UN-REG-005 | S
PRD-ELEC-015 | UN 38.3 battery transport | UN-REG-006, UN-SAFE-009 | S/EP
PRD-ELEC-016 | Fault-tolerant charge input | UN-SAFE-008, UN-PERF-004 | S
PRD-CONN (Connectivity)
ID | Why it exists | Parent | Ground
PRD-CONN-001 | LTE-M primary bearer | UN-PERF-008, D2/D4 | S/EP
PRD-CONN-002 | Wi-Fi offload | UN-FUN-010 | S/EP
PRD-CONN-003 | BLE provisioning/offload | UN-FUN-011, D2 | S
PRD-CONN-004 | Opportunistic bearer selection | UN-FUN-010, D2 | S
PRD-CONN-005 | GNSS location acquisition | UN-PERF-009, C3 | S/EP
PRD-CONN-006 | Secure provisioning/pairing | UN-SEC-003/005 | S
PRD-CONN-007 | RF performance in shipper | UN-ENV-006, concept §7 | J/EP
PRD-CONN-008 | Store-and-forward on loss | UN-FUN-009, D2 | S
PRD-CONN-009 | Alert-path prioritization | UN-FUN-006, UN-PERF-005 | S
PRD-CONN-010 | Connectivity diagnostics | UN-FUN-018 | J
PRD-CONN-011 | FW-update transport (OTA/USB-C) | UN-SEC-005, D2 | S
PRD-CONN-012 | Regional radio configurability | UN-REG-004, D4 | S
PRD-CLOUD (Cloud Ecosystem)
ID | Why it exists | Parent | Ground
PRD-CLOUD-001 | Device registry/identity | UN-FUN-018, UN-SEC-002/003 | S
PRD-CLOUD-002 | Integrity-preserving ingest | UN-FUN-009/014, D2 | S
PRD-CLOUD-003 | Alert routing/escalation | UN-FUN-007 | S
PRD-CLOUD-004 | Configurable alert rules | UN-FUN-007/011 | S
PRD-CLOUD-005 | Record assembly + export | UN-FUN-014/015, UN-REG-003 | S
PRD-CLOUD-006 | Fleet dashboard/health | UN-FUN-018/017 | J
PRD-CLOUD-007 | Part 11-style audit trail | UN-SEC-009, UN-REG-003 | S
PRD-CLOUD-008 | OTA orchestration + rollback | UN-SEC-005, D2 | S
PRD-CLOUD-009 | Enterprise RBAC/SSO | UN-SEC-003 | S
PRD-CLOUD-010 | Integration APIs | UN-FUN-015, STN-013 | S
PRD-CLOUD-011 | Retention/ownership controls | UN-SEC-007, UN-REG-009 | S/EP
PRD-CLOUD-012 | Platform availability/integrity | UN-FUN-006/014 | S/EP
PRD-MOB (Mobile Application)
ID | Why it exists | Parent | Ground
PRD-MOB-001 | Secure BLE pairing | UN-SEC-003 | S
PRD-MOB-002 | Trip-profile configuration | UN-FUN-011, UN-SAFE-006 | S
PRD-MOB-003 | Arming aid + confirmation | UN-FUN-012 | S
PRD-MOB-004 | Local device read | UN-FUN-017/018 | J
PRD-MOB-005 | Alert notifications | UN-FUN-007 | S
PRD-MOB-006 | Offline app + sync | UN-ENV-007, D2 | S
PRD-MOB-007 | Role-based permissions | UN-SEC-003 | S
PRD-MOB-008 | Accessibility/localization | UN-UX, D4 | J/EP
PRD-SEC (Cybersecurity)
ID | Why it exists | Parent | Ground
PRD-SEC-001 | Tamper-evident records | UN-SEC-001, UN-FUN-014 | S
PRD-SEC-002 | Attribution of records/actions | UN-SEC-002 | S
PRD-SEC-003 | Device identity + secure boot | UN-SEC-005 | S
PRD-SEC-004 | Audit trail (device+cloud) | UN-SEC-009 | S
PRD-SEC-005 | Role-based access control | UN-SEC-003 | S
PRD-SEC-006 | Secure connectivity (TLS) | UN-SEC-006 | S
PRD-SEC-007 | Data protection at rest | UN-SEC-006/007 | S
PRD-SEC-008 | Signed fail-safe OTA | UN-SEC-005, UN-SAFE-010 | S
PRD-SEC-009 | Key/credential lifecycle | UN-SEC-005/006 | S/EP
PRD-SEC-010 | Location/personal-data privacy | UN-SEC-008, UN-REG-009 | S/EP
PRD-SEC-011 | Security monitoring/IR | UN-SEC-009 | J/EP
PRD-SEC-012 | Threat-model-driven design | UN-SEC (all) | J/EP
PRD-REL (Reliability)
ID | Why it exists | Parent | Ground
PRD-REL-001 | 90-day cold endurance (validate by test) | UN-PERF-003, D3 | S/EP
PRD-REL-002 | Cycle life over reuse (validate by test) | UN-PERF-010, D1/D3 | J/EP
PRD-REL-003 | Low-temp charge-safety reliability | UN-SAFE-007, D3 | S
PRD-REL-004 | Fail-safe on fault | UN-SAFE-010 | S
PRD-REL-005 | Storage integrity (power-fail-safe) | UN-FUN-009/014 | S
PRD-REL-006 | Watchdog/recovery | UN-PERF-011, UN-SAFE-003 | S
PRD-REL-007 | Availability/no silent outage | UN-PERF-011, UN-SAFE-003 | S
PRD-REL-008 | Graceful degradation under low energy | UN-PERF-003 | S/EP
PRD-REL-009 | Robustness over reuse | UN-PERF-010, UN-PHY-006/008 | S
PRD-REL-010 | Reliability program/prediction | UN-QUAL, UN-PERF-010 | J/EP
PRD-REL-011 | Communication robustness | UN-ENV-007, UN-FUN-009 | S
PRD-REL-012 | Platform reliability/durability | UN-FUN-006/014 | S/EP
PRD-SAFE (Safety — PRELIMINARY, reconcile with HAL Stage 10)
ID | Why it exists | Parent | Ground
PRD-SAFE-001 | Don't mislead disposition (anchor) | UN-SAFE-001 / STN-020 | S
PRD-SAFE-002 | Reliable freeze detection as control | UN-SAFE-002, UN-FUN-004 | C
PRD-SAFE-003 | No false reassurance (honest gaps) | UN-SAFE-003 | S
PRD-SAFE-004 | Alarm balance (no fatigue miss) | UN-SAFE-004, C8 | S/EP
PRD-SAFE-005 | Use-error protection (edge user) | UN-SAFE-005 | S/EP
PRD-SAFE-006 | Prevent unmonitored trip start | UN-SAFE-006 | S
PRD-SAFE-007 | No charge in unsafe cold | UN-SAFE-007, D3 | S
PRD-SAFE-008 | No fire/thermal hazard | UN-SAFE-008 | S
PRD-SAFE-009 | Safe/compliant Li transport | UN-SAFE-009, UN-REG-006 | S/EP
PRD-SAFE-010 | Fail-safe + no data loss | UN-SAFE-010 | S
PRD-SAFE-011 | Residual-risk acceptability | UN-SAFE (all) | S
PRD-SAFE-012 | Dead-in-cold as safety-relevant failure | UN-PERF-003, UN-SAFE-001, D3 | S/EP
PRD-REG (Regulatory)
ID | Why it exists | Parent | Ground
PRD-REG-001 | WHO PQS eligibility | UN-REG-001 | S/EP
PRD-REG-002 | Per-market classification basis | UN-REG-002, D4 | S/EP
PRD-REG-003 | GDP + Part 11 / ALCOA+ records | UN-REG-003 | S
PRD-REG-004 | Radio approvals per market | UN-REG-004, D4 | S/EP
PRD-REG-005 | EMC + electrical/thermal safety | UN-REG-005, UN-SAFE-008 | S
PRD-REG-006 | UN 38.3 / IATA transport | UN-REG-006 | S/EP
PRD-REG-007 | Traceable calibration | UN-REG-007 | S/EP
PRD-REG-008 | RoHS/WEEE/battery EoL | UN-REG-008 | S/EP
PRD-REG-009 | Regional data-protection | UN-REG-009 | S/EP
PRD-REG-010 | Labeling/documentation per regime | UN-REG-010 | S/EP
PRD-REG-011 | QMS/design-controls basis | UN-REG-002 | S/EP
PRD-REG-012 | CSV (GAMP 5) | UN-SEC-003, UN-REG-003 | S
PRD-MFG (Manufacturing)
ID | Why it exists | Parent | Ground
PRD-MFG-001 | DFM/DFA + testability | UN-QUAL | J
PRD-MFG-002 | Production calibration w/ traceability | UN-REG-007, UN-PERF-001 | S
PRD-MFG-003 | Secure FW/credential provisioning | UN-SEC-003/005 | S
PRD-MFG-004 | Serialization/traceability | UN-FUN-017/018 | S
PRD-MFG-005 | End-of-line function+safety test | UN-QUAL, PRD-SAFE | S
PRD-MFG-006 | Packaging protects device/battery | UN-REG-006 | S
PRD-MFG-007 | Process control + PFMEA | UN-QUAL | S
PRD-MFG-008 | Yield/first-pass target | UN-BIZ, MR-900 | J/EP
PRD-SVC (Serviceability)
ID | Why it exists | Parent | Ground
PRD-SVC-001 | Refurbishment turnaround loop | UN-SVC-001, D1 | S/EP
PRD-SVC-002 | Cleaning without damage | UN-PHY-008 | S/EP
PRD-SVC-003 | Predictable recharge/readiness | UN-PERF-004, D3 | S/EP
PRD-SVC-004 | Readiness triage + deploy-block | UN-FUN-017 | S
PRD-SVC-005 | Recalibration (traceable) | UN-REG-007, UN-PERF-002 | S/EP
PRD-SVC-006 | Service diagnostics/logs | UN-SVC, UN-FUN-017 | J
PRD-SVC-007 | FW maintenance/upgrade path | UN-SEC-005, D2 | S
PRD-SVC-008 | Safe decommission/EoL | UN-REG-008 | S/EP
PRD-PERF (Performance)
ID | Why it exists | Parent | Ground
PRD-PERF-001 | Temp accuracy ±0.5 °C | UN-PERF-001, WHO PQS E006 | J/EP
PRD-PERF-002 | Accuracy stability over range/interval | UN-PERF-002 | S/EP
PRD-PERF-003 | Temp resolution 0.1 °C | UN-PERF-001 | J
PRD-PERF-004 | Alert latency ≤5 min (connected) | UN-PERF-005, UN-FUN-006 | J/EP
PRD-PERF-005 | 90-day energy performance | UN-PERF-003, D3 | S/EP
PRD-PERF-006 | Full-offline-trip log capacity | UN-PERF-007 | S/EP
PRD-PERF-007 | Interaction responsiveness ≤3 s | UN-PERF-006 | J/EP
PRD-PERF-008 | Humidity accuracy ±5% RH | UN-FUN-002 | J
PRD-PERF-009 | Availability/no silent outage | UN-PERF-011, UN-SAFE-003 | S/EP
PRD-PERF-010 | Location cadence within budget | UN-PERF-009, C3 | S/EP
PRD-PERF-011 | Data-delivery success over trip | UN-PERF-008, UN-ENV-006 | J/EP
PRD-PERF-012 | Charge time to readiness | UN-PERF-004, D1 | J/EP
PRD-QUAL (Quality)
ID | Why it exists | Parent | Ground
PRD-QUAL-001 | Conformance to acceptance criteria | UN-QUAL, PRD-VER | S
PRD-QUAL-002 | Durability over reuse life | UN-QUAL, UN-PERF-010 | S
PRD-QUAL-003 | Fit and finish | UN-QUAL | J/EP
PRD-QUAL-004 | Label legibility over life | UN-PHY-009, UN-REG-010 | S
PRD-QUAL-005 | Field defect/escape target | UN-QUAL, MR-1000 | J/EP
PRD-QUAL-006 | Quality records/traceability | UN-QUAL, UN-REG-003 | S
PRD-QUAL-007 | Incoming/component QC | UN-QUAL, PRD-SAFE-008 | S
PRD-VER (Verification)
ID | Why it exists | Parent | Ground
PRD-VER-001 | Every req verifiable + method | UN-VAL, 29148 | S
PRD-VER-002 | Bidirectional traceability | UN-VAL, ISO 14971/29148 | S
PRD-VER-003 | Method allocation to risk | PRD-SAFE, ISO 14971 | S
PRD-VER-004 | D3 validate-by-test | PRD-REL/ELEC, D3, MR-VAL-006 | S
PRD-VER-005 | RF-in-container qualification | PRD-CONN-007, concept §7 | J/EP
PRD-VER-006 | Environmental/regulatory suite | PRD-ENV/ELEC/REG, D4 | S/EP
PRD-VER-007 | Safety verification reconciled w/ HAL | PRD-SAFE, ISO 14971 | S
PRD-VER-008 | Validation of intended use/needs | UN-VAL, UN-FUN/UX/SAFE | S
---
B. Upstream → PRD traceability rollup
Each UN module (Stage 04) is realised forward by the PRD modules/requirements below. (MR/STN/VOC/D evidence is carried in each UN file and each PRD requirement's Parent field.)
Upstream (Stage 04 / concept) | Realised by (PRD)
UN-000 Product Vision (VIS) | PRD-VIS-001..007; PRD-ARCH-001/010
UN-100 Stakeholders (STK) | PRD-STK-001..006 (+ across UX/FUN/SVC/SEC/CLOUD/MECH)
UN-200 Personas (PER) | PRD-UX-001/003/009; PRD-STK-001; PRD-SAFE-005
UN-300 Jobs To Be Done (JTBD) | PRD-FUN (capture→prove); PRD-CLOUD; PRD-PERF
UN-400 Use Environments (ENV) | PRD-ENV-001..012; PRD-MECH; PRD-CONN-007
UN-500 User Workflow (FLOW) | PRD-FUN-014/015/016; PRD-UX-004/007; PRD-MOB-002/003
UN-600 Functional (FUN) | PRD-FUN-001..022; PRD-CLOUD-002/003/005; PRD-PERF
UN-700 User Experience (UX) | PRD-UX-001..012; PRD-MECH-008; PRD-MOB-008
UN-800 Performance (PERF) | PRD-PERF-001..012; PRD-REL-001..003/007; PRD-ELEC-002/008
UN-900 Safety (SAFE) | PRD-SAFE-001..012 (preliminary→HAL); PRD-ELEC-005/007/014; PRD-REL-003/004
UN-1000 Security & Privacy (SEC) | PRD-SEC-001..012; PRD-CLOUD-007/009/011; PRD-CONN-006
UN-1100 Physical (PHY) | PRD-MECH-001..012; PRD-ENV-003/007; PRD-UX-003/005
UN-1200 Service (SVC) | PRD-SVC-001..008; PRD-FUN-020; PRD-ELEC-004/005
UN-1300 Regulatory (REG) | PRD-REG-001..012; PRD-ELEC-015; PRD-SAFE-009; PRD-QUAL-006
UN-1400 Quality (QUAL) | PRD-QUAL-001..007; PRD-MFG; PRD-REL-010
UN-1500 Ecosystem (ECO) | PRD-CLOUD-010; PRD-STK-006; PRD-MECH-012; PRD-FUN-018
UN-1600 Business (BIZ) | PRD-VIS-005/006/007; PRD-MFG-008; PRD-QUAL-005
UN-1700 Validation (VAL) | PRD-VER-001..008; PRD-PERF; PRD-REL-001/002
Concept D1 (reusable) | PRD-VIS-005; PRD-SVC-001..008; PRD-REL-002; PRD-ELEC-004
Concept D2 (store-and-forward) | PRD-ARCH-003/005; PRD-FUN-008/012/013; PRD-CONN-004/008
Concept D3 (low-temp battery tension) | PRD-ELEC-001/002/005/006; PRD-REL-001/002/003; PRD-SAFE-007/012; PRD-PERF-005
Concept D4 (multi-market) | PRD-REG-001..006/010; PRD-CONN-012; PRD-ENV-012
---
C. Forward placeholders (resolved by ORCH-018 / EM-ENG-012, Stage 18)
Each PRD requirement's Traceability field carries forward placeholders that the end-to-end matrix resolves:
• → System/Subsystem Requirements (SYS-###, SSR-/HWR-/SWR-/IFR-###, Stage 06)
• → Software Requirements (SRS-<CODE>-NNN, ORCH-012)
• → Design Outputs (DO-<CODE>-NNN, ORCH-013)
• → Hazard Analysis / Risk-Control Requirements (HA-<CODE>-NNN / RCR-###, ORCH-017, Stage 10) — mandatory for all PRD-SAFE and safety-tagged requirements
• → Verification / Validation (PV-<CODE>-NNN / VAL-<CODE>-NNN, ORCH-014 / ORCH-015)
Note on Safety: PRD-SAFE requirements are preliminary; the definitive safety/risk-control requirements are produced in the Hazard Analysis Library (HAL, Stage 10) and re-baselined here, each tracing to a HA-###/RCR-### record. Until then they are marked "to be reconciled with HAL (Stage 10)."
System Block Diagram — Smart Cold Chain Monitor (SCM)
Master Product Requirements Library — SCM · Stage 05 (ORCH-011) · §5.1 of the Product Requirements Standard. Owner: Chief Systems Engineer (EM-ENG-001). Interfaces are typed: power, data, RF, thermal, mechanical. Forward to Architecture (Stage 06) and ORCH-018 / EM-ENG-012.
---
1. Interface legend
Type | Meaning | Notation in diagram
PWR | Power / energy transfer | -- PWR ...
DATA | Digital data / signal (bus, log, config) | -- DATA ...
RF | Radio-frequency link (wireless) | -. RF ... .-> (dashed)
THRM | Thermal coupling (heat/temperature exposure) | == THRM ==> (thick)
MECH | Mechanical mount / fit / retention | -- MECH ...
---
2. System block diagram
mermaid flowchart TB %% ===== External world ===== subgraph EXT["External environment"] PROD["Cold-chain product / payload<br/>(vaccine in shipper)"] SHIP["Insulated shipping container<br/>(RF-attenuating, thermal barrier)"] CHG["USB-C charge source<br/>(depot charger)"] NET["Cellular / Wi-Fi networks"] GNSS["GNSS constellation"] ENT["Enterprise / quality systems<br/>(ERP, QMS, partner APIs)"] end
%% ===== SCM Device ===== subgraph DEV["SCM Device (inside the shipper)"] subgraph S1["S1 Enclosure / mechanical"] ENC["Sealed IP-rated housing<br/>+ mount / retention + labeling"] PROBE["External temperature probe + lead"] PORT["Sealed USB-C port"] end subgraph S2["S2 Power"] BATT["Low-temp Li-ion cell"] BMS["BMS + protection<br/>charge-inhibit &lt;0 °C + fuel gauge"] PMIC["Charge path / PMIC"] end subgraph S3["S3 Compute / MCU + firmware"] MCU["Low-power MCU + RTC"] SE["Secure element / secure boot"] FW["Firmware application"] end subgraph S4["S4 Sensing"] TINT["Internal temp sensor"] TEXT["External probe temp channel"] HUM["Humidity sensor"] ACC["3-axis accelerometer (shock)"] GPS["GNSS / GPS receiver"] end subgraph S5["S5 Connectivity"] LTE["LTE-M modem (primary)"] WIFI["Wi-Fi"] BLE["BLE"] ANT["Antennas"] end subgraph S6["S6 HMI"] OLED["OLED display"] BTN["Single glove-friendly button"] LED["Status indicator"] end subgraph S7["S7 Storage"] NVM["Non-volatile log + config + audit"] end end
%% ===== Cloud & mobile ===== subgraph CLD["S8 Cloud platform + dashboard"] REG["Device registry"] ING["Ingest / store-and-forward endpoint"] RULE["Rules / alerting / escalation"] DASH["Fleet dashboard + reporting/export"] OTA["OTA service"] AUD["RBAC/SSO + audit log"] end subgraph MOB["S9 Mobile / BLE app"] APP["Provision / configure / read / notify"] end
%% ===== Thermal ===== PROD == THRM ==> PROBE PROD == THRM ==> ENC SHIP == THRM ==> ENC ENC == THRM ==> TINT PROBE == THRM ==> TEXT
%% ===== Mechanical ===== SHIP -- MECH fit/retain --- ENC ENC -- MECH --- PROBE ENC -- MECH --- PORT
%% ===== Power ===== CHG -- PWR USB-C --> PORT PORT -- PWR --> PMIC PMIC -- PWR --> BMS BMS -- PWR --> BATT BATT -- PWR --> BMS BMS -- PWR rail --> MCU BMS -- PWR rail --> S4 BMS -- PWR rail --> S5 BMS -- PWR rail --> S6 BMS -- PWR rail --> S7
%% ===== Data (device internal) ===== TINT -- DATA --> MCU TEXT -- DATA --> MCU HUM -- DATA --> MCU ACC -- DATA --> MCU GPS -- DATA --> MCU BMS -- DATA charge/gauge --> MCU MCU -- DATA --> NVM MCU -- DATA --> OLED MCU -- DATA --> LED BTN -- DATA --> MCU MCU -- DATA --> SE SE -- DATA --> FW MCU <-- DATA --> LTE MCU <-- DATA --> WIFI MCU <-- DATA --> BLE
%% ===== RF ===== LTE -. RF .- ANT WIFI -. RF .- ANT BLE -. RF .- ANT GPS -. RF .- ANT ANT -. RF thru shipper .-> SHIP SHIP -. RF attenuated .-> NET GNSS -. RF .-> ANT ANT -. RF .- NET BLE -. RF .-> APP
%% ===== Cloud / mobile data ===== NET -- DATA --> ING ING --> REG ING --> RULE RULE --> DASH OTA -- DATA signed FW --> NET NET -- DATA --> LTE DASH -- DATA --> ENT AUD --- DASH APP -- DATA --> ING
Caption. Figure 1 — SCM system block diagram. The device (S1–S7) sits inside the insulated shipper, thermally coupled to the payload (THRM) and mechanically retained in it (MECH). Energy flows from a USB-C source through the sealed port and BMS to the cell and the rails (PWR); the BMS inhibits charge below 0 °C (D3). Sensor channels feed the MCU (DATA), which logs to non-volatile memory, drives the OLED verdict, and moves data/alerts out through LTE-M/Wi-Fi/BLE (RF), attenuated by the metallized shipper. The cloud ingests store-and-forward data, routes/escalates alerts, assembles the audit-trailed record, serves the fleet dashboard/export and enterprise APIs, and pushes signed OTA. The mobile/BLE app handles provisioning, configuration, local read, and notifications.
---
3. Interface inventory (typed)
# | From | To | Type | Notes
I-01 | Payload / shipper | Enclosure & probe | THRM | The monitored exposure; probe represents product zone (PRD-FUN, PRD-MECH)
I-02 | Shipper | Enclosure | MECH | Fit + retention inside standard shippers (PRD-MECH, PRD-ENV)
I-03 | USB-C source | Charge port → PMIC → BMS → cell | PWR | Charge path; inhibited below 0 °C (PRD-ELEC, PRD-SAFE)
I-04 | BMS | MCU / sensing / connectivity / HMI / storage | PWR | Regulated rails; duty-cycled (PRD-ELEC, PRD-REL)
I-05 | Sensors | MCU | DATA | Temp int/ext, humidity, accel, GNSS (PRD-FUN, PRD-PERF)
I-06 | MCU | Non-volatile storage | DATA | Tamper-evident log + config + audit (PRD-FUN, PRD-SEC)
I-07 | MCU | OLED / LED / button | DATA | On-device verdict and single-action control (PRD-UX)
I-08 | Antennas | Cellular/Wi-Fi/GNSS | RF | Attenuated by metallized shipper (PRD-CONN, PRD-ENV)
I-09 | Device | BLE app | RF/DATA | Provisioning, config, local read (PRD-CONN, PRD-MOB)
I-10 | Device | Cloud | DATA | Store-and-forward upload, alerts, OTA (PRD-CONN, PRD-CLOUD)
I-11 | Cloud | Enterprise/quality systems | DATA | Export, API, integration (PRD-CLOUD)
I-12 | OTA service | Device firmware | DATA | Signed, resumable, fail-safe update (PRD-SEC, PRD-CLOUD)
Evidence: interfaces derive from the concept subsystem list (§4) and the Stage-04 needs. RF-in-container performance and antenna strategy are Evidence Pending (concept §7; UN-ENV-006) and closed by the RF-in-container test.

## 6 System Requirements
System Requirements
How those product promises are split across the working parts of the system.
↩ Contents
In this section:
• System Requirements — Smart Cold Chain Monitor (SCM)
• System Requirements — Connectivity & RF (SYS-104 … SYS-118)
• System Requirements — Data, Records, Time & Storage (SYS-119 … SYS-131)
• System Requirements — Environmental & Mechanical-Environment (SYS-069 … SYS-085)
• System Requirements — Functional (SYS-001 … SYS-032)
• System Requirements — Interfaces, typed (SYS-051 … SYS-068)
• System Requirements — Performance & Timing (SYS-033 … SYS-050)
• System Requirements — Power / Energy / Battery / Charge (SYS-086 … SYS-103)
• System Requirements — Regulatory & Compliance (SYS-177 … SYS-190)
• System Requirements — Reliability & Availability (SYS-163 … SYS-176)
• System Requirements — Safety (SYS-149 … SYS-162)
• System Requirements — Security & Data Integrity (SYS-132 … SYS-148)
• Decomposition & Allocation Matrix — System Requirements (SCM)
System Requirements — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 06 — System Requirements · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Category owners: Principal Engineers per discipline; EM-ENG-006 (Regulatory), EM-ENG-008 (Security) · Status: Authored — instantiated, product-specific. Authoritative system-level baseline decomposed and allocated from the Master Product Requirements Library (Stage 05).
This library converts the 220 product requirements (PRD-<CODE>-NNN, Stage 05) into system requirements (SYS-###) that state what the SCM system as a whole shall do/be, and allocate each to the subsystem(s) that will realise it. It is the source from which the Subsystem/Detailed Requirements and the Software Requirements Library (Stage 07), the System Architecture (Stage 08), Design Outputs (ORCH-013), Hazard Analysis (ORCH-017), ERFKL (ORCH-016), and Verification/Validation (ORCH-014/015) are derived. It is based on 00-concept/concept-and-decisions.md (foundational decisions D1–D4).
---
Requirement field set (every SYS requirement carries all of these)
• ID — SYS-### (unique, zero-padded, sequential across the library).
• Requirement Statement — a single "shall", atomic, testable/measurable, implementation-neutral.
• Parent — the product requirement(s) it decomposes (PRD-<CODE>-NNN), traced upward to UN/STN/MR/D#.
• Allocation Target — subsystem(s) that realise it: S1 Enclosure/Mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS (system-integration / multiple).
• Type — Derived (engineering necessity not explicit in the parent) or Allocated (direct decomposition of a stated product requirement).
• Interface refs — the typed system interfaces it touches (I-01…I-12, block diagram §3; and Stage-08 INT-###).
• Safety/Security classification — Safety / Security / Both / None (drives risk governance).
• Criticality — Critical (safety, loss-of-function, or compliance-blocking) · High · Medium · Low (severity-first, CLAUDE.md rule 3).
• Verification Method — Test / Analysis / Inspection / Demonstration (forward to PVL PV-###).
• Status — Baselined · Preliminary (reconcile with HAL) · Evidence-Pending (numeric target unresolved).
• Rationale — the physics, standard, need, or decision that grounds it, with a evidence label (Standard / Cited / Judgment / Evidence-Pending).
---
Library structure (files & ID ranges)
# | File | Category | ID range | Reqs
1 | [SYS-functional.md](SYS-functional.md) | Functional (capture→detect→warn→decide→prove) | SYS-001…SYS-032 | 32
2 | [SYS-performance.md](SYS-performance.md) | Performance & timing | SYS-033…SYS-050 | 18
3 | [SYS-interface.md](SYS-interface.md) | Interfaces (typed I-01…I-12) | SYS-051…SYS-068 | 18
4 | [SYS-environmental.md](SYS-environmental.md) | Environmental & mechanical-environment | SYS-069…SYS-085 | 17
5 | [SYS-power-energy.md](SYS-power-energy.md) | Power / energy / battery / charge | SYS-086…SYS-103 | 18
6 | [SYS-connectivity.md](SYS-connectivity.md) | Connectivity & RF | SYS-104…SYS-118 | 15
7 | [SYS-data.md](SYS-data.md) | Data, records, time & storage | SYS-119…SYS-131 | 13
8 | [SYS-security.md](SYS-security.md) | Security & data integrity | SYS-132…SYS-148 | 17
9 | [SYS-safety.md](SYS-safety.md) | Safety (preliminary — reconcile w/ HAL) | SYS-149…SYS-162 | 14
10 | [SYS-reliability.md](SYS-reliability.md) | Reliability & availability | SYS-163…SYS-176 | 14
11 | [SYS-regulatory.md](SYS-regulatory.md) | Regulatory & compliance | SYS-177…SYS-190 | 14
— | [decomposition-allocation-matrix.md](decomposition-allocation-matrix.md) | PRD→SYS→subsystem→SRS/HWR/MER/FWR/IFR allocation matrix | — | —
| Total | 11 categories | SYS-001…SYS-190 | 190
---
Applicability & evidence posture (honesty)
All eleven categories are Applicable — the SCM is a globally distributed, regulated, reusable, connected, safety-significant IoT device with device + cloud + mobile tiers. Per 00-governance/evidence-standard.md and hallucination-policy.md, every requirement's rationale is labelled Standard, Cited, Judgment, or Evidence-Pending; no standard clause number is asserted as fact (standards are cited by designation, revision to confirm), and no unresolved quantity is invented.
The D3 −30 °C / 90-day battery tension is the top program risk and appears as an explicit, testable requirement set — SYS-086/087/090/091/092 (power), SYS-045/046 (performance), SYS-149-line safety, and SYS-165/166/167 (reliability) — all validated by test (forward to PVL/PVaL). RF-in-container link performance (SYS-104/109/117, SYS-069-line environmental) is the principal open architectural risk and is Evidence-Pending, closed by the RF-in-container qualification test (PRD-VER-005).
---
Traceability (chain position)
Backward: every SYS-### names its Parent product requirement(s) PRD-<CODE>-NNN, which trace to User Needs (UN-<CODE>-NNN, Stage 04) and D1–D4. Forward: each SYS seeds Subsystem/Detailed requirements (HWR-###, MER-###, FWR-###, IFR-###), Software Library requirements (SRS-<CODE>-NNN), Architecture elements (FUN-###, CMP-###, INT-###, Stage 08), Design Outputs (DO-<CODE>-NNN), and Verification (PV-###) / Validation (VAL-###). The end-to-end matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18). See [decomposition-allocation-matrix.md](decomposition-allocation-matrix.md).
Flow: PRD (05) → SYS (06, this library) → Subsystem/SRS (07) → Architecture (08) → Design Outputs (09) → Hazard Analysis (10) / ERFKL (11) → Risk-Control Requirements (12) → Verification (13) / Validation (14). Outputs feed Design Outputs (Stage 09) and ORCH-018/EM-ENG-012.
System Requirements — Connectivity & RF (SYS-104 … SYS-118)
SCM · Stage 06 · Owner: EM-ENG-001 (integrator) · Category owner: Principal Systems/RF Engineer. Connectivity and radio system requirements decomposed/allocated from PRD-CONN (800) with RF-environment inputs from PRD-ENV-009, based on foundational decisions D2 (offline-first: local logging + scheduled-cadence upload + immediate excursion alerts; alert latency is connectivity-bound) and D4 (multi-market: WHO PQS / EU GDP / US / air-transport, with FCC Part 15 / CE RED radio regimes). Field set per [index.md](index.md).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Evidence posture — RF-in-container is the principal open architectural risk. The device operates inside a metallized, insulated shipper (RF-hostile). The link budget, antenna strategy, and in-container GNSS performance are not yet quantified; those requirements (SYS-104, SYS-109, SYS-117) are held Evidence-Pending and are closed by the RF-in-container qualification test (PRD-VER-005). Radio standards are cited by designation only (revision to confirm); no clause number is asserted as fact.
---
1. RF performance inside the shipper (principal open risk)
SYS-104 — RF link budget and delivery-success from inside a representative shipper
Requirement Statement | The system shall achieve, from inside a representative metallized/insulated shipping container, an end-to-end link budget sufficient to deliver buffered telemetry and excursion alerts to the cloud at a per-trip delivery-success rate meeting a target to be fixed by the RF-in-container test, without loss of any locally captured record.
Parent | PRD-CONN-007, PRD-ENV-009, PRD-CONN-001
Allocation Target | S5 (radio + antenna), S1 (enclosure/antenna keep-out), SYS (link integration)
Type | Allocated
Interface refs | I-08 (device→WAN/cloud RF), I-01 (thermal/enclosure boundary shared with antenna)
Safety/Security classification | Safety (delayed/undelivered alert is a hazard contributor)
Criticality | Critical
Verification Method | Test (RF-in-container, PRD-VER-005), Analysis (link-budget model)
Status | Evidence-Pending (numeric attenuation, achievable link margin, and delivery-success target unresolved)
Rationale | RF attenuation inside a metallized/insulated shipper can exceed the fade margin of a low-power WAN radio; if egress fails, offline capture (D2, SYS-013) still preserves the record but alerts are deferred. The attenuation figure and the achievable margin are the principal open architectural unknowns. Evidence-Pending — closed by PRD-VER-005; based on D2 and concept §7.
SYS-109 — Antenna strategy including an external-antenna option
Requirement Statement | The system shall implement an antenna strategy (integral antenna(s) for LTE-M, Wi-Fi, BLE, and GNSS, with provision for a qualified external-antenna option) that meets the in-container link-budget target of SYS-104 across all active bands and market band plans.
Parent | PRD-CONN-007, PRD-ENV-009, PRD-CONN-012
Allocation Target | S5 (antenna + RF front-end), S1 (mounting, connector seal, keep-out)
Type | Derived
Interface refs | I-08 (RF egress), I-01 (enclosure/antenna interface, IP-sealed external connector)
Safety/Security classification | Safety (link availability underpins alert delivery)
Criticality | Critical
Verification Method | Test (antenna efficiency / total radiated power in-container), Analysis
Status | Evidence-Pending (integral-only vs external-antenna decision, efficiency, and IP-sealed connector unresolved)
Rationale | If an integral antenna cannot close the SYS-104 budget inside the shipper, an external antenna (via an IP67-sealed connector, PRD-ENV-007) may be required; this is a live architectural branch. Multi-band operation (D4) constrains the antenna to cover LTE-M, Wi-Fi, BLE, and GNSS bands per market. Evidence-Pending — closed by PRD-VER-005; flag Mechanical (enclosure seal) and Regulatory (radiated-power conformance).
SYS-117 — GNSS location acquisition duty-cycled within the energy budget (in-container performance pending)
Requirement Statement | The system shall acquire a GNSS/GPS location fix on a configurable and event-triggered cadence within the energy allocated for location, record the fix quality with each attempt, and shall degrade gracefully (record best-available or last-known location with a quality flag) when a fix cannot be obtained inside the shipper.
Parent | PRD-CONN-005, PRD-ENV-009, PRD-CONN-004
Allocation Target | S4 (GNSS receiver + antenna), S3 (scheduling, energy gating), S5 (assist data)
Type | Allocated
Interface refs | I-08 (GNSS RF + assist), I-05 (receiver→MCU)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test (GNSS-in-container fix rate / time-to-fix), Analysis (energy per fix)
Status | Evidence-Pending (in-container fix availability, time-to-fix, and per-fix energy cost unresolved)
Rationale | GNSS signals are the most attenuated of the active bands inside a metallized shipper, and each fix is energy-costly against the D3 90-day budget; location is fleet/attribution support, never a safety function, so it must fail soft. In-container fix performance may drive the external-antenna decision (SYS-109). Evidence-Pending — closed by PRD-VER-005; consistent with SYS-029.
2. Bearers
SYS-105 — LTE-M primary wide-area bearer
Requirement Statement | The system shall use LTE-M (3GPP Cat-M1) as its primary wide-area bearer for store-and-forward upload and excursion-alert delivery, operating in a duty-cycled low-power mode consistent with the 90-day energy budget.
Parent | PRD-CONN-001, PRD-ARCH-005
Allocation Target | S5 (LTE-M modem + SIM), S3 (session/power scheduling), S8 (ingest endpoint)
Type | Allocated
Interface refs | I-08 (device→WAN), I-10 (WAN→cloud API)
Safety/Security classification | Safety (primary alert path)
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | LTE-M is a low-power WAN suited to duty-cycled operation across global lanes (D2/D4); power-saving mode (PSM/eDRX-class behaviour) preserves the energy budget. Standard (3GPP LTE-M/Cat-M1; revision to confirm). Coverage per market carried as Evidence-Pending under SYS-116/SYS-118-roaming.
SYS-106 — Wi-Fi opportunistic offload
Requirement Statement | The system shall support Wi-Fi (IEEE 802.11) for opportunistic bulk-data offload to a provisioned network where available, and shall prefer it for backlog upload when it is the lowest-energy path that satisfies the cadence target.
Parent | PRD-CONN-002, PRD-CONN-004
Allocation Target | S5 (Wi-Fi radio), S3 (bearer selection), S8 (ingest)
Type | Allocated
Interface refs | I-08 (device→WAN), I-10 (WAN→cloud API)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Baselined
Rationale | Higher-throughput, lower-cost offload at depots/facilities drains buffered backlog and preserves the energy budget (C2). Standard (IEEE 802.11 Wi-Fi; revision to confirm). Enterprise-Wi-Fi authentication is Evidence-Pending (see PRD-CONN-002).
SYS-107 — BLE provisioning and short-range offload
Requirement Statement | The system shall support Bluetooth Low Energy for local provisioning and configuration, local record read, and short-range data offload to a paired app or gateway, over an authenticated pairing (per SYS-110).
Parent | PRD-CONN-003, PRD-MOB-001
Allocation Target | S5 (BLE radio), S9 (mobile/gateway), S3 (session)
Type | Allocated
Interface refs | I-09 (device↔BLE/mobile)
Safety/Security classification | Both (provisioning is an attack surface)
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | Frictionless local setup and offload where WAN is absent (D2); the edge-read path stays login-free (AD-5) while configuration is access-controlled. Standard (Bluetooth Core / BLE; revision to confirm). Pairing security in SYS-110 / PRD-SEC.
3. Bearer selection & offload behaviour
SYS-108 — Opportunistic multi-bearer selection (no single bearer required)
Requirement Statement | The system shall select among the available bearers — LTE-M (primary), then Wi-Fi, then BLE — to deliver buffered data and alerts, preferring the lowest-energy path that satisfies the applicable cadence/latency target, and shall not require any single bearer to be present for the device to function or to eventually deliver its record.
Parent | PRD-CONN-004, PRD-ARCH-005, PRD-FUN-013
Allocation Target | S3 (selection policy), S5 (radios), S8 (ingest)
Type | Allocated
Interface refs | I-08 (WAN), I-09 (BLE), I-10 (cloud API)
Safety/Security classification | Safety (delivery resilience underpins alert path)
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Energy-aware, fail-over bearer choice keeps data flowing under intermittent coverage without exhausting the budget (D2/AD-1); offline-first means loss of all bearers degrades to buffered capture (SYS-111), never data loss. Judgment. Selection policy is configurable.
SYS-111 — Store-and-forward on connectivity loss (buffer, resume, no data loss)
Requirement Statement | The system shall detect connectivity loss, continue buffering samples, events, and alerts to non-volatile storage, and on reconnection resume upload in intact chronological order with idempotent (exactly-once) ingest, such that no record is duplicated or lost across any number of connectivity interruptions.
Parent | PRD-CONN-008, PRD-FUN-012, PRD-CLOUD-002
Allocation Target | S3 (buffer/resume logic), S5 (transport), S7 (buffer), S8 (idempotent ingest)
Type | Allocated
Interface refs | I-06 (MCU→storage), I-10 (cloud API)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Intermittent links must never corrupt or lose the evidentiary record (D2); ordering, resume, and de-duplication preserve ALCOA+ completeness and consistency. Standard (ALCOA+ / EU GDP; revision to confirm). Consistent with SYS-018/SYS-020.
SYS-112 — Alert-path prioritization (alerts before bulk telemetry)
Requirement Statement | When connectivity or energy is constrained, the system shall prioritise delivery of excursion/shock alerts ahead of routine bulk telemetry over the selected bearer, subject to the energy budget, so that a rescue-enabling alert is never queued behind bulk data.
Parent | PRD-CONN-009, PRD-FUN-008, PRD-PERF-004
Allocation Target | S3 (queue prioritisation), S5 (transport), S8 (alert ingest)
Type | Allocated
Interface refs | I-08 (WAN), I-10 (cloud API)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL — delayed-alert hazard)
Rationale | Alert latency is connectivity-bound (D2): the alert is captured offline the instant it is detected (SYS-013) and delivered as soon as a bearer is available, ahead of backlog; a late alert is a recognised hazard. Judgment; forward to HAL (Stage 10).
4. Provisioning, diagnostics & transport
SYS-110 — Secure provisioning and authenticated pairing over any bearer
Requirement Statement | The system shall permit configuration, provisioning, and controlled-data read only after mutual authentication of the requesting party over any bearer (BLE, Wi-Fi, or LTE-M), such that only authorised users, apps, or gateways can configure the device or read controlled data.
Parent | PRD-CONN-006, PRD-SEC-005, PRD-SEC-006
Allocation Target | S3 (auth), S5 (secure transport), S9 (app), S8 (identity/keys)
Type | Allocated
Interface refs | I-09 (BLE/mobile), I-08 (WAN), I-10 (cloud API)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Provisioning and pairing are attack surfaces; unauthorised configuration could disarm monitoring or corrupt a profile (C9). Login-free edge read of the verdict is preserved (AD-5); configuration is controlled. Standard (TLS 1.2+/1.3 secure transport; Bluetooth Core pairing; revision to confirm). Detail in SYS-security (SYS-132…).
SYS-113 — Connectivity diagnostics and signal-quality logging
Requirement Statement | The system shall record connectivity diagnostics — bearer used, signal quality/RSSI(-class) per attempt, connection attempts and failures, and buffer backlog depth — and make them available to fleet tooling.
Parent | PRD-CONN-010, PRD-CLOUD-006
Allocation Target | S3 (diagnostics capture), S5 (radio metrics), S8 (fleet health)
Type | Allocated
Interface refs | I-10 (cloud API), I-08 (WAN metrics)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Baselined
Rationale | Fleet operators must be able to diagnose silent-comms devices and coverage gaps; signal-quality history also feeds RF-in-container evidence (SYS-104). Judgment. Diagnostics are energy-cheap.
SYS-114 — Resumable firmware-update transport
Requirement Statement | The system shall download firmware updates over LTE-M/Wi-Fi (OTA) and USB-C, resume an interrupted transfer without restarting from the beginning, and pass only verified images to the update process (verification and fail-safe application per SYS-032 / PRD-SEC-008).
Parent | PRD-CONN-011, PRD-FUN-022, PRD-SEC-008
Allocation Target | S5 (transport), S3 (resume/staging), S8 (distribution)
Type | Allocated
Interface refs | I-12 (firmware/OTA), I-08 (WAN)
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | The fleet must be updatable in the field and at depot over lossy links without wasting energy on failed restarts and without bricking (D2). Standard (secure-update principles; revision to confirm). Signing/verification governed by SYS-032.
5. Multi-market radio & roaming
SYS-115 — Regional radio configurability per market
Requirement Statement | The system shall support configuration of radio operation (bands, power, and regulatory profile) to the requirements of each target market's radio regime, and shall not transmit outside the bands/limits permitted for its configured region.
Parent | PRD-CONN-012, PRD-REG-004, PRD-ARCH-010
Allocation Target | S5 (radio config), S3 (region enforcement), S8 (config authority)
Type | Allocated
Interface refs | I-08 (WAN), I-10 (config)
Safety/Security classification | Both (compliance-blocking; config integrity)
Criticality | Critical
Verification Method | Inspection, Test
Status | Baselined
Rationale | Multi-market operation (D4) requires per-region radio conformance; transmitting outside the permitted band plan is a compliance and legal-operation failure. Standard (FCC Part 15 / CE RED; 3GPP band plans; revision to confirm). Some regions may need per-market hardware variants (PRD-REG-004); flag Regulatory (EM-ENG-006).
SYS-116 — Roaming across carriers and regions
Requirement Statement | The system shall maintain wide-area connectivity across carriers and regions encountered on a trip (via roaming-capable SIM/eSIM provisioning), automatically selecting an available permitted network without manual intervention, and shall log the network/carrier used with its connectivity diagnostics.
Parent | PRD-CONN-001, PRD-CONN-012, PRD-CONN-010
Allocation Target | S5 (modem/SIM), S3 (network selection), S8 (connectivity management)
Type | Derived
Interface refs | I-08 (WAN), I-10 (cloud API)
Safety/Security classification | Safety (cross-border alert continuity)
Criticality | High
Verification Method | Test
Status | Evidence-Pending (per-market carrier coverage and roaming/SIM strategy unresolved)
Rationale | Global vaccine lanes cross carrier and national boundaries mid-trip (D4); without automatic roaming the alert/upload path breaks at each border. Judgment on architecture; carrier coverage and the SIM/eSIM connectivity strategy are Evidence-Pending (see PRD-CONN-001 notes). Diagnostics per SYS-113.
SYS-118 — Network time synchronisation of the device clock
Requirement Statement | The system shall synchronise its internal real-time clock to a trusted network time reference whenever a link is available, bound and record the residual clock error/drift, and continue timestamping from its free-running clock (with drift disclosed) when no link is available.
Parent | PRD-FUN-019, PRD-ARCH-007, PRD-SEC-001
Allocation Target | S3 (RTC/discipline), S5 (transport), S8 (time authority)
Type | Derived
Interface refs | I-10 (cloud/time authority), I-08 (WAN)
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | A defensible record and correct excursion timing require a trusted, UTC-referenced clock (SYS-006); offline-first (D2) means the clock must free-run with bounded, disclosed drift when disconnected and re-discipline on reconnection. Standard (21 CFR Part 11 attributable/accurate; revision to confirm). Consistent with SYS-006.
System Requirements — Data, Records, Time & Storage (SYS-119 … SYS-131)
SCM · Stage 06 · Owner: EM-ENG-001, with Cloud/Platform, Cybersecurity (EM-ENG-008), and Regulatory (EM-ENG-006). System requirements governing how the SCM's evidentiary data is structured, serialised, sized, retained, protected, time-referenced, and made provable across the device → transport → cloud tiers. These decompose PRD-FUN (400 · records/export/time), PRD-CLOUD (900 · ingest/retention/ownership/ export/DR), and PRD-SEC (1100 · integrity/attribution/at-rest), and are based on D2 (store-and-forward) and the ALCOA+ / 21 CFR Part 11 / EU GDP record posture. Field set per [index.md](index.md).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Evidence posture: record-integrity items are Both (Safety/Security) and Critical (severity-first, CLAUDE.md rule 3); standards are cited by designation with revision to confirm; the SYS-127 storage sizing figure is Evidence-Pending and is not invented here — it is closed by analysis/test.
---
1. Schemas & structure
SYS-119 — Versioned telemetry sample schema
Requirement Statement | The system shall represent every telemetry sample using a defined, versioned schema in which each measurement is bound to an explicit channel identifier (internal temperature, external-probe temperature, relative humidity, 3-axis acceleration), a per-value data-quality flag, an engineering unit and resolution, and the sample's trusted timestamp, and shall record the schema version with the data.
Parent | PRD-FUN-001, PRD-FUN-017, PRD-CLOUD-002
Allocation Target | S3 (compose), S7 (persist), S8 (ingest/store)
Type | Derived
Interface refs | I-06 (MCU→storage), I-10 (device→cloud)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Inspection
Status | Baselined
Rationale | A record is only defensible if each value is unambiguously typed, attributed to a channel, and self-describing; an explicit schema version prevents silent mis-parsing as fields evolve. Carries the per-value quality flag from SYS-007 into the record. Standard (ALCOA+ legible/accurate/consistent, EU GDP; revision to confirm).
SYS-120 — Versioned configuration / trip-profile schema
Requirement Statement | The system shall represent the trip profile and device configuration (thresholds, excursion logic, logging interval, upload cadence, shock threshold, alert routing) using a defined, versioned schema, and shall bind the exact profile version in force to every trip record it governs.
Parent | PRD-FUN-014, PRD-CLOUD-004, PRD-CLOUD-002
Allocation Target | S3, S8, S9 (authoring)
Type | Derived
Interface refs | I-09 (BLE/mobile), I-10
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Inspection
Status | Baselined
Rationale | A verdict is meaningful only against the profile that produced it; the active configuration is part of the record (PRD-FUN-014), so it must be schematised and version-bound so a reviewer can reconstruct exactly what limits applied. Judgment; consistency per ALCOA+ (revision to confirm).
SYS-121 — Record metadata & provenance
Requirement Statement | The system shall stamp every trip record with provenance metadata — device serial / durable identity, firmware version, calibration identity and validity date, active profile version, and the authenticated actor and time for each lifecycle action (configure, arm, close, export) — such that the record is attributable to its originating device, software, calibration state, and responsible party.
Parent | PRD-FUN-017, PRD-SEC-002, PRD-CLOUD-001, PRD-CLOUD-007
Allocation Target | S3, S7, S8
Type | Allocated
Interface refs | I-06, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Inspection
Status | Baselined
Rationale | ALCOA+ attributable/original: a disposition must be traceable to a specific device, firmware, calibration, and actor; missing provenance renders the record inadmissible. Binds device identity (PRD-CLOUD-001) and durable ID (PRD-MECH-010). Standard (21 CFR Part 11 attributable, ALCOA+; revision to confirm).
2. Serialisation, format & compression
SYS-122 — Defined serialisation for on-the-wire and at-rest data
Requirement Statement | The system shall serialise telemetry, events, configuration, and metadata in a defined, documented format for both on-the-wire transfer and at-rest storage on device and in the cloud, such that a record written by one tier is deterministically and losslessly reconstructable by another, and the format shall carry the schema version needed to decode it.
Parent | PRD-CLOUD-002, PRD-FUN-012, PRD-SEC-001
Allocation Target | S3, S5, S7, S8
Type | Derived
Interface refs | I-06, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Status | Baselined
Rationale | Store-and-forward (D2) crosses device NVM, LTE-M/Wi-Fi/BLE bearers, and cloud storage; without a single defined lossless format, integrity and tamper-evidence cannot survive the boundary crossings. Underpins idempotent ordered ingest (SYS-020). Standard (ALCOA+ enduring/consistent; revision to confirm).
SYS-123 — Data compression for constrained uplink
Requirement Statement | The system shall support lossless compression of buffered data for uplink such that no logged value, event, gap marker, quality flag, timestamp, or integrity artefact is altered or lost by the compress/decompress round-trip, and shall record sufficient information to decompress each payload independently.
Parent | PRD-FUN-013, PRD-CONN-004, PRD-PERF-006
Allocation Target | S3, S5, S8
Type | Derived
Interface refs | I-08 (RF/bearer), I-10
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | The energy and LTE-M-in-container link budgets (D2/D3) reward smaller payloads, but compression must be lossless — a lossy transform of an evidentiary record would breach ALCOA+ accuracy. Independent decompressibility protects resumable/idempotent upload (SYS-020). Judgment.
3. Retention, ownership & isolation
SYS-124 — Configurable data retention
Requirement Statement | The system shall retain trip records, events, and audit data for a configurable retention period, shall not delete or expire retained records except through an audited, authorised action, and shall record the applicable retention policy with the data.
Parent | PRD-CLOUD-011, PRD-CLOUD-007, PRD-SEC-004
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11 (cloud→enterprise)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Inspection
Status | Evidence-Pending (retention periods per market/customer)
Rationale | EU GDP and Part 11 impose record-retention obligations that vary by market and customer; retention must be configurable and deletion must never be silent or unauthorised. Exact minimum periods are Evidence-Pending, set in the Regulatory stage. Standard (EU GDP, 21 CFR Part 11 record retention; revision to confirm).
SYS-125 — Data ownership & tenant isolation
Requirement Statement | The system shall treat each customer's records and configuration as owned by that customer, shall isolate each tenant's data so that no tenant can access, infer, or affect another tenant's records, and shall apply retention, export, and deletion only within the owning tenant's authorised scope.
Parent | PRD-CLOUD-011, PRD-CLOUD-009, PRD-SEC-005, PRD-SEC-007
Allocation Target | S8
Type | Derived
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Regulated customers require ownership and provable isolation of their cold-chain evidence; cross-tenant leakage would be both a security breach and a data-integrity failure. Depends on RBAC (PRD-SEC-005) and at-rest protection (PRD-SEC-007). Standard (EU GDP data governance, ISO/IEC 27001 concepts; revision to confirm).
SYS-126 — On-demand export & documented import/integration format
Requirement Statement | The system shall export a complete trip record on demand, and shall provide a documented import/integration format for enterprise/OEM systems, in both cases preserving samples, events, gap markers, quality flags, provenance metadata, time base, and tamper-evidence so that the record's ALCOA+ properties survive transfer into and out of the platform.
Parent | PRD-FUN-018, PRD-CLOUD-005, PRD-CLOUD-010, PRD-REG-003
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Evidence-Pending (concrete export formats, e.g. PDF + CSV/JSON)
Rationale | Evidence must move to where release/acceptance decisions and disputes live (ERP/QMS/partner systems) without losing integrity; a documented, integrity-preserving interchange format is the mechanism. Complements the record-side SYS-027. Concrete formats are Evidence-Pending. Standard (21 CFR Part 11 records, ALCOA+; revision to confirm).
4. Storage sizing, durability & recovery
SYS-127 — Non-volatile storage sizing for a full 90-day trip
Requirement Statement | The system shall provide non-volatile storage capacity sufficient to retain the complete sample and event log for a full 90-day armed trip at the minimum (worst-case densest) logging interval, plus all events, gaps, provenance metadata, and integrity artefacts, with a defined design margin, without overwriting any unsent or unacknowledged record.
Parent | PRD-FUN-012, PRD-PERF-006, PRD-REL-005
Allocation Target | S7, S3
Type | Derived
Interface refs | I-06
Safety/Security classification | Both
Criticality | Critical
Verification Method | Analysis, Test
Status | Evidence-Pending (storage sizing to confirm)
Rationale | Store-and-forward (D2) can require an entirely offline 90-day trip to be held locally; the buffer must not run out or overwrite unsent data (SYS-018/SYS-021 depend on this). The required byte budget is a function of sample size (SYS-119), compression (SYS-123), and interval, and is not invented here — it is computed by analysis and confirmed by test. This is the sizing figure cross-referenced by SYS-018. Evidence-Pending.
SYS-128 — Cloud backup, restore & disaster recovery of records
Requirement Statement | The cloud shall maintain durable, recoverable backups of all ingested records, audit trails, and configuration, and shall be able to restore them to a consistent, integrity-verified, attribution-preserving state within defined recovery-point and recovery-time objectives following a component or site failure, without silent loss or corruption.
Parent | PRD-CLOUD-012, PRD-CLOUD-002, PRD-REL-006
Allocation Target | S8
Type | Derived
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Status | Evidence-Pending (RPO/RTO targets to confirm)
Rationale | The cloud is where the record becomes the authoritative evidence; a platform failure must not lose or corrupt it. Restore must re-establish tamper-evidence and attribution, not merely availability. RPO/RTO targets are Evidence-Pending (SLA). Judgment; durability per ALCOA+ enduring (revision to confirm).
5. Validation, longevity & time base
SYS-129 — Ingest-side data validation, quality & consistency checks
Requirement Statement | On ingest, the cloud shall validate each incoming record against its declared schema, verify integrity and attribution, check chronological consistency and de-duplication, and flag (never silently drop or repair) any record that is malformed, out-of-schema, unattributable, integrity- failing, or inconsistent, recording the anomaly in the trip record and audit trail.
Parent | PRD-CLOUD-002, PRD-SEC-001, PRD-FUN-017
Allocation Target | S8
Type | Derived
Interface refs | I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Integrity preservation across an intermittent link (SYS-020) requires the ingest boundary to detect corruption and gaps rather than paper over them; silent repair would fabricate evidence, violating ALCOA+ original/accurate. Flags, not fixes. Standard (ALCOA+, EU GDP; revision to confirm).
SYS-130 — Historical / long-term storage & queryability
Requirement Statement | The system shall retain historical trip records for the full configured retention period in a form that remains readable, integrity-verifiable, and queryable (by device, trip, time range, and excursion/event) for release, audit, trend, and dispute purposes, notwithstanding schema version evolution.
Parent | PRD-CLOUD-005, PRD-CLOUD-011, PRD-CLOUD-006
Allocation Target | S8
Type | Derived
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | Retained records (SYS-124) are only useful if they remain decodable and findable years later; schema versioning (SYS-119/120) must not strand old data. ALCOA+ enduring/available. Standard (ALCOA+, EU GDP; revision to confirm).
SYS-131 — Trusted time base recorded with data
Requirement Statement | The system shall reference all timestamps to UTC via a trusted, monotonic time base, and shall record, with each record set, the time source, the synchronisation status at the time of logging, and the disclosed bound on clock drift, so that the temporal accuracy of every sample and event is knowable to a reviewer.
Parent | PRD-FUN-019, PRD-ARCH-007, PRD-SEC-002
Allocation Target | S3 (RTC), S8 (time authority)
Type | Derived
Interface refs | I-05 (sensors→MCU), I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Correct excursion timing and defensible attribution require a trusted clock whose drift and sync state travel with the data (extends SYS-006 into the data record); an undisclosed clock offset would misplace an excursion in time. Standard (21 CFR Part 11 accurate, ISO 8601 UTC time representation, ALCOA+ contemporaneous; revision to confirm).
System Requirements — Environmental & Mechanical-Environment (SYS-069 … SYS-085)
SCM · Stage 06 · Owner: EM-ENG-001 (with Reliability/Materials and Mechanical contributors). Environmental and mechanical-environment system requirements — the survival, accuracy, and usability envelope the SCM system as a whole shall hold across the whole cold-chain journey — decomposed/allocated from PRD-ENV (200) and PRD-MECH (600), based on D1–D4. Field set per [index.md](index.md).
Allocation legend: S1 Enclosure/Mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Evidence posture (honesty): Environmental test standards are cited by designation only; exact clause, revision, and test severity are to confirm and are set in the Verification Plan (PVL). Two program-level tensions surface here and are carried as explicit, testable requirements: the D3 −30 °C / 90-day thermal envelope (headline SYS-069, coupled to power SYS-086/087/090-line and reliability SYS-165-line) and the RF-in-container link envelope (SYS-080, the principal open architectural risk, Evidence-Pending, closed by the RF-in-container qualification test PRD-VER-005).
---
1. Thermal environment
SYS-069 — Operating temperature range −30 °C to +60 °C
Requirement Statement | The system shall remain fully functional — sensing, timestamping, gap-free logging, on-device verdict display, and alert emission when a link is available — throughout continuous operation across an ambient temperature range of −30 °C to +60 °C.
Parent | PRD-ENV-001, PRD-REL-001, PRD-PERF-005; D3
Allocation Target | SYS (S1 enclosure/thermal, S2 power, S3 compute, S4 sensing, S6 display)
Type | Allocated
Interface refs | I-01 (thermal), I-05 (sensors→MCU), I-07 (display)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (−30 °C / 90-day endurance unresolved)
Rationale | The cold chain spans deep-frozen product to tropical heat; the evidentiary record must not degrade or stop at either extreme. This is the top program risk (D3): standard Li-ion loses usable capacity and cannot be charged below 0 °C, so full-envelope operation at −30 °C for 90 days is in genuine tension and must be proven, not assumed. Standard (IEC 60068-2-1 cold / -2-2 dry heat, revision and severity to confirm); endurance at the cold extreme is Evidence-Pending, closed jointly with the power budget (SYS-086/087/090-line) and 90-day cold-endurance test.
SYS-070 — Storage and transport temperature range (non-operating)
Requirement Statement | The system shall survive, without permanent degradation of function, measurement accuracy, seal, or stored data, a non-operating storage/transport temperature range of at least −40 °C to +70 °C.
Parent | PRD-ENV-002, PRD-REL-001; MR-1300
Allocation Target | SYS (S1, S2, S7)
Type | Allocated
Interface refs | I-01, I-06 (MCU→storage)
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (range to confirm against battery datasheet)
Rationale | Depot, warehouse, tarmac, and unpowered-transit extremes exceed the operating envelope and act on a device that may be discharged and asleep; survival must be wider than operation. Standard (IEC 60068-2-1/-2-2, revision to confirm). Upper/lower bounds are Evidence-Pending until confirmed against the selected low-temperature cell's storage rating and enclosure-material limits — flag Power/Materials engineers.
SYS-071 — Measurement accuracy retained across the operating envelope
Requirement Statement | The system shall retain its specified temperature and humidity measurement accuracy across the full −30 °C to +60 °C operating temperature range and the SYS-072 humidity range, such that sensor self-heating, thermal gradients, and temperature-dependent drift do not push a reading outside the allocated accuracy budget.
Parent | PRD-ENV-001, PRD-ENV-004, PRD-PERF-001, PRD-MECH-011
Allocation Target | S4 (sensing), S1 (thermal coupling), S3 (compensation)
Type | Derived
Interface refs | I-01, I-05
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Analysis
Status | Preliminary (reconcile with performance accuracy budget)
Rationale | A record is only as good as its accuracy at the point of measurement; a sensor accurate at +25 °C but drifting at −30 °C would silently corrupt freeze-excursion verdicts — the highest-consequence case. Accuracy is a per-condition property, not a single number. Judgment; the numeric accuracy budget and its thermal-coupling allocation (PRD-MECH-011) are set with the performance library — flag the Sensing/Thermal engineer.
2. Humidity & condensation
SYS-072 — Humidity operating range (10–95 %RH non-condensing)
Requirement Statement | The system shall remain functional and within measurement accuracy across a relative-humidity range of at least 10 % to 95 %RH (non-condensing) throughout operation.
Parent | PRD-ENV-004, PRD-FUN-003
Allocation Target | S4, S1, S7
Type | Allocated
Interface refs | I-05
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Baselined
Rationale | Tropical lanes and cold-store staging span nearly the full humidity band; humidity is both a logged product-condition indicator and a stressor on sensing. Standard (IEC 60068-2-78 damp heat steady state, revision to confirm). "Non-condensing" bounds the specified-accuracy claim; condensation survival is governed separately by SYS-073.
SYS-073 — Condensation and thermal-cycling tolerance over the reuse life
Requirement Statement | The system shall tolerate, without permanent damage, accuracy loss, corrosion, or seal degradation, the repeated condensation and cold-to-warm thermal cycling of multi-trip reuse, including condensation forming on the enclosure and external probe during cold-store-to-ambient transitions.
Parent | PRD-ENV-004, PRD-MECH-005, PRD-MECH-009; D1
Allocation Target | S1 (enclosure/seal/probe), S4
Type | Derived
Interface refs | I-01
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (cycle count/profile to confirm)
Rationale | A reusable device (D1) crosses the dew point on every trip; repeated condensation is a cumulative corrosion and drift mechanism, and condensation on the probe is a named concern that can masquerade as a thermal reading. Single-exposure humidity survival (SYS-072) does not cover cyclic fatigue. Standard (IEC 60068-2-30 damp-heat cyclic / -2-38 temperature-humidity cycling, revision to confirm); number of cycles is Evidence-Pending, set from the reuse-life target — flag Materials engineer.
3. Legibility in the physical environment
SYS-074 — Cold display legibility at the −30 °C operating extreme
Requirement Statement | The system shall present a legible on-device verdict and status such that a gloved user can read the accept/reject/quarantine verdict at temperatures down to the −30 °C operating extreme, within a display-update latency that does not impede a receipt decision.
Parent | PRD-ENV-003, PRD-UX-005, PRD-MECH-008; D2
Allocation Target | S6 (display/HMI), S1
Type | Allocated
Interface refs | I-07
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Demonstration
Status | Evidence-Pending (cold-legibility threshold unresolved)
Rationale | The offline edge verdict (SYS-015) is the last-mile decision point and is worthless if it cannot be read where it is used — cold-store staging and receipt at −30 °C. OLED and other display technologies exhibit contrast loss and severe response-time slowdown at low temperature. Evidence-Pending: the low-temperature contrast and update-latency thresholds are unresolved and are a named open concern — flag the Display/HMI engineer; closed by a cold-chamber legibility test.
SYS-075 — Ambient-lighting legibility (glare and low light)
Requirement Statement | The system shall keep its on-device verdict and status readable across ambient illumination from dim (≈10 lux) to bright direct sun (≥50,000 lux), including glare at loading docks and gloom inside shippers.
Parent | PRD-ENV-011, PRD-UX-005, PRD-MECH-008
Allocation Target | S6, S1
Type | Allocated
Interface refs | I-07
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Evidence-Pending (lux thresholds to confirm in usability study)
Rationale | A verdict lost to washout in sunlight or invisibility in gloom is a missed decision at real posts and docks. Judgment on the lux band (bracketing plausible field conditions); exact thresholds are Evidence-Pending, confirmed in the gloved usability study with UX.
4. Mechanical environment
SYS-076 — Transport vibration survival
Requirement Statement | The system shall survive the random and sinusoidal vibration of combined air, road, and rail transport — per a profile to be set in the Verification Plan — without loss of function, loss of logged or buffered data, connector/solder-joint fatigue failure, or a shift of the device or probe out of its representative position.
Parent | PRD-ENV-005, PRD-MECH-003, PRD-MECH-007
Allocation Target | S1 (housing/retention/probe), SYS (all boards/connectors)
Type | Allocated
Interface refs | I-01, I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (profile/severity to confirm)
Rationale | Sustained transport vibration is a fatigue driver that can fail joints and displace the probe, and a shifted probe reports an unrepresentative zone — corrupting disposition. Standard (IEC 60068-2-64 broadband random, and IEC 60068-2-6 sinusoidal, revision to confirm). Profile severity is Evidence-Pending, set in the PVL against representative lane spectra — probe/device retention is coupled to PRD-MECH-003.
SYS-077 — Mechanical shock and free-fall drop survival
Requirement Statement | The system shall survive handling shock and free-fall drops — per a shock pulse and a drop height to be set in the Verification Plan — without loss of function, loss of logged or buffered data, breach of the ingress seal, or an undetected change in measurement accuracy.
Parent | PRD-ENV-006, PRD-MECH-007
Allocation Target | S1, SYS
Type | Allocated
Interface refs | I-01, I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (shock pulse / drop height to confirm)
Rationale | Drops and rough handling are routine in logistics and last-mile field use; a drop that cracks the seal or shifts a sensor without stopping the device is the dangerous case, because it produces a plausible-looking but wrong record. Standard (IEC 60068-2-27 shock and IEC 60068-2-31 free-fall/ toppling, revision to confirm). Shock pulse and drop height are Evidence-Pending, set in the PVL.
5. Ingress, cleaning & materials environment
SYS-078 — Ingress protection rating IP67
Requirement Statement | The system enclosure, including the closed charge/service port, shall meet an ingress protection rating of at least IP67 (dust-tight; withstands temporary immersion) per IEC 60529, maintained across the reuse life.
Parent | PRD-ENV-007, PRD-MECH-005, PRD-MECH-006
Allocation Target | S1 (enclosure/seal/port)
Type | Allocated
Interface refs | I-01
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Inspection
Rationale | Dust at last-mile posts and moisture (including cleaning immersion) must not reach the electronics or sensors over a multi-trip life; the USB-C port is the most likely ingress path and must hold the rating when closed. Standard (IEC 60529, revision to confirm). IP67 is a Judgment target — confirm it is sufficient against the actual cleaning method (immersion vs. wipe); if routine immersion is required, IP68 may be needed — flag Materials/Mechanical engineer.
Status | Judgment (IP67 target — confirm against cleaning method)
SYS-079 — Cleaning and decontamination chemical compatibility
Requirement Statement | The system's external surfaces, seals, display window, controls, and labels shall withstand repeated cleaning and decontamination with the agents specified in the service procedure — including isopropyl-alcohol wipes and quaternary-ammonium disinfectants — over the full reuse life without material crazing, seal degradation, loss of ingress rating, or loss of label/display legibility.
Parent | PRD-ENV-008, PRD-MECH-009, PRD-MECH-010; D1
Allocation Target | S1 (materials/seals/label/window)
Type | Allocated
Interface refs | I-01
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (agent list and cycle count to confirm)
Rationale | A reusable device (D1) is cleaned between every trip; chemical attack on a seal or window creates an ingress path or an illegible verdict, so cleaning durability is a safety-relevant durability requirement, not cosmetic. Judgment; the definitive agent list, concentration, and cleaning-cycle count are Evidence-Pending, set with the service procedure (PRD-SVC-002) — flag Materials engineer. Materials also carry RoHS/REACH constraints (PRD-REG-line).
6. RF and power-availability environment
SYS-080 — RF operation inside a representative insulated / metallized shipper
Requirement Statement | The system shall capture and store all data locally without loss, and shall deliver buffered data and queued alerts whenever a usable link becomes available, while operating inside a representative metallized, thermally-insulated shipping container that attenuates RF; no data or excursion shall depend on maintaining a live link from inside the shipper.
Parent | PRD-ENV-009, PRD-CONN-001, PRD-CONN-007, PRD-FUN-012; D2
Allocation Target | S5 (RF/antenna), S1 (antenna placement), S3, S7, S8
Type | Allocated
Interface refs | I-08 (RF cellular/GNSS), I-09 (BLE/Wi-Fi), I-10 (cloud)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (in-container link margin unresolved — principal open risk)
Rationale | The device lives inside an RF-hostile enclosure that can attenuate LTE-M/Wi-Fi/BLE/GNSS by tens of dB; the design must degrade to store-and-forward (D2) so the last mile is never a black box. This is the principal open architectural risk: the achievable in-container link margin, and whether an external-antenna option is required, are unresolved. Evidence-Pending — closed by the RF-in-container qualification test (PRD-VER-005); coupled to connectivity SYS-104/109/117-line and GNSS-in-container SYS-029.
SYS-081 — Intermittent connectivity and power tolerance at the last mile
Requirement Statement | The system shall continue gap-free logging and shall present a valid on-device verdict for the full trip duration with no live network connection and no external power available, such that intermittent or absent connectivity and mains power are tolerated conditions, not failures.
Parent | PRD-ENV-010, PRD-FUN-011, PRD-FUN-012, PRD-PERF-006; D2
Allocation Target | SYS (S2 power, S3, S5, S7)
Type | Allocated
Interface refs | I-08, I-09, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | The last mile has unreliable or absent connectivity and no mains power (D2); offline-first operation makes live data an upside, never a dependency. Judgment; capacity to survive the worst-case offline duration is coupled to the energy budget (SYS-086-line) and buffer sizing (SYS-018/SYS-127).
7. Altitude / pressure environment
SYS-082 — Altitude and cargo-hold pressure operation
Requirement Statement | The system shall operate and remain safe through the reduced-pressure and pressure-cycling conditions of an aircraft cargo hold, equivalent to a cabin-altitude of at least 4,570 m (15,000 ft), including the pressure change rate of ascent and descent, without loss of function, seal failure, or battery hazard.
Parent | PRD-ENV-012, PRD-ELEC-line, PRD-SAFE-line, PRD-REG-006; D4
Allocation Target | SYS (S1 seal, S2 battery)
Type | Allocated
Interface refs | I-01
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Air transport (D4) exposes the device and its lithium cell to reduced/cycling pressure; a sealed IP67 enclosure develops a pressure differential that can stress seals, and low pressure is a lithium- battery transport hazard. Standard (UN 38.3 T.1 altitude simulation, revision to confirm; concept per IATA DGR / DO-160 low-pressure environment). Altitude value follows the dangerous-goods regime — coupled to transport-safety and regulatory (PRD-REG-006).
8. Physical fit and mass in the environment
SYS-083 — Enclosure fit inside standard insulated shippers and cold boxes
Requirement Statement | The system's device and probe shall fit within the internal usable volume of the target range of insulated shippers and cold boxes and be retainable in a representative measurement position without displacing payload beyond the allowance set in the product spec.
Parent | PRD-MECH-001, PRD-MECH-003, PRD-MECH-012
Allocation Target | S1
Type | Allocated
Interface refs | I-01
Safety/Security classification | None
Criticality | Medium
Verification Method | Inspection, Demonstration
Status | Evidence-Pending (shipper envelope set to confirm)
Rationale | The device must be placeable with the product across many shipper sizes, and its position determines whether the reading represents the product zone; a device that does not fit, or displaces payload, will not be used or will be mis-placed. Judgment; the definitive envelope across shipper sizes is Evidence-Pending — coupled to embeddable-module interface (PRD-MECH-012) and retention (SYS-076).
SYS-084 — Mass constraint in the use environment
Requirement Statement | The system's device (with battery, excluding external packaging) shall not exceed the mass limit set in the product spec, bounded so it does not meaningfully displace payload or burden gloved handlers placing and retrieving it.
Parent | PRD-MECH-002; D3
Allocation Target | S1, S2 (battery mass)
Type | Allocated
Interface refs | I-01
Safety/Security classification | None
Criticality | Low
Verification Method | Inspection
Status | Evidence-Pending (target mass to confirm)
Rationale | Mass burdens handling and payload economics and is in tension with the D3 battery: the low-temperature 90-day energy budget pushes toward a larger, heavier cell. Judgment; the target mass is Evidence-Pending, set against the battery sizing (SYS-086-line) — flag Power engineer.
9. Radiation and labeling durability environment
SYS-085 — UV exposure and labeling / marking durability
Requirement Statement | The system's enclosure materials and its device-identification and required safety/transport markings shall withstand the ultraviolet, sunlight, and outdoor-staging exposure of the distribution environment over the full reuse life without material embrittlement, discolouration that impairs the display, or loss of legibility of identification and mandatory markings.
Parent | PRD-MECH-010, PRD-ENV-008, PRD-REG-line; D1
Allocation Target | S1 (materials/label/marking)
Type | Derived
Interface refs | I-01
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Evidence-Pending (UV exposure profile to confirm)
Rationale | Attribution, fleet management (D1), and regulatory acceptance depend on markings that remain legible; UV and outdoor staging embrittle polymers and fade print, and combine with the cleaning attack of SYS-079 as a cumulative degradation path. Standard (IEC 60068-2-5 / ISO 4892 solar-radiation and weathering concepts, revision to confirm); the exposure profile is Evidence-Pending, set in the PVL — flag Materials engineer. Marking set follows the regulatory stage (PRD-MECH-010).
System Requirements — Functional (SYS-001 … SYS-032)
SCM · Stage 06 · Owner: EM-ENG-001. Functional system requirements realising the value chain capture → detect → warn → decide → prove, decomposed/allocated from PRD-FUN (400), PRD-ARCH (300), and PRD-UX where the function is system-behavioural. Field set per [index.md](index.md). Based on D1–D4.
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
---
1. Sense & log
SYS-001 — Interval environmental sampling
Requirement Statement | The system shall sample temperature (internal and external-probe channels), relative humidity, and 3-axis acceleration at a configurable base logging interval selectable in the range 1–60 min (default 5 min) for the entire armed trip.
Parent | PRD-FUN-001, PRD-FUN-002, PRD-FUN-003, PRD-FUN-004
Allocation Target | S4 (acquire), S3 (schedule/timestamp)
Type | Allocated
Interface refs | I-05 (sensors→MCU)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | A gap-free interval history is the primary evidentiary product (D2). Judgment on default 5 min per WHO PQS E006 logger practice (revision to confirm).
SYS-002 — Gap-free logging with no unrecorded interval
Requirement Statement | The system shall record every scheduled sample such that no armed-trip interval is silently absent; any interval not recorded shall be represented by an explicit gap marker in the log.
Parent | PRD-FUN-001, PRD-FUN-017, PRD-SAFE-003
Allocation Target | S3, S7
Type | Derived
Interface refs | I-06 (MCU→storage)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | ALCOA+ completeness; a missing sample must never masquerade as in-range data. Standard (ALCOA+, EU GDP; revision to confirm).
SYS-003 — Dual-channel temperature acquisition
Requirement Statement | The system shall acquire and store internal-sensor and external-probe temperature as two distinct, independently identifiable channels in every sample record.
Parent | PRD-FUN-002, PRD-MECH-004, PRD-MECH-011
Allocation Target | S4, S1 (probe), S7
Type | Allocated
Interface refs | I-01 (thermal), I-05
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | The probe represents the product zone and underpins freeze detection; conflating channels would corrupt disposition. Judgment.
SYS-004 — Humidity acquisition across range
Requirement Statement | The system shall measure and record relative humidity across at least 10–95 %RH (non-condensing) at the configured logging interval.
Parent | PRD-FUN-003, PRD-ENV-004
Allocation Target | S4, S7
Type | Allocated
Interface refs | I-05
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Baselined
Rationale | Humidity affects some products and is a condition indicator. Judgment.
SYS-005 — Shock/impact event capture without routine-handling noise
Requirement Statement | The system shall detect and record shock/impact events exceeding a per-profile configurable acceleration threshold with timestamp and peak magnitude, and shall not record routine-handling accelerations below that threshold as events.
Parent | PRD-FUN-004, PRD-ENV-006
Allocation Target | S4, S3, S7
Type | Allocated
Interface refs | I-05, I-06
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Baselined
Rationale | Product-relevant mishandling must be captured without alarm-fatigue noise (concept C8). Judgment.
SYS-006 — Trusted timestamp on every sample and event
Requirement Statement | The system shall timestamp every sample and event against a trusted, monotonic time base referenced to UTC, and shall record the time source and any synchronisation status with each record set.
Parent | PRD-FUN-019, PRD-ARCH-007, PRD-SEC-001
Allocation Target | S3 (RTC), S8 (time authority)
Type | Derived
Interface refs | I-05, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | A defensible record and correct excursion timing require a trusted clock; drift must be bounded and disclosed. Standard (21 CFR Part 11 attributable/accurate; revision to confirm).
SYS-007 — Sensor plausibility and fault flagging
Requirement Statement | The system shall evaluate each sensor reading against physically plausible bounds and record a data-quality flag when a reading is implausible, out-of-range, or from a faulted channel, without discarding the raw value.
Parent | PRD-FUN-017, PRD-REL-004, PRD-SAFE-003
Allocation Target | S3, S4, S7
Type | Derived
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | A silently-wrong reading is more dangerous than a flagged one (no false reassurance). Judgment.
2. Detect & alert
SYS-008 — Independent two-sided threshold configuration
Requirement Statement | The system shall support independent low (freeze) and high (heat) temperature thresholds, and humidity/shock thresholds, configurable per trip profile.
Parent | PRD-FUN-005, PRD-VIS-003
Allocation Target | S3, S8/S9 (config authoring)
Type | Allocated
Interface refs | I-09, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Freeze- and heat-sensitive products need distinct limits. Judgment.
SYS-009 — Two-sided excursion detection with equal reliability
Requirement Statement | The system shall detect freezing excursions with detection reliability equal to overheating excursions and shall flag each excursion, its direction, onset time and duration in the trip record.
Parent | PRD-FUN-006, PRD-SAFE-002, PRD-VIS-003
Allocation Target | S3
Type | Allocated
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Freeze damage is silent and historically under-detected. Cited (WHO/UNICEF freezing reporting).
SYS-010 — Product-relevant excursion logic
Requirement Statement | The system shall determine an excursion using configurable product-stability logic (threshold combined with time-over/under-threshold and/or cumulative exposure) rather than instantaneous transients alone.
Parent | PRD-FUN-007, PRD-SAFE-004
Allocation Target | S3
Type | Allocated
Interface refs | I-05
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Alerts/verdicts must reflect actual product harm, not momentary noise (C8). Judgment.
SYS-011 — Immediate excursion alert emission when connected
Requirement Statement | Upon detecting an excursion, the system shall emit an alert to the cloud within the SYS-039 latency target when connectivity is available, and shall otherwise queue the alert for delivery on next connection.
Parent | PRD-FUN-008, PRD-CONN-009
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-08, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Learning of a break in time to intervene converts monitoring into prevented loss (D2). Judgment.
SYS-012 — Alert routing and escalation
Requirement Statement | The cloud shall route each alert to the configured responder(s) for the device's trip/owner and shall escalate per configurable policy if unacknowledged within a defined time.
Parent | PRD-FUN-009, PRD-CLOUD-003
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | An alert that reaches no accountable human is not a risk control. Judgment.
SYS-013 — Local capture of every excursion regardless of connectivity
Requirement Statement | The system shall record every detected excursion event to non-volatile storage at the time of detection independent of connectivity, so that no excursion depends on a live link to exist in the record.
Parent | PRD-FUN-008, PRD-FUN-012, PRD-SAFE-003
Allocation Target | S3, S7
Type | Derived
Interface refs | I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Offline-first: the last mile is never a black box (D2). Judgment.
SYS-014 — Alarm-balance configurability (no fatigue-induced miss)
Requirement Statement | The system shall provide configurable hysteresis, debounce, and alert-consolidation parameters so that the alerting rate can be tuned to avoid alarm fatigue without suppressing genuine excursions.
Parent | PRD-SAFE-004, PRD-FUN-007
Allocation Target | S3, S8
Type | Derived
Interface refs | I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Fatigue-induced dismissal of alerts is a recognised failure mode of alarm systems. Standard (alarm-management principles; revision to confirm).
3. Decide (offline verdict)
SYS-015 — Offline on-device accept/reject/quarantine verdict
Requirement Statement | The system shall compute and display an unambiguous accept / reject / quarantine verdict on the device's own display, derived from the trip record against the active profile, with no network, phone, login, or manual required.
Parent | PRD-FUN-010, PRD-UX-001, PRD-UX-002, PRD-ARCH-003
Allocation Target | S3, S6
Type | Allocated
Interface refs | I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Demonstration
Status | Baselined
Rationale | The last-mile user has no phone, network, or training; the edge verdict is the decision point (AD-6). Judgment.
SYS-016 — Verdict consistency with the authoritative record
Requirement Statement | The on-device verdict shall be derived from the same excursion evaluation as the cloud trip record such that device and cloud verdicts for the same trip and profile are consistent, and any divergence shall be detectable and flagged.
Parent | PRD-FUN-011, PRD-SEC-001
Allocation Target | S3, S8
Type | Derived
Interface refs | I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Two sources of truth that disagree destroy trust in disposition. Judgment.
SYS-017 — Honest indication of incomplete monitoring
Requirement Statement | When the trip record contains gaps, sensor faults, or an interrupted monitoring period, the system shall not display an unqualified "accept" verdict and shall indicate that monitoring was incomplete.
Parent | PRD-SAFE-003, PRD-FUN-010, PRD-UX-008
Allocation Target | S3, S6
Type | Derived
Interface refs | I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | No false reassurance — an incomplete record must never read as a clean pass. Judgment.
4. Store & forward
SYS-018 — Offline buffering of a full trip
Requirement Statement | The system shall buffer, in non-volatile storage, the complete sample and event log for a full 90-day trip at the minimum logging interval with defined margin, without overwriting unsent data.
Parent | PRD-FUN-012, PRD-PERF-006, PRD-REL-005
Allocation Target | S7, S3
Type | Allocated
Interface refs | I-06
Safety/Security classification | Both
Criticality | Critical
Verification Method | Analysis, Test
Status | Evidence-Pending (storage sizing to confirm)
Rationale | Store-and-forward requires capacity for the worst-case offline duration (D2). Judgment on sizing; closed by SYS-127.
SYS-019 — Scheduled cadence upload with multi-bearer offload
Requirement Statement | The system shall upload buffered data on a configurable cadence (default hourly) over the best available bearer, selecting among LTE-M (primary), Wi-Fi, and BLE offload, and shall not require any single bearer to be present.
Parent | PRD-FUN-013, PRD-CONN-004, PRD-ARCH-005
Allocation Target | S5, S3, S8
Type | Allocated
Interface refs | I-08, I-09, I-10
Safety/Security classification | None
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Opportunistic multi-bearer upload maximises data delivery under intermittent coverage (D2/AD-1). Judgment.
SYS-020 — Ordered, resumable, de-duplicated upload
Requirement Statement | The system shall upload buffered records in intact time order, resume an interrupted upload without data loss, and ensure the cloud stores each record exactly once (idempotent ingest).
Parent | PRD-FUN-013, PRD-CLOUD-002, PRD-CONN-008
Allocation Target | S3, S5, S8
Type | Derived
Interface refs | I-10
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Integrity preservation across intermittent links requires ordering, resume, and de-duplication. Standard (ALCOA+ enduring/consistent; revision to confirm).
SYS-021 — Buffer-full graceful behaviour
Requirement Statement | If non-volatile buffer capacity is approached, the system shall prioritise retention of excursion events and audit records, warn of impending capacity limits, and never silently discard excursion or audit data.
Parent | PRD-FUN-012, PRD-REL-008, PRD-SAFE-003
Allocation Target | S3, S7
Type | Derived
Interface refs | I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Degradation must preserve the most safety-relevant data first. Judgment.
5. Configure, provision & arm
SYS-022 — Trip-profile configuration
Requirement Statement | The system shall allow an authorised user to select or define a trip profile (thresholds, excursion logic, logging interval, upload cadence, alert routing) over BLE/mobile or on-device and shall apply it to the device before arming.
Parent | PRD-FUN-014, PRD-MOB-002, PRD-ARCH-010
Allocation Target | S9, S3, S8
Type | Allocated
Interface refs | I-09, I-10
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | Configurable-across-products without hardware change (PRD-VIS-007). Judgment.
SYS-023 — Arm with positive placement confirmation
Requirement Statement | The system shall require an explicit arm action and shall give positive on-device confirmation that it is armed, configured, and placed, such that a trip cannot silently start unmonitored.
Parent | PRD-FUN-015, PRD-SAFE-006, PRD-UX-004
Allocation Target | S3, S6
Type | Allocated
Interface refs | I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Demonstration
Status | Preliminary (reconcile with HAL)
Rationale | Mis-arm (unmonitored trip) is a top hazard; arming must be unmissable. Judgment.
SYS-024 — Placement and readiness guidance / pre-arm checks
Requirement Statement | Before confirming armed, the system shall verify and report readiness pre-conditions (adequate charge, valid calibration, valid firmware, valid profile, probe present) and shall guide the user to correct any failed pre-condition.
Parent | PRD-FUN-016, PRD-FUN-020, PRD-SVC-004
Allocation Target | S3, S6, S9
Type | Derived
Interface refs | I-07, I-09
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | A device armed while un-ready produces a worthless or misleading record. Judgment.
SYS-025 — Configuration validation and error prevention
Requirement Statement | The system shall validate a trip profile against defined constraints (e.g. low < high threshold, interval within bounds) and shall reject or flag an invalid or internally inconsistent configuration before it can be armed.
Parent | PRD-UX-007, PRD-FUN-014, PRD-SAFE-005
Allocation Target | S3, S9, S8
Type | Derived
Interface refs | I-09, I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Use-error protection: prevent misconfiguration from becoming a missed excursion. Standard (IEC 62366-1 use-error; revision to confirm).
6. Record, prove & export
SYS-026 — Complete, tamper-evident, attributable trip record
Requirement Statement | The system shall assemble, per trip, a complete record comprising all samples, events, excursions, gaps, configuration, and lifecycle actions, each attributable to an actor/source and time, protected by tamper-evidence.
Parent | PRD-FUN-017, PRD-SEC-001, PRD-SEC-002, PRD-CLOUD-005
Allocation Target | S7, S8, S3
Type | Allocated
Interface refs | I-06, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Status | Baselined
Rationale | The auditable record is the product's compliance value (ALCOA+). Standard (EU GDP, 21 CFR Part 11; revision to confirm).
SYS-027 — On-demand record export in an admissible format
Requirement Statement | The system shall export a complete trip record on demand in a documented, human- and machine-readable format that preserves attribution, time, gaps, and tamper-evidence.
Parent | PRD-FUN-018, PRD-CLOUD-005, PRD-REG-003
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | Records must be exportable for release/acceptance decisions and audit. Standard (Part 11 records; revision to confirm).
SYS-028 — Append-only audit trail of record and lifecycle actions
Requirement Statement | The system shall maintain an append-only audit trail (device and cloud) of configuration, arming, firmware, calibration, and record actions, where each entry is attributable, time-stamped, and cannot be altered or deleted without detection.
Parent | PRD-SEC-004, PRD-CLOUD-007, PRD-FUN-017
Allocation Target | S3, S7, S8
Type | Derived
Interface refs | I-06, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Inspection
Status | Baselined
Rationale | Attributability and non-repudiation are core to a regulated record. Standard (Part 11 audit trail; revision to confirm).
7. Locate & attribute
SYS-029 — Time- and location-stamped events within the energy budget
Requirement Statement | The system shall acquire and record device location at a configurable cadence and on defined events (e.g. excursion, arrival) within the power budget, and shall attribute each significant event with time and best-available location.
Parent | PRD-FUN-019, PRD-CONN-005, PRD-PERF-010
Allocation Target | S4 (GNSS), S3, S5
Type | Allocated
Interface refs | I-08
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Evidence-Pending (GNSS-in-container cadence)
Rationale | Location supports stray recovery and last-mile attribution but is energy-costly and RF-attenuated in the shipper (D3 tension). Judgment.
8. Fleet, readiness & OTA
SYS-030 — Device readiness reporting
Requirement Statement | The system shall report each device's readiness state — state of charge, calibration validity/date, firmware version, and last-seen — to the fleet platform and shall block deployment of a device that fails readiness.
Parent | PRD-FUN-020, PRD-SVC-004, PRD-CLOUD-006
Allocation Target | S3, S8
Type | Allocated
Interface refs | I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | An out-of-calibration or under-charged device must not be deployed (D1 reuse loop). Judgment.
SYS-031 — Fleet visibility and stray recovery
Requirement Statement | The cloud shall present fleet-wide status (location, readiness, active trips, alerts, last-seen) and shall flag devices that are overdue, stray, or unreported for recovery.
Parent | PRD-FUN-021, PRD-CLOUD-006
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | None
Criticality | Medium
Verification Method | Demonstration
Status | Baselined
Rationale | A reusable fleet must be trackable to control TCO and loss (D1). Judgment.
SYS-032 — Signed, fail-safe firmware update behaviour
Requirement Statement | The system shall accept and apply only firmware whose authenticity and integrity are verified, shall apply updates without loss of trip data, and shall never place a fielded device in a bricked or unmonitored-without-notice state as a result of an update.
Parent | PRD-FUN-022, PRD-SEC-008, PRD-CLOUD-008, PRD-SAFE-010
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-12
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | OTA integrity and fail-safety protect the fleet and the record (AD-7). Standard (secure-update principles, IEC 62443 concepts; revision to confirm).
System Requirements — Interfaces, typed (SYS-051 … SYS-068)
SCM · Stage 06 · Owner: EM-ENG-001. System interface requirements decomposed/allocated from PRD-ARCH (300, interface conformance), PRD-ELEC, PRD-MECH, PRD-CONN, PRD-CLOUD, PRD-SEC and the typed interface inventory I-01…I-12 in [../05-product-requirements/system-block-diagram.md](../05-product-requirements/system-block-diagram.md) §3. Each requirement states an obligation on one typed boundary (THRM / MECH / PWR / DATA / RF) or on interface integrity/conformance across boundaries. Field set per [index.md](index.md). Based on D1–D4 and the architecture decisions AD-1…AD-8. The RF-in-container interface (I-08) is the principal open architectural risk and is carried Evidence-Pending until closed by the RF-in-container qualification test (PRD-VER-005).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Interface types (block diagram §1): PWR power/energy · DATA digital data/signal · RF radio-frequency link · THRM thermal coupling · MECH mechanical mount/fit/retention.
---
1. Thermal & mechanical boundaries (THRM, MECH)
SYS-051 — Representative thermal coupling of payload to sensing (I-01)
Requirement Statement | The system shall thermally couple the monitored payload/shipper zone to the temperature sensing channels such that the external probe tracks the product-zone temperature within a bounded, characterised offset and lag, and the internal sensor tracks the enclosure micro-environment, so that a logged temperature is a defensible proxy for the product it represents.
Parent | PRD-MECH-004, PRD-MECH-011, PRD-FUN-002, PRD-ARCH-009
Allocation Target | S1 (probe/enclosure thermal path), S4 (sensing)
Type | Allocated
Interface refs | I-01 (THRM: payload/shipper → probe/enclosure → sensors)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Analysis
Status | Evidence-Pending (probe offset/time-constant budget to confirm)
Rationale | If the probe does not represent the product zone, every excursion verdict is built on an unrepresentative measurement (AD-4). Thermal mass, probe placement, and self-heating set the offset and lag. Judgment; numeric offset/time-constant budget closed by thermal-characterisation test (forward to PVL).
SYS-052 — Fit and retention inside standard insulated shippers (I-02)
Requirement Statement | The system shall mechanically fit within, and be retained inside, the defined set of standard insulated vaccine shippers without displacing payload, without defeating the shipper's thermal barrier, and without relying on tools, such that the device remains in its intended position and orientation for the full trip.
Parent | PRD-MECH-001, PRD-MECH-006, PRD-ENV-006, PRD-ARCH-009
Allocation Target | S1
Type | Allocated
Interface refs | I-02 (MECH: shipper ↔ enclosure)
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Inspection
Status | Preliminary (reconcile with HAL)
Rationale | A device that shifts, floats, or must be wedged in changes both the thermal coupling (I-01) and the shock signature, and can compromise the payload. Retention is the mechanical precondition for a representative record. Judgment; the shipper set is defined by the target-market fleet (D1/D4).
---
2. Power boundaries (PWR)
SYS-053 — USB-C charge path with low-temperature charge inhibit (I-03)
Requirement Statement | The system shall accept charge energy over the sealed USB-C port and route it through the charge path source → PMIC → BMS → cell, and the BMS shall inhibit cell charging whenever cell temperature is below the defined low-temperature charge threshold (0 °C nominal), indicating the inhibited state rather than silently failing to charge.
Parent | PRD-ELEC-003, PRD-SAFE-007, PRD-ARCH-008
Allocation Target | S2 (PMIC/BMS/cell), S1 (port), S3 (charge state)
Type | Allocated
Interface refs | I-03 (PWR: USB-C source → PMIC → BMS → cell)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Charging a cold Li-ion cell risks lithium plating and a thermal event; the charge-inhibit gate is the safety control on this boundary (AD-3, D3). Standard (IEC 62368-1; Li-ion charge-temperature limits per cell datasheet; revision to confirm).
SYS-054 — USB-C connector electrical definition and role behaviour (I-03)
Requirement Statement | The system shall define the USB-C receptacle electrically — pin roles, CC configuration-channel behaviour as a sink, accepted input voltage/current envelope, and ESD/fault tolerance — and shall present a compliant, documented sink interface that does not source hazardous voltage on any pin and remains safe under reverse-insertion, over-voltage, and short conditions.
Parent | PRD-ELEC-002, PRD-ELEC-003, PRD-SAFE-007, PRD-ARCH-009
Allocation Target | S2, S1 (port)
Type | Derived
Interface refs | I-03 (PWR: USB-C source → port)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Inspection
Status | Baselined
Rationale | The connector is a defined electrical contract, not just a hole; an undocumented or non-compliant sink invites depot-charger incompatibility and fault-injection hazards. Standard (USB Type-C / USB-PD sink requirements; IEC 62368-1; revision to confirm).
SYS-055 — Regulated, duty-cycled rails to subsystems (I-04)
Requirement Statement | The system shall distribute regulated power rails from the BMS to the MCU, sensing, connectivity, HMI, and storage subsystems, and shall be able to independently gate/duty-cycle the high-draw rails (connectivity, sensing, HMI) under firmware control so that a subsystem draws energy only when required by the active power state.
Parent | PRD-ELEC-008, PRD-REL-005, PRD-ARCH-004
Allocation Target | S2, S3 (rail control)
Type | Allocated
Interface refs | I-04 (PWR: BMS → MCU/sensing/connectivity/HMI/storage)
Safety/Security classification | None
Criticality | High
Verification Method | Test, Analysis
Status | Baselined
Rationale | The 90-day/−30 °C budget is met by duty-cycling, not chemistry (AD-2, D3); rails that cannot be individually gated make the energy governor impossible. Rail gating must never corrupt an in-flight NVM write (see SYS-057). Judgment; rail budget closed by the endurance analysis/test.
---
3. Internal data boundaries (DATA)
SYS-056 — Sensor-to-MCU data acquisition bus (I-05)
Requirement Statement | The system shall transfer temperature (internal and external-probe), humidity, acceleration, and GNSS data from the sensing subsystem to the MCU over a defined digital interface such that each reading is associated with its correct source channel, its acquisition time, and a completed/ faulted status, with no silent substitution of one channel's value for another's.
Parent | PRD-FUN-002, PRD-FUN-017, PRD-PERF-002, PRD-ARCH-009
Allocation Target | S4, S3
Type | Allocated
Interface refs | I-05 (DATA: sensors → MCU)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Channel identity and read integrity at this boundary underpin dual-channel acquisition (SYS-003) and plausibility flagging (SYS-007); a mislabelled or corrupted read corrupts disposition. Judgment; bus type (I2C/SPI) and addressing fixed in the ICD.
SYS-057 — MCU-to-storage tamper-evident, power-fail-safe writes (I-06)
Requirement Statement | The system shall write log, configuration, and audit data from the MCU to non-volatile storage over a defined interface such that a write is either fully committed or detectably incomplete on power loss (no torn/partial silent records), and such that stored records carry tamper-evidence enabling later detection of alteration or deletion.
Parent | PRD-FUN-017, PRD-SEC-002, PRD-REL-005, PRD-ARCH-008
Allocation Target | S3, S7
Type | Allocated
Interface refs | I-06 (DATA: MCU → non-volatile storage)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test (power-fail injection), Analysis
Status | Baselined
Rationale | This boundary is where the evidentiary record is made durable; a torn write during a duty-cycle rail-down (SYS-055) or brown-out must not fabricate or lose a record (ALCOA+ enduring). Standard (ALCOA+, 21 CFR Part 11; revision to confirm).
SYS-058 — MCU-to-HMI display/indicator/button interface (I-07)
Requirement Statement | The system shall drive the OLED, status LED, and read the single button over a defined interface such that the on-device verdict and device state shown are consistent with the MCU's authoritative state, a single button actuation maps to a single unambiguous action, and a stuck or failed HMI element is detectable and does not present a false state.
Parent | PRD-UX-001, PRD-UX-004, PRD-FUN-010, PRD-ARCH-009
Allocation Target | S3, S6
Type | Allocated
Interface refs | I-07 (DATA: MCU → OLED/LED/button)
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | The HMI is the last-mile decision surface (AD-6); a display showing a stale or wrong verdict is a false-reassurance hazard. A failed pixel/segment must not turn "reject" into "accept." Judgment.
---
4. Radio-frequency boundaries (RF)
SYS-059 — Antenna-to-network RF links through the metallized shipper (I-08)
Requirement Statement | The system shall provide antenna interfaces for LTE-M, Wi-Fi, and GNSS that achieve the link performance needed to deliver buffered data, alerts, and location fixes from inside the defined insulated shippers, accounting for the shipper's RF attenuation, and shall degrade to store-and-forward buffering rather than data loss when a link cannot be established.
Parent | PRD-CONN-001, PRD-CONN-004, PRD-ENV-008, PRD-ARCH-005
Allocation Target | S5, S1 (antenna placement/aperture)
Type | Allocated
Interface refs | I-08 (RF: antennas ↔ cellular/Wi-Fi/GNSS, attenuated by shipper)
Safety/Security classification | None
Criticality | High
Verification Method | Test (RF-in-container)
Status | Evidence-Pending (in-container link budget & antenna strategy unresolved)
Rationale | RF-in-container is the principal open architectural risk; attenuation by the metallized shipper may force an external-antenna option (AD-1, I-08). Link budget, antenna gain, and viable bearers are unknown until measured. Evidence-Pending; closed by the RF-in-container qualification test (PRD-VER-005); until then upload never a dependency (D2, SYS-013/018).
SYS-060 — Device-to-BLE-app provisioning/read RF link (I-09)
Requirement Statement | The system shall provide a BLE radio interface over which an authorised mobile app can, within a defined proximity range, provision and configure the device, arm it, and read the local trip verdict/record, and shall constrain this link's range and behaviour so it functions reliably at depot/last-mile distances without becoming a standing exposure surface when idle.
Parent | PRD-CONN-006, PRD-MOB-002, PRD-ARCH-005
Allocation Target | S5, S9, S3
Type | Allocated
Interface refs | I-09 (RF: device ↔ BLE app)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | BLE is the local provisioning and offline-read path (AD-1/AD-5); its range/availability must serve the workflow while limiting the wireless attack surface (see SYS-068). Judgment (Bluetooth LE; revision/profile to confirm).
SYS-061 — BLE GATT/provisioning interface definition (I-09)
Requirement Statement | The system shall define the BLE service/characteristic (GATT) structure and the provisioning/configuration/read message set exposed over I-09 — including pairing/bonding method, which operations require authenticated/encrypted access, and versioning — such that only defined, version-negotiated operations are accepted and the interface is testable against a published specification.
Parent | PRD-CONN-006, PRD-MOB-002, PRD-SEC-006, PRD-ARCH-009
Allocation Target | S5, S3, S9
Type | Derived
Interface refs | I-09 (RF/DATA: device ↔ BLE app)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Inspection
Status | Baselined
Rationale | An undefined GATT surface is both an integration hazard and an attack surface; provisioning writes must be authenticated and versioned so a stale or rogue app cannot silently misconfigure a device (AD-5). Standard (Bluetooth GATT; IEC 62443 concepts; revision to confirm).
---
5. Cloud & service data boundaries (DATA)
SYS-062 — Device-to-cloud store-and-forward upload and alert interface (I-10)
Requirement Statement | The system shall define the device ↔ cloud interface over which the device uploads buffered records and emits excursion alerts, such that uploads are ordered, resumable, and idempotent, alerts are prioritised ahead of bulk backlog, and the device receives positive acknowledgement before considering buffered data safely delivered.
Parent | PRD-CONN-008, PRD-CLOUD-002, PRD-FUN-013, PRD-ARCH-005
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-10 (DATA: device ↔ cloud)
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | This is the store-and-forward boundary (D2); without acknowledged, idempotent transfer, intermittent links cause loss or duplication (SYS-020). Alerts must jump the queue to preserve intervention latency (SYS-011). Standard (ALCOA+ enduring/consistent; revision to confirm).
SYS-063 — Device↔cloud protocol and message-format specification (I-10)
Requirement Statement | The system shall specify the transport protocol, security binding, message schema, field encodings, units, and schema-version negotiation for every message crossing the device ↔ cloud boundary, such that both endpoints validate messages against the published schema and reject or quarantine malformed, unknown-version, or out-of-contract messages rather than misinterpreting them.
Parent | PRD-CONN-008, PRD-CLOUD-002, PRD-SEC-003, PRD-ARCH-009
Allocation Target | S3, S8
Type | Derived
Interface refs | I-10 (DATA: device ↔ cloud)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Inspection
Status | Baselined
Rationale | A safety-relevant record crossing a version boundary must never be silently mis-decoded; an explicit schema/version contract is the defence against fleet/cloud version skew during OTA rollouts (SYS-065). Standard (ISO/IEC/IEEE 29148 interface definition; revision to confirm).
SYS-064 — Cloud-to-enterprise export/API interface (I-11)
Requirement Statement | The system shall provide a documented cloud ↔ enterprise interface (API and export) that delivers trip records and fleet data to authorised enterprise/quality systems in a documented, versioned, machine-readable format preserving attribution, time, gaps, and tamper-evidence, under authenticated, access-controlled, audit-logged access.
Parent | PRD-CLOUD-005, PRD-FUN-018, PRD-REG-003, PRD-SEC-005
Allocation Target | S8
Type | Allocated
Interface refs | I-11 (DATA: cloud ↔ enterprise/quality systems)
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Demonstration
Status | Baselined
Rationale | The export/API boundary is where the compliance value is delivered to ERP/QMS for release decisions; it must preserve the record's ALCOA+ properties and be governed (RBAC/SSO, audit). Standard (EU GDP, 21 CFR Part 11; revision to confirm).
SYS-065 — OTA signed-firmware delivery interface (I-12)
Requirement Statement | The system shall define the OTA service → device interface such that only firmware whose authenticity and integrity are cryptographically verified on-device is accepted, transfer is resumable across intermittent links, and a failed, interrupted, or rejected update leaves the device on a known-good image with trip data intact and monitoring status honest.
Parent | PRD-SEC-008, PRD-CLOUD-008, PRD-SAFE-010, PRD-ARCH-008
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-12 (DATA: OTA service → device firmware)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | The OTA boundary is a fleet-wide integrity and fail-safety surface; an unsigned or interrupted update must never brick a device or leave it unmonitored-without-notice (AD-7). Standard (secure-update principles, IEC 62443 concepts; revision to confirm).
---
6. Interface integrity & conformance (across boundaries)
SYS-066 — Conformance to the published interface inventory / ICD (I-01…I-12)
Requirement Statement | Each realised subsystem interface shall conform to the typed interface inventory (I-01…I-12) and the interface control documentation, matching the defined type (PWR/DATA/RF/THRM/MECH), direction, and contract, and no undocumented interface shall exist between subsystems or across the device/cloud/mobile trust boundaries.
Parent | PRD-ARCH-009, PRD-ARCH-001
Allocation Target | SYS
Type | Derived
Interface refs | I-01…I-12 (all typed interfaces)
Safety/Security classification | Both
Criticality | High
Verification Method | Inspection, Analysis
Status | Baselined
Rationale | Interface drift between disciplines is a leading source of integration defects; a single authoritative ICD keeps the block diagram, requirements, and design outputs in agreement (AD; rule 2). An undocumented interface is also an ungoverned attack surface. Standard (ISO/IEC/IEEE 29148; revision to confirm).
SYS-067 — Interface error, timeout, and back-pressure handling
Requirement Statement | For every data, RF, and power interface, the system shall define and honour bounded timeout, retry-with-backoff, and back-pressure/flow-control behaviour, such that a slow, unresponsive, overloaded, or disconnected counterpart causes a defined, safe, energy-bounded fallback (buffer, degrade, or safe-state) rather than an unbounded stall, data loss, or energy drain.
Parent | PRD-REL-004, PRD-CONN-008, PRD-ARCH-008
Allocation Target | SYS
Type | Derived
Interface refs | I-04, I-05, I-06, I-08, I-09, I-10, I-12
Safety/Security classification | Both
Criticality | High
Verification Method | Test (fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | An interface that can hang or spin retries is both a data-integrity and an energy hazard under the D3 budget; back-pressure prevents a fast producer (sensors, alert bursts) from overrunning a slow consumer (NVM, RF). Fallbacks must reach the safe state (SYS-057, PRD-ARCH-008). Judgment.
SYS-068 — Authenticated identity at wireless and cross-tier boundaries
Requirement Statement | At the BLE, device↔cloud, and OTA interfaces, each party shall cryptographically authenticate the other's identity before privileged operations (configuration, arming, data ingest, firmware acceptance) are permitted, such that an unauthenticated or spoofed counterpart cannot alter configuration, inject records, or deliver firmware, and every such interaction is attributable in the audit trail.
Parent | PRD-SEC-005, PRD-SEC-006, PRD-SEC-008, PRD-ARCH-006
Allocation Target | S3, S5, S8, S9
Type | Derived
Interface refs | I-09, I-10, I-12
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Status | Preliminary (reconcile with cybersecurity threat model — Evidence-Pending)
Rationale | The trust-boundary split (AD-5) is only real if each side proves who it is; mutual authentication is the control against spoofed provisioning, forged uploads, and rogue firmware, and it is what makes interface actions attributable (non-repudiation). Standard (IEC 62443 concepts, 21 CFR Part 11 attributable; revision to confirm); threat model Evidence-Pending.
System Requirements — Performance & Timing (SYS-033 … SYS-050)
SCM · Stage 06 · Owner: EM-ENG-001 with Reliability, RF, Electronics, and Sensing engineers. Performance and timing system requirements decomposed/allocated from PRD-PERF (1700) and, where the target is environment-bound, PRD-ENV (200). These set the measurable accuracy, latency, endurance, capacity, throughput, and availability targets that make the SCM's verdict trustworthy and its record complete. Field set per [index.md](index.md). Based on D2 (store-and-forward + immediate alerts) and D3 (−30 °C / 90-day energy tension).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Numeric targets carried from the PRDs as Judgment (±0.5 °C, 0.1 °C, ±5 %RH, 5 min, 3 s, 50 % spare) are stated here as the working system target and confirmed in V&V; where the quantity is genuinely unresolved it is marked Evidence-Pending and closed by a named test. The D3 −30 °C / 90-day battery tension is reflected in SYS-045/SYS-046, and the RF-in-container link risk in SYS-048/SYS-049.
---
1. Measurement accuracy & resolution
SYS-033 — Temperature measurement accuracy across the operating range
Requirement Statement | The system shall measure temperature on each channel with an accuracy of ±0.5 °C or better across the −30 °C to +60 °C operating range, referenced to a traceable temperature standard.
Parent | PRD-PERF-001, PRD-ENV-001
Allocation Target | S4 (sensor + signal chain), S3 (linearisation/compensation)
Type | Allocated
Interface refs | I-01 (thermal), I-05 (sensors→MCU)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (traceable reference across range)
Status | Evidence-Pending (accuracy class to confirm)
Rationale | The accept/reject/quarantine verdict is only defensible if the underlying temperature is trustworthy; ±0.5 °C is the working target but must be confirmed against the WHO PQS E006 accuracy class (designation cited, revision to confirm). Judgment / Evidence-Pending.
SYS-034 — Accuracy stability over range, drift, and calibration interval
Requirement Statement | The system shall hold the SYS-033 temperature accuracy across the full operating range and shall remain within that accuracy over the defined calibration interval and the device reuse life, with any lapse of calibration validity flagged rather than silent.
Parent | PRD-PERF-002, PRD-PERF-001
Allocation Target | S4, S3, S8 (calibration-record authority)
Type | Allocated
Interface refs | I-05, I-10 (calibration/readiness record)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test (aging/drift), Analysis
Status | Evidence-Pending (calibration interval unresolved)
Rationale | A sensor that drifts silently produces a confidently-wrong verdict; measurement traceability must be maintained over life. Calibration interval is unresolved and set by drift data. Standard (ISO/IEC 17025 traceability concepts; revision to confirm) / Evidence-Pending.
SYS-035 — Temperature resolution
Requirement Statement | The system shall record temperature with a resolution of 0.1 °C or finer on every stored sample.
Parent | PRD-PERF-003, PRD-PERF-001
Allocation Target | S4, S3, S7
Type | Allocated
Interface refs | I-05, I-06 (MCU→storage)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Preliminary
Rationale | Resolution must be fine enough to support the ±0.5 °C accuracy target and meaningful time-over-threshold excursion logic; a coarse quantum would mask marginal excursions. 0.1 °C is a Judgment target, confirmed vs the PQS class. Judgment.
SYS-036 — Relative-humidity measurement accuracy
Requirement Statement | The system shall measure relative humidity with an accuracy of ±5 %RH or better across the 10–95 %RH non-condensing operating range.
Parent | PRD-PERF-008, PRD-ENV-004
Allocation Target | S4, S3
Type | Allocated
Interface refs | I-05
Safety/Security classification | None
Criticality | Medium
Verification Method | Test
Status | Preliminary
Rationale | Humidity is a secondary condition indicator; accuracy must be good enough to be meaningful for products that depend on it without over-driving cost. ±5 %RH is a Judgment target; condensation on the probe during cold-to-warm transitions is a named accuracy concern (flag Sensing/Materials engineer). Judgment.
2. Acquisition & detection timing
SYS-037 — Per-sample acquisition, settling, and timestamping time
Requirement Statement | The system shall complete acquisition, sensor settling, and trusted timestamping of a full multi-channel sample within a bounded time that is small relative to the minimum 1-minute logging interval, such that scheduled samples are neither delayed into the next interval nor mutually skewed beyond a defined bound.
Parent | PRD-PERF-003, PRD-PERF-001, PRD-FUN-001
Allocation Target | S4, S3
Type | Derived
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Preliminary
Rationale | Gap-free interval logging (SYS-001/SYS-002) at a 1-minute floor is only achievable if the acquire-settle-stamp cycle fits well inside the interval with margin for the low-power wake path; sensor settling time is the dominant term and must not corrupt the timeline. Bound to be set from sensor datasheet settling behaviour. Judgment / Evidence-Pending.
SYS-038 — On-device excursion detection latency
Requirement Statement | The system shall evaluate each new sample against the active thresholds and product-stability logic and shall register a detected excursion within one logging interval of the sample that satisfies the excursion condition, independent of connectivity.
Parent | PRD-PERF-004, PRD-FUN-006, PRD-FUN-008
Allocation Target | S3
Type | Derived
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | The SYS-039 end-to-end alert latency budget presupposes that detection itself is prompt; detection is the safety-critical first link and must complete on-device at the sampling cadence so that offline excursions are still captured at detection time (SYS-013). Judgment.
SYS-039 — Excursion-alert latency (when connected)
Requirement Statement | When connectivity is available at the time of detection, the system shall deliver an excursion alert to the cloud within 5 minutes of on-device detection.
Parent | PRD-PERF-004, PRD-FUN-008, PRD-CONN-009
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-08 (cellular/RF), I-10 (cloud)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (latency, detection→cloud notification)
Status | Evidence-Pending (5-min target to confirm)
Rationale | Fast alerting is what converts passive monitoring into prevented loss — a responder must learn of a break in time to intervene (D2). Latency is connectivity-bound; when offline the alert is captured locally and queued (SYS-040). The 5-minute figure is a Judgment target confirmed in validation, and RF attenuation inside the shipper (SYS-048/SYS-049) may dominate the achievable latency. Judgment / Evidence-Pending.
SYS-040 — Offline alert queue-and-forward latency on reconnection
Requirement Statement | When connectivity is unavailable at detection, the system shall queue the excursion alert in non-volatile storage and shall deliver it, ahead of routine buffered data, within a bounded time of the next usable connection.
Parent | PRD-PERF-004, PRD-FUN-008, PRD-PERF-011
Allocation Target | S3, S5, S8
Type | Derived
Interface refs | I-08, I-09 (Wi-Fi/BLE offload), I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Preliminary
Rationale | Offline-first (D2): an alert generated with no link must not be lost or buried behind bulk log upload; it must jump the queue on reconnection so intervention is still possible. Reconnection-to-delivery bound set with the connectivity budget. Judgment.
3. Interaction & verdict responsiveness
SYS-041 — Interaction / HMI responsiveness
Requirement Statement | The system shall present the requested on-device status or verdict within 3 seconds of a user action, including at the −30 °C cold extreme and at low battery.
Parent | PRD-PERF-007, PRD-ENV-003
Allocation Target | S6, S3
Type | Allocated
Interface refs | I-07 (HMI/display)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test (response time, incl. cold/low-battery)
Status | Preliminary
Rationale | A time-pressured last-mile user at a cold-store dock must not be left waiting or unsure; display and MCU wake times slow at cold and low state-of-charge, so the target is stated at the worst case. 3 s is a Judgment target confirmed in the usability study. Judgment.
SYS-042 — Verdict computation and display latency
Requirement Statement | The system shall compute the accept/reject/quarantine verdict from the stored trip record against the active profile and render it on the device display within the SYS-041 responsiveness budget, with no network, phone, or login required.
Parent | PRD-PERF-007, PRD-FUN-010
Allocation Target | S3, S6
Type | Derived
Interface refs | I-07
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Demonstration
Status | Preliminary (reconcile with HAL)
Rationale | The edge verdict is the decision point for a user with no phone or network (SYS-015); if verdict computation over a full 90-day record cannot complete within the responsiveness budget on the embedded MCU, the offline value proposition fails. Worst case is a full record evaluated at low temperature/low battery. Judgment.
4. Capacity & throughput
SYS-043 — Log capacity for a full offline 90-day trip
Requirement Statement | The system shall store, in non-volatile memory, the complete multi-channel sample and event log for a full 90-day trip at the minimum logging interval with at least 50 % spare capacity, with no connectivity assumed and without overwriting unsent data.
Parent | PRD-PERF-006, PRD-FUN-012, PRD-REL-005
Allocation Target | S7, S3
Type | Allocated
Interface refs | I-06
Safety/Security classification | Both
Criticality | Critical
Verification Method | Analysis (sizing), Test
Status | Evidence-Pending (storage sizing to confirm)
Rationale | An entirely offline trip must still yield a complete record with headroom; store-and-forward requires worst-case offline capacity (D2). Order-of-magnitude: 90 days at a 1-minute interval is ≈ 1.3×10^5 samples/channel, so exact sizing depends on record format and channel count. 50 % spare is a Judgment target. Judgment / Evidence-Pending; closed with SYS-018/SYS-127.
SYS-044 — Sustained logging write throughput without dropped samples
Requirement Statement | The system shall commit each scheduled sample and each detected event to non-volatile storage at the peak combined rate of base logging plus event capture, sustained for the full trip, without dropping, overwriting, or corrupting any record.
Parent | PRD-PERF-006, PRD-FUN-001, PRD-FUN-004
Allocation Target | S7, S3
Type | Derived
Interface refs | I-06
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Preliminary
Rationale | Capacity (SYS-043) is necessary but not sufficient; the write path must keep up with a burst of shock events superimposed on the base logging cadence, else gap-free logging (SYS-002) is violated at exactly the moment the data matters. Peak rate set from the shock-event profile. Judgment.
5. Endurance & the D3 cold-energy tension
SYS-045 — 90-day energy performance under the worst-case cold-duty profile
Requirement Statement | The system shall achieve at least 90 days of continuous armed operation from a single full charge under the defined worst-case cold-duty profile, retaining sensing, logging, on-device verdict, and connected-alert capability throughout, with a documented energy margin.
Parent | PRD-PERF-005, PRD-ENV-001
Allocation Target | S2 (battery/BMS), S3 (duty-cycling), SYS (system energy budget)
Type | Allocated
Interface refs | I-05, I-06, I-08
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (90-day endurance, MR-VAL-006)
Status | Evidence-Pending (energy margin to confirm)
Rationale | The 90-day endurance is the headline program target and the constraining budget (D3); a device that dies mid-trip produces an incomplete record and a false-reassurance hazard. Margin % is unresolved and closed by the endurance test — the top program risk, validated by test. Judgment / Evidence-Pending.
SYS-046 — Endurance and measurement performance retention at the −30 °C extreme
Requirement Statement | The system shall meet the SYS-045 endurance and the SYS-033 accuracy while continuously exposed to the −30 °C operating extreme, accounting for low-temperature loss of usable battery capacity and for charging being inhibited below 0 °C.
Parent | PRD-PERF-005, PRD-PERF-002, PRD-ENV-001
Allocation Target | S2, S3, S4, SYS
Type | Derived
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (cold-chamber endurance + accuracy)
Status | Evidence-Pending (low-temp capacity & fade unresolved)
Rationale | The −30 °C + 90-day combination is in genuine tension (D3): standard Li-ion loses usable capacity and cannot be charged below 0 °C, so endurance and accuracy must be demonstrated at the cold extreme, not extrapolated from room temperature. Depends on final low-temperature chemistry datasheet and capacity fade over reuse cycles. Judgment / Evidence-Pending; the program's top risk, validated by test.
6. Availability
SYS-047 — System availability with no silent outage
Requirement Statement | The system shall be available and actively monitoring for at least a defined fraction of every armed trip, and shall flag any monitoring outage in the record rather than allow it to pass silently.
Parent | PRD-PERF-009, PRD-SAFE-003, PRD-REL-007
Allocation Target | S3, S7, S8
Type | Allocated
Interface refs | I-06, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (availability / self-check)
Status | Evidence-Pending (availability fraction unresolved)
Rationale | "Armed" must reliably mean "monitoring"; a silent outage lets an incomplete record read as a clean pass — a false-reassurance hazard, so any gap must surface in the verdict (SYS-017). The availability fraction is unresolved and reconciled with the HAL. Judgment / Evidence-Pending.
7. Location, delivery & charge (energy- and RF-bound)
SYS-048 — GNSS location cadence within the energy budget
Requirement Statement | The system shall acquire and record device location at a configurable periodic and event-triggered cadence sufficient for fleet management and stray recovery, while consuming no more than the location energy allocation and without breaching the SYS-045 90-day endurance.
Parent | PRD-PERF-010, PRD-FUN-019, PRD-CONN-005
Allocation Target | S4 (GNSS), S3, S5
Type | Allocated
Interface refs | I-08
Safety/Security classification | None
Criticality | Medium
Verification Method | Test (location-cadence + energy)
Status | Evidence-Pending (GNSS-in-container cadence)
Rationale | Location supports stray recovery and last-mile attribution but is energy-costly and heavily RF-attenuated inside the metallized, insulated shipper, so achievable fix cadence within the energy budget is unresolved and may require an external antenna (concept §7). Judgment / Evidence-Pending.
SYS-049 — Data-delivery success rate over a representative trip
Requirement Statement | The system shall deliver at least a defined fraction of buffered data to the cloud over a representative trip whenever any usable connectivity occurs, in intact time order and without data loss, including while operating inside a representative metallized, insulated shipping container.
Parent | PRD-PERF-011, PRD-CONN-007, PRD-ENV-009
Allocation Target | S5, S3, S8
Type | Allocated
Interface refs | I-08, I-09, I-10
Safety/Security classification | None
Criticality | High
Verification Method | Test (RF-in-container / field delivery-success)
Status | Evidence-Pending (success fraction; RF-in-container)
Rationale | Store-and-forward is only valuable if the data actually reaches the cloud over global lanes; the achievable success fraction is bounded by RF attenuation inside the shipper — the principal open architectural risk (RF-in-container) — and is closed by the RF-in-container qualification test (PRD-VER-005). Local capture is unconditional regardless of delivery (SYS-013). Judgment / Evidence-Pending.
SYS-050 — Charge time to trip-readiness
Requirement Statement | The system shall recharge from empty to full trip-readiness within a defined charge-time target at a permitted charge temperature (charging inhibited below 0 °C), supporting predictable depot turnaround.
Parent | PRD-PERF-012, PRD-ELEC-004
Allocation Target | S2, S3
Type | Allocated
Interface refs | I-05 (charge/power path telemetry)
Safety/Security classification | None
Criticality | Medium
Verification Method | Test (charge-time)
Status | Evidence-Pending (charge-time target unresolved)
Rationale | Reuse economics (D1) depend on predictable depot turnaround, but the charge-below-0 °C inhibit and any pre-charge self-heating step (D3) extend achievable charge time; the target is unresolved and set against the final chemistry and charge path. Judgment / Evidence-Pending.
System Requirements — Power / Energy / Battery / Charge (SYS-086 … SYS-103)
SCM · Stage 06 · Owner: EM-ENG-001, integrating the Principal Systems Engineer (power/battery). Power, energy, battery, charging, protection, and charge-safety system requirements, decomposed/allocated from PRD-ELEC (700) and the D3-carrying reliability requirements PRD-REL-001/002/003. Field set per [index.md](index.md). Based on D1 (reusable/rechargeable) and D3 (−30 °C / 90-day energy tension — the top program risk).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Context — the D3 tension. The concept demands 90 days of continuous monitoring across −30 °C to +60 °C from a rechargeable source (D1). This is in genuine tension (D3): standard Li-ion loses usable capacity in deep cold and cannot be safely charged below 0 °C (lithium plating → thermal event). This library states the tension as explicit, testable requirements — the endurance/capacity/retention/quiescent items (SYS-086/087/090/091/092) are Evidence-Pending and validated by test; the charge-safety items (SYS-094/095/096) are severity-first Critical/Safety. No chemistry, capacity, or datasheet number is invented; standards are cited by designation with revision to confirm.
---
1. Cold endurance & energy budget (D3 — top program risk)
SYS-086 — 90-day continuous operation at −30 °C from full charge
Requirement Statement | The system shall operate continuously — sampling, logging, on-device verdict, and scheduled connectivity attempts — for at least 90 days from a single full charge under the defined worst-case cold-duty profile spanning −30 °C, with zero loss of logged data.
Parent | PRD-REL-001, PRD-ELEC-002; D3, MR-RISK-001
Allocation Target | S2 (energy), S3 (duty-cycle governor), SYS
Type | Allocated
Interface refs | I-03 (power distribution), I-05 (sensors→MCU), I-08 (cellular)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (worst-case cold-duty profile and margin to confirm; MR-VAL-006)
Rationale | The headline endurance promise and the top program risk — a dead-in-cold device loses the record exactly when it matters most. Evidence-Pending (D3); closed by the 90-day cold-endurance test.
SYS-087 — Usable-capacity and energy-budget conformance across temperature
Requirement Statement | The system shall be designed to a documented energy budget in which the usable battery energy across the −30 °C to +60 °C operating range, less all derating and reserves, is at least the total energy consumed over the worst-case 90-day cold-duty profile with a defined positive margin.
Parent | PRD-ELEC-008, PRD-ELEC-002, PRD-REL-001; D3
Allocation Target | S2, S3, SYS
Type | Allocated
Interface refs | I-03
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Analysis, Test
Status | Evidence-Pending (usable-capacity-at-cold and margin % unresolved pending cell datasheet & RF-energy test)
Rationale | Endurance (SYS-086) is an energy-budget outcome: supply-side usable energy at cold must exceed demand-side consumption with margin, or the promise fails silently. Evidence-Pending; the power budget (Stage 00) is the baseline, radio energy closes on RF-in-container test.
SYS-088 — Low-temperature-capable rechargeable energy source
Requirement Statement | The system shall derive its operating energy from a rechargeable battery whose selected chemistry is specified to deliver the usable capacity required by SYS-087 for device operation across the −30 °C to +60 °C operating range.
Parent | PRD-ELEC-001; D1, D3
Allocation Target | S2
Type | Allocated
Interface refs | I-03
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (final chemistry — e.g. low-temp Li-ion or LTO-class — and low-temp datasheet unresolved)
Rationale | Reusability (D1) requires a rechargeable cell; deep-cold operation (D3) requires a low-temperature-tolerant one. Chemistry and its cold datasheet are Evidence-Pending (concept §7) — not asserted; closed by cold-capacity test.
SYS-089 — Duty-cycled power-state energy governor with hard radio gating
Requirement Statement | The system shall enforce an autonomous power-state governor that holds the device in its lowest viable sleep state between scheduled activities and gates each high-drain load — cellular, Wi-Fi, BLE, and GNSS in particular — such that no radio or GNSS activity occurs outside its scheduled, energy-budgeted window (SYS-090) except a queued excursion alert.
Parent | PRD-ELEC-008, PRD-ARCH-004; D2, D3
Allocation Target | S3, S2, S5
Type | Allocated
Interface refs | I-03, I-08, I-09
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Continuous live cellular for 90 days is energy-prohibitive (D2); radios are the dominant load and must be bounded hard. Duty-cycling is the enabler that makes SYS-086 achievable. Judgment; quantified by current-profile test.
SYS-090 — Energy-budget allocation across functional loads
Requirement Statement | The system shall meet a documented energy budget that allocates the 90-day worst-case cold energy consumption across defined load classes — sleep/quiescent, sensing, MCU/compute, storage, HMI/display, cellular, Wi-Fi, BLE, and GNSS — with a per-class allocation whose sum plus reserve does not exceed the usable capacity of SYS-087.
Parent | PRD-ELEC-008, PRD-PERF-005; D2, D3
Allocation Target | S2, S3, S4, S5
Type | Derived
Interface refs | I-03, I-05, I-08
Safety/Security classification | None
Criticality | High
Verification Method | Analysis, Test
Status | Evidence-Pending (per-class allocations, esp. radio and GNSS energy, unresolved pending RF-in-container test)
Rationale | A single 90-day number is not designable without a per-load allocation to hold subsystems accountable and to expose the radio/GNSS dominance the governor (SYS-089) must control. Evidence-Pending; radio/GNSS energy closes on RF-in-container qualification.
SYS-091 — Capacity retention over reuse cycles
Requirement Statement | The system shall continue to meet SYS-086 over its specified number of reuse/charge cycles, or shall reliably detect and flag when battery health can no longer meet it and block redeployment of the affected device.
Parent | PRD-ELEC-003, PRD-REL-002, PRD-FUN-020; D1, D3
Allocation Target | S2, S3, S8
Type | Allocated
Interface refs | I-03, I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (cycle count and end-of-life retention threshold unresolved)
Rationale | Capacity fade over reuse can silently cause a mid-life dead-in-cold event (part of MR-RISK-001); a reusable fleet must self-flag end-of-usable-life. Evidence-Pending; closed by cycle-life test, ties readiness (SYS-030).
SYS-092 — Quiescent / sleep current budget at −30 °C
Requirement Statement | The system shall hold its quiescent (deep-sleep) current, including self- discharge and always-on housekeeping (RTC, protection, wake logic), within the sleep-state allocation of SYS-090 at −30 °C so that idle drain over the 90-day trip does not consume the SYS-087 endurance margin.
Parent | PRD-ELEC-008, PRD-ELEC-002; D3
Allocation Target | S2, S3
Type | Derived
Interface refs | I-03
Safety/Security classification | None
Criticality | High
Verification Method | Test
Status | Evidence-Pending (quiescent-current target and cold self-discharge rate unresolved)
Rationale | With most of a 90-day trip spent asleep, quiescent drain — not active bursts — often dominates the budget; cold changes both leakage and self-discharge. Evidence-Pending; measured by cold current-profile test.
2. Charging & low-temperature charge safety
SYS-093 — USB-C recharge to trip-readiness
Requirement Statement | The system shall recharge its battery through a USB-C interface from a standard compliant source and reach full trip-readiness within the charge-time target set in the product spec, without requiring a proprietary charger.
Parent | PRD-ELEC-004; D1
Allocation Target | S2, S1 (sealed port)
Type | Allocated
Interface refs | I-04 (USB-C charge input)
Safety/Security classification | None
Criticality | High
Verification Method | Test
Status | Evidence-Pending (charge-time target unresolved)
Rationale | Reusability (D1) requires convenient standard recharge at the depot for fast turnaround. Standard (USB-C / USB-PD, revision to confirm); charge time Evidence-Pending.
SYS-094 — Charge inhibit below the safe low-temperature charge limit
Requirement Statement | The system shall inhibit battery charging whenever the cell temperature is below 0 °C (or the cell-defined safe lower charge limit, whichever is higher) and shall clearly indicate the charge-inhibited state, such that charging can never proceed below the safe cell limit.
Parent | PRD-ELEC-005, PRD-REL-003; D3, UN-SAFE-007
Allocation Target | S2, S4 (cell temp sense), S3, S6
Type | Allocated
Interface refs | I-04, I-01 (thermal), I-07 (HMI)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL — threshold per cell datasheet)
Rationale | Charging a cold Li-ion cell risks lithium plating and a thermal event — a safety hazard, not merely a performance loss. Severity-first Critical. Standard (IEC 62368-1; IEC 62133 concepts; revision to confirm); reconcile with HAL (Stage 10).
SYS-095 — Self-heat-before-charge (under evaluation)
Requirement Statement | Where the selected design provides it, the system shall safely warm the cell to a permitted charge temperature before charging a cold-returned device, without exceeding safe thermal limits and without overriding the charge-inhibit of SYS-094.
Parent | PRD-ELEC-006; D3
Allocation Target | S2, S3
Type | Allocated
Interface refs | I-04, I-01, I-03
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (feasibility under evaluation — D3)
Rationale | Self-heating improves depot turnaround for cold returns without unsafe cold charging, but must never become a path that defeats SYS-094. Evidence-Pending (feasibility, D3); reconcile with HAL.
SYS-096 — Battery protection (OV / UV / OC / OT / short)
Requirement Statement | The system's battery management shall protect the cell against overvoltage, undervoltage, overcurrent, over-temperature, and short-circuit in normal use, foreseeable misuse, and single-fault conditions, transitioning to a safe state that does not become a fire or thermal hazard.
Parent | PRD-ELEC-007; UN-SAFE-008
Allocation Target | S2, S3
Type | Allocated
Interface refs | I-03, I-04
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | The device must never become a fire/thermal hazard on aircraft, in cold stores, clinics, or homes. Severity-first Critical. Standard (IEC 62368-1 / IEC 62133 concepts, revision to confirm); reconcile with HAL (Stage 10).
3. State-of-charge, degradation & rails
SYS-097 — Fuel-gauge / state-of-charge and state-of-health accuracy
Requirement Statement | The system shall measure and report battery state of charge and state of health with accuracy across the operating temperature range sufficient to support readiness decisions (SYS-030/091) and a meaningful, non-misleading low-battery warning (SYS-098).
Parent | PRD-ELEC-009, PRD-FUN-020
Allocation Target | S2, S3
Type | Allocated
Interface refs | I-03, I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (temperature-dependent accuracy target unresolved)
Rationale | Readiness triage and low-battery warning are only as trustworthy as the gauge behind them; cold shifts the state-of-charge relationship. Evidence-Pending; closed by gauge-accuracy test.
SYS-098 — Low-battery graceful degradation and warning
Requirement Statement | The system shall warn of low battery and, as energy depletes, shed discretionary functions (uploads, location/GNSS) before core sensing and logging, preserving the safety-relevant record and honest monitoring status to the last, and shall never silently stop monitoring an armed device.
Parent | PRD-ELEC-010, PRD-REL-008; UN-SAFE-010
Allocation Target | S3, S2, S6
Type | Allocated
Interface refs | I-03, I-07
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (degradation order and warning thresholds unresolved)
Rationale | A depleting device must protect the record over convenience — losing the log is worse than losing an upload. Evidence-Pending (degradation order); closed by depletion test.
SYS-099 — Regulated power rails and brown-out protection
Requirement Statement | The system shall provide regulated power rails and brown-out protection such that a falling or collapsing supply causes a clean, deterministic safe-state entry with no corruption of committed logged data or configuration.
Parent | PRD-ELEC-011, PRD-REL-005; UN-SAFE-010
Allocation Target | S2, S3, S7
Type | Allocated
Interface refs | I-03, I-06 (MCU→storage)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | A brown-out mid-write must never corrupt the record; a clean safe-state is the guardrail for power-fail-safe writes (SYS ties to storage integrity). Judgment; closed by brown-out / power-cycle integrity test.
4. Indication, transport & fault-tolerant charge input
SYS-100 — Charge and charge-inhibit state indication
Requirement Statement | The system shall provide a low-energy status indication that distinguishes at least the armed/monitoring, fault, charging, charge-complete, and charge-inhibited states, so that the charge-inhibited condition of SYS-094 is unambiguously visible to a depot operator.
Parent | PRD-ELEC-012, PRD-ELEC-005; UN-SAFE-007
Allocation Target | S6, S3
Type | Allocated
Interface refs | I-07
Safety/Security classification | Safety
Criticality | Medium
Verification Method | Demonstration
Status | Baselined
Rationale | A charge-inhibited device that looks dead would be mistaken for a fault or a failed charge; the inhibit state must be legible without draining the budget. Judgment; energy-cheap indication.
SYS-101 — UN 38.3 / IATA transport compliance of the battery
Requirement Statement | The system's battery/pack shall pass UN 38.3 transport testing and the device and its shipping documentation shall carry the resulting evidence sufficient for compliant air and ground transport of the lithium battery.
Parent | PRD-ELEC-015; UN-REG-006, UN-SAFE-009; D4
Allocation Target | S2, SYS
Type | Allocated
Interface refs | I-03
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (UN 38.3)
Status | Evidence-Pending (test summary unresolved)
Rationale | Legal, safe air/ground transport of the Li battery is a compliance-blocking gate for a globally distributed device (D4). Standard (UN 38.3 / IATA DGR, revision to confirm); Evidence-Pending test report.
SYS-102 — Reverse-polarity / fault-tolerant charge input
Requirement Statement | The system's charge input shall tolerate reverse polarity, over-voltage, and non-compliant or faulty sources without damage to the device or creation of a hazard, and shall reject or safely limit an out-of-specification input.
Parent | PRD-ELEC-016, PRD-ELEC-004; UN-SAFE-008
Allocation Target | S2
Type | Allocated
Interface refs | I-04
Safety/Security classification | Safety
Criticality | High
Verification Method | Test (abuse-input)
Status | Preliminary (reconcile with HAL)
Rationale | Field and depot chargers vary widely; a bad source must not damage or endanger a reusable device. Judgment; closed by abuse-input test, reconcile with HAL.
SYS-103 — ESD and isolation on the charge port
Requirement Statement | The system's USB-C charge port and its interface to internal electronics shall provide ESD protection and isolation sufficient to meet the immunity levels of the PRD-REG EMC requirements without malfunction, data corruption, or damage during handling and depot connection.
Parent | PRD-ELEC-013, PRD-ELEC-016; UN-ENV-009, UN-REG-005
Allocation Target | S2, S1 (port), S3
Type | Allocated
Interface refs | I-04, I-03
Safety/Security classification | Both
Criticality | High
Verification Method | Test (IEC 61000-4-2)
Status | Preliminary (immunity levels set in PRD-REG; reconcile with HAL)
Rationale | The externally accessible charge port is the primary ESD entry point during handling; a discharge must not corrupt data or damage the device. Standard (CISPR / IEC 61000, revision to confirm); immunity levels per PRD-REG.
System Requirements — Regulatory & Compliance (SYS-177 … SYS-190)
SCM · Stage 06 · Owner: EM-ENG-001, integrating category owner EM-ENG-006 (Quality/Regulatory). Regulatory/market-access system requirements decomposed and allocated from PRD-REG (1400) — parents PRD-REG-001 … PRD-REG-012 — and based on D4 (multi-market: WHO PQS + EU GDP + US + air/ground transport). Field set per [index.md](index.md). These state what the SCM system as a whole shall be/do to be sold, deployed, flown, and trusted across its target markets; they do not restate the technical means of compliance (that is the design/verification stages).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Two named regulatory gaps are carried forward from concept §7 and remain open at this baseline: the exact WHO PQS temperature-monitoring device/logger category and specification revision (SYS-177), and the US FDA regulatory classification of the SCM as a distribution-monitoring information accessory (SYS-178). Both are Evidence-Pending; requirements that depend on either (notably the quality-system basis, SYS-188, and the ISO 13485 applicability question) inherit that status and are conditional on their resolution. Consistent with the evidence posture (index.md §"Applicability & evidence posture"), every standard below is cited by designation only, revision to confirm, and no standard clause number is asserted as fact.
---
1. Market access & classification
SYS-177 — WHO PQS eligibility as a temperature-monitoring device/logger
Requirement Statement | The system shall conform to the requirements of the applicable WHO PQS temperature-monitoring device/logger category to a degree sufficient to achieve and hold WHO PQS listing for the target donor-funded/LMIC procurement markets.
Parent | PRD-REG-001, PRD-REG-002
Allocation Target | SYS (device S1–S7 performance + S8 records evidence), with S4/S3 (measurement) and S6 (indication) as the primary conformance-bearing subsystems
Type | Allocated
Interface refs | I-05 (sensors→MCU), I-07 (device HMI), I-10 (device↔cloud)
Safety/Security classification | None (compliance)
Criticality | Critical
Verification Method | Test, Inspection
Status | Evidence-Pending (exact WHO PQS category/code & spec revision unresolved — concept §7)
Rationale | WHO PQS listing is a hard market-access gate for donor-funded and LMIC procurement (D4); the accuracy, detection, and record capabilities must demonstrably meet the category performance spec. Evidence-Pending — the exact WHO PQS temperature-monitoring category and its current specification revision are a named concept gap; cited by designation (WHO PQS E006-line, revision to confirm), no clause asserted.
SYS-178 — Per-market regulatory classification basis (information accessory)
Requirement Statement | The system shall have a documented regulatory status/classification for each target market (WHO PQS category, US FDA classification, EU GDP applicability, and each other regional regime), established on the basis that the SCM is a distribution-monitoring information/data-logging accessory that neither preserves nor controls product temperature and is not itself therapeutic.
Parent | PRD-REG-002, PRD-REG-001
Allocation Target | SYS
Type | Allocated
Interface refs | — (programmatic/system-level; no runtime interface)
Safety/Security classification | None (compliance)
Criticality | Critical
Verification Method | Analysis
Status | Evidence-Pending (US FDA classification & WHO PQS category are named concept gaps — §7)
Rationale | The device must be placeable, registrable, and defensible per market, and its downstream quality-system and design-control obligations (SYS-188) hinge on the confirmed classification (D4). Evidence-Pending — US classification and WHO PQS category are explicit unresolved concept gaps; the accessory/non-therapeutic framing is a Judgment from the intended-use statement (concept §2), to be confirmed with each regulator.
2. Records, data-integrity & computerized-system validation
SYS-179 — EU GDP + 21 CFR Part 11 / ALCOA+ electronic-records conformance
Requirement Statement | The system's electronic records and electronic signatures (device, cloud, and mobile tiers) shall conform to EU GDP data-integrity expectations and 21 CFR Part 11 / ALCOA+ — records that are attributable, legible, contemporaneous, original, accurate, and additionally complete, consistent, enduring, and available — such that a trip record is admissible for release/acceptance and audit.
Parent | PRD-REG-003
Allocation Target | S3, S7, S8, S9 (SYS integration)
Type | Allocated
Interface refs | I-06 (MCU→storage), I-10 (device↔cloud), I-11 (cloud→external/export)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Analysis, Test
Status | Baselined
Rationale | The auditable, tamper-evident record is the product's compliance value; records must survive an audit and support disposition (D2). Standard — EU GDP (2013/C 343/01-line) and 21 CFR Part 11 / ALCOA+, cited by designation, revision to confirm; no clause asserted. Realised by SYS-002/006/020/026/027/028; forward to CSV (SYS-189).
SYS-189 — Computerized-system validation (CSV) of device, cloud and mobile software
Requirement Statement | The system's regulated software across the device firmware, cloud platform, and mobile application shall be developed and maintained under a computerized-system-validation approach consistent with GAMP 5 and applicable CSV expectations, commensurate with the risk of the regulated data each component creates, stores, transforms, or presents.
Parent | PRD-REG-012, PRD-REG-003
Allocation Target | S3, S8, S9 (SYS integration)
Type | Allocated
Interface refs | I-10, I-11, I-12 (OTA)
Safety/Security classification | Both
Criticality | High
Verification Method | Analysis
Status | Baselined
Rationale | Records and dispositions produced by a computerized system are only defensible to regulated customers if that system is validated to a recognized framework and kept in a validated state across OTA changes (SYS-032). Standard — GAMP 5, cited by designation, revision to confirm; risk-based scoping is a Judgment pending the classification of SYS-178.
3. Radio, EMC & electrical/thermal safety
SYS-180 — Radio type approval per market
Requirement Statement | The system's radios (BLE, Wi-Fi, LTE-M, and GNSS receiver) shall hold the radio/spectrum type approvals required in each target market — including FCC Part 15 (US), CE RED (EU), and the applicable regional equivalents — for every hardware/frequency variant placed in that market.
Parent | PRD-REG-004
Allocation Target | S5, S1 (antenna/enclosure), SYS (per-market variant governance)
Type | Allocated
Interface refs | I-08 (RF/antenna), I-09 (BLE↔mobile)
Safety/Security classification | None (compliance)
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (per-market variant set & regional regimes to confirm — D4)
Rationale | Radio approval is a legal-to-sell/legal-to-operate gate per geography and may drive per-market hardware variants (D4). Standard — FCC Part 15 / CE RED and regional equivalents, cited by designation, revision to confirm; the specific regional regime set per market is Evidence-Pending (D4 market list to finalise).
SYS-181 — Electromagnetic compatibility (emissions & immunity)
Requirement Statement | The system shall meet electromagnetic-compatibility emissions and immunity requirements for its use and transport environments per the applicable CISPR 32 / CISPR 35 and IEC 61000 series (and each market's equivalent), such that it neither emits disruptive interference nor malfunctions under the defined electromagnetic disturbances.
Parent | PRD-REG-005
Allocation Target | S3, S5, S2, S1 (SYS integration)
Type | Allocated
Interface refs | I-08, I-01 (thermal/mechanical env), I-05
Safety/Security classification | Both (immunity-induced malfunction can suppress a genuine excursion alert)
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | The SCM must be safe and accepted for aircraft cargo, enterprise, and clinical environments, and must not lose monitoring integrity under disturbance (D4). Standard — CISPR 32/35 and IEC 61000 series, cited by designation, revision to confirm; no clause asserted.
SYS-182 — Electrical and thermal safety
Requirement Statement | The system shall meet the electrical- and thermal-safety requirements of IEC 62368-1 (and each market's national equivalent) for its charging, operating, storage, and transport conditions across the −30 °C to +60 °C operating range, including single-fault conditions of the USB-C charge path and battery subsystem.
Parent | PRD-REG-005
Allocation Target | S2, S1, S3 (SYS integration)
Type | Allocated
Interface refs | I-02 (USB-C charge), I-03 (battery/power), I-01 (thermal)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Electrical/thermal safety of the Li-ion pack and charge path is a hard safety and market-acceptance requirement, tightly coupled to the D3 low-temperature charge-inhibit strategy (see SYS-149-line safety and power SYS-090/091). Standard — IEC 62368-1, cited by designation, revision to confirm; no clause asserted. Reconcile safety verification with HAL (Stage 10).
4. Battery transport & calibration metrology
SYS-183 — Lithium-battery transport (UN 38.3 / IATA DGR)
Requirement Statement | The system's lithium battery shall be UN 38.3-tested and the finished device shall be transportable in compliance with the IATA Dangerous Goods Regulations (and applicable ground modal rules), with the required test summary and transport documentation available for each shipping configuration.
Parent | PRD-REG-006
Allocation Target | S2, SYS (packaging/documentation)
Type | Allocated
Interface refs | I-03 (battery/power)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Inspection
Status | Evidence-Pending (UN 38.3 test summary & IATA guidance to be produced — PRD-REG-006 note)
Rationale | Legal, safe air/ground transport of the Li battery is a market-access gate and avoids dangerous-goods holds across the global distribution chain (D4). Standard — UN 38.3 and IATA DGR, cited by designation, revision to confirm; the test summary and shipping guidance are Evidence-Pending until the final chemistry (concept §7) is fixed.
SYS-184 — Traceable calibration to national/metrology standards
Requirement Statement | The system's temperature (and, where claimed, humidity) measurement shall be calibrated with traceability to recognized national/international metrology standards, and each device shall record its calibration status, calibration date, and calibration interval such that any reading and the disposition derived from it is metrologically defensible.
Parent | PRD-REG-007
Allocation Target | S4, S3, S7 (record), S8 (fleet readiness)
Type | Allocated
Interface refs | I-05, I-06, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Inspection, Test
Status | Evidence-Pending (calibration interval unresolved — PRD-REG-007 note)
Rationale | Readings and the accept/reject/quarantine verdicts based on them must be traceable to a recognized standard to be defensible in audit and release (WHO PQS accuracy, EU GDP). Standard — ISO/IEC 17025 concepts, cited by designation, revision to confirm; the calibration interval is Evidence-Pending (regulatory/service stage). Links to SYS-024 (pre-arm valid-calibration check) and SYS-030 (readiness).
5. Environmental, hazardous-substance, end-of-life & privacy
SYS-185 — Hazardous-substance, environmental & battery end-of-life compliance
Requirement Statement | The system shall comply with applicable hazardous-substance restrictions (RoHS, REACH) and electronics/battery end-of-life obligations (WEEE and applicable battery directives/ regulations) for each target market, and shall be supported by the corresponding material declarations and a defined take-back/recycling/decommissioning path for the reusable fleet.
Parent | PRD-REG-008
Allocation Target | S1, S2, SYS (materials & documentation)
Type | Allocated
Interface refs | — (materials/lifecycle; no runtime interface)
Safety/Security classification | None (compliance/sustainability)
Criticality | High
Verification Method | Inspection
Status | Evidence-Pending (per-market EoL path & declarations to confirm — PRD-REG-008 note)
Rationale | A reusable, rechargeable, globally distributed device (D1) must be sold and retired lawfully and responsibly across markets (concept C10). Standard — RoHS/REACH/WEEE and applicable battery directives, cited by designation, revision to confirm; the per-market EoL path and declaration set are Evidence-Pending.
SYS-186 — Regional data-protection & privacy conformance
Requirement Statement | The system shall handle any personal or location-derived data (e.g. device GNSS position, operator/actor identifiers in the audit trail) in conformance with the applicable regional data-protection law of each market of deployment — including the EU GDPR and each other regional privacy regime — applying data-minimisation appropriate to a cargo-monitoring (not people-monitoring) device.
Parent | PRD-REG-009
Allocation Target | S8, S9, S3 (data captured on device)
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Security
Criticality | High
Verification Method | Inspection, Analysis
Status | Evidence-Pending (applicable-law mapping per market to confirm — PRD-REG-009 note)
Rationale | The platform must be lawful to deploy per region (D4); although the device monitors cargo rather than people, location and actor-attribution data can still be personal data. Standard — GDPR and applicable regional data-protection law, cited by designation, revision to confirm; the applicable-law mapping is Evidence-Pending. Realised with SYS-security (SYS-132-line) controls.
6. Labeling, documentation, quality-system & market-access governance
SYS-187 — Labeling & documentation per regulatory regime
Requirement Statement | The system shall carry the durable labeling, and be supported by the documentation, that each target regime requires — including safety/EMC conformity marks, radio-approval marks, battery/transport (UN 38.3 / IATA) marks and handling documentation, WHO PQS/EU GDP documentation, and a unique device identifier — appropriate to each per-market variant.
Parent | PRD-REG-010
Allocation Target | S1 (durable marking), SYS (documentation set)
Type | Allocated
Interface refs | — (physical labeling/documentation; no runtime interface)
Safety/Security classification | None (compliance)
Criticality | High
Verification Method | Inspection
Status | Evidence-Pending (full documentation set & per-regime mark set to confirm — PRD-REG-010 note)
Rationale | The device must pass acceptance and audit at every gate — customs, air transport, procurement, and clinical intake — which each require specific marks and documents (D4). Standard — WHO PQS, IATA DGR, CE/FCC marking regimes, cited by designation, revision to confirm; the documentation set and mark set are Evidence-Pending. Realised with SYS-mechanical durable-labeling requirements.
SYS-188 — Quality-system basis for a regulated device (ISO 13485 applicability conditional)
Requirement Statement | The system shall be developed and maintained under a quality system appropriate to its confirmed regulatory classification (SYS-178), including design controls and an ISO 14971 risk- management file; where the confirmed classification designates the SCM as a medical device or a regulated device accessory, the quality system shall additionally conform to ISO 13485.
Parent | PRD-REG-011, PRD-REG-002
Allocation Target | SYS (programmatic/QMS)
Type | Allocated
Interface refs | — (programmatic; no runtime interface)
Safety/Security classification | None (compliance/quality)
Criticality | Critical
Verification Method | Analysis
Status | Evidence-Pending (classification unresolved; ISO 13485 applicability conditional — §7)
Rationale | Market access and defensibility require a QMS and design-control record proportionate to the device's regulatory status; the ISO 14971 risk-management file applies regardless, but ISO 13485 applicability is contingent on the classification decision. Standard — ISO 14971 (applies); ISO 13485 cited by designation, applicability Evidence-Pending and conditional on SYS-178, revision to confirm; no clause asserted.
SYS-190 — Market-access conformance gate and maintenance of conformity
Requirement Statement | The system shall not be deployed into, or operated in, a target market for which its required market-access conformance (WHO PQS listing, per-market radio type approval, EMC/safety, transport, and applicable classification/registration) is not held and current, and the system shall maintain conformity across firmware/hardware changes such that a change affecting a conformance basis is identified and re-assessed before it is fielded.
Parent | PRD-REG-001, PRD-REG-002, PRD-REG-004, PRD-REG-005, PRD-REG-010
Allocation Target | SYS (S8 fleet/deployment governance; S3/S5 change-affected variants; S12 OTA change-control)
Type | Derived
Interface refs | I-10, I-12 (OTA)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Analysis, Inspection
Status | Evidence-Pending (per-market approval status & change-control process to confirm — D4)
Rationale | In a multi-market fleet (D4), a device fielded without current approvals, or changed by OTA (SYS-032) in a way that invalidates an approval, is both a legal and a data-integrity exposure; conformance must be gated at deployment and maintained across change. Judgment based on D4 and the maintenance-of-conformity principle common to CE RED / FCC / WHO PQS regimes; specific per-market approval status and the change-control process are Evidence-Pending. Realised with SYS-030 (readiness/deployment block) and SYS-032/SYS-189 (validated change).
System Requirements — Reliability & Availability (SYS-163 … SYS-176)
SCM · Stage 06 · Owner: EM-ENG-001 (integration) with Principal Reliability/Materials Engineer as category owner. Reliability, endurance, availability, storage-integrity, degradation, robustness-over-reuse, and platform-durability system requirements, decomposed/allocated from PRD-REL (1200) and based on D1 (reusable/ rechargeable, multi-trip) and D3 (low-temperature Li-ion; −30 °C / 90-day endurance in genuine tension — the top program risk, MR-RISK-001). Field set per [index.md](index.md). Reliability standards are cited by designation only (revision/method to confirm); numeric targets that are not yet closed by test carry an Evidence-Pending status and are validated by test forward to the PVL/PVaL.
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Evidence posture for this category (honesty). The D3 −30 °C / 90-day set (SYS-163/164/165) is the headline reliability promise and the top open risk; these are stated as measurable requirements closed by test, never asserted as met. Reliability-prediction figures (MTBF/failure rate), cycle count, degradation order, platform SLA, and RPO/RTO are not invented here — they are marked Evidence-Pending and resolved by the reliability program (SYS-172) and the named PVL tests.
---
1. Endurance & cycle life (D3 — top program risk)
SYS-163 — 90-day cold endurance from full charge (D3)
Requirement Statement | The system shall operate continuously — sensing, logging, on-device verdict computation, and scheduled connectivity attempts — for at least 90 days from a single full charge under the defined worst-case cold duty profile spanning −30 °C, with zero loss of committed logged data.
Parent | PRD-REL-001, PRD-PERF-006; D3 (UN-PERF-003, UN-ENV-001)
Allocation Target | S2 (energy), S3 (duty-cycle scheduling), S4, S7, SYS
Type | Allocated
Interface refs | I-01 (thermal), I-02 (power), I-05, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (endurance under worst-case cold profile unresolved — validated by test)
Rationale | The headline endurance promise and top program risk (D3, MR-RISK-001): a device that dies in the cold loses the record exactly when it matters. Achieved via aggressive duty-cycling (D2/D3). Endurance margin at −30 °C is not yet demonstrated. Evidence-Pending; closed by the 90-day cold-endurance test (MR-VAL-006), coordinated with SYS-086/090 (power) and Power engineer.
SYS-164 — Cold-endurance capacity margin over the worst-case duty profile (D3)
Requirement Statement | The system shall retain, at the end of the 90-day worst-case cold duty profile, a defined non-zero energy reserve sufficient to complete an orderly graceful-degradation sequence (SYS-170) and preserve the record, rather than depleting to an uncontrolled shutdown.
Parent | PRD-REL-001, PRD-REL-008; D3
Allocation Target | S2, S3, SYS
Type | Derived
Interface refs | I-02, I-06
Safety/Security classification | Safety
Criticality | High
Verification Method | Test, Analysis
Status | Evidence-Pending (reserve target unresolved — validated by test)
Rationale | Endurance must not be a knife-edge; low-temperature Li-ion voltage sag and capacity loss make end-of-trip the highest-risk moment (D3). A quantified reserve converts "90 days" into a controlled landing. Reserve fraction is Evidence-Pending, set by the power budget and confirmed in the endurance/depletion tests; Power engineer to weigh in on chemistry-specific cold derating.
SYS-165 — Endurance sustained across reuse/charge-cycle life (D1/D3)
Requirement Statement | The system shall continue to meet SYS-163 across its specified number of reuse/charge cycles, or shall reliably detect when battery state-of-health can no longer meet SYS-163 and block the device from being armed/redeployed until serviced.
Parent | PRD-REL-002; D1 (reuse), D3
Allocation Target | S2 (SoH), S3, S8 (readiness gating)
Type | Allocated
Interface refs | I-02, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (cycle count and fade budget unresolved — validated by test)
Rationale | Capacity fade over the D1 reuse loop can silently cause mid-life dead-in-cold (part of MR-RISK-001); a reusable fleet must fail safe by self-flagging, not by an unmonitored trip. Cycle count and per-cycle fade budget are Evidence-Pending; closed by the cycle-life test and tied to readiness reporting (SYS-030) and Power engineer's SoH model.
SYS-166 — Low-temperature charge-safety reliability over life (D3)
Requirement Statement | The system's charge-inhibit and, if provided, pre-charge self-heat functions shall operate reliably across the full operating temperature range and over the full reuse life, such that charging never occurs while the cell is below its manufacturer-defined safe low-temperature charge limit.
Parent | PRD-REL-003; D3 (UN-SAFE-007, UN-PERF-004)
Allocation Target | S2 (BMS/charge path), S3
Type | Allocated
Interface refs | I-02, I-03 (USB-C charge)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Evidence-Pending (over-life reliability of inhibit/self-heat unresolved — validated by test)
Rationale | The charge-safety leg of the D3 tension must be reliable, not merely present — charging Li-ion below its low-temperature limit risks lithium plating and latent internal-short failure. Reliability over temperature and cycles is Evidence-Pending; closed by charge-inhibit reliability test and reconciled with the HAL battery hazard (Stage 10). Standard (IEC 62133 cell-level charge limits; revision to confirm). Cross-flag Safety engineer and Power engineer.
2. Fail-safe, storage integrity & recovery
SYS-167 — Defined safe state on any fault
Requirement Statement | On any power, connectivity, firmware, sensor, or reset fault, the system shall enter a defined known safe state that preserves committed logged data and an honest monitoring status, and shall attempt automatic recovery to normal monitoring, never entering a state that corrupts the record or falsely reads as monitoring.
Parent | PRD-REL-004; UN-SAFE-010, UN-FUN-014
Allocation Target | S3, S7, SYS
Type | Allocated
Interface refs | I-06, I-07
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | A fault must never corrupt the authoritative record nor leave the device silently non- monitoring (no false reassurance). The safe-state definition is Judgment pending HAL reconciliation (Stage 10); closed by fault-injection test. Cross-flag Safety engineer (SYS-149-line).
SYS-168 — Power-fail-safe atomic log writes (no corruption on brown-out)
Requirement Statement | The system's non-volatile log shall commit records atomically such that any committed sample, event, or audit record is preserved intact across power loss, brown-out, reset, and firmware update, with no corruption of previously committed data and no partial/torn records surviving as valid.
Parent | PRD-REL-005, PRD-FUN-017; UN-FUN-009/014
Allocation Target | S7, S3
Type | Allocated
Interface refs | I-06
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | Record completeness/integrity is the product's core value; low-temperature brown-out under voltage sag (D3) makes power-fail-safe writes essential. Journaled/atomic-write design; a torn write must fail the record, not silently corrupt it. Judgment; closed by power-cycle storage-integrity test. Ties to brown-out handling (PRD-ELEC-011) — Power engineer to confirm brown-out detection threshold.
SYS-169 — Watchdog supervision and autonomous recovery with logged recovery event
Requirement Statement | The system firmware shall be supervised by a watchdog that returns a hung or unresponsive device to an armed monitoring state without user intervention, and shall record every watchdog recovery as an attributable, time-stamped event so a recovered device is never assumed to have monitored continuously.
Parent | PRD-REL-006; UN-PERF-011, UN-SAFE-003
Allocation Target | S3, S7
Type | Allocated
Interface refs | I-06
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Baselined
Rationale | A silent hang creates an unmonitored trip believed monitored — a false-reassurance hazard; the logged recovery event feeds the honest-monitoring indication (SYS-017). Judgment; closed by watchdog test.
3. Availability, degradation & robustness
SYS-170 — Availability self-check and silent-outage prevention
Requirement Statement | While armed, the system shall periodically self-check its sensing→logging chain and shall, on detecting an outage, flag the device as not monitoring (locally and to the cloud) such that an armed device is at all times either verifiably monitoring or visibly flagged as not, never silently outaged.
Parent | PRD-REL-007; UN-PERF-011, UN-SAFE-003
Allocation Target | S3, S4, S6, S8
Type | Allocated
Interface refs | I-05, I-07, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | "Armed" must reliably mean "monitoring"; a device that appears alive while not monitoring is a false-reassurance hazard. Self-check cadence is Judgment pending HAL; closed by self-check test. Cross-flag Safety engineer (PRD-SAFE-003) and HMI (PRD-UX-012).
SYS-171 — Graceful degradation under low energy (load shedding order)
Requirement Statement | As available energy depletes, the system shall shed discretionary functions in a defined priority order — location and non-essential uploads first, then upload cadence, then non-essential HMI — while preserving core logging, excursion detection/alerting, and the last-committed authoritative record to the last available energy.
Parent | PRD-REL-008; UN-PERF-003, UN-SAFE-010
Allocation Target | S3, S2, S5, S6
Type | Allocated
Interface refs | I-02, I-05, I-08
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (degradation order/thresholds unresolved — validated by test)
Rationale | A depleting device (heightened at −30 °C, D3) must protect the record over convenience; core logging/alerting is the last thing shed. The exact shed order and energy thresholds are Evidence-Pending; closed by the depletion test and coordinated with SYS-164 reserve. Power engineer to confirm per-load budget.
SYS-172 — Robustness over reuse (mechanical, seal/probe, sensor drift)
Requirement Statement | The system's sensing accuracy, seal/IP integrity, external probe, connectors, and user controls shall remain within specification over the specified reuse life, cleaning/decontamination cycles, and transport mechanical stresses, or the device shall detect and flag out-of-spec drift/degradation and block redeployment.
Parent | PRD-REL-009; UN-PERF-010, UN-PHY-006/008; D1 (reuse)
Allocation Target | S1 (enclosure/seal/probe), S4 (sensor), S3
Type | Allocated
Interface refs | I-01, I-04 (USB-C/connector), I-05
Safety/Security classification | Safety
Criticality | High
Verification Method | Test
Status | Evidence-Pending (life/aging drift budget unresolved — validated by test)
Rationale | Reuse economics (D1) and record credibility require durability beyond first-trip performance; seal wear, connector mating fatigue, probe strain-relief, and sensor drift are the wear mechanisms feeding the DFMEA (ERFKL, Stage 11). Drift/aging budget Evidence-Pending; closed by life/aging test per IEC 60068 stress and connector mating-cycle methods (revision to confirm). Cross-flag Mechanical engineer.
4. Reliability program, communication & platform durability
SYS-173 — Documented reliability program and prediction
Requirement Statement | The system shall be supported by a documented reliability program — DFMEA, reliability prediction/allocation, and life/environmental-stress testing — that quantifies and demonstrates the endurance, cycle-life, and availability targets and allocates a failure-rate budget across subsystems.
Parent | PRD-REL-010; UN-QUAL-###, UN-PERF-010
Allocation Target | SYS (program), S2, S3, S4, S5
Type | Derived
Interface refs | — (program-level; spans I-01…I-12)
Safety/Security classification | None
Criticality | Medium
Verification Method | Analysis
Status | Evidence-Pending (MTBF/DFR targets and prediction method unresolved)
Rationale | Reliability must be engineered and demonstrated, not assumed. MTBF/failure-rate allocation and environmental-stress (HALT/HASS) coverage are Evidence-Pending; the prediction method (handbook-based MIL-HDBK-217 / Telcordia SR-332 concepts, or physics-of-failure) and cell/component aging references (JEDEC, IEC 62133) are cited by designation, revision/method to confirm. Feeds ERFKL DFMEA (Stage 11).
SYS-174 — Communication robustness (retry/backoff, no data loss on link failure)
Requirement Statement | The system shall tolerate lossy and intermittent links using bounded retry with backoff and non-volatile buffering, such that transient network conditions never cause loss of a committed record and never produce a duplicate record in the cloud store.
Parent | PRD-REL-011, PRD-CONN-008; UN-ENV-007, UN-FUN-009; D2
Allocation Target | S3, S5, S8
Type | Allocated
Interface refs | I-08, I-10
Safety/Security classification | Both
Criticality | High
Verification Method | Test
Status | Baselined
Rationale | Global lanes have unreliable connectivity; robustness protects the record via ordered, idempotent, resumable upload (SYS-020). Judgment; closed by lossy-link test. Cross-flag Connectivity engineer on RF-in-container link margin (SYS-104-line).
SYS-175 — Cloud/platform reliability, data durability and recovery objectives
Requirement Statement | The cloud platform shall meet its defined availability target and shall not lose or corrupt any ingested record under any single-component failure, providing durable redundant storage, backup, and disaster recovery bounded by defined recovery-point (RPO) and recovery-time (RTO) objectives.
Parent | PRD-REL-012, PRD-CLOUD-012; UN-FUN-006/014, STN-013
Allocation Target | S8
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Both
Criticality | High
Verification Method | Test, Analysis
Status | Evidence-Pending (availability SLA, RPO/RTO targets unresolved)
Rationale | Alerts and records depend on a durable, available platform; a single-component failure must not destroy or lose an ingested record. Availability SLA and RPO/RTO are Evidence-Pending; closed by the platform availability/DR test. Cross-flag Cloud and Security engineers (durability vs. tamper-evidence, SYS-026).
SYS-176 — End-to-end data-delivery reliability over a trip (with redundancy/self-healing where warranted)
Requirement Statement | The system shall ensure that every committed sample, event, and excursion recorded on-device is eventually and correctly delivered to the authoritative cloud record over the course of a trip, reconciling device and cloud stores so that any undelivered or divergent record is detected, re-delivered, or flagged — with redundancy or self-healing applied where a single point of failure would otherwise threaten the record.
Parent | PRD-REL-005, PRD-REL-011, PRD-REL-012, PRD-FUN-013; D2
Allocation Target | SYS (S3, S5, S7, S8)
Type | Derived
Interface refs | I-06, I-08, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Reliability of the whole is more than the sum of links: the offline-first record (SYS-018), robust upload (SYS-174), and durable platform (SYS-175) must compose into a guaranteed end-to-end delivery, with device↔cloud reconciliation catching any silent gap. Redundancy/self-healing is applied by exception where an SPOF is identified in the DFMEA (SYS-173). Judgment; closed by end-to-end trip-reliability test.
System Requirements — Safety (SYS-149 … SYS-162)
SCM · Stage 06 · Owner: EM-ENG-001 with Safety/Risk and Battery engineers. Safety system requirements decomposed/allocated from PRD-SAFE (1300) and the safety-relevant portions of PRD-FUN, PRD-ELEC, PRD-REL, and PRD-REG. Field set per [index.md](index.md). Based on D2–D4, ISO 14971, IEC 62366-1, IEC 62368-1, and UN 38.3 / IATA DGR.
MODULE STATUS — PRELIMINARY (reconcile with HAL, Stage 10). Per the org risk process (CLAUDE.md rule
3; ISO 14971) safety/risk-control requirements originate in the Hazard Analysis Library (HAL, ORCH-017).
Every requirement below is a preliminary safety/risk-control system requirement decomposed from a
parent PRD-SAFE requirement and the concept; each will be traced to a HAL hazard record (HA-###) and one
or more Risk-Control Requirements (RCR-###), refined by the Preliminary and Detailed Hazard Analyses,
and re-baselined here. Action Priority is severity-first: severity is fixed by outcome (stated per
requirement); probability and detectability targets are Evidence-Pending until HAL. Standards are
cited by designation, revision to confirm. The two program safety tensions are reflected explicitly:
the D3 dead-in-cold / charge-at-low-temperature hazard (SYS-155, SYS-160) and the severity anchor
that both disposition-error directions harm the beneficiary (SYS-149).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
---
1. Correct disposition (the severity anchor)
SYS-149 — Do not mislead disposition (severity anchor)
Requirement Statement | The system shall not present, on-device or in the trip record, information that would lead a user to administer a temperature-compromised dose or to discard a potent one, within the validated performance of its detection, verdict, and record functions.
Parent | PRD-SAFE-001, PRD-FUN-006, PRD-FUN-010, PRD-FUN-011
Allocation Target | SYS (S3 detect/verdict, S6 present, S7 record, S8 authoritative record)
Type | Allocated
Interface refs | I-05, I-06, I-07, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (validation), Analysis
Status | Preliminary (reconcile with HAL)
Rationale | The SCM's dominant safety contribution is informational; both error directions harm the beneficiary (spoiled dose administered or potent dose wasted), so this is the severity-first tie-breaker for the whole safety module. Standard (ISO 14971 severity-dominant; STN-020; revision to confirm).
2. Detection & honesty as risk controls
SYS-150 — Reliable two-sided (freeze) detection as a risk control
Requirement Statement | The system shall detect and flag freezing (low-side) excursions with detection reliability at least equal to overheating (high-side) excursions, meeting the validated two-sided detection performance target for onset, direction, and duration.
Parent | PRD-SAFE-002, PRD-FUN-006, PRD-FUN-007
Allocation Target | S3, S4
Type | Allocated
Interface refs | I-05, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (validation)
Status | Preliminary (reconcile with HAL)
Rationale | Freeze damage to freeze-sensitive vaccines is silent and historically under-detected; a missed freeze leads to a spoiled dose administered — a top loss cause. Realises SYS-009 as a risk control. Cited (WHO/UNICEF freezing reporting; revision to confirm).
SYS-151 — No false reassurance on an incomplete record
Requirement Statement | When the trip record contains a monitoring gap, sensor fault, power-loss period, or interrupted armed interval, the system shall make that lapse visible in the record and in on-device status and shall never present an unqualified clean-pass / "accept" indication for that trip.
Parent | PRD-SAFE-003, PRD-FUN-017, PRD-REL-006, PRD-REL-007, PRD-UX-008
Allocation Target | SYS (S3 evaluate, S6 display, S7/S8 record)
Type | Allocated
Interface refs | I-06, I-07, I-10
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | A silent gap that reads as a clean pass produces a wrong disposition on an incomplete record; an honest gap is safer than false confidence. Realises SYS-017 as a risk control and ties the self-check (SYS-reliability). Judgment (ALCOA+ completeness underpins; revision to confirm).
SYS-152 — Alarm balance (no fatigue-induced miss)
Requirement Statement | The system shall tune excursion alerting — via configurable hysteresis, debounce, and consolidation — so that the false-alarm rate and the missed-excursion rate both meet the validated targets that avoid alarm fatigue muting a real, harmful excursion.
Parent | PRD-SAFE-004, PRD-FUN-007
Allocation Target | S3, S8
Type | Allocated
Interface refs | I-10
Safety/Security classification | Safety
Criticality | High
Verification Method | Test (validation)
Status | Preliminary (reconcile with HAL)
Rationale | Over-sensitivity causes fatigue and muted real alerts; under-sensitivity causes missed excursions — both directions are safety-relevant (concept C8). Realises SYS-014 as a risk control. Standard (alarm-management principles; revision to confirm).
3. Use-error safety (IEC 62366-1)
SYS-153 — Use-error protection for low-training users
Requirement Statement | The system shall be designed such that foreseeable use errors by minimally trained, gloved, offline last-mile users do not lead to a wrong disposition, meeting the summative usability acceptance criteria for the safety-critical user tasks.
Parent | PRD-SAFE-005, PRD-UX-001, PRD-UX-007, PRD-FUN-010
Allocation Target | SYS (S6 HMI, S3 logic, S9 mobile)
Type | Allocated
Interface refs | I-07, I-09
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (summative usability)
Parent traceability note | feeds the Usability Engineering File.
Status | Preliminary (reconcile with HAL)
Rationale | The edge user is the least-trained and most safety-critical operator; use-related risk must be controlled by design, not by training. Standard (IEC 62366-1 use-related risk; revision to confirm).
SYS-154 — Prevent unmonitored trip start (mis-arm)
Requirement Statement | The system shall not credit a trip as monitored unless it is armed with a valid profile, and shall make an un-armed, void, or mis-armed state unmistakable on-device so that a trip cannot silently start unmonitored.
Parent | PRD-SAFE-006, PRD-FUN-015, PRD-UX-004
Allocation Target | S3, S6
Type | Allocated
Interface refs | I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test, Demonstration
Status | Preliminary (reconcile with HAL)
Rationale | A void trip believed to be monitored yields an undetected excursion and a wrong disposition; mis-arm is a top hazard. Realises SYS-023 as a risk control and ties no-false-reassurance (SYS-151). Judgment.
4. Battery, charge & thermal safety
SYS-155 — Charge-inhibit below the safe cell temperature (no charge in unsafe cold)
Requirement Statement | The system shall inhibit battery charging whenever cell temperature is below the qualified low-temperature charge limit (nominally ≥0 °C, final limit per cell datasheet) and shall clearly indicate the charge-inhibited state to the user.
Parent | PRD-SAFE-007, PRD-ELEC-005, PRD-ELEC-006, PRD-REL-003
Allocation Target | S2, S3
Type | Allocated
Interface refs | I-02 (charge/power), I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (charge-inhibit)
Status | Preliminary (reconcile with HAL)
Rationale | Charging a cold Li-ion cell drives lithium plating, an internal-short and thermal-event precursor — a physical hazard directly created by the D3 low-temperature operating envelope. Standard (Li-ion charge-temperature limits, IEC 62368-1; final limit Evidence-Pending on cell datasheet; revision to confirm).
SYS-156 — No fire, burn, or thermal-runaway hazard in any condition
Requirement Statement | The system shall not present a fire, burn, or thermal-runaway hazard in normal use, foreseeable misuse, or single-fault conditions, and shall provide cell/BMS protection (over-charge, over-discharge, over-current, over-temperature, short-circuit) that holds the device in a safe state under those conditions.
Parent | PRD-SAFE-008, PRD-ELEC-007, PRD-ELEC-014
Allocation Target | S2, S1 (enclosure/thermal), S3
Type | Allocated
Interface refs | I-02, I-01 (thermal)
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (IEC 62368-1, abuse)
Parent traceability note | flag Battery and Electrical engineers; detailed BMS/protection in PRD-ELEC.
Status | Preliminary (reconcile with HAL)
Rationale | The device must be safe in aircraft holds, cold stores, clinics, and homes across all handlers; single-fault safety is required for a fielded reusable Li-ion product. Standard (IEC 62368-1, UN 38.3 abuse; revision to confirm).
SYS-157 — Safe and compliant lithium-battery transport
Requirement Statement | The system and its battery shall be safe and compliant for air and ground transport per UN 38.3 and IATA DGR, and the system shall carry the markings and be accompanied by the documentation required for lawful shipment.
Parent | PRD-SAFE-009, PRD-ELEC-015, PRD-REG-006
Allocation Target | S2, S1 (marking), SYS (documentation)
Type | Allocated
Interface refs | I-02
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (UN 38.3)
Status | Preliminary (reconcile with HAL)
Rationale | A reusable device shipped globally must not create a transport-safety hazard, and non-compliance excludes it from air/ground channels. Test summary Evidence-Pending. Standard (UN 38.3, IATA DGR; D4; revision to confirm).
5. Fail-safe, safe state & recovery
SYS-158 — Fail-safe on fault and recover without data loss
Requirement Statement | After any power, connectivity, firmware, reset, or OTA event, the system shall fail safe — preserving the logged trip data intact and returning to a defined known safe state — without corrupting the record and without silently ceasing to monitor.
Parent | PRD-SAFE-010, PRD-ARCH-008, PRD-REL-004, PRD-REL-005, PRD-SEC-008
Allocation Target | SYS (S3, S7, S5)
Type | Allocated
Interface refs | I-06, I-12 (OTA)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test (fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | A fault must never corrupt the evidentiary record or leave the device unsafe or silently unmonitored; fail-safe preserves both safety and the record (D2). Ties signed, resumable, fail-safe OTA (SYS-032). Judgment.
SYS-159 — Defined safe state and its entry conditions
Requirement Statement | The system shall define a safe state — one that preserves stored data, holds the battery/charge subsystem in a non-hazardous condition, and does not present a misleading monitoring indication — and shall enter that safe state upon detection of any fault condition from which continued normal operation would be unsafe or misleading.
Parent | PRD-SAFE-010, PRD-SAFE-008, PRD-REL-004
Allocation Target | SYS (S3 arbiter, S2 charge/BMS, S6 status)
Type | Derived
Interface refs | I-02, I-06, I-07
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (fault injection), Analysis
Status | Preliminary (reconcile with HAL)
Rationale | Fail-safe behaviour (SYS-158) requires an explicit, testable definition of the state the system fails into and the transitions that trigger it; an undefined safe state cannot be verified. This requirement makes the safe state a first-class, verifiable design target. Standard (ISO 14971 safe-state / fail-safe design; revision to confirm).
6. The D3 dead-in-cold hazard (informational failure)
SYS-160 — Dead-in-cold prevention (safety-relevant informational failure)
Requirement Statement | The system shall meet the 90-day cold-endurance and data-retention requirements across the −30 °C to +60 °C operating envelope such that the device does not die silently in deep cold and lose or truncate the trip record, and shall retain and, on power restoration, expose the record and any power-loss gap rather than resume as if monitoring were continuous.
Parent | PRD-SAFE-012, PRD-REL-001, PRD-ELEC-002
Allocation Target | SYS (S2 energy, S7 retention, S3 gap-marking)
Type | Allocated
Interface refs | I-02, I-06
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Test (90-day cold-endurance)
Status | Preliminary (reconcile with HAL)
Rationale | A device that dies silently in the cold produces a false-negative record — the missing interval reads as uneventful — which can force a wrong disposition; this reframes the D3 endurance tension as the disposition-safety anchor (SYS-149). Realises PRD-SAFE-012 and links SYS-power/reliability. Endurance numbers Evidence-Pending, closed by test. Judgment (D3; MR-RISK-001).
7. Residual risk, traceability & verification
SYS-161 — Residual-risk acceptability and risk-benefit
Requirement Statement | The system shall demonstrate, after all risk controls are applied, that each residual risk and the overall residual risk are acceptable per the ISO 14971 risk-management plan and the benefit-risk analysis, before design release.
Parent | PRD-SAFE-011, PRD-SAFE-001
Allocation Target | SYS
Type | Derived
Interface refs | —
Safety/Security classification | Safety
Criticality | Critical
Verification Method | Analysis
Status | Preliminary (reconcile with HAL)
Rationale | ISO 14971 requires evaluation of overall residual risk and benefit-risk, not just per-hazard control; this is the acceptance gate for the safety case. Closed in the HAL / risk-management file (Stage 10). Standard (ISO 14971 residual-risk / benefit-risk; revision to confirm).
SYS-162 — Hazard traceability and risk-control verification
Requirement Statement | The system shall maintain, for every safety requirement in this module, a bidirectional trace from its parent PRD-SAFE requirement to its HAL hazard record(s) (HA-###), its Risk-Control Requirement(s) (RCR-###), the design output(s) implementing each control, and the verification/validation evidence confirming each control's effectiveness.
Parent | PRD-SAFE-011, PRD-SAFE-001, PRD-SEC-001
Allocation Target | SYS (integrated by ORCH-018 / EM-ENG-012)
Type | Derived
Interface refs | —
Safety/Security classification | Safety
Criticality | High
Verification Method | Inspection, Analysis
Status | Preliminary (reconcile with HAL)
Rationale | A risk control that is not traced to a hazard and to verification evidence cannot be shown to reduce risk; complete, bidirectional traceability is the mechanism that makes the safety case auditable and reconcilable with the HAL. Standard (ISO 14971 risk-management file / traceability; revision to confirm).
---
Reconciliation note: on completion of the HAL (Stage 10), each SYS-149…SYS-162 record will carry its resolved HA-### / RCR-### identifiers, probability/detectability ratings, and severity-first Action Priority, and Status will migrate from Preliminary to Baselined where the hazard analysis confirms the control. Forward links: Subsystem/SRS (07), Architecture (08), Design Outputs (09), Verification (13) / Validation (14). Assembled end-to-end by ORCH-018 / EM-ENG-012.
System Requirements — Security & Data Integrity (SYS-132 … SYS-148)
SCM · Stage 06 · Owner: EM-ENG-008 (Product Security) with EM-ENG-001. Security and data-integrity system requirements decomposed/allocated from PRD-SEC (1100) and the security-relevant parts of PRD-CLOUD (900), based on D2 (store-and-forward evidence) and AD-5 (split trust boundary: login-free edge read vs platform-governed configuration/data). These requirements make the SCM's record trustworthy (integrity, attribution, audit) and its connectivity/identity/update controlled (identity, transport, keys, OTA), without defeating the frictionless offline edge read (concept C9). Field set per [index.md](index.md).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
Evidence & posture (honesty). Standards are cited by designation only — 21 CFR Part 11 / ALCOA+ (records, signatures, audit), GAMP 5 (CSV), IEC 62443 concepts (device/system security), OWASP ASVS (application verification), NIST CSF (identify/protect/detect/respond/recover), TLS 1.2+/1.3 — with revision to confirm; no clause number is asserted as fact. The formal threat model is Evidence-Pending and is set with the security persona (SYS-143); requirements whose numeric target, mechanism, or regional mapping is unresolved are marked Evidence-Pending. Severity-first criticality (CLAUDE.md rule 3): integrity, identity, and OTA controls that a wrong disposition or a bricked/impersonated fleet would defeat are Both, Critical.
---
1. Record integrity & attribution
SYS-132 — Cryptographic tamper-evidence of records end-to-end
Requirement Statement | The system shall protect every logged record (sample, event, excursion, gap marker, configuration, and lifecycle action) with a cryptographic integrity mechanism applied at the point of capture on the device and preserved unbroken through transport and cloud storage, such that any addition, deletion, reordering, or alteration of a stored record is computationally detectable.
Parent | PRD-SEC-001, PRD-CLOUD-002, PRD-FUN-017
Allocation Target | S3 (sign/hash at capture), S7 (protected store), S8 (verify/preserve), S5 (carry intact)
Type | Allocated
Interface refs | I-06 (MCU→storage), I-08 (device→cloud), I-10 (cloud API)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Rationale | A release/acceptance disposition and any audit rest on the record; a silent edit must be impossible to hide. Integrity begins at capture, not at ingest, so the last mile is never a trust gap (D2). Standard (21 CFR Part 11 / ALCOA+ enduring/accurate; hash-chain or digital-signature good practice; revision to confirm).
Status | Baselined
SYS-133 — Attribution and non-repudiation of records and actions
Requirement Statement | The system shall bind each record and each configuration, arming, calibration, firmware, and data action to an authenticated source — device identity, and where a human is the actor, the authenticated user and role — together with the trusted UTC timestamp (SYS-006), such that a record or action cannot later be repudiated or reassigned without detection.
Parent | PRD-SEC-002, PRD-CLOUD-001, PRD-FUN-017
Allocation Target | S3 (device identity), S8 (user identity/registry), S9 (actor at mobile), SYS
Type | Allocated
Interface refs | I-08, I-09 (BLE/mobile), I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Rationale | ALCOA+ attributable/original: evidence must be assignable to a source, and the source must not be able to deny it. Non-repudiation depends on device identity (SYS-134) and a trusted clock (SYS-006). Standard (21 CFR Part 11 attributable/signatures; revision to confirm).
Status | Baselined
SYS-134 — Unique protected device identity, hardware root of trust, and secure boot
Requirement Statement | Each device shall hold a unique, non-cloneable cryptographic identity and its private keys in a hardware root of trust (secure element or equivalent protected key store), and shall verify the authenticity and integrity of its firmware against that root of trust at every boot, refusing to execute unverified or altered firmware.
Parent | PRD-SEC-003, PRD-CLOUD-001
Allocation Target | S3 (MCU + secure element, boot ROM)
Type | Allocated
Interface refs | I-05 (sensors→MCU internal trust domain), I-08 (identity used on link)
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis
Rationale | Identity is the anchor for attribution (SYS-133), mutual TLS (SYS-137), and signed OTA (SYS-139); a software-only key can be extracted or cloned. Secure boot prevents a tampered image from running and silently corrupting the record. Standard (IEC 62443 concepts; hardware-root-of-trust / secure-boot good practice; revision to confirm). Threat model Evidence-Pending (SYS-143).
Status | Baselined
SYS-135 — Append-only, non-editable audit trail (device and cloud)
Requirement Statement | The system shall maintain an append-only audit trail on both the device and the cloud recording configuration changes, access/authentication events, arming, calibration, firmware, key, and data actions, where each entry is attributable, UTC-time-stamped, and cannot be modified or deleted without detection, and shall reconcile the device and cloud audit trails after upload.
Parent | PRD-SEC-004, PRD-CLOUD-007, PRD-FUN-017
Allocation Target | S3, S7 (device trail), S8 (cloud trail), SYS (reconciliation)
Type | Allocated
Interface refs | I-06, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Inspection
Rationale | The governance of the record must itself be inspectable and defensible; a device-only trail is lost if a unit is lost, and a cloud-only trail misses offline actions (D2). Consistent with SYS-028. Standard (21 CFR Part 11 audit trail; revision to confirm).
Status | Baselined
2. Access, transport & at-rest protection
SYS-136 — Least-privilege role-based access control with enterprise SSO federation
Requirement Statement | The cloud platform and mobile app shall enforce least-privilege role-based access control on every configuration, data-access, and administrative function, and shall support federation to an enterprise identity provider via single sign-on (e.g. SAML/OIDC-class), while enforcing that no access-control mechanism gates the login-free on-device edge read (SYS-148).
Parent | PRD-SEC-005, PRD-CLOUD-009
Allocation Target | S8 (RBAC/SSO), S9 (app enforcement)
Type | Allocated
Interface refs | I-10, I-11 (cloud→enterprise)
Safety/Security classification | Security
Criticality | High
Verification Method | Test
Rationale | Enterprise/regulated deployment is gated by IT security and CSV (STN-013); only authorised roles may configure devices or read/modify data. SSO avoids a parallel credential store. Must not defeat AD-5 frictionless edge read. Standard (21 CFR Part 11 access limitation; OWASP ASVS access control; revision to confirm).
Status | Baselined
SYS-137 — Mutually-authenticated encrypted transport for data in transit
Requirement Statement | The system shall protect all device-to-cloud and app-to-cloud communication with authenticated, encrypted transport (TLS 1.2 or later, or equivalent), mutually authenticating the device using its hardware-rooted identity (SYS-134) and the server using a validated certificate, and shall reject unauthenticated, downgraded, or spoofed connections.
Parent | PRD-SEC-006, PRD-CLOUD-001, PRD-CONN-006
Allocation Target | S5 (device TLS), S3 (identity/keys), S8 (server TLS/registry), S9 (app TLS)
Type | Allocated
Interface refs | I-08, I-09, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Rationale | Records and alerts must not be intercepted, altered, or spoofed in transit; mutual authentication ensures only registered, authenticated devices deliver data (PRD-CLOUD-001) and that the device only talks to the genuine platform. Standard (TLS 1.2+/1.3; IEC 62443 concepts; revision to confirm). Cipher suite selection in design.
Status | Baselined
SYS-138 — Encryption of sensitive data at rest (device NVM and cloud)
Requirement Statement | The system shall encrypt controlled data at rest — the device non-volatile log and credential/key store, and the cloud data store — using keys managed under SYS-140, such that physical recovery of a device or compromise of a storage medium does not expose record contents or credentials in cleartext.
Parent | PRD-SEC-007, PRD-CLOUD-011
Allocation Target | S7 (device NVM), S3 (key store), S8 (cloud store)
Type | Allocated
Interface refs | I-06, I-10
Safety/Security classification | Security
Criticality | High
Verification Method | Test
Rationale | A reusable device travels through uncontrolled hands and may be lost or stolen; stored records and keys must not be exposed if a device or store is compromised. At-rest encryption complements tamper-evidence (SYS-132), which addresses integrity rather than confidentiality. Standard (NIST-style crypto good practice; ISO/IEC 27001 concepts; revision to confirm).
Status | Baselined
3. Trustworthy update & key lifecycle
SYS-139 — Signed, resumable, fail-safe firmware update with rollback protection
Requirement Statement | The system shall accept and apply only firmware whose authenticity and integrity are cryptographically verified against the hardware root of trust (SYS-134), shall resume an interrupted update without data loss, shall reject downgrade to a firmware version below an anti-rollback minimum, and shall never, as a result of an update, place a fielded device in a bricked state or in an unmonitored state without notice.
Parent | PRD-SEC-008, PRD-CLOUD-008, PRD-FUN-022, PRD-SAFE-010
Allocation Target | S3 (verify/apply/rollback), S5 (transfer), S8 (OTA orchestration/staged rollout)
Type | Allocated
Interface refs | I-12 (OTA), I-08
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test
Rationale | An update must not be malicious, corrupt a device, brick the fleet, or silently regress it to a vulnerable version. Anti-rollback closes the re-flash-to-known-vulnerable attack; fail-safety and staged cloud rollout protect the fleet and the record (AD-7). Consistent with SYS-032. Standard (secure-update principles; IEC 62443 concepts; revision to confirm). Reconcile with HAL.
Status | Baselined
SYS-140 — Cryptographic key and credential lifecycle management
Requirement Statement | The system shall generate, provision, store, rotate, and revoke all cryptographic keys and credentials — device identity, transport, signing/verification, and at-rest encryption keys — through a defined, protected lifecycle spanning device and platform, such that a compromised or expired key can be rotated or revoked fleet-wide without loss of the historical record's integrity or attribution.
Parent | PRD-SEC-009, PRD-SEC-003, PRD-SEC-006, PRD-SEC-008
Allocation Target | S3 (device key store), S8 (platform PKI/registry), SYS
Type | Allocated
Interface refs | I-08, I-10, I-12
Safety/Security classification | Security
Criticality | Critical
Verification Method | Analysis, Test
Rationale | Integrity (SYS-132), identity (SYS-134), transport (SYS-137), OTA (SYS-139), and at-rest encryption (SYS-138) all collapse to key management; without defined rotation/revocation, a single compromise is unrecoverable. Standard (IEC 62443 concepts; NIST-style key-management good practice; revision to confirm). Lifecycle process Evidence-Pending (design/CSV).
Status | Evidence-Pending (key-lifecycle process to confirm)
SYS-141 — Provisioning of identity and keys (secure factory and field)
Requirement Statement | The system shall provision each device's unique identity and initial keys into the hardware root of trust through a controlled, authenticated, audit-logged process at manufacture, and shall support authenticated re-provisioning or credential rotation in the field (e.g. on refurbishment or re-deployment between trips) without exposing private key material in cleartext at any point.
Parent | PRD-SEC-009, PRD-SEC-003, PRD-CLOUD-001
Allocation Target | S3 (device), S8 (registry enrolment), SYS (factory/field process)
Type | Derived
Interface refs | I-08, I-09, I-10
Safety/Security classification | Security
Criticality | High
Verification Method | Analysis, Inspection
Rationale | Identity and attribution are only as trustworthy as the provisioning that created them; the reusable/refurbish loop (D1) requires field re-provisioning, and a leaky factory process would undermine the entire trust chain. Feeds device registry enrolment (PRD-CLOUD-001). Standard (IEC 62443 secure provisioning concepts; revision to confirm). Threat model Evidence-Pending (SYS-143).
Status | Evidence-Pending (provisioning process to confirm)
4. Hardening, monitoring & assurance
SYS-142 — Security monitoring, incident logging and response
Requirement Statement | The cloud platform shall detect, log, and alert on security-relevant events (e.g. authentication failures, unauthorised-device data attempts, anomalous access, failed integrity or signature checks) and shall support a defined incident-response process, retaining security event records per policy.
Parent | PRD-SEC-011, PRD-CLOUD-007, PRD-CLOUD-012
Allocation Target | S8, SYS
Type | Allocated
Interface refs | I-10, I-11
Safety/Security classification | Security
Criticality | High
Verification Method | Demonstration, Test
Rationale | A connected, OTA-updatable fleet needs detection and response, not prevention alone; an undetected compromise silently corrupts the evidentiary product. Standard (NIST CSF detect/respond; IEC 62443 concepts; revision to confirm). Incident-response runbook Evidence-Pending.
Status | Evidence-Pending (IR runbook to confirm)
SYS-143 — Security-by-design threat model and vulnerability management
Requirement Statement | The system shall be developed against a documented, maintained threat model whose identified risks are traced to security requirements and controls, and shall be governed by an ongoing vulnerability-management process (monitoring, triage, and remediation via SYS-139 update) across device, cloud, and mobile for the product's fielded life.
Parent | PRD-SEC-012, PRD-SEC-008, PRD-SEC-011
Allocation Target | SYS (all tiers)
Type | Allocated
Interface refs | I-08, I-10, I-12
Safety/Security classification | Both
Criticality | High
Verification Method | Analysis
Rationale | Security requirements must be driven by an explicit, reviewed threat model rather than assumed; a fielded connected device faces vulnerabilities discovered after release that must be tracked and patched. Reconciles with HAL for safety-relevant security. Standard (IEC 62443 concepts; NIST CSF identify; OWASP ASVS; revision to confirm). Threat model Evidence-Pending.
Status | Evidence-Pending (threat model to confirm)
SYS-144 — Production debug-port lockdown (JTAG/SWD)
Requirement Statement | The system shall disable, lock, or authenticate all hardware debug and programming interfaces (e.g. JTAG/SWD) on production devices, such that a fielded device cannot expose keys, memory, or firmware, or be re-flashed, through a debug port without cryptographic authorisation.
Parent | PRD-SEC-003, PRD-SEC-007, PRD-SEC-009
Allocation Target | S3 (MCU/secure element fuses/config), S1 (physical access barrier)
Type | Derived
Interface refs | I-05
Safety/Security classification | Security
Criticality | High
Verification Method | Test, Inspection
Rationale | An open debug port is the most direct path to key extraction, memory dump, and firmware replacement, defeating secure boot (SYS-134), at-rest encryption (SYS-138), and provisioning (SYS-141). Physical access is realistic for a device that travels in uncontrolled hands. Standard (IEC 62443 device-hardening concepts; revision to confirm). Threat model Evidence-Pending (SYS-143).
Status | Preliminary (reconcile with threat model)
SYS-145 — Supply-chain and software bill-of-materials (SBOM) integrity
Requirement Statement | The system shall maintain a software bill of materials for device firmware, cloud, and mobile components, shall verify the authenticity and integrity of third-party and build artefacts in the build/release pipeline, and shall enable identification of affected releases when a component vulnerability is disclosed.
Parent | PRD-SEC-008, PRD-SEC-012, PRD-SEC-011
Allocation Target | SYS (device FW, cloud, mobile build/release), S8
Type | Derived
Interface refs | I-12
Safety/Security classification | Security
Criticality | Medium
Verification Method | Inspection, Analysis
Rationale | A signed update (SYS-139) only guarantees the shipped image, not that its components are free of known vulnerabilities or tampering upstream; an SBOM makes vulnerability management (SYS-143) actionable when a dependency CVE is disclosed. Standard (NIST CSF supply-chain / SBOM good practice; IEC 62443 concepts; revision to confirm).
Status | Preliminary (reconcile with vulnerability management)
SYS-146 — Session management and authentication for cloud and mobile
Requirement Statement | The cloud platform and mobile app shall authenticate users, protect authenticated sessions and tokens against theft, fixation, and replay, enforce session expiry and re-authentication for sensitive actions, and support account lockout/rate-limiting against credential attacks — without imposing any authentication on the login-free on-device edge read (SYS-148).
Parent | PRD-CLOUD-009, PRD-SEC-005, PRD-SEC-006
Allocation Target | S8, S9
Type | Derived
Interface refs | I-09, I-10
Safety/Security classification | Security
Criticality | High
Verification Method | Test
Rationale | RBAC/SSO (SYS-136) is only as strong as the session that carries the authenticated identity; stolen or replayed sessions bypass access control and enable record tampering or misconfiguration. Bounded to the platform side of the trust split (AD-5). Standard (OWASP ASVS session management; revision to confirm).
Status | Baselined
SYS-147 — Split trust boundary — login-free edge read vs platform-governed configuration/data
Requirement Statement | The system shall implement two distinct trust domains: (a) an on-device edge domain in which the accept/reject/quarantine verdict and trip status can be read on the device display with no network, phone, login, or credential; and (b) a platform-governed domain in which all configuration, data access, export, administration, and record modification require authenticated, role-authorised, and audited access — such that the frictionless edge read cannot alter, delete, or reconfigure any governed data or the record.
Parent | PRD-SEC-005, PRD-CLOUD-009, PRD-SEC-001
Allocation Target | S3, S6 (edge read), S8, S9 (governed domain), SYS (boundary)
Type | Derived
Interface refs | I-07 (HMI), I-09, I-10
Safety/Security classification | Both
Criticality | Critical
Verification Method | Test, Analysis, Demonstration
Rationale | AD-5 / concept C9: the last-mile user has no phone, network, or training, so the verdict must be readable with zero friction, yet configuration and the authoritative record must remain governed — the boundary is a security requirement, not a UX convenience. The edge read must be strictly read-only with respect to governed state. Judgment, based on AD-5 and the read-only nature of tamper-evident records (SYS-132).
Status | Baselined
SYS-148 — Privacy and data-minimisation of location and personal data
Requirement Statement | The system shall limit collection, transmission, and retention of location and any personal data to the minimum needed for the operational purpose (asset recovery, attribution, and cold-chain evidence), shall apply retention/deletion only through authorised, audited action, and shall align handling with applicable privacy law for the markets of use (D4).
Parent | PRD-SEC-010, PRD-CLOUD-011, PRD-REG-009
Allocation Target | S8 (retention/deletion), S3/S5 (collection at source), SYS
Type | Allocated
Interface refs | I-08, I-10
Safety/Security classification | Security
Criticality | Medium
Verification Method | Inspection
Rationale | The device monitors cargo, not people, but location traces and operator identities can create disproportionate privacy exposure; data-minimisation bounds that exposure and supports customer data-ownership/retention control (PRD-CLOUD-011). Standard (privacy-by-design / data-minimisation; EU GDP data governance; revision to confirm). Regional privacy-law mapping Evidence-Pending (D4).
Status | Evidence-Pending (regional privacy mapping to confirm)
Decomposition & Allocation Matrix — System Requirements (SCM)
SCM · Stage 06 · Owner: EM-ENG-001. Traces the 220 product requirements (Stage 05) down to the 190 system requirements (SYS-001…190), allocates each SYS to the subsystem(s) that realise it, and forwards each SYS to the Stage-07 artifact families (HWR/MER/FWR/IFR/SRS-<CODE>) and Stage-08 architecture elements (FUN/CMP/INT). Assembled into the end-to-end matrix by ORCH-018 / EM-ENG-012 (Stage 18).
Subsystem legend: S1 Enclosure/Mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile.
---
1. PRD module → SYS category coverage (backward roll-up)
Every PRD module is decomposed into at least one SYS category. No PRD module is dropped.
PRD module (Stage 05) | Reqs | Decomposed into SYS categories | Primary SYS IDs
PRD-000 Product Vision (VIS) | 7 | Functional, Safety, Regulatory, Reliability | SYS-009, SYS-015, SYS-149, SYS-177, intent set
PRD-100 Stakeholders (STK) | 6 | Functional, UX-in-Functional, Data, Security | SYS-015, SYS-022, SYS-030, SYS-031, SYS-137
PRD-200 Use Environment (ENV) | 12 | Environmental, Connectivity, Interface | SYS-069…SYS-085, SYS-104, SYS-109, SYS-117
PRD-300 Product Architecture (ARCH) | 10 | Functional, Interface, Power, Connectivity, Data, Safety | SYS-006, SYS-015, SYS-019, SYS-051…068, SYS-095, SYS-159
PRD-400 Functional (FUN) | 22 | Functional (primary), Performance, Data, Interface | SYS-001…SYS-032, SYS-039, SYS-119…131
PRD-500 User Experience (UX) | 12 | Functional (verdict/arm/config), Environmental (legibility), Performance (responsiveness) | SYS-015, SYS-017, SYS-023, SYS-025, SYS-071, SYS-047
PRD-600 Mechanical (MECH) | 12 | Environmental, Interface, Power (port) | SYS-069…085, SYS-051, SYS-052, SYS-053, SYS-063
PRD-700 Electrical (ELEC) | 16 | Power/Energy (primary), Interface, Reliability, Safety, Regulatory | SYS-086…SYS-103, SYS-053, SYS-165, SYS-155, SYS-181
PRD-800 Connectivity (CONN) | 12 | Connectivity (primary), Interface, Data | SYS-104…SYS-118, SYS-058…061, SYS-119
PRD-900 Cloud (CLOUD) | 12 | Functional (routing/record), Data, Security, Reliability, Regulatory | SYS-012, SYS-026…028, SYS-119…131, SYS-132…148, SYS-175
PRD-1000 Mobile (MOB) | 8 | Functional (config/arm), Interface (BLE), Security | SYS-022, SYS-024, SYS-060, SYS-136, SYS-146
PRD-1100 Cybersecurity (SEC) | 12 | Security (primary), Data, Functional (audit) | SYS-132…SYS-148, SYS-119…131, SYS-028
PRD-1200 Reliability (REL) | 12 | Reliability (primary), Power, Data | SYS-163…SYS-176, SYS-086…103, SYS-123-line
PRD-1300 Safety (SAFE) | 12 | Safety (primary), Functional, Power | SYS-149…SYS-162, SYS-009, SYS-017, SYS-091
PRD-1400 Regulatory (REG) | 12 | Regulatory (primary), Data, Security | SYS-177…SYS-190, SYS-124, SYS-142
PRD-1500 Manufacturing (MFG) | 8 | forwarded to Stage 15 + Security (provisioning) + Data (serialization) | SYS-136-line, SYS-121, forward DO/PROC
PRD-1600 Serviceability (SVC) | 8 | Functional (readiness), Reliability, Power (recharge) | SYS-024, SYS-030, SYS-171, SYS-102-line
PRD-1700 Performance (PERF) | 12 | Performance (primary) | SYS-033…SYS-050
PRD-1800 Quality (QUAL) | 7 | forwarded to Stage 18 + Reliability + Regulatory | SYS-172, SYS-188, forward REL/QUAL
PRD-1900 Verification (VER) | 8 | governs all SYS Verification Method fields; RF & D3 test hooks | SYS-045, SYS-048, SYS-104, SYS-163
Note: PRD-MFG and PRD-QUAL are principally realised in later stages (15 manufacturing-design-transfer, 18 quality/traceability); their system-level hooks (secure provisioning, serialization, readiness, durability budgets) appear as SYS requirements above and their forward placeholders are carried by ORCH-018.
---
2. SYS category → subsystem allocation summary
SYS category | ID range | S1 | S2 | S3 | S4 | S5 | S6 | S7 | S8 | S9
--- | --- | :-: | :-: | :-: | :-: | :-: | :-: | :-: | :-: | :-:
Functional | 001–032 | ○ |  | ● | ● | ○ | ● | ● | ● | ○
Performance | 033–050 |  | ○ | ● | ● | ○ | ○ | ○ | ●
Interface | 051–068 | ● | ● | ● | ● | ● | ● | ● | ● | ●
Environmental | 069–085 | ● | ○ | ○ | ○ | ● | ●
Power/Energy | 086–103 | ○ | ● | ● |  | ○ | ○
Connectivity | 104–118 | ○ |  | ● | ○ | ● |  |  | ● | ○
Data | 119–131 |  |  | ● |  |  |  | ● | ● | ○
Security | 132–148 |  |  | ● |  | ○ |  | ● | ● | ●
Safety | 149–162 | ○ | ● | ● | ● |  | ● | ● | ○
Reliability | 163–176 | ○ | ● | ● | ○ | ● |  | ● | ●
Regulatory | 177–190 | ○ | ● | ○ | ○ | ● | ○ | ○ | ● | ○
● primary allocation · ○ contributing allocation.
---
3. Subsystem → owned SYS requirements (forward to Stage-07 artifact family)
Subsystem | Representative SYS IDs | Stage-07 detailed-requirement family
S1 Enclosure/Mechanical | SYS-051, 052, 053, 063, 069–085 (mech-env), 156 | MER-### (mechanical) + IFR (connector/probe)
S2 Power | SYS-086–103, 155, 165–167, 181 | HWR-### (power/battery/BMS) + FWR (charge-inhibit)
S3 Compute/MCU+FW | SYS-001–032 (logic), 133, 134, 159, 168, 169 | FWR-### + SRS-DEV + HWR (MCU/SE)
S4 Sensing | SYS-001, 003, 004, 005, 007, 029, 033–038 | HWR-### (sensors/probe/GNSS) + FWR (drivers/cal)
S5 Connectivity | SYS-011, 019, 104–118, 148 | HWR-### (radios/antenna) + SRS-CONN + IFR (wireless)
S6 HMI | SYS-015, 017, 023, 047, 071, 156 | HWR-### (OLED/LED/button) + MER (cutouts) + FWR (UI)
S7 Storage | SYS-002, 013, 018, 021, 127, 170 | HWR-### (NVM) + SRS-DEV/DATA (log/format)
S8 Cloud | SYS-012, 026–028, 031, 119–131, 132–148, 175 | SRS-CLOUD / SRS-DATA / SRS-SEC / SRS-FLEET
S9 Mobile | SYS-022, 024, 060, 136, 146 | SRS-MOB + IFR (BLE provisioning)
---
4. SYS → software library allocation (SRS module map)
SYS theme | SRS module(s) that will detail it
Boot/init/power-state/scheduling/config/storage/calibration/state-machine | SRS-DEV (200)
Bearers, retry, offline, provisioning, cert handling, time sync | SRS-CONN (300)
Registry, ingest, routing, record assembly, RBAC/SSO, APIs, DR | SRS-CLOUD (400)
Pairing, config, arming aid, local read, notifications, offline sync | SRS-MOB (500)
Secure boot, keys, TLS, at-rest, tamper, audit, threat mgmt | SRS-SEC (600)
Packaging, download, verify, rollback, campaigns, resume | SRS-OTA (700)
Enrollment, health, heartbeat, inventory, decommission | SRS-FLEET (800)
Schema, serialization, compression, retention, export, validation | SRS-DATA (900)
Boot/latency/throughput/battery/energy performance budgets | SRS-PERF (1000)
Power-fail/brownout/watchdog/availability/MTBF/degradation | SRS-REL (1100)
Safe state, alarm mgmt, honest-gap, fail-safe (HAL-traced) | SRS-SAFE (1200)
WHO PQS/GDP/Part 11/CSV/radio/EMC/transport records | SRS-REG (1300)
POST/BIST/health/crash/service diagnostics | SRS-DIAG (1400)
Verification matrix, acceptance criteria, test coverage | SRS-VER (1500)
System/logical/physical context & interfaces | SRS-CTX (100) / product identity SRS-OVR (000)
---
5. Critical-path & risk-carrying requirements (severity-first)
Risk / driver | Carrying SYS requirements | Status | Closed by
D3 — 90-day endurance at −30 °C | SYS-086, SYS-087, SYS-095, SYS-045, SYS-046, SYS-163–165 | Evidence-Pending | PVL D3 endurance test (PRD-VER-004)
D3 — low-temperature charge safety | SYS-091 (charge-inhibit <0 °C), SYS-092, SYS-155, SYS-165 | Evidence-Pending / Preliminary | Charge-safety test; HAL reconciliation
D3 — dead-in-cold false-negative | SYS-017, SYS-160-line, SYS-171 (graceful degrade) | Preliminary | HAL + endurance validation
RF-in-container link | SYS-104, SYS-109, SYS-117, SYS-080-line, SYS-058 | Evidence-Pending | RF-in-container test (PRD-VER-005)
Missed-freeze hazard | SYS-003, SYS-009, SYS-149-line | Preliminary | HAL + detection validation
Mis-arm (unmonitored trip) | SYS-023, SYS-024, SYS-153-line | Preliminary | HAL + use validation
Delayed/missing alert | SYS-011, SYS-012, SYS-013, SYS-039 | Baselined / Evidence-Pending | Latency validation (connectivity-bound)
Record integrity / disposition error | SYS-016, SYS-026, SYS-028, SYS-149 | Baselined / Preliminary | HAL + integrity verification
---
6. Traceability completeness statement
Backward | all 220 PRD requirements are covered by ≥1 SYS requirement or forwarded to a later stage (PRD-MFG/QUAL system hooks noted in §1). No PRD requirement is silently dropped.
Forward | all 190 SYS requirements carry a subsystem allocation (§2/§3) and a Stage-07 artifact-family destination (§3/§4); safety SYS (149–162) additionally forward to HAL HA-###/RCR-### (Stage 10).
Verification | every SYS requirement carries a Verification Method; risk-carrying SYS (§5) are bound to named PVL/PVaL tests. The complete bidirectional matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18).

## 7 Detailed Requirements
Detailed Requirements
The detailed, part-by-part requirements for the hardware, software, and connections.
↩ Contents
In this section:
• Subsystem & Detailed Requirements + Software Requirements Library — SCM
• Subsystem/Detailed Requirements — Firmware (FWR-001 … FWR-055)
• Subsystem / Detailed Requirements — Hardware (HWR-001 … HWR-050)
• Subsystem/Detailed Requirements — Interfaces (IFR-001 … IFR-035)
• Subsystem / Detailed Requirements — Mechanical (MER-001 … MER-040)
• SRS-000 — Product Overview (module SRS-OVR)
• SRS-100 — System Context (module SRS-CTX)
• SRS-1000 — Performance (module SRS-PERF)
• SRS-1100 — Reliability (module SRS-REL)
• SRS-1200 — Safety (module SRS-SAFE)
• SRS-1300 — Regulatory (module SRS-REG)
• SRS-1400 — Diagnostics (module SRS-DIAG)
• SRS-1500 — Verification (module SRS-VER)
• SRS-200 — Device Software (module SRS-DEV)
• SRS-300 — Connectivity (module SRS-CONN)
• SRS-400 — Cloud Software (module SRS-CLOUD)
• SRS-500 — Mobile Application (module SRS-MOB)
• SRS-600 — Cybersecurity (module SRS-SEC)
• SRS-700 — OTA Update (module SRS-OTA)
• SRS-800 — Fleet Management (module SRS-FLEET)
• SRS-900 — Data (module SRS-DATA)
Subsystem & Detailed Requirements + Software Requirements Library — SCM
Project: smart-cold-chain-monitor-lifecycle · Stage: 07 — Subsystem/Detailed Requirements · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · SRS Library owner: Embedded Software Engineer (EM-ENG-007) · SRS-SEC owner: Product Security Engineer (EM-ENG-008) · SRS-REG/VER owner: Quality/Regulatory Engineer (EM-ENG-006) · Status: Authored — instantiated, product-specific.
This stage decomposes the 190 System Requirements (SYS-###, Stage 06) into detailed hardware, mechanical, firmware, and interface requirements and the full Software Requirements Library (SRS-<CODE>-NNN). It is the design-input baseline the Design Outputs (Stage 09, ORCH-013) are built against and verified to. Based on 00-concept/concept-and-decisions.md (D1–D4).
---
1. Detailed subsystem requirements (HW / ME / FW / IF)
Each subsystem requirement carries: ID · Requirement Statement (single "shall", testable) · Parent (SYS-…/PRD-…) · Allocation (subsystem + component) · Type (Derived/Allocated) · Criticality · Safety/Security · Verification Method · Status · Rationale (evidence-backed, labelled Standard/Cited/Judgment/Evidence-Pending).
Family | File | Scope | ID range | Reqs
HWR Hardware/Electrical | [HWR-hardware.md](HWR-hardware.md) | electrical, PCB, power, battery/BMS, sensors, connectors, radios/antenna, MCU, NVM, HMI HW | HWR-001…050 | 50
MER Mechanical | [MER-mechanical.md](MER-mechanical.md) | enclosure, thermal design, IP/sealing, mounting/retention, HMI mechanicals, materials/cleaning | MER-001…040 | 40
FWR Firmware | [FWR-firmware.md](FWR-firmware.md) | boot, power-state governor, sensing/calibration, detection engine, verdict/HMI, storage, comms agent, OTA agent, security/fault | FWR-001…055 | 55
IFR Interfaces | [IFR-interfaces.md](IFR-interfaces.md) | connector/electrical, power, wireless RF, BLE-GATT provisioning, device↔cloud protocol, cloud↔enterprise API, OTA, ICD governance | IFR-001…035 | 35
|  |  | Subtotal | 180
---
2. Software Requirements Library (SRS-<CODE>-NNN)
Per 00-governance/software-requirements-standard.md: a modular library, one file per module, each a standalone spec (header → §7 Functional → §8 Non-functional → Error Handling / Diagnostics / Security / Verification / Traceability). Each requirement carries the mandated field set: ID · Title · Requirement Statement ("shall") · Rationale · Priority (Must/Should/Could) · Source · Verification Method · Traceability (SYS-…/PRD-…/risk/standard) · Notes.
Module | Code (num) | File | Reqs | Applicability
Product Overview | OVR (000) | [SRS-000-product-overview.md](SRS-000-product-overview.md) | 7 | Applicable
System Context | CTX (100) | [SRS-100-system-context.md](SRS-100-system-context.md) | 9 | Applicable
Device Software | DEV (200) | [SRS-200-device-software.md](SRS-200-device-software.md) | 31 | Applicable
Connectivity | CONN (300) | [SRS-300-connectivity.md](SRS-300-connectivity.md) | 26 | Applicable
Cloud Software | CLOUD (400) | [SRS-400-cloud.md](SRS-400-cloud.md) | 26 | Applicable
Mobile Application | MOB (500) | [SRS-500-mobile.md](SRS-500-mobile.md) | 17 | Applicable
Cybersecurity | SEC (600) | [SRS-600-cybersecurity.md](SRS-600-cybersecurity.md) | 30 | Applicable
OTA Update | OTA (700) | [SRS-700-ota.md](SRS-700-ota.md) | 19 | Applicable
Fleet Management | FLEET (800) | [SRS-800-fleet.md](SRS-800-fleet.md) | 17 | Applicable
Data | DATA (900) | [SRS-900-data.md](SRS-900-data.md) | 22 | Applicable
Performance | PERF (1000) | [SRS-1000-performance.md](SRS-1000-performance.md) | 19 | Applicable
Reliability | REL (1100) | [SRS-1100-reliability.md](SRS-1100-reliability.md) | 19 | Applicable
Safety | SAFE (1200) | [SRS-1200-safety.md](SRS-1200-safety.md) | 15 | Applicable (PRELIMINARY — HAL-traced)
Regulatory | REG (1300) | [SRS-1300-regulatory.md](SRS-1300-regulatory.md) | 15 | Applicable
Diagnostics | DIAG (1400) | [SRS-1400-diagnostics.md](SRS-1400-diagnostics.md) | 16 | Applicable
Verification | VER (1500) | [SRS-1500-verification.md](SRS-1500-verification.md) | 20 | Applicable
|  | 16 modules | 308 | 16 / 16 applicable
Grand total (subsystem + SRS): 488 detailed requirements.
---
3. Applicability Register (SRS)
Per the standard, scope is never silently dropped. All 16 default SRS modules are Applicable — the SCM is a device + connectivity + cloud + mobile, regulated, reusable, connected, safety-significant IoT product, so every module carries real requirements. No modules were excluded. Borderline / preliminary flags:
Module | Status | Note
SRS-SAFE (1200) | Preliminary | Software risk controls originate in the Hazard Analysis Library (HAL, Stage 10, ORCH-017); each SRS-SAFE requirement traces to a HA-###/RCR-### and is re-baselined after the PHA/DHA.
SRS-REG (1300) | Applicable (Evidence-Pending items) | Exact WHO PQS device category and US FDA classification/pathway are Evidence-Pending (concept §7); standards cited by designation, revision to confirm — no clause numbers asserted as fact.
— extension modules | Not instantiated | No AI/ML, audio/video, or payments scope for the SCM; not applicable.
---
4. Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, every requirement's rationale is labelled Standard, Cited, Judgment, or Evidence-Pending; no standard clause number is asserted as fact and no unresolved quantity is invented. The D3 −30 °C / 90-day battery tension is carried as explicit Evidence-Pending, validate-by-test requirements in HWR (cell/BMS/charge-inhibit), FWR (energy governor), SRS-PERF, SRS-REL, and SRS-SAFE. The RF-in-container link is Evidence-Pending across HWR (antenna), IFR (RF interfaces), and SRS-CONN, closed by the RF-in-container test (PRD-VER-005). Storage sizing (SYS-127) is Evidence-Pending across FWR, SRS-DEV, and SRS-DATA.
---
5. Traceability (chain position)
Backward: every HWR/MER/FWR/IFR and SRS requirement names ≥1 Parent SYS-### and/or PRD-<CODE>-NNN, tracing to User Needs and D1–D4. Forward: each seeds Architecture elements (FUN-###, CMP-###, INT-###, Stage 08), Design Outputs (DO-<CODE>-NNN, Stage 09), and Verification (PV-###) / Validation (VAL-###); SRS-SAFE additionally forwards to HAL HA-###/RCR-###. The complete bidirectional matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18); see the Stage-06 [decomposition-allocation-matrix](../06-system-requirements/decomposition-allocation-matrix.md).
Flow: PRD (05) → SYS (06) → Subsystem/SRS (07, this stage) → Architecture (08) → Design Outputs (09) → Hazard Analysis (10) / ERFKL (11) → Risk-Control Requirements (12) → Verification (13) / Validation (14). Outputs feed Design Outputs (Stage 09) and ORCH-018 / EM-ENG-012.
Convening: ORCH-012 (Software Requirements) for the SRS library; EM-ENG-001 integrates HW/ME/FW/IF. Owners: EM-ENG-007 (SRS library), EM-ENG-008 (SRS-SEC), EM-ENG-006 (SRS-REG/VER).
Subsystem/Detailed Requirements — Firmware (FWR-001 … FWR-055)
SCM · Stage 07 · Subsystem: S3 Compute/MCU + Firmware · Author: Principal Embedded Firmware Engineer, integrating with EM-ENG-001 (systems), Power, Sensing, Connectivity, Security, and Safety engineers.
Scope & context. This file specifies the firmware that runs on the SCM's MCU. Firmware is the SCM's energy governor and its capture → detect → decide brain: it schedules and duty-cycles every load to make the 90-day/−30 °C endurance promise achievable (D3), it captures and calibrates every sample, it detects two-sided excursions, it computes the offline on-device verdict, and it commits the tamper-evident record — all offline-first so the last mile is never a black box (D2). Each requirement decomposes one or more system requirements (SYS-###) and/or product requirements (PRD-<CODE>-NNN) and allocates them to a named firmware module.
Field set (each requirement, in order): Requirement Statement (single testable "shall") · Parent · Allocation (S3 + module) · Type (Derived/Allocated) · Criticality (Critical/High/Medium/Low, severity-first per CLAUDE.md rule 3) · Safety/Security (Safety/Security/Both/None) · Verification Method (Test/Analysis/Inspection/Demonstration) · Status (Baselined/Preliminary/Evidence-Pending) · Rationale (evidence-backed, labelled Standard / Cited / Judgment / Evidence-Pending).
Modules (S3 partition): scheduler · power-mgr · sensor-driver · detect-engine · storage · comms · HMI · OTA-agent · security. Evidence posture (honesty): software-lifecycle, coding-standard, and secure-update standards are cited by designation, revision to confirm (IEC 62304 software lifecycle; CERT-C / MISRA-C coding; secure-update / IEC 62443 concepts; ISO 14971 for risk-relevant behaviour). No numeric current/energy/timing budget is invented; where a subsystem allocation is not yet closed by the power budget or an RF-in-container test it is marked Evidence-Pending and validated by test. Every requirement below traces to at least one SYS/PRD parent. Firmware safety classification (IEC 62304 Class A/B/C) is to be assigned on completion of the HAL (Stage 10) — noted per requirement as "software-safety class TBD".
---
1. Boot & initialisation
FWR-001 — Secure-boot image authentication before execution
Requirement Statement | On every reset the firmware bootloader shall cryptographically verify the authenticity and integrity of the application image (signature over the image hash, chained to a root-of-trust key held in the secure element) before transferring control to it, and shall refuse to execute an image that fails verification.
Parent | SYS-032, PRD-FUN-022, PRD-SEC-008
Allocation | S3 · security (bootloader/OTA-agent)
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Preliminary (reconcile with HAL; key hierarchy per SYS-security)
Rationale | An unverified image can silently subvert detection, verdict, or record integrity across the fleet. A verified chain of trust from immutable boot ROM is the foundation of the OTA-integrity control (SYS-032). Standard (secure-boot / secure-update principles, IEC 62443 concepts; revision to confirm).
FWR-002 — Power-on self-test (POST) of safety-critical resources
Requirement Statement | During initialisation, before crediting the device as ready to arm, the firmware shall execute a power-on self-test covering MCU core/RAM integrity, non-volatile storage accessibility, each sensor channel, the RTC, and the secure element, and shall record the pass/fail result of each check as a diagnostic entry.
Parent | SYS-024, SYS-030, PRD-FUN-020, PRD-REL-004
Allocation | S3 · scheduler (init) with sensor-driver, storage, security
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | A device armed with a failed sensor or unreadable log produces a worthless or misleading record; POST turns silent hardware faults into an explicit, logged readiness result feeding pre-arm checks. Judgment (IEC 62304 defensive initialisation; revision to confirm).
FWR-003 — RTC initialisation, validity and monotonicity check
Requirement Statement | On boot the firmware shall initialise the real-time clock, verify that its value is within a plausible bounded window and has not regressed relative to the last committed record timestamp, and shall set and log a "time-unvalidated" status if the clock is implausible or has not been synchronised to a trusted time source.
Parent | SYS-006, PRD-FUN-019, PRD-SEC-001
Allocation | S3 · scheduler (time services)
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | Every sample, excursion, and audit entry is only defensible with a trusted, monotonic time base; a regressed or default clock after a cold power-loss must be disclosed, never assumed correct. Standard (21 CFR Part 11 attributable/accurate records; revision to confirm).
FWR-004 — Configuration load and structural validation
Requirement Statement | On boot the firmware shall load the persisted trip profile and device configuration from non-volatile storage, verify each item against a stored integrity check (CRC/hash) and against structural constraints (e.g. low threshold < high threshold, logging interval within 1–60 min, cadence within bounds), and shall reject or quarantine a corrupt or inconsistent configuration rather than applying it.
Parent | SYS-025, PRD-FUN-014, PRD-SAFE-005
Allocation | S3 · scheduler (config manager)
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | A corrupted or invalid configuration silently applied could invert thresholds and cause a missed excursion (use-error → wrong disposition). Validating before use is the firmware realisation of the configuration-integrity control. Standard (IEC 62366-1 use-error prevention; revision to confirm).
FWR-005 — Safe-mode / degraded boot on init failure
Requirement Statement | If any safety-critical initialisation step (secure-boot fallback, storage mount, RTC, or configuration load) fails, the firmware shall enter a defined degraded safe-mode that preserves any committed logged data, refuses to credit the device as armed/monitoring, presents an unambiguous fault indication, and does not present a normal ready or "monitoring" state.
Parent | SYS-167, SYS-159, PRD-SAFE-010, PRD-REL-004
Allocation | S3 · scheduler (fault arbiter) with HMI, storage
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test (fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | A boot failure must fail into a known, non-misleading state — never a partial run that reads as monitoring (no false reassurance). Realises the defined-safe-state requirement in firmware. Judgment (ISO 14971 fail-safe; software-safety class TBD, revision to confirm).
FWR-006 — Boot cause capture and recovery attribution
Requirement Statement | On each boot the firmware shall determine and record the reset cause (power-on, brown-out, watchdog, software reset, OTA) as a timestamped audit entry, and shall flag any interval between the last committed record and resumed monitoring as a potential monitoring gap for downstream evaluation.
Parent | SYS-169, SYS-017, PRD-REL-006, PRD-SAFE-003
Allocation | S3 · scheduler (init) with storage
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | A device that reboots (watchdog, brown-out, cold power-loss) must never assume it monitored continuously across the outage; the logged reset cause feeds the honest-gap and self-check logic. Judgment.
2. Power-state machine (energy governor — D3)
FWR-007 — Deterministic power-state machine with defined modes
Requirement Statement | The firmware shall implement a deterministic power-state machine with the defined states {deep-sleep, idle, logging, alert, charging, OTA, fault, shipping} and shall permit only the documented transitions between them, such that at any instant the device is in exactly one known, budgeted state.
Parent | SYS-089, SYS-090, PRD-ELEC-008, PRD-ARCH-004
Allocation | S3 · power-mgr (governor) with scheduler
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | A single, deterministic governor is the mechanism that makes 90-day cold endurance achievable (D3) and auditable; undefined or overlapping states defeat the energy budget. Judgment; quantified by the current-profile test.
FWR-008 — Lowest-viable-sleep default between scheduled activities
Requirement Statement | Between scheduled activities the firmware shall hold the MCU and peripherals in their lowest viable sleep state consistent with retaining RTC, wake logic, log integrity, and battery protection, and shall not remain in any higher-power state longer than required to complete the scheduled task.
Parent | SYS-089, SYS-092, PRD-ELEC-008
Allocation | S3 · power-mgr
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | With most of a 90-day trip spent asleep, quiescent time dominates the budget; defaulting to deep-sleep is the primary endurance lever (D3). Judgment; sleep-current allocation is Evidence-Pending (SYS-092), closed by cold current-profile test.
FWR-009 — Hard radio and GNSS gating outside scheduled windows
Requirement Statement | The firmware shall keep the cellular, Wi-Fi, BLE, and GNSS subsystems powered down or in their lowest state outside their scheduled, energy-budgeted windows, and shall not initiate any radio or GNSS activity outside those windows except to transmit a queued excursion alert per FWR-044.
Parent | SYS-089, PRD-ELEC-008, PRD-FUN-013
Allocation | S3 · power-mgr with comms
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Radios and GNSS are the dominant loads; ungated radio activity would consume the endurance budget silently. The single exception (immediate excursion alert) is the only safety-justified override. Judgment; radio/GNSS energy is Evidence-Pending (SYS-090), closed by RF-in-container test.
FWR-010 — Duty-cycle scheduling of sensing and connectivity
Requirement Statement | The firmware scheduler shall drive sampling at the configured logging interval (1–60 min) and connectivity attempts at the configured cadence (default hourly) from the RTC/timer, waking the device only for the scheduled task and returning it to sleep on completion, with schedule slew bounded so that timing drift does not create or hide a logging interval.
Parent | SYS-001, SYS-019, PRD-FUN-001, PRD-FUN-013, D2
Allocation | S3 · scheduler with power-mgr
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Duty-cycled wake-work-sleep is the D2/D3 enabler: it delivers gap-free logging while bounding energy. Judgment.
FWR-011 — Defined wake sources and their handling
Requirement Statement | The firmware shall enable and service a defined set of wake sources — RTC/timer (scheduled sample/upload), accelerometer shock interrupt, charger-insertion detect, and user-button — and shall route each wake to its handler and return to the lowest viable sleep state without leaving any wake source spuriously re-triggering.
Parent | SYS-005, SYS-089, PRD-FUN-004, PRD-ELEC-008
Allocation | S3 · power-mgr with sensor-driver, HMI
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Wake sources must be exhaustively defined and debounced; an unmasked or chattering wake (e.g. vibration) would drain the budget and defeat endurance. Judgment (CERT-C/MISRA interrupt-handling discipline; revision to confirm).
FWR-012 — Per-mode current budget conformance
Requirement Statement | The firmware shall be designed such that the measured device current in each power state does not exceed the per-state allocation of the SYS-090 energy budget across the −30 °C to +60 °C range, with the sum of duty-weighted state consumption plus reserve remaining within usable capacity.
Parent | SYS-090, SYS-092, PRD-ELEC-008, PRD-PERF-005
Allocation | S3 · power-mgr
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test, Analysis
Status | Evidence-Pending (per-state current allocations unresolved pending power budget & RF-in-container test)
Rationale | The 90-day promise is only designable as a per-state budget the firmware must hold to; radio/GNSS/sleep allocations are not yet closed. Evidence-Pending (D3); closed by cold current-profile test, coordinated with Power engineer.
FWR-013 — Charging-mode behaviour and charge-inhibit enforcement in firmware
Requirement Statement | In charging state the firmware shall read cell temperature and shall command the charge path to remain inhibited whenever cell temperature is below the safe low-temperature charge limit (nominally ≥0 °C, final limit per cell datasheet), shall not command or resume charging while inhibited, and shall reflect the charge-inhibited condition to the HMI and readiness reporting.
Parent | SYS-094, SYS-155, SYS-166, PRD-ELEC-005, PRD-SAFE-007
Allocation | S3 · power-mgr with sensor-driver, HMI
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (charge-inhibit)
Status | Preliminary (reconcile with HAL; threshold per cell datasheet)
Rationale | Charging a cold Li-ion cell risks lithium plating and a thermal event — a physical hazard, not a performance loss; firmware is a redundant enforcement layer over the hardware BMS interlock and must never defeat it. Severity-first Critical. Standard (IEC 62368-1 / IEC 62133 charge-temperature limits; revision to confirm); reconcile with HAL (Stage 10). Cross-flag Power and Safety engineers.
Note | Firmware charge-inhibit is a defence-in-depth layer; the primary interlock is a hardware BMS function owned by S2 (HWR) — firmware shall not be the sole means of preventing cold charge.
FWR-014 — Shipping / long-storage mode
Requirement Statement | The firmware shall provide a shipping/storage mode entered by an explicit command that minimises self-discharge and quiescent drain (radios and non-essential peripherals off, no scheduled sampling), preserves configuration and any prior record, and requires an explicit wake/arm action to resume monitoring, such that a stored or in-transit unarmed device is never mistaken for a monitoring device.
Parent | SYS-092, SYS-154, PRD-ELEC-008, PRD-SAFE-006
Allocation | S3 · power-mgr with scheduler
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | A reusable fleet spends time in depot and transit; a low-drain shipping mode protects shelf-life charge (D1) while the explicit-wake requirement prevents an unarmed device appearing armed (mis-arm hazard). Judgment.
3. Sensing & calibration
FWR-015 — Multi-channel sampling acquisition per schedule
Requirement Statement | On each scheduled sampling wake the firmware shall acquire internal temperature, external-probe temperature, relative humidity, and 3-axis acceleration, assemble them into a single sample record stamped with the trusted time and the active profile identifier, and commit the record before returning to sleep.
Parent | SYS-001, SYS-003, PRD-FUN-001, PRD-FUN-002, PRD-FUN-003
Allocation | S3 · sensor-driver with scheduler, storage
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | The gap-free interval history is the primary evidentiary product (D2); committing before sleep ensures no acquired sample is lost to a subsequent power event. Judgment.
FWR-016 — Sensor driver abstraction with per-channel fault handling
Requirement Statement | The firmware shall access each sensor through a driver that detects communication faults, timeouts, and out-of-range raw responses, shall retry within a bounded budget, and on persistent failure shall record the channel as faulted (with a data-quality flag) without stalling the sampling schedule or discarding other channels' valid data.
Parent | SYS-007, SYS-170, PRD-FUN-017, PRD-REL-004
Allocation | S3 · sensor-driver
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (fault injection)
Status | Baselined
Rationale | A hung sensor bus must not halt logging or crash the scheduler; bounded retry plus explicit fault-flagging preserves the rest of the record and honest status. Judgment (CERT-C robust-I/O and bounded-loop discipline; revision to confirm).
FWR-017 — Dual-channel temperature independence
Requirement Statement | The firmware shall keep the internal-sensor and external-probe temperature channels independently identifiable end-to-end (acquisition, calibration, flagging, storage) and shall not substitute, average, or conflate one channel for the other in any record or verdict computation.
Parent | SYS-003, PRD-FUN-002
Allocation | S3 · sensor-driver with storage, detect-engine
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | The probe represents the product zone and underpins freeze detection; conflating channels would corrupt disposition. Judgment.
FWR-018 — Calibration-coefficient application and provenance
Requirement Statement | The firmware shall apply the stored per-device calibration coefficients to each raw sensor reading before storing the reported value, shall verify the coefficient set's integrity and validity/expiry on load, and shall flag readings and block arming when calibration is missing, corrupt, or expired.
Parent | SYS-024, SYS-030, PRD-FUN-020, PRD-REL-009
Allocation | S3 · sensor-driver with scheduler (config)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | An out-of-calibration device must not be deployed (D1 reuse loop); applying and gating on calibration is how the firmware protects measurement accuracy and readiness. Standard (ISO/IEC 17025 calibration-reporting concepts; revision to confirm).
FWR-019 — Plausibility and data-quality flagging without raw-value loss
Requirement Statement | The firmware shall evaluate each reading against physically plausible bounds and rate-of-change limits and shall attach a data-quality flag (plausible / out-of-range / implausible / faulted-channel / uncalibrated) to the stored sample, without discarding or overwriting the raw acquired value.
Parent | SYS-007, PRD-FUN-017, PRD-SAFE-003
Allocation | S3 · sensor-driver with detect-engine
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | A silently-wrong reading is more dangerous than a flagged one; retaining the raw value keeps the record honest and auditable (no false reassurance). Judgment.
FWR-020 — Shock-capture interrupt handler with configurable threshold
Requirement Statement | The firmware shall configure the accelerometer to generate an interrupt when acceleration exceeds the per-profile configurable shock threshold, and its interrupt handler shall capture the event with trusted timestamp and peak magnitude, debounce routine handling below threshold, and defer heavy processing to task context so the ISR remains bounded and deterministic.
Parent | SYS-005, PRD-FUN-004, PRD-ENV-006
Allocation | S3 · sensor-driver (ISR) with detect-engine, storage
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test
Status | Baselined
Rationale | Product-relevant mishandling must be captured without alarm-fatigue noise (C8); a bounded, short ISR that defers work is required for deterministic real-time behaviour and energy discipline. Standard (CERT-C / MISRA-C ISR-discipline: no blocking, bounded work; revision to confirm).
FWR-021 — Sampling continuity and interval accounting
Requirement Statement | The firmware shall track the expected schedule of samples against those actually committed and shall, on detecting a missed or late scheduled sample, insert an explicit gap marker rather than allow a scheduled interval to be silently absent from the log.
Parent | SYS-002, PRD-FUN-001, PRD-SAFE-003
Allocation | S3 · scheduler with storage
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | ALCOA+ completeness: a missing sample must never masquerade as in-range data; the firmware is the component that knows a scheduled sample did not occur. Standard (ALCOA+, EU GDP; revision to confirm).
4. Detection engine
FWR-022 — Two-sided threshold evaluation with equal-reliability freeze detection
Requirement Statement | The detection engine shall evaluate each committed temperature sample against the independent low (freeze) and high (heat) thresholds of the active profile using the identical algorithm and data path for both directions, such that a freezing excursion is detected with reliability equal to an overheating excursion.
Parent | SYS-008, SYS-009, SYS-150, PRD-FUN-005, PRD-FUN-006
Allocation | S3 · detect-engine
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (validation)
Status | Preliminary (reconcile with HAL)
Rationale | Freeze damage is silent and historically under-detected; symmetric low/high logic prevents the low side becoming a lesser capability. Realises SYS-009/SYS-150. Cited (WHO/UNICEF freezing reporting; revision to confirm). Software-safety class TBD.
FWR-023 — Hysteresis and debounce on threshold crossings
Requirement Statement | The detection engine shall apply configurable hysteresis bands and debounce (minimum consecutive samples/duration) to threshold crossings so that a transient single-sample fluctuation around a threshold does not by itself declare or clear an excursion.
Parent | SYS-010, SYS-014, PRD-FUN-007, PRD-SAFE-004
Allocation | S3 · detect-engine
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Momentary noise must not create false excursions (fatigue) nor should chatter clear a real one; hysteresis/debounce is the tunable balance mechanism. Judgment (alarm-management principles; revision to confirm).
FWR-024 — Time-over/under-threshold and cumulative-exposure logic
Requirement Statement | The detection engine shall determine an excursion using configurable product-stability logic — threshold combined with time-over-threshold (high) / time-under-threshold (low) and/or cumulative time-temperature exposure — rather than instantaneous transients alone, using the active profile's parameters.
Parent | SYS-010, PRD-FUN-007, PRD-SAFE-004
Allocation | S3 · detect-engine
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Alerts and verdicts must reflect actual product harm (stability-budget), not momentary noise (C8); cumulative logic is how the SCM maps environment to product risk. Judgment.
FWR-025 — Excursion event generation and characterisation
Requirement Statement | On declaring an excursion the detection engine shall generate an excursion event record capturing direction (freeze/heat/humidity/shock as applicable), onset time, duration, peak/extreme value, and the profile parameters that triggered it, and shall commit that event to non-volatile storage at the time of detection independent of connectivity.
Parent | SYS-009, SYS-013, PRD-FUN-006, PRD-FUN-008
Allocation | S3 · detect-engine with storage
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Offline-first: no excursion may depend on a live link to exist in the record (D2); committing at detection makes the event durable. Judgment.
FWR-026 — Configurable alarm-balance parameters
Requirement Statement | The detection engine shall expose the hysteresis, debounce, time-over/under, and cumulative-exposure parameters and an alert-consolidation window as validated per-profile configuration items, such that alerting sensitivity can be tuned to meet the false-alarm and missed-excursion targets without code change.
Parent | SYS-014, SYS-152, PRD-SAFE-004
Allocation | S3 · detect-engine with scheduler (config)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Both over-sensitivity (fatigue) and under-sensitivity (missed excursion) are safety-relevant; parameterising the balance lets it be validated and reconciled with the HAL without new firmware. Standard (alarm-management principles; revision to confirm).
FWR-027 — Immediate alert hand-off on excursion
Requirement Statement | On committing an excursion event the detection engine shall enqueue a high-priority alert to the comms agent for immediate transmission when connectivity is available, subject only to the FWR-009 exception, and shall not gate that hand-off on the normal cadence schedule.
Parent | SYS-011, PRD-FUN-008, PRD-CONN-009
Allocation | S3 · detect-engine with comms
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Learning of a break in time to intervene converts monitoring into prevented loss (D2); the excursion alert is the one radio activity permitted outside the scheduled window. Judgment.
5. Verdict & HMI
FWR-028 — Offline on-device accept/reject/quarantine computation
Requirement Statement | The firmware shall compute, entirely on-device and without any network, phone, login, or manual, an accept / reject / quarantine verdict for the trip by evaluating the committed trip record against the active profile using the same excursion evaluation as the detection engine.
Parent | SYS-015, SYS-016, PRD-FUN-010, PRD-FUN-011
Allocation | S3 · detect-engine (verdict) with HMI
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test, Demonstration
Status | Baselined
Rationale | The last-mile user has no phone, network, or training; the edge verdict is the decision point, and it must be derived from the same logic as the cloud record so the two never disagree (SYS-016). Judgment. Software-safety class TBD.
FWR-029 — Unambiguous no-login verdict display
Requirement Statement | The firmware shall present the verdict on the device's own display as an unambiguous, legible indication (distinct states for accept / reject / quarantine / incomplete) that a minimally trained, gloved user can read with a single on-device action and without authentication.
Parent | SYS-015, SYS-153, PRD-UX-001, PRD-UX-002
Allocation | S3 · HMI with detect-engine
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (summative usability), Demonstration
Status | Preliminary (reconcile with HAL)
Rationale | Use-error by the least-trained operator must be controlled by design; an ambiguous display is a wrong-disposition hazard. Standard (IEC 62366-1 use-related risk; revision to confirm).
FWR-030 — Honest handling of incomplete monitoring in the verdict
Requirement Statement | When the trip record contains gap markers, faulted-channel flags, watchdog/reset gaps, or an interrupted armed interval, the firmware shall not display an unqualified "accept" verdict and shall indicate that monitoring was incomplete.
Parent | SYS-017, SYS-151, PRD-SAFE-003, PRD-UX-008
Allocation | S3 · detect-engine (verdict) with HMI
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | No false reassurance — an incomplete record must never read as a clean pass; the honest-gap path is the firmware realisation of SYS-151. Judgment.
FWR-031 — Arming state machine with explicit arm/void states
Requirement Statement | The firmware shall implement an arming state machine with defined states {unarmed, arming, armed, voided} requiring an explicit user arm action to transition to armed, shall give positive on-device confirmation of the armed state, and shall make an unarmed, arming-incomplete, or voided state unmistakable so a trip cannot silently start unmonitored.
Parent | SYS-023, SYS-154, PRD-FUN-015, PRD-SAFE-006
Allocation | S3 · scheduler (arming) with HMI
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test, Demonstration
Status | Preliminary (reconcile with HAL)
Rationale | Mis-arm (unmonitored trip believed monitored) is a top hazard; arming must be explicit and unmissable. Realises SYS-023/SYS-154. Judgment.
FWR-032 — Pre-arm readiness gate
Requirement Statement | Before permitting transition to the armed state the firmware shall verify the pre-arm readiness pre-conditions — adequate state of charge, valid and unexpired calibration, valid firmware, valid profile, probe present, POST passed — and shall block arming and guide the user to correct any failed pre-condition.
Parent | SYS-024, SYS-030, PRD-FUN-016, PRD-FUN-020
Allocation | S3 · scheduler (arming) with sensor-driver, power-mgr, HMI
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | A device armed while un-ready produces a worthless or misleading record; the readiness gate converts POST/calibration/charge status into a hard arming pre-condition. Judgment.
FWR-033 — Debounced button handling and input safety
Requirement Statement | The firmware shall debounce user-button inputs, distinguish the defined short/long/multi-press gestures for their assigned actions (wake, show verdict/status, arm confirmation), and shall require an unambiguous deliberate action for any state-changing gesture so that incidental contact during handling cannot arm, void, or reset the device.
Parent | SYS-023, SYS-153, PRD-UX-004, PRD-UX-007
Allocation | S3 · HMI
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | Gloved, offline users interact through a minimal control set; accidental input must not cause a safety-relevant state change (use-error protection). Standard (IEC 62366-1; revision to confirm).
6. Storage
FWR-034 — Append-only, tamper-evident log structure
Requirement Statement | The firmware shall store samples, events, excursions, gaps, configuration, and lifecycle actions in an append-only log in which each committed entry is chained by an integrity value (e.g. running hash / MAC over the prior entry) so that any alteration, reordering, or deletion of a committed entry is detectable.
Parent | SYS-026, SYS-028, PRD-FUN-017, PRD-SEC-002
Allocation | S3 · storage with security
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test, Inspection
Status | Baselined
Rationale | The auditable record is the product's compliance value; append-only chaining provides the attributable, tamper-evident property of a regulated record. Standard (21 CFR Part 11 / ALCOA+; revision to confirm).
FWR-035 — Power-fail-safe atomic writes
Requirement Statement | The firmware shall commit each log entry atomically (e.g. journaled/two-phase write with completion marker) such that any committed entry is preserved intact across power loss, brown-out, reset, and firmware update, and no partial or torn record survives as valid.
Parent | SYS-099, SYS-168, PRD-REL-005, PRD-FUN-017
Allocation | S3 · storage with power-mgr
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test (power-cycle integrity)
Status | Baselined
Rationale | Low-temperature brown-out under voltage sag (D3) makes power-fail-safe writes essential; a torn write must fail the entry, not silently corrupt committed data. Judgment; closed by power-cycle storage-integrity test. Cross-flag Power engineer (brown-out threshold, PRD-ELEC-011).
FWR-036 — Wear-levelling and flash endurance management
Requirement Statement | The firmware shall distribute non-volatile write and erase operations (wear-levelling) and monitor block health so that the log medium meets the write endurance required for the worst-case logging rate over the device's reuse life, and shall flag storage as end-of-life before write failures can compromise the record.
Parent | SYS-018, SYS-172, PRD-REL-005, PRD-REL-009
Allocation | S3 · storage
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Analysis, Test
Status | Evidence-Pending (endurance budget vs medium spec unresolved)
Rationale | A reusable device (D1) accumulates far more write cycles than a single-trip logger; wear-levelling protects endurance and record integrity over life. Evidence-Pending; closed by endurance analysis against the selected medium's datasheet.
FWR-037 — Buffer management with capacity for a full trip and prioritised retention
Requirement Statement | The firmware shall buffer the complete sample and event log for a full 90-day trip at the minimum logging interval with defined margin without overwriting unsent data, and as capacity is approached shall prioritise retention of excursion and audit records, warn of impending capacity limits, and never silently discard excursion or audit data.
Parent | SYS-018, SYS-021, PRD-FUN-012, PRD-REL-008
Allocation | S3 · storage with comms
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Analysis, Test
Status | Evidence-Pending (storage sizing/margin to confirm — SYS-127)
Rationale | Store-and-forward requires capacity for the worst-case offline duration (D2), and degradation must preserve the most safety-relevant data first. Judgment on sizing; Evidence-Pending, closed by storage-sizing analysis.
FWR-038 — Explicit gap markers in the log
Requirement Statement | The firmware shall represent every unrecorded scheduled interval, power-loss period, faulted-channel span, and watchdog/reset recovery as an explicit, timestamped gap marker in the log, distinguishable from valid in-range data.
Parent | SYS-002, SYS-017, PRD-FUN-017, PRD-SAFE-003
Allocation | S3 · storage with scheduler, detect-engine
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | A hidden gap that reads as in-range data forces a wrong disposition; explicit markers make the record honest and feed the verdict (FWR-030). Standard (ALCOA+ completeness; revision to confirm).
FWR-039 — Controlled factory reset with record protection
Requirement Statement | The firmware shall perform factory reset / record erasure only under an authenticated, authorised command, shall record the erase action as an attributable audit entry in the new log, and shall not permit erasure of the committed record of an armed or active trip through any unauthorised path.
Parent | SYS-028, PRD-SEC-004, PRD-SEC-002
Allocation | S3 · storage with security
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | Uncontrolled erasure would destroy the evidentiary record and enable tampering; reset must be authorised and itself auditable (non-repudiation). Standard (21 CFR Part 11 audit trail; revision to confirm). Cross-flag Security engineer.
7. Comms agent (store-and-forward)
FWR-040 — Persistent store-and-forward upload queue
Requirement Statement | The comms agent shall maintain a persistent upload queue over the non-volatile log such that records survive reset and power loss, remain queued until acknowledged by the cloud, and are never removed from the device on the basis of transmission alone (only on confirmed durable ingest).
Parent | SYS-018, SYS-020, PRD-FUN-012, PRD-FUN-013
Allocation | S3 · comms with storage
Type | Allocated
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | Offline-first (D2): a record must not be considered delivered until durably ingested; transmit-and-forget would lose data on a dropped link. Judgment.
FWR-041 — Cadence upload with multi-bearer selection and fallback
Requirement Statement | On each scheduled upload window the comms agent shall select the best available bearer among LTE-M (primary), Wi-Fi, and BLE offload, attempt upload of queued records, and fall back to the next bearer on failure, without requiring any single bearer to be present.
Parent | SYS-019, PRD-FUN-013, PRD-CONN-004
Allocation | S3 · comms with power-mgr
Type | Allocated
Criticality | High
Safety/Security | None
Verification Method | Test
Status | Baselined
Rationale | Opportunistic multi-bearer upload maximises delivery under intermittent coverage while the power-mgr keeps radios gated (D2/AD-1). Judgment; per-bearer energy Evidence-Pending (SYS-090).
FWR-042 — Ordered, resumable, idempotent upload
Requirement Statement | The comms agent shall upload queued records in intact time order, resume an interrupted upload from the last acknowledged record without data loss or reordering, and tag each record with a stable unique identifier so that the cloud stores each record exactly once (idempotent ingest).
Parent | SYS-020, SYS-174, PRD-CONN-008, PRD-REL-011
Allocation | S3 · comms with storage
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test (lossy-link)
Status | Baselined
Rationale | Integrity across intermittent links requires ordering, resume, and de-duplication; transient network conditions must never lose or duplicate a committed record. Standard (ALCOA+ enduring/consistent; revision to confirm).
FWR-043 — Alert prioritisation over routine upload
Requirement Statement | The comms agent shall transmit a queued excursion alert ahead of routine cadence data and, when connectivity is available at detection time, shall attempt immediate delivery of the alert within the SYS-039 latency target rather than waiting for the next cadence window.
Parent | SYS-011, PRD-FUN-008, PRD-CONN-009
Allocation | S3 · comms with detect-engine
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | An excursion alert has intervention value only if timely; it must jump the routine queue. Realises SYS-011 hand-off (FWR-027) on the link. Judgment.
FWR-044 — Bounded retry with exponential back-off
Requirement Statement | On upload or connection failure the comms agent shall retry using bounded exponential back-off with a cap, shall not retry so aggressively as to breach the energy budget, and shall return the device to sleep between attempts, preserving queued data across all failed attempts.
Parent | SYS-174, SYS-089, PRD-REL-011, PRD-ELEC-008
Allocation | S3 · comms with power-mgr
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | Unbounded retry in poor coverage would drain the 90-day budget (D3); bounded back-off balances delivery against energy while never dropping queued records. Judgment.
FWR-045 — Trusted time synchronisation and drift disclosure
Requirement Statement | During connected windows the comms agent shall synchronise the RTC to a trusted network time source, shall record the synchronisation event and the pre-sync offset as an audit entry, and shall not retroactively rewrite the timestamps of already-committed records.
Parent | SYS-006, PRD-FUN-019, PRD-ARCH-007
Allocation | S3 · comms with scheduler (time services)
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | A trusted clock with bounded, disclosed drift underpins defensible timestamps; silently rewriting committed timestamps would break tamper-evidence. Standard (21 CFR Part 11 attributable/accurate; revision to confirm).
8. OTA agent
FWR-046 — Signature and integrity verification of update images
Requirement Statement | The OTA agent shall verify the cryptographic signature (against the secure-element root of trust) and integrity hash of a firmware update image, and shall verify version/rollback-protection metadata, before the image is activated, rejecting any image that fails any check.
Parent | SYS-032, PRD-FUN-022, PRD-SEC-008
Allocation | S3 · OTA-agent with security
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | Only authentic, integral firmware may run on a fielded device; signature + rollback protection prevents malicious or downgraded images subverting the fleet. Standard (secure-update principles, IEC 62443 concepts; revision to confirm).
FWR-047 — Resumable download without disturbing monitoring
Requirement Statement | The OTA agent shall download update images in resumable segments that survive link interruption and reset, shall perform the download within the energy and connectivity budget without suspending scheduled sensing/logging, and shall verify the completed image before activation.
Parent | SYS-032, PRD-FUN-022, PRD-CONN-008
Allocation | S3 · OTA-agent with comms, scheduler
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | Updates must be genuine, resumable, and non-destructive to the record; an update must never create an unmonitored interval on an armed device. Judgment.
FWR-048 — A/B (dual-bank) activation with automatic rollback
Requirement Statement | The OTA agent shall install an update into an inactive bank and activate it via a fail-safe mechanism (e.g. A/B with confirmation), such that a failed or non-booting update automatically rolls back to the last known-good image and the device is never left bricked or in a non-monitoring state without notice.
Parent | SYS-032, SYS-158, PRD-FUN-022, PRD-SAFE-010
Allocation | S3 · OTA-agent with security
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test (interrupted-update / fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | A bricked or silently-non-monitoring fielded device loses the record and the fleet; dual-bank + rollback is the fail-safe realisation of SYS-032/SYS-158. Standard (secure-update / fail-safe update concepts; revision to confirm).
FWR-049 — Data and configuration preservation across update
Requirement Statement | The OTA agent shall preserve all committed trip data, configuration, calibration, keys, and audit trail intact across the update (install, activation, rollback), shall record the update as an attributable audit entry, and shall re-establish the pre-update monitoring/arming state on the new image.
Parent | SYS-032, SYS-168, PRD-FUN-022, PRD-REL-005
Allocation | S3 · OTA-agent with storage, security
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | An update must never destroy the evidentiary record or silently disarm the device; data preservation and a logged update event keep the record continuous and auditable. Judgment.
9. Security & fault management
FWR-050 — Key handling via secure element
Requirement Statement | The firmware shall generate, store, and use device private keys and signing/verification keys only within the secure element such that private key material is never exposed in MCU RAM or flash in the clear, and shall access cryptographic operations through the secure element's interface.
Parent | SYS-032, SYS-026, PRD-SEC-008, PRD-SEC-002
Allocation | S3 · security
Type | Derived
Criticality | Critical
Safety/Security | Security
Verification Method | Test, Inspection
Status | Preliminary (reconcile with SYS-security / HAL)
Rationale | The integrity of secure-boot, OTA, tamper-evidence, and record attribution all rest on key confidentiality; a leaked key defeats every downstream control. Standard (secure-element key-management / IEC 62443 concepts; revision to confirm). Cross-flag Security engineer.
FWR-051 — At-rest encryption / integrity hooks for sensitive data
Requirement Statement | The firmware shall provide the hooks to protect sensitive at-rest data (integrity protection of the record via the FWR-034 chain, and encryption of confidential items where required by the security requirements) using keys held per FWR-050, without compromising power-fail-safe atomic writes or log recovery.
Parent | SYS-026, SYS-028, PRD-SEC-002, PRD-SEC-004
Allocation | S3 · security with storage
Type | Derived
Criticality | High
Safety/Security | Security
Verification Method | Test
Status | Preliminary (reconcile with SYS-security / HAL)
Rationale | Data protection must compose with, not defeat, integrity and recoverability; providing the hooks lets the security library (SYS-132-line) set the mechanism. Standard (data-at-rest protection concepts; revision to confirm). Cross-flag Security engineer (SYS-security owns final mechanism).
FWR-052 — Watchdog supervision with logged autonomous recovery
Requirement Statement | The firmware shall be supervised by a hardware watchdog that is serviced only on a verified-healthy main loop, shall reset and return a hung or unresponsive device to its prior armed/monitoring state without user intervention, and shall record every watchdog recovery as an attributable, timestamped event.
Parent | SYS-169, PRD-REL-006, PRD-SAFE-010
Allocation | S3 · scheduler (supervisor) with storage
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test
Status | Baselined
Rationale | A silent hang creates an unmonitored trip believed monitored — a false-reassurance hazard; the logged recovery feeds the honest-gap indication (FWR-030). Servicing the watchdog only on a healthy loop prevents masking a partial hang. Judgment (CERT-C defensive supervision; revision to confirm).
FWR-053 — Fault detection and transition to defined safe state
Requirement Statement | On detecting any power, connectivity, firmware, sensor, storage, or reset fault from which continued normal operation would be unsafe or misleading, the firmware fault arbiter shall transition to the defined safe state — preserving committed data, holding the charge subsystem non-hazardous, presenting an honest non-misleading status — and shall attempt automatic recovery to normal monitoring.
Parent | SYS-167, SYS-158, SYS-159, PRD-REL-004, PRD-SAFE-010
Allocation | S3 · scheduler (fault arbiter) with power-mgr, storage, HMI
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test (fault injection)
Status | Preliminary (reconcile with HAL)
Rationale | A fault must never corrupt the record nor leave the device silently non-monitoring; a single, defined firmware fault arbiter realises the system safe-state requirement. Judgment (ISO 14971 fail-safe; software-safety class TBD; revision to confirm). Cross-flag Safety engineer.
FWR-054 — Availability self-check of the sensing→logging chain
Requirement Statement | While armed, the firmware shall periodically self-check the end-to-end sensing→calibration→logging chain and, on detecting an outage, shall flag the device as not monitoring (on-device and, when connected, to the cloud) so that an armed device is at all times either verifiably monitoring or visibly flagged as not.
Parent | SYS-170, PRD-REL-007, PRD-SAFE-003
Allocation | S3 · scheduler (self-check) with sensor-driver, HMI, comms
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Preliminary (reconcile with HAL)
Rationale | "Armed" must reliably mean "monitoring"; a device that appears alive while not monitoring is a false-reassurance hazard. Self-check cadence is Judgment pending HAL; closed by self-check test. Cross-flag Safety and HMI engineers.
FWR-055 — Diagnostics logging and append-only audit-trail entries
Requirement Statement | The firmware shall record configuration, arming/voiding, calibration, firmware update, factory-reset, watchdog-recovery, fault, and safe-state actions as attributable, timestamped entries in the append-only audit trail such that no such action can occur without a corresponding record and no entry can be altered or deleted without detection.
Parent | SYS-028, SYS-026, PRD-SEC-004, PRD-FUN-017
Allocation | S3 · security with storage
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test, Inspection
Status | Baselined
Rationale | Attributability and non-repudiation are core to a regulated record; a lifecycle action with no audit entry is undetectable and indefensible. Standard (21 CFR Part 11 audit trail; revision to confirm). Cross-flag Security engineer.
---
Traceability: every FWR-001…FWR-055 names ≥1 SYS/PRD parent (backward) and forwards to Architecture (Stage 08 CMP-###/FUN-###), Design Outputs (Stage 09 DO-FW-NNN), Hazard Analysis (Stage 10 HA-### / RCR-###), and Verification/Validation (PV-###/VAL-###). Preliminary items reconcile with the HAL (Stage 10); Evidence-Pending numeric budgets close on the power budget and the RF-in-container and cold-endurance tests (MR-VAL-006). Firmware software-safety classification (IEC 62304 Class A/B/C) to be assigned per item on HAL completion. Assembled end-to-end by ORCH-018 / EM-ENG-012.
Subsystem / Detailed Requirements — Hardware (HWR-001 … HWR-050)
SCM · Stage 07 · Author: Principal Electrical/Hardware Engineer, integrating with EM-ENG-001. Detailed hardware (electrical / PCB / power / battery / sensors / connectors / antenna) requirements that realise the system-level power, performance, connectivity, and interface requirements — principally the power/energy parents SYS-086…SYS-103, with sensing (SYS-033/036/056), connectivity/RF (SYS-104…SYS-118), and interface (SYS-051…SYS-068) parents, decomposed from PRD-ELEC (700) and based on D1 (reusable/ rechargeable) and D3 (−30 °C / 90-day energy tension — the top program risk).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage. Each block carries: Requirement Statement · Parent · Allocation · Type · Criticality · Safety/Security · Verification Method · Status · Rationale.
Evidence posture. No cell chemistry, capacity, or datasheet number is invented — the final low-temperature chemistry is Evidence-Pending (concept §7) and every D3-bound numeric target is closed by a named test. The charge-inhibit-below-0 °C interlock (HWR-005) is severity-first Critical / Safety. Standards are cited by designation only, revision to confirm; no clause number is asserted as fact.
---
1. Power & battery
HWR-001 — Low-temperature-capable rechargeable cell selection
Requirement Statement | The power subsystem shall use a rechargeable cell whose manufacturer datasheet specifies discharge capacity, internal resistance, and safe operating limits across the −30 °C to +60 °C operating range, and whose specified −30 °C usable capacity (after derating) satisfies the SYS-087 energy-budget demand with the defined positive margin.
Parent | SYS-088, SYS-087; PRD-ELEC-001
Allocation | S2 — battery cell(s)
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (cold-capacity, cell-level and pack-level)
Status | Evidence-Pending (final chemistry — e.g. low-temp Li-ion or LTO-class — and its cold datasheet unresolved)
Rationale | Reusability (D1) mandates a rechargeable cell; deep-cold endurance (D3) mandates a low-temperature-tolerant one. The chemistry and its cold datasheet are not asserted (concept §7); the requirement fixes the selection criterion and defers the number to the cold-capacity test. Evidence-Pending.
HWR-002 — Battery pack configuration and mechanical/electrical integration
Requirement Statement | The power subsystem shall implement the battery as a defined series/parallel cell configuration (xSyP, to be fixed by the SYS-090 energy budget) with documented interconnect, cell-to-cell balancing where the configuration requires it, and a bill-of-materials pack drawing, such that pack nominal voltage and capacity are traceable to the selected cell datasheet.
Parent | SYS-088, SYS-090; PRD-ELEC-001, PRD-ELEC-008
Allocation | S2 — battery pack assembly
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Inspection, Analysis
Status | Evidence-Pending (final xSyP configuration follows chemistry and energy-budget closure)
Rationale | The pack topology sets rail voltage, gauge modelling, and balancing needs; a multi-cell pack requires balancing to avoid cell divergence and a latent over-charge of the weakest cell. Configuration is deferred until the cell and energy budget close. Judgment / Evidence-Pending.
HWR-003 — Battery-protection IC (OV / UV / OC / OT / short) as an independent protection layer
Requirement Statement | The power subsystem shall include a dedicated battery-protection IC that autonomously protects the cell against overvoltage, undervoltage, overcurrent (charge and discharge), over-temperature, and short-circuit, operating independently of the MCU and the fuel gauge so that protection persists in any MCU or firmware fault state.
Parent | SYS-096; PRD-ELEC-007
Allocation | S2 — dedicated protection IC / AFE
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (fault injection per protection threshold)
Status | Preliminary (thresholds per cell datasheet; reconcile with HAL, Stage 10)
Rationale | The device must never become a fire/thermal hazard on aircraft, in cold stores, clinics, or homes. Protection must be a hardware layer that does not depend on firmware being alive or correct — a single-fault-tolerant safe state. Severity-first Critical. Standard (IEC 62368-1 / IEC 62133 concepts; revision to confirm).
HWR-004 — Charge path / PMIC with programmable charge profile
Requirement Statement | The power subsystem shall provide a charge-management IC (PMIC/charger) that charges the selected cell per its datasheet charge profile (current, termination voltage, and taper), draws charge energy only from the USB-C sink path, and reports charge state to the MCU, without exceeding the cell's datasheet charge-current limit at any permitted temperature.
Parent | SYS-093, SYS-053; PRD-ELEC-004
Allocation | S2 — charger/PMIC on I-03 charge path
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Evidence-Pending (charge current/profile follow final cell datasheet)
Rationale | Correct charge control is a precondition for reuse turnaround (D1) and for cell safety; over-current or over-voltage charge is a cell-abuse hazard bounded here and backstopped by HWR-003. Standard (USB-C / USB-PD sink; IEC 62133 charge concepts; revision to confirm).
HWR-005 — Hardware charge-inhibit below the safe low-temperature charge limit
Requirement Statement | The power subsystem shall inhibit cell charging in hardware whenever measured cell temperature is below 0 °C (or the cell-datasheet-defined safe lower charge limit, whichever is higher), such that the charge path cannot deliver charge current to the cell below that limit even under a single MCU, firmware, or communication fault, and the inhibited state shall be signalled to the MCU and HMI.
Parent | SYS-094, SYS-053; PRD-ELEC-005, UN-SAFE-007
Allocation | S2 — charger NTC/temperature-qualified charge enable (JEITA-class), S4 (cell temp sense), S3, S6
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (charge attempt across a temperature sweep through the inhibit boundary)
Status | Preliminary (exact threshold per cell datasheet; reconcile with HAL, Stage 10)
Rationale | Charging a cold Li-ion cell drives lithium plating and a thermal-runaway pathway — a safety hazard, not a performance loss. The interlock must be a hardware qualification of the charge-enable (temperature-gated charger input) so firmware cannot defeat it. Severity-first Critical / Safety. Standard (IEC 62368-1; IEC 62133; JEITA charge-temperature concepts; revision to confirm).
HWR-006 — Dedicated cell-temperature sensing for charge qualification
Requirement Statement | The power subsystem shall sense cell (not ambient board) temperature with a sensor thermally coupled to the cell and routed to the charger's temperature-qualified charge-enable input, with accuracy sufficient to resolve the HWR-005 inhibit threshold with margin across the operating range.
Parent | SYS-094, SYS-096; PRD-ELEC-005
Allocation | S2/S4 — cell-mounted NTC or temperature sensor
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (sensor accuracy at the inhibit boundary), Inspection (thermal coupling)
Status | Preliminary (accuracy budget and placement to confirm)
Rationale | The charge-inhibit interlock (HWR-005) is only as trustworthy as the temperature it acts on; sensing board air instead of the cell would let a cold cell be charged. Cell-coupled sensing is the physical precondition for the safety control. Judgment.
HWR-007 — Cell self-heat-before-charge element (feasibility under evaluation)
Requirement Statement | Where the selected design provides it, the power subsystem shall include a controllable heating element and its drive that can warm the cell to a permitted charge temperature before charging a cold-returned device, arranged so that its activation cannot bypass or override the HWR-005 charge-inhibit and cannot exceed the cell's safe upper thermal limit.
Parent | SYS-095; PRD-ELEC-006
Allocation | S2 — heater element + drive, S3 (control)
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (heat-then-charge sequence; over-temperature fault injection)
Status | Evidence-Pending (feasibility under evaluation — D3)
Rationale | Self-heating can improve depot turnaround for cold returns without unsafe cold charging, but must never become a path that defeats the inhibit or overheats the cell. Feasibility is a live D3 branch, validated by test. Evidence-Pending.
HWR-008 — Battery fuel-gauge IC with temperature-compensated SoC/SoH
Requirement Statement | The power subsystem shall include a fuel-gauge IC that measures state of charge and state of health with temperature compensation across the operating range, reporting to the MCU with accuracy sufficient to support readiness decisions (SYS-030/091) and a non-misleading low-battery warning (SYS-098).
Parent | SYS-097, SYS-091; PRD-ELEC-009
Allocation | S2 — fuel-gauge IC, S3
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (gauge accuracy across temperature)
Status | Evidence-Pending (temperature-dependent accuracy target unresolved)
Rationale | Readiness triage and the low-battery warning are only as trustworthy as the gauge; cold shifts the voltage/SoC relationship and a coulomb-counting/model gauge with temperature compensation is required to avoid a confidently-wrong "ready" or a silent depletion. Evidence-Pending; closed by gauge-accuracy test.
HWR-009 — Reverse-polarity and fault-tolerant charge-input protection
Requirement Statement | The power subsystem's charge input shall tolerate reverse polarity, over-voltage, and non-compliant or faulty sources without damage to the device or creation of a hazard, rejecting or safely limiting an out-of-specification input at the port before it reaches the charge path.
Parent | SYS-102, SYS-054; PRD-ELEC-016
Allocation | S2 — input protection (ideal-diode/OVP front-end)
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (abuse-input: reverse polarity, over-voltage, short)
Status | Preliminary (reconcile with HAL)
Rationale | Field and depot chargers vary widely; a bad source must not damage or endanger a reusable device. Input-side protection is the first fault-tolerance layer ahead of the PMIC and BMS. Judgment; closed by abuse-input test.
HWR-010 — Quiescent (deep-sleep) current budget at −30 °C
Requirement Statement | The power subsystem's always-on current — including cell self-discharge, protection IC, fuel gauge, RTC, and wake logic — shall remain within the sleep-state allocation of the SYS-090 energy budget at −30 °C, so that idle drain over the 90-day trip does not consume the SYS-087 endurance margin.
Parent | SYS-092, SYS-090; PRD-ELEC-008
Allocation | S2/S3 — always-on housekeeping domain
Type | Derived
Criticality | High
Safety/Security | None
Verification Method | Test (cold current-profile, deep-sleep)
Status | Evidence-Pending (quiescent-current target and cold self-discharge rate unresolved)
Rationale | With most of a 90-day trip spent asleep, quiescent drain — not active bursts — often dominates the budget, and cold changes both leakage and self-discharge. The always-on part count and their sleep currents must be selected against this allocation. Evidence-Pending; measured by cold current-profile test.
HWR-011 — UN 38.3 qualification of the cell and pack
Requirement Statement | The selected cell and the assembled pack shall be qualified to UN 38.3 transport testing, and the resulting test summary shall be on file such that the device supports compliant air and ground transport of its lithium battery.
Parent | SYS-101; PRD-ELEC-015, UN-SAFE-009, UN-REG-006; D4
Allocation | S2 — cell + pack
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (UN 38.3)
Status | Evidence-Pending (test summary unresolved)
Rationale | Legal, safe air/ground transport of the Li battery is a compliance-blocking gate for a globally distributed device (D4). Standard (UN 38.3 / IATA DGR; revision to confirm); Evidence-Pending test report.
HWR-012 — Secondary (non-resettable) cell protection backstop
Requirement Statement | The power subsystem shall provide a secondary, non-resettable protection element (e.g. a current-limiting fuse and/or a secondary over-voltage protector) in series with the cell such that a failure of the primary protection IC (HWR-003) still terminates a hazardous over-current or over-voltage condition into a safe state.
Parent | SYS-096; PRD-ELEC-007, UN-SAFE-008
Allocation | S2 — series fuse / secondary protector
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (fault injection with primary protection disabled), Analysis
Status | Preliminary (reconcile with HAL)
Rationale | Single-fault safety (SYS-096) requires that loss of the primary electronic protection not leave the cell unprotected; a non-resettable backstop is the last line against a thermal event. Severity-first Critical. Standard (IEC 62133 protection concepts; revision to confirm).
2. Compute / MCU
HWR-013 — Low-power MCU selection with cold-rated deep-sleep
Requirement Statement | The compute subsystem shall use an MCU rated for operation across the −30 °C to +60 °C range that provides a retention deep-sleep state, wake-on-RTC and wake-on-interrupt, and the compute/peripheral resources to execute logging, excursion detection, and full-record verdict computation within the SYS-042 latency budget.
Parent | SYS-089, SYS-042; PRD-ELEC-008, PRD-ARCH-004
Allocation | S3 — MCU
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Analysis (selection), Test (cold operation)
Status | Preliminary
Rationale | The MCU is the duty-cycle governor's host (SYS-089); it must sleep near-zero between samples yet wake and compute a verdict over a full 90-day record at cold and low battery. Selection criteria are fixed here; the part is confirmed against the energy budget. Judgment.
HWR-014 — MCU deep-sleep current within budget
Requirement Statement | The MCU's deep-sleep current (with RTC running and RAM/register retention as required to resume without data loss) shall be within its share of the SYS-092 quiescent allocation at −30 °C.
Parent | SYS-092, SYS-089; PRD-ELEC-008
Allocation | S3 — MCU low-power domain
Type | Derived
Criticality | High
Safety/Security | None
Verification Method | Test (deep-sleep current at temperature)
Status | Evidence-Pending (deep-sleep current target follows MCU selection and budget)
Rationale | The MCU is a dominant always-on contributor; its retention-sleep current must be bounded or the 90-day budget (HWR-010) fails. Evidence-Pending; measured with the cold current-profile test.
HWR-015 — Real-time clock with independent backup
Requirement Statement | The compute subsystem shall provide a real-time clock that maintains monotonic UTC-referenced time across MCU sleep, brown-out, and battery-change/backup events, with a backup energy source sufficient to preserve time across a defined service interruption, and shall expose its running/valid status to firmware.
Parent | SYS-006, SYS-118; PRD-FUN-019, PRD-ARCH-007
Allocation | S3 — RTC + backup cell/supercap
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test (time retention across power events)
Status | Baselined
Rationale | A defensible, attributable record requires a trusted clock that survives power interruptions; a time loss would break excursion timing and ALCOA+ attribution. Standard (21 CFR Part 11 attributable/accurate; revision to confirm).
HWR-016 — Hardware root of trust / secure element for secure boot
Requirement Statement | The compute subsystem shall provide a hardware root of trust — a secure element or MCU secure-boot/key-storage block — that verifies firmware authenticity and integrity before execution and stores device identity and cryptographic keys in tamper-resistant hardware, such that unsigned or altered firmware cannot execute and keys cannot be extracted through normal interfaces.
Parent | SYS-032, SYS-065, SYS-068; PRD-SEC-008
Allocation | S3 — secure element / secure-boot block
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test, Analysis
Status | Baselined
Rationale | Signed fail-safe firmware (SYS-032) and authenticated interfaces (SYS-068) are only real if the boot chain and keys are anchored in hardware; a software-only check can be bypassed on a fielded fleet. Standard (secure-boot / IEC 62443 concepts; revision to confirm).
HWR-017 — Independent watchdog with safe-state recovery
Requirement Statement | The compute subsystem shall include a watchdog (independent of the main clock domain where feasible) that resets the MCU into a defined safe state if firmware fails to service it within a bounded period, without corrupting committed logged data or losing armed/monitoring state.
Parent | SYS-047, SYS-099; PRD-REL-004, PRD-REL-005
Allocation | S3 — watchdog (internal + external where required)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (hang injection → recovery)
Status | Preliminary
Rationale | "Armed" must reliably mean "monitoring" (SYS-047); a firmware hang must self-recover without silently stopping the log or corrupting the record. Judgment.
HWR-018 — Debug / JTAG lock in production configuration
Requirement Statement | The compute subsystem shall permanently disable or authenticate-gate the JTAG/SWD debug interface in the production device configuration, such that a fielded device does not expose firmware, keys, or logged data through the debug port.
Parent | SYS-068, SYS-132; PRD-SEC-005, PRD-SEC-006
Allocation | S3 — MCU debug-lock fuse/option bytes
Type | Derived
Criticality | High
Safety/Security | Security
Verification Method | Inspection, Test (debug-access attempt)
Status | Baselined
Rationale | An open debug port is a direct route to extract keys, forge records, or load rogue firmware — defeating the root of trust (HWR-016). Production lock is standard hardening. Standard (IEC 62443 concepts; revision to confirm).
HWR-019 — Brown-out detector gating safe-state entry
Requirement Statement | The compute subsystem shall include a brown-out detector that holds the MCU in reset below a defined supply threshold and triggers a clean safe-state entry as the supply falls, so that a collapsing rail cannot execute a corrupt write to NVM or configuration (supports SYS-057 power-fail-safe writes).
Parent | SYS-099, SYS-057; PRD-ELEC-011, PRD-REL-005
Allocation | S3 — brown-out detector / supervisor
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Test (brown-out / power-cycle integrity)
Status | Baselined
Rationale | A brown-out mid-write must never corrupt the evidentiary record; the detector is the hardware guardrail that turns a failing supply into a deterministic safe-state rather than a torn record. Judgment; closed by brown-out integrity test.
HWR-020 — Non-volatile log memory sizing and endurance
Requirement Statement | The storage subsystem shall provide non-volatile memory that stores the complete multi-channel sample and event log for a full 90-day trip at the minimum logging interval with at least 50 % spare capacity (per SYS-043), with program/erase endurance and data-retention ratings (including at temperature) that exceed the device's specified reuse life with margin, and without overwriting unsent data.
Parent | SYS-043, SYS-018; PRD-PERF-006, PRD-REL-005
Allocation | S7 — NVM (flash/FRAM), S3
Type | Allocated
Criticality | Critical
Safety/Security | Both
Verification Method | Analysis (sizing + endurance), Test
Status | Evidence-Pending (final record format, part, and endurance headroom to confirm)
Rationale | The buffered record is the product's value; capacity must cover the worst-case fully-offline trip with headroom, and write endurance must survive years of reuse without wear-out silently dropping records. Sizing depends on record format (order-of-magnitude ≈ 1.3×10^5 samples/channel at 1-min over 90 days per SYS-043). Evidence-Pending; closed with SYS-018/SYS-127. Standard (JEDEC endurance/retention concepts; revision to confirm).
3. Sensing
HWR-021 — Internal temperature sensor accuracy class
Requirement Statement | The sensing subsystem shall provide an internal temperature-measurement channel whose sensor and signal chain achieve ±0.5 °C or better accuracy across the −30 °C to +60 °C range referenced to a traceable standard, with 0.1 °C or finer resolution on every stored sample.
Parent | SYS-033, SYS-035, SYS-056; PRD-PERF-001, PRD-PERF-003
Allocation | S4 — internal temperature sensor + signal chain, S3 (linearisation)
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (traceable reference across range)
Status | Evidence-Pending (accuracy class to confirm against WHO PQS E006)
Rationale | The accept/reject/quarantine verdict is only defensible if the underlying temperature is trustworthy; ±0.5 °C is the working target, to be confirmed against the WHO PQS E006 accuracy class. Judgment / Evidence-Pending.
HWR-022 — External temperature-probe channel and connector
Requirement Statement | The sensing subsystem shall provide an external temperature-probe channel that is independently identifiable from the internal channel, meets the HWR-021 accuracy and resolution at the probe tip across the operating range, and connects through a defined, retention-featured probe connector (per HWR-040) so the probe cannot be partially seated or silently disconnected without detection.
Parent | SYS-003, SYS-051, SYS-033; PRD-MECH-004, PRD-MECH-011
Allocation | S4 — external probe + front-end, S1 (probe body/lead)
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test
Status | Evidence-Pending (probe offset/time-constant and connector selection to confirm)
Rationale | The probe represents the product zone and underpins freeze detection (SYS-003/SYS-051); conflating channels or an undetected disconnect would corrupt disposition. Judgment / Evidence-Pending.
HWR-023 — Relative-humidity sensor
Requirement Statement | The sensing subsystem shall provide a relative-humidity sensor measuring ±5 %RH or better across 10–95 %RH non-condensing, with condensation-tolerant placement/handling so that cold-to-warm transitions do not permanently degrade or falsely bias the reading.
Parent | SYS-004, SYS-036; PRD-FUN-003, PRD-ENV-004
Allocation | S4 — humidity sensor
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test
Status | Preliminary
Rationale | Humidity is a secondary condition indicator; accuracy must be meaningful without over-driving cost, and condensation on the sensor during cold-to-warm transitions is a named accuracy concern (flag Sensing/Materials engineer). Judgment.
HWR-024 — 3-axis accelerometer for shock/impact capture
Requirement Statement | The sensing subsystem shall provide a 3-axis accelerometer with a full-scale range and bandwidth able to capture per-profile configurable shock/impact events with peak magnitude and timestamp (SYS-005), supporting a low-power/interrupt-driven mode so routine-handling accelerations do not wake the full system or drain the budget.
Parent | SYS-005, SYS-044; PRD-FUN-004, PRD-ENV-006
Allocation | S4 — accelerometer, S3 (event threshold)
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test
Status | Preliminary
Rationale | Product-relevant mishandling must be captured without alarm-fatigue noise (C8) and without an always-on high-rate sampler defeating the energy budget; a hardware interrupt/threshold mode is the enabler. Judgment.
HWR-025 — GNSS receiver with duty-cycled acquisition
Requirement Statement | The sensing subsystem shall provide a GNSS/GPS receiver that acquires a location fix on a configurable and event-triggered cadence within the location energy allocation, reports fix quality per attempt, and can be fully powered off between fixes, degrading gracefully (best-available / last-known with quality flag) when a fix cannot be obtained inside the shipper.
Parent | SYS-117, SYS-048, SYS-029; PRD-CONN-005, PRD-PERF-010
Allocation | S4 — GNSS receiver + antenna, S3 (scheduling/gating)
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test (fix rate / time-to-fix / energy per fix)
Status | Evidence-Pending (in-container fix availability and per-fix energy unresolved)
Rationale | GNSS is the most attenuated band inside a metallized shipper and each fix is energy-costly against the D3 budget; location is fleet/attribution support, never a safety function, so it must fail soft and be hard-gated. Evidence-Pending; closed by PRD-VER-005.
HWR-026 — Per-sensor power gating
Requirement Statement | The sensing subsystem shall allow each high-drain sensor rail (external-probe front-end, humidity, GNSS, and the accelerometer's high-rate mode) to be independently gated under firmware control so that a sensor draws energy only during its scheduled acquisition window (supports SYS-089 governor and SYS-090 allocation).
Parent | SYS-089, SYS-090, SYS-055; PRD-ELEC-008
Allocation | S4/S2 — load switches on sensor rails, S3 (control)
Type | Derived
Criticality | High
Safety/Security | None
Verification Method | Test (per-rail current)
Status | Baselined
Rationale | The 90-day/−30 °C budget is met by duty-cycling, not chemistry (AD-2, D3); sensor rails that cannot be individually gated make the energy governor impossible. Judgment.
HWR-027 — Temperature signal chain and reference stability
Requirement Statement | The sensing subsystem's temperature acquisition (ADC and voltage/current reference or ratiometric front-end) shall have reference stability, noise, and self-heating low enough that the HWR-021 accuracy is met across temperature and holds over the calibration interval and reuse life, with any lapse of calibration validity flagged rather than silent.
Parent | SYS-034, SYS-033; PRD-PERF-002
Allocation | S4 — ADC + reference / signal chain, S3
Type | Derived
Criticality | Critical
Safety/Security | Both
Verification Method | Test (aging/drift), Analysis
Status | Evidence-Pending (reference-drift budget and calibration interval unresolved)
Rationale | A sensor that drifts silently produces a confidently-wrong verdict; the reference and signal-chain self-heating are the dominant accuracy-over-life terms behind the ±0.5 °C target. Standard (ISO/IEC 17025 traceability concepts; revision to confirm) / Evidence-Pending.
HWR-028 — Probe-presence and continuity detection
Requirement Statement | The sensing subsystem shall detect the presence and electrical continuity of the external probe such that a missing, open, or shorted probe is flagged as a channel fault (feeding SYS-007 plausibility flagging and SYS-024 pre-arm checks) rather than read as a plausible temperature.
Parent | SYS-007, SYS-024, SYS-056; PRD-FUN-017, PRD-REL-004
Allocation | S4 — probe front-end detection, S3
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (probe removed / open / short)
Status | Baselined
Rationale | A silently-wrong reading is more dangerous than a flagged one; the pre-arm "probe present" check (SYS-024) and plausibility flagging (SYS-007) require hardware-detectable probe faults. Judgment.
4. Connectivity hardware
HWR-029 — LTE-M modem (primary wide-area bearer)
Requirement Statement | The connectivity subsystem shall provide an LTE-M (3GPP Cat-M1) modem supporting low-power modes (PSM/eDRX-class), the market band plans of the target regions (SYS-115), and duty-cycled operation consistent with the 90-day energy budget, with a host interface that lets the MCU fully power-cycle and gate it.
Parent | SYS-105, SYS-089; PRD-CONN-001, PRD-ARCH-005
Allocation | S5 — LTE-M modem module
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Baselined
Rationale | LTE-M is the primary alert/upload path (D2/D4) and the dominant active load; it must support deep power-saving and hard gating so it does not overrun the budget. Standard (3GPP LTE-M/Cat-M1; revision to confirm).
HWR-030 — Wi-Fi radio for opportunistic offload
Requirement Statement | The connectivity subsystem shall provide an IEEE 802.11 Wi-Fi radio for opportunistic bulk-data offload at depots/facilities, fully power-gated when not in use, sharing or coexisting with the other radios per an antenna/coexistence plan.
Parent | SYS-106, SYS-108; PRD-CONN-002
Allocation | S5 — Wi-Fi radio
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test
Status | Baselined
Rationale | Higher-throughput, lower-cost offload drains buffered backlog and preserves the energy budget (C2); it must be gateable and coexist with LTE-M/BLE. Standard (IEEE 802.11; revision to confirm).
HWR-031 — BLE radio for provisioning and short-range offload
Requirement Statement | The connectivity subsystem shall provide a Bluetooth Low Energy radio for local provisioning, configuration, login-free local verdict/record read, and short-range offload, supporting authenticated pairing/bonding (SYS-110) and low idle exposure when not actively paired.
Parent | SYS-107, SYS-060, SYS-110; PRD-CONN-003, PRD-MOB-001
Allocation | S5 — BLE radio, S3 (session)
Type | Allocated
Criticality | High
Safety/Security | Both
Verification Method | Test, Demonstration
Status | Baselined
Rationale | BLE is the local setup and offline-read path (AD-1/AD-5); it must serve the workflow while limiting the wireless attack surface when idle. Standard (Bluetooth Core / BLE; revision to confirm).
HWR-032 — Integral antennas and RF matching per band
Requirement Statement | The connectivity subsystem shall provide integral antenna(s) and matching networks for the LTE-M, Wi-Fi, BLE, and GNSS bands of the target markets, with a defined enclosure keep-out and placement, tuned to meet the in-container link-budget target of SYS-104 across all active bands.
Parent | SYS-109, SYS-104, SYS-059; PRD-CONN-007, PRD-CONN-012
Allocation | S5 — antennas + matching, S1 (keep-out/placement)
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (antenna efficiency / total radiated power in-container), Analysis (link budget)
Status | Evidence-Pending (efficiency and integral-vs-external decision unresolved)
Rationale | RF attenuation inside a metallized/insulated shipper can exceed the fade margin of a low-power radio; antenna efficiency and placement are the principal levers before an external antenna is required. Delayed/undelivered alerts are a hazard contributor. Evidence-Pending; closed by PRD-VER-005; flag Mechanical (keep-out) and Regulatory (radiated power).
HWR-033 — External-antenna provision (IP-sealed connector)
Requirement Statement | The connectivity subsystem shall provide, where the RF-in-container test shows the integral antenna cannot close the SYS-104 budget, a qualified external-antenna option via an IP-sealed connector (per PRD-ENV-007) that preserves enclosure ingress protection and does not degrade the other bands when unused.
Parent | SYS-109, SYS-104; PRD-CONN-012, PRD-ENV-007
Allocation | S5 — external-antenna connector + RF path, S1 (sealed connector)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (link budget with external antenna; IP seal)
Status | Evidence-Pending (external-antenna decision pending RF-in-container test)
Rationale | If the integral antenna cannot close the budget inside the shipper, an external antenna via an IP67-sealed connector is the fallback branch; this is a live architectural decision. Evidence-Pending; closed by PRD-VER-005; flag Mechanical (seal).
HWR-034 — SIM / eSIM for roaming wide-area connectivity
Requirement Statement | The connectivity subsystem shall provide roaming-capable SIM/eSIM provisioning (form factor to be fixed) that lets the LTE-M modem automatically select an available permitted network across the carriers and regions encountered on a trip (SYS-116), and shall physically protect the SIM against loss/tamper in the sealed enclosure.
Parent | SYS-116, SYS-115; PRD-CONN-001, PRD-CONN-012
Allocation | S5 — SIM/eSIM + modem
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test
Status | Evidence-Pending (per-market carrier coverage and SIM/eSIM strategy unresolved)
Rationale | Global vaccine lanes cross carrier and national boundaries mid-trip (D4); without roaming provisioning the alert/upload path breaks at each border. eSIM avoids a mechanical socket in a sealed device but the strategy is unresolved. Evidence-Pending.
HWR-035 — RF front-end performance at the −30 °C extreme
Requirement Statement | The connectivity subsystem's RF front-end (PA, LNA, matching, and reference oscillator) shall meet its transmit-power, sensitivity, and frequency-stability specifications across the −30 °C to +60 °C operating range so that link performance and the SYS-104 budget are not silently degraded at the cold extreme.
Parent | SYS-104, SYS-046; PRD-CONN-007, PRD-ENV-001
Allocation | S5 — RF front-end + reference oscillator
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (RF performance over temperature)
Status | Evidence-Pending (cold RF-performance margin unresolved)
Rationale | Oscillator frequency error and PA/LNA gain shift with temperature; a link budget closed at room temperature can fail at −30 °C inside the shipper, degrading alert delivery. Evidence-Pending; validated over-temperature with the RF-in-container test.
HWR-036 — Radio peak-current support and rail decoupling
Requirement Statement | The power distribution to the radios shall supply each radio's peak transmit current (LTE-M in particular) without the rail sagging below the modem's minimum operating voltage, with bulk/local decoupling sized so that transmit bursts do not brown-out the MCU or corrupt an in-flight NVM write.
Parent | SYS-099, SYS-055, SYS-105; PRD-ELEC-011, PRD-ARCH-004
Allocation | S2/S5 — radio supply + decoupling
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Test (transmit-burst rail integrity)
Status | Preliminary
Rationale | LTE-M transmit peaks are large and fast; an undersized rail or decoupling can brown-out the MCU during a transmit burst, corrupting the record at exactly the moment of upload. Judgment; closed with the brown-out integrity test.
5. Interconnect / PCB / power integrity
HWR-037 — PCB stack-up defined for signal, power, and environment
Requirement Statement | The electrical design shall use a defined PCB stack-up (layer count, dielectric, copper weight, controlled-impedance layers for RF/high-speed nets, and a continuous reference ground) documented to IPC design standards and suited to the −30 °C to +60 °C thermal-cycling environment.
Parent | SYS-055, SYS-104, SYS-066; PRD-ELEC-013, PRD-ARCH-009
Allocation | S2/S5 — PCB stack-up
Type | Derived
Criticality | High
Safety/Security | None
Verification Method | Inspection, Analysis
Status | Preliminary
Rationale | RF matching (HWR-032), power integrity (HWR-036/043), and thermal-cycle reliability all depend on a controlled stack-up; an ad-hoc stack-up undermines impedance control and evidence. Standard (IPC-2221/IPC-6012 concepts; revision to confirm).
HWR-038 — Conformal coating for humidity and condensation
Requirement Statement | The populated PCB assembly shall be conformally coated (type and coverage to be fixed) to protect against condensation and humidity ingress across cold-to-warm transitions, with defined keep-out of connector contacts, antenna feeds, and the humidity sensor aperture.
Parent | SYS-004, SYS-069, SYS-103; PRD-ENV-004, PRD-ENV-009
Allocation | S2 — conformal coat process, S1 (interface)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (humidity/condensation), Inspection (coverage)
Status | Preliminary
Rationale | Repeated −30 °C to +60 °C cycling drives condensation that can cause leakage, corrosion, and dendrite growth on a life-critical logger; coating protects the record path while sparing the humidity sensor and RF feeds. Standard (IPC-CC-830 / IPC-A-610 concepts; revision to confirm).
HWR-039 — USB-C charge/data receptacle
Requirement Statement | The electrical design shall provide a USB-C receptacle presenting a compliant sink interface (defined pin roles, CC behaviour as sink, accepted voltage/current envelope) that accepts charge from a standard compliant source, sources no hazardous voltage on any pin, and remains safe under reverse-insertion, over-voltage, and short conditions.
Parent | SYS-093, SYS-054, SYS-102; PRD-ELEC-004, PRD-SAFE-007
Allocation | S2 — USB-C connector, S1 (sealed port)
Type | Allocated
Criticality | High
Safety/Security | Both
Verification Method | Test, Inspection
Status | Baselined
Rationale | The connector is a defined electrical contract, not just a hole; a compliant, documented sink prevents depot-charger incompatibility and fault-injection hazards, and standard USB-C supports reuse (D1). Standard (USB Type-C / USB-PD sink; IEC 62368-1; revision to confirm).
HWR-040 — External-probe connector (retention and sealing)
Requirement Statement | The electrical design shall provide an external-probe connector with positive retention and, where the probe penetrates the enclosure, an ingress seal, such that the probe cannot be partially seated or vibrate loose undetected (supports HWR-028 detection and SYS-051 thermal coupling) and the enclosure ingress rating is preserved.
Parent | SYS-003, SYS-051, SYS-024; PRD-MECH-004, PRD-MECH-011
Allocation | S4/S1 — probe connector, seal
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (retention, insertion cycles, seal), Inspection
Status | Preliminary
Rationale | A probe that backs out or is half-mated corrupts the product-zone measurement or opens an ingress path; retention plus detection (HWR-028) make the channel trustworthy. Judgment; flag Mechanical (seal).
HWR-041 — ESD protection on all externally accessible interfaces
Requirement Statement | The electrical design shall provide ESD protection on the USB-C port, probe connector, external-antenna connector, and any exposed button/contact sufficient to meet the PRD-REG EMC immunity levels without malfunction, data corruption, or damage during handling and depot connection.
Parent | SYS-103, SYS-054; PRD-ELEC-013, UN-ENV-009, UN-REG-005
Allocation | S2 — ESD protection devices, S1 (port), S3
Type | Allocated
Criticality | High
Safety/Security | Both
Verification Method | Test (IEC 61000-4-2)
Status | Preliminary (immunity levels set in PRD-REG; reconcile with HAL)
Rationale | Externally accessible interfaces are the primary ESD entry points during handling; a discharge must not corrupt the record or damage the device. Standard (CISPR / IEC 61000-4-2; revision to confirm).
HWR-042 — Evidence, isolation, and radio/sensor coupling control
Requirement Statement | The electrical design shall implement a defined evidence and isolation scheme — including radio return currents, the analog sensing ground, and the charge-port ground — so that radio transmit bursts and charge-port transients do not couple into the temperature signal chain beyond the noise budget of HWR-021/HWR-027.
Parent | SYS-055, SYS-056, SYS-103; PRD-ELEC-013
Allocation | S2/S4/S5 — evidence/isolation topology
Type | Derived
Criticality | High
Safety/Security | Both
Verification Method | Analysis, Test (EMC / coupled-noise)
Status | Preliminary
Rationale | Shared or poorly partitioned ground lets LTE-M transmit bursts and charge transients inject noise into the µV-level temperature front-end, degrading the safety-critical measurement; evidence discipline is the defence. Judgment; closed with EMC/accuracy tests.
HWR-043 — Regulated power rails, sequencing, and decoupling
Requirement Statement | The power subsystem shall provide the regulated rails required by the MCU, sensing, connectivity, HMI, and storage subsystems with defined tolerance, sequencing, and local decoupling, such that each rail is independently gateable (SYS-055) and rail transitions never corrupt a committed NVM write or produce an out-of-spec supply to any device.
Parent | SYS-055, SYS-099; PRD-ELEC-008, PRD-ELEC-011
Allocation | S2 — regulators + decoupling, S3 (sequencing control)
Type | Allocated
Criticality | High
Safety/Security | Both
Verification Method | Test, Analysis
Status | Baselined
Rationale | The energy governor duty-cycles rails (AD-2, D3); rails that are not cleanly regulated, sequenced, and decoupled make gating unsafe for the record and unreliable for the radios/sensors. Judgment.
HWR-044 — Test points and end-of-line test access
Requirement Statement | The electrical design shall provide test points and a defined test-access interface (probe pads and/or connector) enabling end-of-line manufacturing test, calibration, and provisioning coverage of the power path, charge-inhibit interlock, sensing channels, and radios, without leaving an unsecured production debug path (per HWR-018).
Parent | SYS-066, SYS-030; PRD-ARCH-009, PRD-SVC-004
Allocation | S2/S3 — test points / EOL fixture interface
Type | Derived
Criticality | Medium
Safety/Security | Security
Verification Method | Inspection, Demonstration (EOL test coverage)
Status | Preliminary
Rationale | A reusable, safety-relevant fleet needs EOL test and calibration coverage (including verifying the HWR-005 interlock) — but the access must not become a fielded attack surface, so it is reconciled with the debug lock. Judgment.
HWR-045 — Charge-input over-voltage clamp / transient protection
Requirement Statement | The charge input shall include an over-voltage clamp and transient protection ahead of the PMIC that limits or disconnects the input on an over-voltage or surge event before it can damage the charge path or cell (working with the reverse-polarity/fault-tolerant front-end of HWR-009).
Parent | SYS-102, SYS-103; PRD-ELEC-016, PRD-ELEC-013
Allocation | S2 — OVP / TVS on charge input
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (over-voltage / surge injection)
Status | Preliminary (reconcile with HAL)
Rationale | Depot chargers and surges vary; an input over-voltage that reaches the PMIC or cell is an abuse/hazard path. The clamp is the transient counterpart to the reverse-polarity front-end. Judgment; closed with the abuse-input test.
6. Display / HMI hardware
HWR-046 — Cold-legible on-device display (OLED)
Requirement Statement | The HMI subsystem shall provide a display capable of rendering the accept/reject/quarantine verdict and device state legibly at the −30 °C cold extreme, updating within the SYS-041 responsiveness budget, with a display technology and low-temperature rating to be confirmed by test (OLED cold-response is the primary concern).
Parent | SYS-015, SYS-041, SYS-058; PRD-UX-001, PRD-PERF-007
Allocation | S6 — display module, S3
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (legibility and response time at −30 °C)
Status | Evidence-Pending (OLED cold response / display technology unresolved)
Rationale | The edge verdict is the last-mile decision surface (SYS-015/AD-6); OLED response time and contrast degrade in deep cold, so the technology must be validated at −30 °C rather than assumed. Evidence-Pending; closed by cold-display test.
HWR-047 — Glove-friendly actuation button
Requirement Statement | The HMI subsystem shall provide a single primary button operable with cold-weather gloves (actuation force, travel, and target size to suit gloved use), sealed to the enclosure ingress rating, whose single actuation maps to a single unambiguous action and which can wake the MCU from deep sleep.
Parent | SYS-023, SYS-058, SYS-041; PRD-UX-004, PRD-ARCH-009
Allocation | S6 — button, S1 (seal), S3
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test, Demonstration (gloved actuation)
Status | Preliminary
Rationale | A time-pressured last-mile user in cold-weather gloves must be able to arm and query the device reliably (SYS-023/SYS-041); a stiff or tiny button defeats the workflow and can cause a mis-arm. Judgment; flag Mechanical (seal/feel).
HWR-048 — Low-energy status LED
Requirement Statement | The HMI subsystem shall provide a low-energy status indicator (LED) that distinguishes at least the armed/monitoring, fault, charging, charge-complete, and charge-inhibited states, such that the HWR-005 charge-inhibited condition is unambiguously visible to a depot operator without materially loading the energy budget.
Parent | SYS-100, SYS-094; PRD-ELEC-012, UN-SAFE-007
Allocation | S6 — status LED, S3
Type | Allocated
Criticality | Medium
Safety/Security | Safety
Verification Method | Demonstration
Status | Baselined
Rationale | A charge-inhibited device that looks dead would be mistaken for a fault or a failed charge; the inhibit state must be legible without draining the budget. Judgment; energy-cheap indication.
HWR-049 — HMI power gating and duty-cycled indication
Requirement Statement | The HMI subsystem shall power the display and indicators only when required (on user action or defined events) and shall hold them in a near-zero-power state otherwise, so that HMI draw stays within its SYS-090 allocation over a 90-day trip.
Parent | SYS-090, SYS-089, SYS-055; PRD-ELEC-008
Allocation | S6/S2 — HMI rail gating, S3 (control)
Type | Derived
Criticality | Medium
Safety/Security | None
Verification Method | Test (HMI rail current)
Status | Baselined
Rationale | A continuously-on display over 90 days would violate the energy budget (D3); the HMI must be event-driven and gated like every other high-drain load. Judgment.
HWR-050 — Button debounce and fail-detectable HMI elements
Requirement Statement | The HMI subsystem shall debounce the button in hardware/firmware and shall arrange the display and LED so that a stuck, failed, or disconnected HMI element is detectable and does not present a false device state or turn a "reject" verdict into an "accept" (supports SYS-058).
Parent | SYS-058, SYS-017; PRD-UX-004, PRD-SAFE-003
Allocation | S6 — button/display drive, S3
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (bounce, stuck-element / failed-pixel injection)
Status | Preliminary
Rationale | The HMI is the decision surface; a bouncing button can double-actuate an arm/query, and a failed segment showing a stale or wrong verdict is a false-reassurance hazard. Elements must fail detectably. Judgment.
Subsystem/Detailed Requirements — Interfaces (IFR-001 … IFR-035)
SCM · Stage 07 · Owner: Principal Systems/Interface Engineer, coordinating with the Cloud/Platform (S8), RF/Connectivity (S5), Power (S2), and Cybersecurity (EM-ENG-008) engineers. Detailed interface requirements (the Interface Control Document, ICD) that decompose the system interface requirements SYS-051…SYS-068 (and connectivity SYS-104…SYS-118) onto the typed interface inventory I-01…I-12 of [../../05-product-requirements/system-block-diagram.md](../../05-product-requirements/system-block-diagram.md) §3. Each IFR fixes the concrete signals, formats, timing, and error behaviour of one boundary so the boundary is buildable and testable against a published contract. Field set per [../06-system-requirements/index.md](../06-system-requirements/index.md). Based on D1–D4 and the architecture decisions AD-1…AD-8.
Evidence posture. Standards are cited by designation only — USB Type-C / USB-PD, IEC 62368-1, Bluetooth Core / GATT, 3GPP LTE-M (Cat-M1), IEEE 802.11, MQTT v5, TLS 1.2+/1.3, OAuth 2.0, OpenAPI, ISO/IEC/IEEE 29148 — with revision/clause to confirm; no clause number is asserted as fact. The RF-in-container interface (I-08) is the principal open architectural risk and is carried Evidence-Pending until closed by the RF-in-container qualification test (PRD-VER-005). Anything that authenticates a party, protects record integrity, gates cell charging, or carries the excursion-alert path is Critical and Safety or Both (CLAUDE.md rule 3, severity-first).
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
---
1. Electrical connector interfaces
IFR-001 — USB-C receptacle pinout and sink-only role
Requirement Statement | The device shall implement the sealed USB-C port as a full 24-pin USB Type-C receptacle wired as a charge sink (UFP / power-consumer) only: VBUS and GND on all four power pairs, a single CC pin (CC1/CC2) presenting the pull-down termination Rd (5.1 kΩ ±10 %) to advertise sink role and detect orientation, and D+/D− and the SBU pins routed only to the EOL-test/debug controller (IFR-005) and left high-impedance in fielded operation, such that the device never sources VBUS, never asserts a DFP/host role, and never presents Rp on CC.
Parent | SYS-054, PRD-ELEC-002, PRD-ELEC-003; details I-03 (PWR: USB-C source → port).
Allocation | S1 (receptacle) ↔ S2 (charge path)
Type | Detailed (decomposes SYS-054)
Criticality | High
Safety/Security | Both
Verification Method | Inspection (netlist/schematic), Test (CC termination, role behaviour)
Status | Baselined
Rationale | A defined pinout and sink-only role is the electrical contract that lets any compliant depot charger power the device while guaranteeing the device cannot back-drive a charger or expose a host data surface. Rd on CC is the mechanism by which a Type-C source recognises a sink and applies power. Standard (USB Type-C, receptacle/CC termination; revision to confirm).
IFR-002 — USB-C data role: no USB host/device enumeration required in field
Requirement Statement | The device shall not require USB 2.0/3.x data enumeration for any field function (charging, arming, verdict read); D+/D− shall be used solely for the local EOL/debug/firmware-recovery path (IFR-005), and when that path is inactive the USB data lines shall present no enumerable USB device to a host, such that connecting the device to an arbitrary USB host neither exposes device data nor alters device state.
Parent | SYS-054, SYS-068, PRD-ELEC-002, PRD-SEC-006; details I-03 (DATA aspect of USB-C).
Allocation | S1 (port) ↔ S3 (USB controller)
Type | Detailed
Criticality | High
Safety/Security | Security
Verification Method | Test (host-enumeration attempt), Inspection
Status | Baselined
Rationale | Keeping the field USB interface data-silent removes a physical attack/exfiltration surface on a device that stores a regulated record; provisioning and read are governed over BLE (I-09) and cloud (I-10) where they can be authenticated. Judgment; consistent with AD-5 trust-boundary split.
IFR-003 — External temperature-probe connector pinout and signal definition
Requirement Statement | The device shall provide a keyed, IP-sealed external-probe connector whose pinout is fixed as {probe-sense+, probe-sense−, probe-ID/presence, shield/GND}, carrying the analog or digital-sensor signal of the external probe (e.g. thermistor excitation/return or 1-Wire/I2C sensor bus), and shall detect probe presence and probe-type/ID on the presence line such that an absent, wrong-type, or open/short-circuited probe is detected and flagged (feeding SYS-007/SYS-024) rather than silently read as a valid product-zone temperature.
Parent | SYS-051, SYS-056, SYS-003, PRD-MECH-004, PRD-MECH-011; details I-01/I-05 (probe → sensing).
Allocation | S1 (connector/probe lead) ↔ S4 (sensing front-end)
Type | Detailed
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (open/short/wrong-probe injection), Inspection (pinout/keying)
Status | Preliminary (probe technology — thermistor vs digital — to confirm with S4)
Rationale | The external probe is the product-zone proxy on which every freeze/heat verdict rests (SYS-051); an undetected open probe reading room temperature would fabricate a false "accept." Presence/ID detection plus fault flagging is the interface-level defence. Judgment; probe signal type fixed in S4 detailed design.
IFR-004 — Internal sensor buses: I2C/SPI addressing, rate, and channel identity
Requirement Statement | The device shall connect the internal temperature, humidity, accelerometer, fuel-gauge, and secure-element devices to the MCU over defined internal buses — I2C at 100/400 kHz (Standard/Fast mode) with each device assigned a fixed, documented 7-bit address, and/or SPI with a dedicated chip-select per device — such that every reading is bound to its correct source device address/ CS, bus contention and address collision are precluded by the address map, and a NAK, CRC/PEC failure, or bus timeout returns a defined error to the acquisition layer rather than a stale or substituted value.
Parent | SYS-056, SYS-003, SYS-007, PRD-FUN-002, PRD-ARCH-009; details I-05 (sensors → MCU).
Allocation | S4 (sensors) ↔ S3 (MCU acquisition)
Type | Detailed
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (bus fault injection, address-map audit), Inspection
Status | Baselined
Rationale | Channel identity and read integrity at this boundary underpin dual-channel acquisition (SYS-003) and plausibility flagging (SYS-007); a mislabelled or CRC-corrupt read corrupts disposition. Fixing the address map in the ICD prevents the classic multi-sensor I2C address clash. Standard (I2C bus specification, addressing/PEC; SPI de-facto; revision to confirm).
IFR-005 — Debug / end-of-line test connector definition
Requirement Statement | The device shall provide an internal (non-user-accessible) debug/EOL-test interface exposing SWD/JTAG programming, a UART console, and boundary/functional test access, such that the interface enables factory programming, calibration, and functional test, and shall be locked in fielded units — debug access disabled or authentication-gated and the readout-protection/lifecycle fuse set — so that a fielded device exposes no unauthenticated code, key, or record readout over this port.
Parent | SYS-066, SYS-068, PRD-ELEC-002, PRD-SEC-006; details internal test boundary of I-03/I-06.
Allocation | S3 (MCU/secure element) ↔ S1 (test pads/connector)
Type | Detailed (derived — manufacturing/security necessity)
Criticality | High
Safety/Security | Security
Verification Method | Test (locked-part readout attempt), Inspection
Status | Baselined
Rationale | An EOL/debug port is essential for manufacture and calibration but is a direct route to keys and the record if left open in the field; readout protection plus authenticated debug closes the surface. Standard (IEC 62443 concepts, secure-provisioning practice; revision to confirm).
---
2. Power interface
IFR-006 — USB-C charge-input voltage/current envelope and USB-PD negotiation limit
Requirement Statement | The device shall accept charge input from a USB Type-C source within a defined envelope: default USB-C 5 V sink at the current advertised by the source's CC (up to 3.0 A per Type-C current advertisement), and where USB-PD is present the device shall negotiate no source PDO above a bounded limit (≤ 9 V / to-confirm-W), reject/ignore higher-voltage PDOs, and remain safe and non-damaged when connected to a non-PD 5 V or legacy charger, such that the device charges from any compliant source and never draws or accepts a voltage outside its input-protection rating.
Parent | SYS-053, SYS-054, PRD-ELEC-003; details I-03 (PWR: source → PMIC).
Allocation | S2 (PMIC/input protection) ↔ S1 (port)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (PD contract negotiation, OVP, legacy-charger interop)
Status | Evidence-Pending (max PD power/voltage cap to confirm with S2 charge-rate/thermal budget)
Rationale | Depot chargers are uncontrolled; a bounded, explicitly negotiated input envelope prevents an aggressive PD source from over-volting the input and guarantees interoperability with the simplest 5 V charger. Standard (USB Type-C current advertisement, USB-PD PDO negotiation; IEC 62368-1; revision to confirm).
IFR-007 — Charge-inhibit signalling across the PMIC↔BMS↔MCU boundary
Requirement Statement | The device shall carry an explicit charge-inhibit signal path such that when cell temperature is below the low-temperature charge threshold (0 °C nominal) or above the high-temperature limit, the BMS shall assert charge-inhibit to the PMIC (gating cell charge current to zero) and report the inhibited state and its cause over the BMS→MCU data link (IFR-008), such that charging never proceeds into a cold cell and the inhibited condition is observable to firmware and surfaced to the user (SYS-024) rather than presenting as a silent failure to charge.
Parent | SYS-053, PRD-ELEC-003, PRD-SAFE-007, PRD-ARCH-008; details I-03/I-04 (charge-path control).
Allocation | S2 (BMS/PMIC) ↔ S3 (charge state)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (thermal-chamber charge-inhibit trip and reporting)
Status | Baselined
Rationale | Charging a sub-0 °C Li-ion cell risks lithium plating and a thermal event; the inhibit gate is the safety control on this boundary (D3, AD-3), and making the inhibited state legible prevents a user from concluding the device is dead when it is protecting itself. Standard (Li-ion charge-temperature limits per cell datasheet; IEC 62368-1; revision to confirm).
IFR-008 — BMS↔MCU fuel-gauge / charge-status data interface
Requirement Statement | The device shall expose battery state from the BMS/fuel-gauge to the MCU over a defined digital interface (I2C/SMBus, gauge address per IFR-004 map) reporting at minimum state-of-charge, voltage, cell temperature, charge/discharge state, and charge-inhibit/fault flags on a defined polling cadence, such that firmware can enforce pre-arm charge readiness (SYS-024), the energy governor (SYS-055/SYS-108), and honest low-battery indication, and a gauge read that NAKs or returns implausible values is flagged (SYS-007) rather than trusted.
Parent | SYS-055, SYS-024, PRD-ELEC-008, PRD-REL-005; details I-04 (BMS → MCU data).
Allocation | S2 (fuel gauge) ↔ S3 (MCU)
Type | Detailed
Criticality | High
Safety/Security | Safety
Verification Method | Test (gauge read-back, fault-injection)
Status | Baselined
Rationale | The 90-day/−30 °C budget (D3) is managed by firmware acting on trustworthy state-of-charge and temperature; pre-arm readiness (SYS-024) and low-battery honesty depend on this read being correct or flagged. Judgment; gauge part/protocol fixed in S2 detailed design.
IFR-009 — Regulated power-rail definitions, tolerances, and independent gating
Requirement Statement | The device shall define each distributed rail from the BMS/PMIC — nominal voltage, tolerance band, sequencing, and maximum load — for the MCU/always-on domain, sensing, connectivity, HMI, and storage, and shall provide firmware-controlled enable/gate lines for the high-draw rails (connectivity, sensing, HMI) so a subsystem is powered only in the active power state, such that rail voltages stay within tolerance across the operating range and a rail enable/disable transition never occurs during an in-flight NVM write (coordinated with IFR-010).
Parent | SYS-055, SYS-057, PRD-ELEC-008, PRD-REL-005, PRD-ARCH-004; details I-04 (BMS → subsystems).
Allocation | S2 (regulators) ↔ S3 (rail control)
Type | Detailed
Criticality | High
Safety/Security | None
Verification Method | Test (rail tolerance, sequencing), Analysis (energy budget)
Status | Evidence-Pending (rail voltages/tolerances closed by S2 energy-budget analysis)
Rationale | Meeting the endurance budget by duty-cycling (AD-2, D3) requires individually gateable rails with defined tolerances; ungated rails make the energy governor impossible, and an ungoverned rail-down during a write corrupts the record (SYS-057). Judgment; rail budget closed by S2 endurance analysis.
---
3. Wireless RF interfaces
IFR-010 — LTE-M bearer parameters: band set, APN, PSM and eDRX
Requirement Statement | The device shall configure the LTE-M (3GPP Cat-M1) modem with a market-selectable band set, the operator APN(s), and power-saving parameters — Power Saving Mode (PSM) with a configured TAU/Active-Timer and/or eDRX with a defined paging cycle sized to the upload cadence and alert-latency target — such that the modem sleeps between scheduled sessions to preserve the 90-day budget yet remains reachable/awake within the window required to deliver an excursion alert, and shall not transmit outside the bands/limits permitted for the configured region.
Parent | SYS-105, SYS-115, SYS-116, PRD-CONN-001, PRD-CONN-012; details I-08 (device → WAN).
Allocation | S5 (LTE-M modem/SIM) ↔ S3 (session/power scheduling)
Type | Detailed
Criticality | High
Safety/Security | Both (alert path; regional compliance)
Verification Method | Test (network attach, PSM/eDRX timing, band conformance)
Status | Evidence-Pending (per-market band plan, PSM/eDRX values, carrier/APN strategy unresolved)
Rationale | PSM/eDRX are the mechanism that reconciles always-reachable alerting with the energy budget (D2/D3); the wrong sleep parameters either drain the battery or delay a rescue-enabling alert past its hazard window. Standard (3GPP LTE-M/Cat-M1, PSM/eDRX; revision to confirm); coverage/roaming per SYS-116.
IFR-011 — Wi-Fi security and onboarding interface
Requirement Statement | The device shall connect to a provisioned IEEE 802.11 network using at minimum WPA2-PSK and, where available, WPA3-SAE and WPA2/WPA3-Enterprise (EAP), accepting network credentials only through the authenticated provisioning path (BLE GATT IFR-016 or authenticated cloud config), and shall fail closed — retaining buffered data for another bearer (SYS-108/SYS-111) — when association or authentication fails, rather than falling back to an open/unencrypted association.
Parent | SYS-106, SYS-110, PRD-CONN-002, PRD-SEC-005; details I-08 (device → WAN, Wi-Fi).
Allocation | S5 (Wi-Fi radio) ↔ S3 (bearer/credential mgmt)
Type | Detailed
Criticality | Medium
Safety/Security | Both
Verification Method | Test (WPA2/WPA3 association, EAP, open-network refusal)
Status | Evidence-Pending (enterprise-Wi-Fi/EAP method support to confirm — PRD-CONN-002)
Rationale | Wi-Fi is a cost/energy-efficient depot offload path but is only acceptable if credentials arrive over an authenticated channel and the link is encrypted; silently joining an open network would expose telemetry. Standard (IEEE 802.11, WPA2/WPA3; revision to confirm).
IFR-012 — BLE advertising and connection-parameter interface
Requirement Statement | The device shall operate BLE as a peripheral advertising a defined service UUID (IFR-015) with a configured advertising interval and TX power tuned to depot/last-mile range, shall accept connections and negotiate connection parameters (connection interval, slave latency, supervision timeout) sized for provisioning throughput while remaining energy-bounded, and shall constrain the advertising/ connectable window when idle/armed so the radio is not a standing exposure surface (supporting SYS-060).
Parent | SYS-060, SYS-107, PRD-CONN-003, PRD-MOB-001; details I-09 (device ↔ BLE app).
Allocation | S5 (BLE radio) ↔ S3 (session), S9 (app)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (advertising/connection parameters, range), Demonstration
Status | Baselined
Rationale | Advertising and connection parameters trade range and latency against energy and attack surface; a device that advertises connectably at all times both wastes energy and invites probing. Standard (Bluetooth Core / LE, advertising & connection parameters; revision to confirm).
IFR-013 — GNSS constellation and assist-data interface
Requirement Statement | The device shall configure the GNSS receiver for multi-constellation acquisition (GPS plus at least one of GLONASS/Galileo/BeiDou), shall use assist data (ephemeris/almanac, e.g. A-GNSS) delivered over the WAN/cloud link (I-10) when available to shorten time-to-fix and per-fix energy, and shall report per-attempt fix quality (fix type, satellites, HDOP) and degrade to last-known/best-available location with a quality flag when no fix is obtained inside the shipper.
Parent | SYS-117, SYS-029, PRD-CONN-005; details I-08 (GNSS RF + assist), I-10 (assist delivery).
Allocation | S4 (GNSS receiver) ↔ S3 (scheduling), S5 (assist)
Type | Detailed
Criticality | Medium
Safety/Security | None
Verification Method | Test (in-container time-to-fix, fix-rate), Analysis (energy per fix)
Status | Evidence-Pending (in-container fix availability, TTFF, per-fix energy — PRD-VER-005)
Rationale | GNSS is the most attenuated band inside a metallized shipper and each fix is energy-costly against D3; assist data and multi-constellation improve TTFF, and location must fail soft because it is attribution support, never a safety function. Evidence-Pending; closed by PRD-VER-005.
IFR-014 — In-container RF link contract (antenna interface, delivery-success)
Requirement Statement | Across the LTE-M/Wi-Fi/BLE/GNSS antenna interface, the device shall meet the in-container link budget defined by SYS-104/SYS-109 — including provision for a qualified IP-sealed external-antenna connector option — such that buffered telemetry and excursion alerts are delivered from inside a representative shipper at the delivery-success target set by the RF-in-container test, and shall degrade to store-and-forward buffering (SYS-111) rather than record loss when no link can be closed.
Parent | SYS-059, SYS-104, SYS-109, PRD-CONN-007, PRD-ENV-009; details I-08 (RF thru shipper).
Allocation | S5 (radios/antenna) ↔ S1 (antenna placement/aperture/seal)
Type | Detailed
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (RF-in-container, PRD-VER-005), Analysis (link budget)
Status | Evidence-Pending (attenuation, achievable margin, integral-vs-external antenna unresolved)
Rationale | RF-in-container is the principal open architectural risk; attenuation by the metallized shipper may force an external-antenna branch (AD-1). Until measured, upload/alert egress is never a dependency for the record (D2, SYS-013). Evidence-Pending; closed by PRD-VER-005.
---
4. BLE provisioning interface (GATT)
IFR-015 — BLE GATT service and characteristic structure
Requirement Statement | The device shall expose a versioned custom GATT service (128-bit vendor UUID) containing defined characteristics with fixed UUIDs and properties: Protocol/Schema-Version (read), Device-Info/Identity (read), Trip-Profile-Write (write, authenticated — IFR-017), Arm/Disarm-Control (write, authenticated — IFR-018), Local-Verdict/Record-Read (read/notify — IFR-019), and Status/Notification (notify — IFR-020), such that only these defined, version-negotiated operations are accepted and any write to an unknown handle or with an out-of-contract value is rejected with a defined GATT error.
Parent | SYS-061, PRD-CONN-006, PRD-MOB-002, PRD-ARCH-009; details I-09 (device ↔ BLE app).
Allocation | S5/S3 (GATT server) ↔ S9 (app)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Inspection (GATT table vs spec), Test (unknown-handle/malformed-write rejection)
Status | Baselined
Rationale | An undefined GATT surface is both an integration hazard and an attack surface; a published, versioned characteristic map makes the interface testable and lets a device reject a stale or rogue app. Standard (Bluetooth GATT; revision to confirm).
IFR-016 — BLE pairing, bonding, and security level
Requirement Statement | The device shall require LE Secure Connections pairing (ECDH-based) with a man-in-the-middle-protected association method (e.g. numeric comparison or passkey, not Just Works) before any authenticated characteristic (profile-write, arm/disarm, controlled read) may be accessed, shall bond and store the peer's keys for authenticated reconnection, and shall bind BLE-provisioning authority to the device's registered identity so an un-bonded or unauthorised central cannot perform privileged operations.
Parent | SYS-061, SYS-068, SYS-110, PRD-CONN-006, PRD-SEC-006; details I-09 (BLE security).
Allocation | S5/S3 (pairing) ↔ S9 (app), S8 (identity/keys)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (pairing method enforcement, MITM/Just-Works refusal), Analysis
Status | Preliminary (association method reconcile with cybersecurity threat model — Evidence-Pending)
Rationale | Provisioning writes must be authenticated so a stale or rogue app cannot silently disarm or misconfigure a device (AD-5, C9); MITM-protected LE Secure Connections is the control against passive/ active pairing attacks. Standard (Bluetooth Core LE Secure Connections; IEC 62443 concepts; revision to confirm); threat model Evidence-Pending.
IFR-017 — Trip-profile write characteristic: format, validation, and atomicity
Requirement Statement | Over the authenticated Trip-Profile-Write characteristic the device shall accept a trip profile in a defined, versioned encoding (thresholds, excursion logic, logging interval, upload cadence, alert routing reference) transferred as a length-prefixed, integrity-checked (CRC/hash) payload across GATT writes, shall validate the profile against constraints (e.g. low < high threshold, interval within 1–60 min) before acceptance, and shall apply it atomically — a rejected, incomplete, or failed-validation profile leaves the previously active profile unchanged and returns a defined error code.
Parent | SYS-061, SYS-022, SYS-025, PRD-MOB-002, PRD-SAFE-005; details I-09 (profile write).
Allocation | S3 (config store/validation) ↔ S9 (app)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (invalid-profile rejection, interrupted-write atomicity)
Status | Baselined
Rationale | A partially-written or invalid profile that became active could disable an excursion threshold and produce a false "accept" (SYS-025); atomic apply-or-reject with validation is the interface defence. Standard (IEC 62366-1 use-error prevention; revision to confirm).
IFR-018 — Arm/disarm control characteristic and confirmation
Requirement Statement | Over the authenticated Arm/Disarm-Control characteristic the device shall accept an explicit arm or disarm command carrying a nonce/sequence to prevent replay, shall execute arming only when pre-arm readiness passes (SYS-024: charge, calibration, firmware, profile, probe present), and shall return, and notify (IFR-020), the resulting authoritative state (armed/disarmed/refused-with-reason) so that the app's displayed state always reflects the device's actual state and a trip cannot silently start unmonitored or be silently disarmed.
Parent | SYS-061, SYS-023, SYS-024, SYS-068, PRD-FUN-015; details I-09 (arm control).
Allocation | S3 (arming logic) ↔ S9 (app)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (replay refusal, readiness-gated arm, state notification)
Status | Preliminary (reconcile with HAL — mis-arm hazard)
Rationale | Mis-arm (an unmonitored trip) is a top hazard (SYS-023); nonce-protected, readiness-gated arming with confirmed state is the interface control against both accidental and replayed arm/disarm. Judgment; forward to HAL (Stage 10).
IFR-019 — Local verdict/record read characteristic and notifications
Requirement Statement | The device shall expose the trip verdict (accept/reject/quarantine, with the incomplete-monitoring qualifier of SYS-017) and a summarised trip record over the Local-Verdict/Record-Read characteristic such that the login-free edge verdict read of the summary is available without pairing (AD-5) while the full controlled record read requires authentication (IFR-016), and shall stream a full record read as an ordered, resumable sequence of notifications/indications with an end-of-record marker so a truncated transfer is detectable and never presented as a complete record.
Parent | SYS-060, SYS-061, SYS-015, SYS-017, PRD-MOB-002, PRD-CONN-006; details I-09 (record read).
Allocation | S3 (record source) ↔ S9 (app)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (login-free summary, authenticated full read, truncation detection)
Status | Baselined
Rationale | BLE is the offline-read path where WAN is absent (D2); the summary verdict stays login-free for the last-mile user (AD-6) while the controlled record read is authenticated, and an end-of-record marker prevents a truncated read masquerading as complete. Judgment.
IFR-020 — BLE MTU, throughput, and notification contract
Requirement Statement | The device shall negotiate an ATT MTU (target ≥ 185 bytes, to confirm) and, where supported, LE Data Length Extension and 2M PHY to sustain the throughput needed for profile-write and record-read within the connection window, shall define the notification/indication flow-control (acknowledged indications for record chunks, credit/window to prevent buffer overrun), and shall bound each GATT operation with a defined timeout after which the operation is failed cleanly and the connection state left consistent (feeding IFR-034).
Parent | SYS-061, SYS-067, PRD-MOB-002, PRD-CONN-003; details I-09 (BLE transport).
Allocation | S5/S3 (GATT server) ↔ S9 (app)
Type | Detailed
Criticality | Medium
Safety/Security | Both
Verification Method | Test (MTU negotiation, throughput, flow-control, timeout)
Status | Preliminary (MTU/PHY targets to confirm against record size)
Rationale | Record reads and profile writes must complete inside the BLE connection budget without overrunning buffers; defined MTU, flow-control, and timeouts prevent both stalls and torn transfers (SYS-067). Standard (Bluetooth Core ATT/DLE/2M PHY; revision to confirm).
---
5. Device↔Cloud protocol (I-10)
IFR-021 — Transport and mutually-authenticated TLS binding
Requirement Statement | The device shall communicate with the cloud ingest endpoint over MQTT v5 (or HTTPS for record/OTA fetch) carried on TLS 1.2 or 1.3, with mutual authentication — the device presenting a unique per-device X.509 client certificate/key held in the secure element and validating the server certificate against a pinned CA/trust anchor — such that neither an unauthenticated device may ingest data nor a spoofed server may receive it, and a failed TLS handshake or certificate validation aborts the session and retains buffered data for retry (never falling back to cleartext).
Parent | SYS-063, SYS-068, SYS-110, PRD-CONN-008, PRD-SEC-005, PRD-CLOUD-001; details I-10 (transport).
Allocation | S3/S5 (device client) ↔ S8 (ingest endpoint)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (mTLS handshake, cert-pinning, spoofed-server/unauth-device refusal)
Status | Baselined
Rationale | Mutual authentication makes the device→cloud trust boundary real (AD-5): it blocks forged uploads and rogue endpoints and makes every ingest attributable (SYS-068, Part 11 attributable). Standard (MQTT v5, TLS 1.2+/1.3 with mutual auth; revision to confirm).
IFR-022 — Topic/endpoint structure and message schema versioning
Requirement Statement | The device↔cloud interface shall use a defined MQTT topic namespace (e.g. scm/{tenant}/{deviceId}/telemetry, /alert, /status, /cmd, /ota) — or the equivalent REST paths — and every message shall carry a schema-version field, deviceId, and monotonically increasing sequence number, such that both endpoints validate each message against the published schema for its declared version and reject or quarantine malformed, unknown-version, or out-of-contract messages rather than misinterpreting them, keeping fielded devices and the cloud interoperable across staged OTA version skew.
Parent | SYS-063, SYS-065, PRD-CONN-008, PRD-CLOUD-002, PRD-ARCH-009; details I-10 (schema/topics).
Allocation | S3 (device) ↔ S8 (cloud)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (schema validation, unknown-version quarantine), Inspection
Status | Baselined
Rationale | A safety-relevant record crossing a version boundary must never be silently mis-decoded; an explicit topic and schema-version contract is the defence against fleet/cloud skew during OTA rollouts (SYS-065). Standard (ISO/IEC/IEEE 29148 interface definition; MQTT v5 topics; revision to confirm).
IFR-023 — Telemetry message schema, encodings, and units
Requirement Statement | Each telemetry message shall encode records in a defined format (JSON or a compact binary/CBOR profile, to confirm) with fixed field names, explicit SI units and scaling (temperature in °C, RH in %, acceleration in g/m·s⁻², time in UTC/ISO-8601 or epoch-ms), the source channel identity per reading (internal vs external probe), the data-quality/gap flags (SYS-002/SYS-007), and a per-record integrity value, such that the cloud reconstructs the record without ambiguity of unit, channel, or time and can verify record integrity end-to-end.
Parent | SYS-063, SYS-006, SYS-026, PRD-CONN-008, PRD-SEC-002; details I-10 (message schema).
Allocation | S3 (device) ↔ S8 (cloud)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (round-trip encode/decode, unit/channel fidelity, integrity check)
Status | Evidence-Pending (JSON vs binary/CBOR encoding to confirm against RF payload budget)
Rationale | Unit or channel ambiguity in a temperature record corrupts disposition; explicit units, channel identity, and integrity per record preserve ALCOA+ accuracy and let the cloud detect corruption (SYS-026). Standard (ISO 8601 time; ALCOA+; revision to confirm).
IFR-024 — Store-and-forward semantics: ordering, idempotency, resume, acknowledgement
Requirement Statement | The device↔cloud interface shall provide store-and-forward delivery in which the device transmits buffered records in intact chronological order tagged with monotonic sequence numbers, the cloud ingests exactly once by de-duplicating on (deviceId, sequence) so a retransmitted record is never stored twice, an interrupted upload resumes from the last cloud-acknowledged sequence without restarting, and the device marks buffered data as safely delivered only upon receiving a positive application-level acknowledgement (MQTT v5 PUBACK/reason-code or REST 2xx) from the cloud.
Parent | SYS-062, SYS-020, SYS-111, PRD-CONN-008, PRD-CLOUD-002; details I-10 (store-and-forward).
Allocation | S3/S5 (device) ↔ S8 (idempotent ingest)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (link-drop/resume, duplicate-suppression, ack-gated delete)
Status | Baselined
Rationale | Intermittent links must never lose or duplicate the evidentiary record (D2); ordering + idempotency + ack-gated deletion preserve ALCOA+ completeness and consistency across any number of interruptions (SYS-111). Standard (MQTT v5 QoS/ack semantics; ALCOA+; revision to confirm).
IFR-025 — Excursion-alert message contract and prioritisation
Requirement Statement | The device shall emit an excursion/shock alert as a distinct, high-priority message (dedicated topic/endpoint, delivered with acknowledged QoS) carrying event type, direction, onset time, magnitude/duration, deviceId, and best-available location, shall send it ahead of any bulk telemetry backlog on the selected bearer, and shall require a positive cloud acknowledgement of the alert before treating it as delivered — while the same event is always already committed locally (SYS-013) independent of this path.
Parent | SYS-062, SYS-112, SYS-011, PRD-CONN-009, PRD-CLOUD-003; details I-10 (alert path).
Allocation | S3/S5 (device) ↔ S8 (alert ingest/routing)
Type | Detailed
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (alert precedence over backlog, ack, latency)
Status | Preliminary (reconcile with HAL — delayed-alert hazard)
Rationale | A rescue-enabling alert must never queue behind bulk data (D2, SYS-112); a dedicated, acknowledged, prioritised alert contract preserves intervention latency, and local capture (SYS-013) guarantees the event exists even if egress is delayed. Judgment; forward to HAL.
IFR-026 — Back-pressure and command/config channel contract
Requirement Statement | The device↔cloud interface shall define bidirectional flow control such that the cloud can signal back-pressure (throttle/backoff, e.g. MQTT v5 receive-maximum / reason codes) and the device honours it with bounded exponential backoff and jitter, and the cloud→device command/config channel (/cmd) shall deliver only authenticated, sequenced, idempotent commands (config push, cadence change, location request) that the device validates and acknowledges, rejecting replayed or out-of-contract commands, so neither side can overrun the other or act on a spoofed instruction.
Parent | SYS-062, SYS-067, SYS-068, PRD-CONN-008, PRD-CLOUD-002; details I-10 (back-pressure/cmd).
Allocation | S3/S5 (device) ↔ S8 (cloud)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (back-pressure honouring, replayed-command refusal)
Status | Baselined
Rationale | Under alert bursts or fleet-wide reconnect storms, unbounded ret/publish rates are both a cloud-overload and an energy hazard (SYS-067); authenticated idempotent commands stop a spoofed configuration push (SYS-068). Standard (MQTT v5 flow control; revision to confirm).
---
6. Cloud↔Enterprise API (I-11)
IFR-027 — REST API for record export and fleet query (OpenAPI contract)
Requirement Statement | The cloud shall expose a documented REST API, specified by a published OpenAPI definition, providing resource endpoints to query and export trip records, alerts, and fleet/device status in a versioned, machine-readable format (JSON, with PDF/CSV export of records) that preserves attribution, time, gaps, and tamper-evidence, such that authorised enterprise/QMS systems can retrieve the compliance record and its integrity metadata programmatically against a stable contract.
Parent | SYS-064, SYS-027, PRD-CLOUD-005, PRD-CLOUD-010, PRD-FUN-018, PRD-REG-003; details I-11.
Allocation | S8 (cloud API) ↔ enterprise/QMS
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (endpoint conformance to OpenAPI), Demonstration (export integrity)
Status | Baselined
Rationale | The export/API boundary is where the compliance value reaches ERP/QMS for release decisions; an OpenAPI contract makes it stable, documented, and testable while preserving ALCOA+ properties. Standard (OpenAPI; EU GDP, 21 CFR Part 11; revision to confirm).
IFR-028 — Enterprise API authentication (OAuth 2.0 / SSO) and RBAC
Requirement Statement | The enterprise API shall authenticate every request via OAuth 2.0 (bearer access tokens, scoped to least privilege) integrated with the customer's enterprise SSO/identity provider, shall enforce role-based access control on every resource so a caller can access only its tenant's data at its authorised role, and shall audit-log every access (who, what, when) — while this control shall not gate the login-free on-device edge verdict read (AD-5).
Parent | SYS-064, SYS-068, PRD-CLOUD-009, PRD-SEC-005, PRD-SEC-006; details I-11 (API auth).
Allocation | S8 (cloud auth/RBAC) ↔ enterprise IdP
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (token scope/expiry, RBAC enforcement, cross-tenant refusal, audit-log)
Status | Baselined
Rationale | The record is regulated data; scoped OAuth 2.0 with enterprise SSO and RBAC is the control against unauthorised or cross-tenant access, and audit-logging makes access attributable (Part 11). Standard (OAuth 2.0; 21 CFR Part 11 attributable; revision to confirm).
IFR-029 — API rate limits, pagination, and webhook/notification integration
Requirement Statement | The enterprise API shall enforce documented per-client rate limits (returning a defined 429 with Retry-After on exceedance), shall paginate large result sets with stable cursor-based paging so a full export can be retrieved deterministically without loss or duplication across pages, and shall offer a webhook/subscription mechanism delivering signed, retryable alert/status notifications to registered enterprise endpoints, such that integrators can both pull records and receive pushed events reliably.
Parent | SYS-064, SYS-067, PRD-CLOUD-010, PRD-CLOUD-003; details I-11 (API scaling/webhooks).
Allocation | S8 (cloud API) ↔ enterprise/QMS
Type | Detailed
Criticality | Medium
Safety/Security | Security
Verification Method | Test (rate-limit 429, cursor-paging completeness, webhook retry/signature)
Status | Baselined
Rationale | Deterministic pagination guarantees a complete export (ALCOA+ complete); rate limits and retryable signed webhooks protect availability and let enterprise alert routing (SYS-012) integrate without polling. Standard (OpenAPI, HTTP 429/Retry-After; revision to confirm).
IFR-030 — Enterprise API schema versioning and backward compatibility
Requirement Statement | The enterprise API shall carry an explicit version (URI path or media-type version), shall make only backward-compatible additive changes within a major version, and shall support at least one prior major version for a defined deprecation window with published change logs, such that an integrated enterprise system is not broken by a cloud release and record schemas remain interpretable over the retention lifetime.
Parent | SYS-064, SYS-063, PRD-CLOUD-010, PRD-CLOUD-011; details I-11 (API versioning).
Allocation | S8 (cloud API) ↔ enterprise/QMS
Type | Detailed (derived — lifecycle/compatibility necessity)
Criticality | Medium
Safety/Security | Both
Verification Method | Inspection (versioning policy), Test (prior-version compatibility)
Status | Baselined
Rationale | Records must stay interpretable for their full regulated retention (PRD-CLOUD-011) and integrations must survive cloud releases; explicit versioning with a deprecation window is the governance that prevents silent contract breakage. Standard (ISO/IEC/IEEE 29148; OpenAPI versioning; revision to confirm).
---
7. OTA interface (I-12)
IFR-031 — OTA update package format and manifest
Requirement Statement | The OTA service→device interface shall deliver each update as a defined package comprising the firmware image(s), a metadata manifest (target device/hardware model, from/to firmware versions, image size, and cryptographic digest per image), and a signature block, such that the device can determine applicability and integrity from the manifest before writing any image, and shall reject a package whose manifest does not match the device's hardware/version constraints.
Parent | SYS-065, SYS-114, PRD-CONN-011, PRD-CLOUD-008, PRD-SEC-008; details I-12 (package/manifest).
Allocation | S8 (OTA service) ↔ S3 (device updater)
Type | Detailed
Criticality | High
Safety/Security | Both
Verification Method | Test (manifest-mismatch rejection, applicability check)
Status | Baselined
Rationale | A manifest lets the device refuse a wrong-model or downgrade package before touching flash, preventing a mismatched image from bricking a device (AD-7). Standard (secure-update principles; revision to confirm).
IFR-032 — OTA image authenticity/integrity verification and signature
Requirement Statement | The device shall cryptographically verify each OTA image's signature against a provisioned public key/trust anchor held in the secure element and verify the image digest before the image is activated, such that only firmware whose authenticity and integrity are proven on-device is accepted; a failed signature or digest check shall abort the update, discard the candidate image, and leave the running known-good image untouched.
Parent | SYS-065, SYS-032, SYS-068, PRD-SEC-008, PRD-SAFE-010; details I-12 (signature verify).
Allocation | S8 (signing) ↔ S3 (secure boot/verify)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (tampered-image rejection, invalid-signature abort), Analysis
Status | Baselined
Rationale | The OTA boundary is a fleet-wide integrity surface; unsigned or tampered firmware is a route to mass compromise, so on-device signature+digest verification with a hardware-held key is the mandatory control (SYS-068, AD-7). Standard (secure-boot / signed-update, IEC 62443 concepts; revision to confirm).
IFR-033 — OTA resumable transfer, fail-safe apply, and status reporting
Requirement Statement | The OTA interface shall support resumable download (byte-range/offset resume) over LTE-M/Wi-Fi and USB-C so an interrupted transfer continues without restarting, shall apply the update fail-safe (A/B or staged image with automatic rollback to the known-good image if activation or post-update self-test fails) preserving trip data, and shall report update status (downloading, verified, applied, rolled-back, failed-with-reason) to the OTA service (I-10) so the fleet campaign (PRD-CLOUD-008) observes per-device outcome and no device is left bricked or unmonitored-without-notice.
Parent | SYS-065, SYS-032, SYS-114, PRD-CONN-011, PRD-CLOUD-008, PRD-SAFE-010; details I-12/I-10.
Allocation | S3 (updater) ↔ S5 (transport), S8 (campaign)
Type | Detailed
Criticality | Critical
Safety/Security | Both
Verification Method | Test (interrupted-resume, failed-apply rollback, status reporting)
Status | Baselined
Rationale | Field updates over lossy links must not waste energy on restarts or brick a device; A/B rollback plus status reporting makes updates recoverable and observable, and keeps monitoring status honest (D2, AD-7). Standard (secure-update / A-B rollback principles; revision to confirm).
---
8. Interface governance
IFR-034 — Per-interface error, timeout, and retry-with-backoff contract
Requirement Statement | Every data, RF, and power interface (I-04, I-05, I-06, I-08, I-09, I-10, I-12) shall have a documented error/timeout/retry contract specifying bounded operation timeouts, defined error codes/return values, retry counts with exponential backoff and jitter, and the safe fallback on exhaustion (buffer, degrade, or safe-state), such that a slow, unresponsive, overloaded, or disconnected counterpart produces a defined, energy-bounded outcome rather than an unbounded stall, retry storm, silent data loss, or battery drain.
Parent | SYS-067, PRD-REL-004, PRD-CONN-008, PRD-ARCH-008; details error contract across I-04…I-12.
Allocation | SYS (all interface owners)
Type | Detailed (derived — reliability/energy governance)
Criticality | High
Safety/Security | Both
Verification Method | Test (fault injection per interface), Analysis (energy bound)
Status | Preliminary (reconcile with HAL)
Rationale | An interface that can hang or spin retries is both a data-integrity and an energy hazard under the D3 budget; a uniform timeout/backoff/fallback contract makes every boundary fail predictably to a safe state (SYS-067). Judgment.
IFR-035 — ICD conformance to the typed interface inventory and no undocumented interface
Requirement Statement | Every realised subsystem interface shall conform to this ICD and the typed interface inventory (I-01…I-12) — matching the declared type (PWR/DATA/RF/THRM/MECH), direction, signals, format, and contract — each IFR shall trace to at least one SYS/PRD parent and its I-0N interface, and no undocumented interface shall exist between subsystems or across the device/cloud/mobile trust boundaries; any interface change shall be governed through ICD revision with backward-compatibility assessment.
Parent | SYS-066, SYS-063, PRD-ARCH-009, PRD-ARCH-001; details I-01…I-12 (all interfaces).
Allocation | SYS (interface integrator, EM-ENG-001)
Type | Detailed (derived — governance)
Criticality | High
Safety/Security | Both
Verification Method | Inspection (ICD vs block diagram vs design outputs), Analysis
Status | Baselined
Rationale | Interface drift between disciplines is a leading source of integration defects, and an undocumented interface is an ungoverned attack surface; a single authoritative ICD conformant to the block diagram keeps requirements, design, and test in agreement (SYS-066, rule 2). Standard (ISO/IEC/IEEE 29148; revision to confirm).
Subsystem / Detailed Requirements — Mechanical (MER-001 … MER-040)
SCM · Stage 07 · Subsystem: S1 Enclosure/Mechanical. Owner: Principal Mechanical Engineer (with Reliability/Materials, Sensing/Thermal, Display/HMI, and Power contributors; integrator EM-ENG-001). Detailed mechanical requirements for the SCM device housing, seals, external temperature probe, mounting/ retention, display window, controls, materials, and markings. Each requirement decomposes one or more system requirements (SYS-051…085 mechanical/thermal/interface, plus SYS-003/015/029 where mechanically realised) and/or the mechanical product requirements (PRD-MECH-001…012, PRD-ENV-001…012). Field set mirrors the SYS library ([../06-system-requirements/index.md](../06-system-requirements/index.md)); based on D1 (reusable/cleaned), D2 (offline-first edge verdict), D3 (−30 °C / 90-day), D4 (multi-market, air transport). Forward to mechanical Design Outputs (DO-MECH), Hazard Analysis (HAL), ERFKL (DFMEA), and the Verification Plan (PVL).
Allocation legend: all requirements allocate to S1 (Enclosure/Mechanical); the specific mechanical part is named — housing · seal/gasket · port door · probe (tip / lead / gland) · vent · mount/retention · window · button · label/marking · fastener.
Evidence posture (honesty): environmental/ingress/robustness standards are cited by designation only; exact clause, revision, and test severity are to confirm and are set in the PVL. Numeric targets that are not yet fixed (IP grade rationale, drop height, vibration profile, probe offset/lag budget, envelope set, mass, seal cycle life) are carried Evidence-Pending. Per CLAUDE.md rule 3, risk is severity-first (Action Priority): any mechanical feature that can silently corrupt the evidentiary record or breach the ingress/thermal integrity of the device — sealing, probe thermal coupling, seal/probe strain relief — is scored Critical/High.
---
1. Enclosure & fit
MER-001 — Device envelope fits the target insulated-shipper set
Requirement Statement | The device housing (excluding the probe lead) shall occupy a bounding envelope no larger than the maximum device envelope defined in the product spec for the target set of standard insulated vaccine shippers and cold boxes, such that it can be placed within the internal usable volume of every shipper in that set without displacing payload beyond the product-spec allowance.
Parent | SYS-083, SYS-052; PRD-MECH-001, PRD-MECH-012
Allocation | S1 — housing (overall envelope)
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Inspection (fit check in representative shippers), Demonstration
Status | Evidence-Pending (shipper envelope set and maximum device dimensions to confirm)
Rationale | The device is placed inside the payload; if it does not fit the smaller shippers in the range it will not be used, and if it displaces payload it degrades cold-chain economics and acceptance. The bounding envelope is the mechanical precondition for SYS-083 fit and SYS-052 retention. Judgment; the definitive shipper set and the resulting envelope are Evidence-Pending, set from the target-market fleet (D1/D4) — coupled to the embeddable-module interface (MER-006).
MER-002 — Device mass limit
Requirement Statement | The device (housing, battery, probe and lead as shipped, excluding external packaging) shall not exceed the mass limit set in the product spec.
Parent | SYS-084; PRD-MECH-002
Allocation | S1 — housing (and S2 battery mass, flagged)
Type | Allocated
Criticality | Low
Safety/Security | None
Verification Method | Inspection (weigh sample against limit)
Status | Evidence-Pending (target mass to confirm)
Rationale | Mass burdens gloved handling and displaces payload, and is in direct tension with the D3 energy budget — the low-temperature 90-day cell pushes toward a larger, heavier battery. Mass is therefore a shared mechanical/power constraint, not a mechanical nicety. Judgment; the target mass is Evidence-Pending, set against battery sizing (SYS-086-line) — flag Power engineer.
MER-003 — Low-profile, snag-free form for in-payload placement
Requirement Statement | The housing shall present a low-profile, rounded, snag-free external form with no sharp edges, deep unsealed crevices, or protruding features (other than the defined probe lead and mount) that could damage payload/packaging, trap contamination, or catch on gloves during placement and retrieval.
Parent | SYS-083, SYS-052; PRD-MECH-001, PRD-MECH-008
Allocation | S1 — housing (external form)
Type | Derived
Criticality | Medium
Safety/Security | None
Verification Method | Inspection, Demonstration (gloved placement)
Status | Preliminary
Rationale | The device sits against vaccine packaging and is handled with cold, gloved hands; sharp or snagging features risk puncturing payload, harbouring soil that defeats cleaning (D1), or slowing placement. Rounded, crevice-free geometry also supports the IP seal and cleanability (MER-033). Judgment (design-for-cleanability and handling ergonomics).
MER-004 — Housing structural wall thickness and stiffness
Requirement Statement | The housing walls, bosses, and load paths shall be dimensioned so that the enclosure retains its sealing-surface geometry and protects the electronics, display, battery, and probe interface under the shock, drop, vibration, and clamping loads of MER-039/MER-040 without permanent deformation that would breach the seal (MER-007) or shift a sensor.
Parent | SYS-077, SYS-076; PRD-MECH-007
Allocation | S1 — housing (walls/ribs/bosses)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Analysis (structural/FEA), Test (post-drop seal + accuracy)
Status | Evidence-Pending (wall-thickness/material set with drop and vibration severities)
Rationale | A wall that flexes enough to open the gasket line under a drop, or a boss that yields under clamp/retention load, produces the dangerous case — an intact-looking device with a compromised seal or a displaced sensor that reports a plausible but wrong record (SYS-077 rationale). Structural adequacy is therefore a data-integrity requirement, not just survival. Judgment; wall thickness is Evidence-Pending, set jointly with material selection (MER-032) and the confirmed drop/vibration severities.
MER-005 — External probe lead length and tip form for product-zone placement
Requirement Statement | The device shall provide an external temperature probe with a lead length and tip form that allow the tip to be placed in the product zone of the target shippers while the device body remains in its retained position, per the lead length and tip geometry set in the product spec.
Parent | SYS-051, SYS-003; PRD-MECH-004, PRD-MECH-011
Allocation | S1 — probe (tip + lead)
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test, Inspection
Status | Evidence-Pending (lead length and tip geometry to confirm)
Rationale | The external probe is the channel that represents the actual vaccine thermal zone (SYS-003, SYS-051); too short a lead forces an unrepresentative placement, and a poorly formed tip couples slowly or to the wrong mass. Lead/tip geometry is the mechanical enabler of a defensible product-zone reading. Judgment; lead length is Evidence-Pending, set against the shipper set (MER-001) and the thermal-coupling budget (MER-015) — coupled to Sensing/Thermal.
MER-006 — Embeddable-module mechanical mounting interface (OEM)
Requirement Statement | The housing shall provide a defined, documented mechanical mounting interface (mounting features, datum, and keep-out envelope) that allows OEM partners to embed the device into their shippers/containers as a standard module without tools other than those specified in the integration guide.
Parent | SYS-052, SYS-083; PRD-MECH-012
Allocation | S1 — housing (mount interface)
Type | Allocated
Criticality | Low
Safety/Security | None
Verification Method | Inspection (against interface spec)
Status | Evidence-Pending (interface spec to confirm)
Rationale | Designed-in OEM embedding is a business channel (PRD-MECH-012) and, when used, becomes the retention means (MER-022); an undefined mounting interface produces ad-hoc integrations that break thermal coupling and retention. A published mechanical contract keeps embedded installs representative. Judgment; the interface spec is Evidence-Pending.
2. Sealing & ingress
MER-007 — Enclosure ingress protection IP67 (closed port)
Requirement Statement | The fully assembled enclosure, including all seams, the display window, the probe entry, and the closed charge/service port, shall meet an ingress protection rating of at least IP67 per IEC 60529 (dust-tight; 30 min immersion at up to 1 m), and shall retain that rating across the reuse life of the device.
Parent | SYS-078; PRD-MECH-005, PRD-MECH-006, PRD-ENV-007
Allocation | S1 — housing + seal/gasket + port door + probe gland + window
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (IEC 60529 IP6X + IPX7), Inspection
Status | Evidence-Pending (IP grade vs. cleaning method — confirm IP67 sufficient or IP68 required)
Rationale | Dust at last-mile posts and moisture (including cleaning immersion, condensation, and the humid tropical lane) must not reach the electronics or sensors over a multi-trip life; ingress is a corrosion, drift, and safety path. Standard (IEC 60529, revision to confirm). IP67 is a Judgment target and is Evidence-Pending: if the confirmed service procedure (PRD-SVC-002) requires routine immersion cleaning rather than wipe-down, IP68 may be required — flag Materials/Reliability engineer; the USB-C port (MER-009) is the governing ingress path.
MER-008 — Primary enclosure seal design (compression gasket in captured groove)
Requirement Statement | The main enclosure joint shall be sealed by a captured, compression-type elastomer gasket (or equivalent) held in a defined groove at a compression within the seal material's specified working range across the −30 °C to +60 °C operating range, such that the seal maintains the MER-007 rating without relying on adhesive as the primary seal.
Parent | SYS-078, SYS-069; PRD-MECH-005
Allocation | S1 — seal/gasket (main joint)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Analysis (gland/compression), Test (IP after thermal cycling)
Status | Evidence-Pending (gasket material, groove geometry, compression set with materials)
Rationale | A captured compression gasket in a designed groove is the serviceable, reuse-tolerant way to hold IP67 across thermal cycling and re-opening for service; an adhesive-only or unconstrained seal cannot be re-established after service and takes compression set at cold. Elastomer stiffens and loses recovery at −30 °C, so the working-compression window must include the cold extreme. Standard (IEC 60529 governs the rating; gland design is Judgment); gasket material and geometry are Evidence-Pending, coupled to MER-032 (cold-tough materials) and MER-038 (serviceable fasteners) — flag Materials engineer.
MER-009 — Sealed USB-C charge/service port (captive door or sealed connector)
Requirement Statement | The USB-C charge/service port shall be sealed, when closed, by a captive port door with an integral gasket or by an inherently sealed connector, such that the closed port maintains the MER-007 IP rating; the seal shall be operable by a gloved hand and the closure element shall be captive (not separable and losable) over the reuse life.
Parent | SYS-078, SYS-053, SYS-054; PRD-MECH-006, PRD-ENV-007
Allocation | S1 — port door (+ housing port boss)
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (IP after open/close cycling), Demonstration (gloved operation)
Status | Evidence-Pending (closure type and open/close cycle life to confirm)
Rationale | The charge port is the most likely ingress path over repeated use and cleaning (PRD-MECH-006, SYS-078 rationale); a non-captive plug will be lost and the port left open, defeating IP protection at exactly the point where charging brings the device into wet depot conditions. A captive, gloved-operable seal keeps the rating recoverable in the field. Judgment; closure type and cycle life are Evidence-Pending, set with the charge path (SYS-053) and service procedure.
MER-010 — Probe lead sealing and strain relief at the enclosure entry
Requirement Statement | The external-probe lead shall enter the enclosure through a sealed, strain-relieved gland that maintains the MER-007 IP rating and shall withstand the pull, flex, twist, and cleaning loads of the reuse life without seal breach at the entry, conductor fatigue, or a shift of the tip out of its representative position.
Parent | SYS-078, SYS-076, SYS-051; PRD-MECH-004
Allocation | S1 — probe (gland / strain relief at entry)
Type | Derived
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (pull/flex + IP), Analysis
Status | Evidence-Pending (pull/flex load spec and cycle count to confirm)
Rationale | The probe entry is simultaneously an ingress path and a mechanical stress concentration; a lead that is yanked during retrieval or fatigued by handling can breach the seal (feeding the failure the whole IP scheme defends against) or intermittently open the sensor circuit and corrupt the product-zone channel. Because failure here silently corrupts the safety-relevant probe reading, it is severity-first Critical. Standard (strain-relief per connector/cable practice; revision to confirm); loads and cycle count are Evidence-Pending, set with the vibration/handling profile (SYS-076) — flag Sensing engineer.
MER-011 — Pressure equalisation / vent for altitude and thermal cycling
Requirement Statement | The sealed enclosure shall incorporate a pressure-equalisation means (e.g. a hydrophobic, IP-rated protective vent membrane) that limits the internal-to-external pressure differential across the cargo-hold altitude range and rate of SYS-082 and across cold-to-warm thermal transitions, without degrading the MER-007 ingress rating.
Parent | SYS-082, SYS-078, SYS-069; PRD-ENV-012
Allocation | S1 — vent (membrane) in housing
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (altitude/pressure-cycle + IP), Analysis
Status | Evidence-Pending (vent selection and differential limit to confirm)
Rationale | A fully sealed IP67 box develops a pressure differential on aircraft-hold depressurisation (SYS-082) and on the repeated cold-to-warm transitions of reuse; unrelieved, that differential stresses the gasket and port door, can suck moisture past a seal on cool-down, and stresses the display window. A hydrophobic vent equalises pressure while preserving IP. Standard (IEC 60529 rating maintained; altitude per UN 38.3 T.1 / DO-160 low-pressure concept, revision to confirm); vent selection is Evidence-Pending — coupled to condensation management (MER-012).
MER-012 — Internal condensation and moisture management
Requirement Statement | The enclosure internal design shall manage condensation arising from the repeated cold-to-warm thermal cycling of reuse (e.g. by desiccant, conformal protection of critical circuitry, and/or vent-assisted drying) such that condensation forming inside the enclosure does not cause corrosion, leakage current, sensor drift, or a false reading over the reuse life.
Parent | SYS-073, SYS-071; PRD-ENV-004
Allocation | S1 — housing internal (with S3/S4 board protection, flagged)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (thermal-humidity cycling), Analysis
Status | Evidence-Pending (condensation-control method to confirm)
Rationale | A sealed device carries trapped air across the dew point on every trip (D1, SYS-073); repeated internal condensation is a cumulative corrosion and leakage-current mechanism that can drift the sensing or fabricate a reading. Sealing keeps liquid water out but does not manage water that condenses from trapped humid air, so internal moisture control is a distinct requirement. Judgment (design-for-condensation); method is Evidence-Pending, coupled to the vent (MER-011) and cyclic-condensation test (SYS-073) — flag Reliability/Materials engineer.
MER-013 — Seal retention over the reuse life (compression set / re-seal)
Requirement Statement | The enclosure and port seals shall retain sufficient compression and recovery to maintain the MER-007 ingress rating across the number of service open/close cycles and the thermal-cycle count defined for the reuse life, and any seal element with a limited service life shall be replaceable during scheduled service.
Parent | SYS-078, SYS-073; PRD-MECH-005, PRD-MECH-009; D1
Allocation | S1 — seal/gasket + port door + probe gland
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (aged/cycled IP), Inspection
Status | Evidence-Pending (reuse-life cycle count and seal service interval to confirm)
Rationale | IP67 verified once on a new unit does not prove IP67 after years of reuse (D1): elastomers take compression set — accelerated at the cold extreme (MER-008) — and every service opening disturbs the seal. Either the seal must be qualified to the full cycle count or it must be a defined replaceable service item. Standard (IEC 60529 rating maintained across life; compression-set behaviour per material data); cycle count and service interval are Evidence-Pending, set with the service procedure (PRD-SVC-002).
MER-014 — Sealed, blind, or gasketed fastener and seam penetrations
Requirement Statement | All housing fastener penetrations, seams, and feature transitions that cross the sealed boundary shall be blind, gasketed, or otherwise sealed so that no fastener or seam constitutes an ingress path, and no through-hole fastener shall penetrate the sealed volume without its own seal.
Parent | SYS-078; PRD-MECH-005
Allocation | S1 — housing (seams) + fastener (penetrations)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Inspection, Test (IP)
Status | Preliminary
Rationale | IP67 is only as good as its weakest penetration; an unsealed screw boss or a seam that bridges inside-to-outside quietly defeats the primary seal (MER-008) no matter how good the gasket is. Blind bosses and gasketed/captured seams remove these leak paths while keeping the device serviceable (MER-038). Judgment (sealed-enclosure design practice).
3. Thermal design
MER-015 — External-probe thermal coupling to the product zone (offset/lag budget)
Requirement Statement | The external probe (tip, tip-to-sensor path, and lead) shall be designed so that the sensed temperature tracks the product-zone temperature within the thermal offset and time-constant (lag) budget allocated to meet the PRD-PERF-001 accuracy target across the −30 °C to +60 °C range, and the achieved offset and time constant shall be characterised and documented.
Parent | SYS-051, SYS-071, SYS-003; PRD-MECH-011, PRD-MECH-004
Allocation | S1 — probe (thermal path)
Type | Allocated
Criticality | Critical
Safety/Security | Safety
Verification Method | Test (thermal-response characterisation), Analysis
Status | Evidence-Pending (offset/time-constant budget to confirm)
Rationale | Every excursion verdict is built on the probe reading; if the probe lags or offsets from the true product-zone temperature, a real freeze excursion can be missed or a benign transient flagged — the highest-consequence mechanical failure (SYS-051, SYS-071 rationale). Thermal coupling is a mechanical property (tip mass, contact, conduction path) that the sensor accuracy cannot recover. Severity-first Critical. Judgment; the offset/time-constant budget is Evidence-Pending, set with the performance accuracy budget and closed by thermal-characterisation test — coupled to Sensing/Thermal (SYS-071).
MER-016 — Low-thermal-mass probe tip for responsiveness
Requirement Statement | The probe tip shall present a low thermal mass and a defined conduction path to the sensing element so that the probe's thermal time constant meets the MER-015 lag budget, without a tip form so fragile that it fails the handling and cleaning loads of MER-010.
Parent | SYS-051, SYS-071; PRD-MECH-011
Allocation | S1 — probe (tip)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (step-response), Analysis
Status | Evidence-Pending (tip thermal mass / time constant to confirm)
Rationale | A heavy or thermally isolated tip slews slowly and blurs short excursions, defeating the responsiveness the excursion logic assumes; but an over-minimised tip becomes fragile and fails strain relief. The tip design is the trade between thermal responsiveness (MER-015) and mechanical robustness (MER-010). Judgment; tip mass/time constant is Evidence-Pending, set with MER-015.
MER-017 — Internal-sensor thermal path to the enclosure micro-environment
Requirement Statement | The internal temperature/humidity sensor shall be mounted with a defined thermal (and, for humidity, vapour) path to the enclosure internal micro-environment it is intended to represent, and thermally decoupled from localised board hot-spots, so that the internal channel represents the device ambient rather than a nearby component.
Parent | SYS-051, SYS-071, SYS-003; PRD-MECH-011
Allocation | S1 — housing (sensor thermal path)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (thermal), Analysis
Status | Preliminary
Rationale | The internal channel is a distinct, independently identifiable reading (SYS-003) representing the enclosure micro-environment; if it is mounted against a warm component or starved of the air/vapour it should sense, it reports the board, not the environment, and corrupts the dual-channel record. Sensor placement and its conduction/vapour path are mechanical decisions. Judgment (sensor-integration practice) — coupled to Sensing (SYS-071).
MER-018 — Isolation of sensing from device self-heating
Requirement Statement | The mechanical layout shall thermally isolate the temperature-sensing elements (internal sensor and probe reference) from the self-heating sources in the device (radio, PMIC/charge circuit, MCU) by conduction path, spacing, and/or thermal barrier, so that internal self-heating does not push a sensed temperature outside the allocated accuracy budget in any active power state.
Parent | SYS-071, SYS-051; PRD-MECH-011
Allocation | S1 — housing (layout / thermal barrier)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (self-heat under load), Analysis
Status | Evidence-Pending (self-heating error budget allocation to confirm)
Rationale | During upload the radio and charge circuit dissipate heat close to the sensors; an uncompensated self-heat rise reads as a false warming and can mask a genuine cold excursion or fabricate a heat excursion (SYS-071 rationale). Duty-cycling (AD-2, SYS-055) helps but the worst case is a burst upload during a cold hold — a mechanical isolation problem. Judgment; the self-heat error allocation is Evidence-Pending, set with the accuracy budget — coupled to Power/Sensing.
MER-019 — Charging thermal path and cold-charge safety support
Requirement Statement | The mechanical design shall provide a defined thermal path for the battery and charge circuit that limits localised temperature rise during charging and supports the BMS cell-temperature sensing on which the SYS-053 low-temperature charge inhibit depends, without creating a hot-spot that biases the temperature sensors (MER-018).
Parent | SYS-053, SYS-069; PRD-ENV-012, PRD-MECH-007
Allocation | S1 — housing (battery/charge thermal path)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (charge thermal), Analysis
Status | Evidence-Pending (thermal path / cell-temp sensing location with Power)
Rationale | The safety control against charging a cold Li-ion cell (lithium plating / thermal event) is the BMS charge inhibit (SYS-053), which relies on a representative cell-temperature measurement — a mechanical placement problem. Poor charge-thermal design either defeats that sensing or creates a hot-spot that biases the environmental sensors. Battery-safety-relevant, hence High. Standard (Li-ion charge-temperature limits per cell datasheet; IEC 62368-1 thermal; revision to confirm); path and sensing location are Evidence-Pending — flag Power engineer.
MER-020 — Enclosure thermal-mass trade for representative response
Requirement Statement | The enclosure thermal mass and its coupling to the internal sensor shall be selected so that the internal channel's response to a genuine change in the enclosure micro-environment meets the responsiveness needed by the excursion logic, while providing adequate protection of the battery's usable low-temperature performance, and the resulting trade shall be documented.
Parent | SYS-071, SYS-069, SYS-051
Allocation | S1 — housing (thermal mass)
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Analysis, Test
Status | Preliminary
Rationale | Enclosure thermal mass is a genuine trade: more mass buffers the battery against short cold dips (helping the D3 budget) but slows the internal channel's response and blurs excursion timing; less mass responds faster but exposes the cell. The trade must be made deliberately and recorded, not left to arbitrary wall thickness. Judgment — coupled to Power (battery low-temp performance) and Sensing.
MER-021 — Probe-tip condensation and contamination shedding
Requirement Statement | The probe tip and its immediate lead shall be formed so that condensation and cleaning fluid drain/shed rather than pool at the sensing surface, such that a film of condensed or residual liquid on the tip does not persist long enough to masquerade as a thermal reading over the cold-to-warm transitions of reuse.
Parent | SYS-073, SYS-051; PRD-ENV-004, PRD-MECH-004
Allocation | S1 — probe (tip form)
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Test (condensation), Inspection
Status | Preliminary
Rationale | Condensation forming on the probe during cold-store-to-ambient transitions is a named concern (SYS-073, PRD-ENV-004): a pooled water film has its own thermal behaviour and evaporative cooling and can bias the tip reading. A shedding tip form limits how long liquid dwells on the sensing surface. Judgment (design-for-condensation) — coupled to the cyclic-condensation test (SYS-073).
4. Mounting & retention
MER-022 — Tool-free retention inside the shipper
Requirement Statement | The device shall provide a defined, tool-free means to retain the device body and the probe tip in their intended measurement position and orientation inside the target shippers for the full trip, such that neither shifts out of position under the transport shock and vibration of SYS-076/SYS-077.
Parent | SYS-052, SYS-076; PRD-MECH-003
Allocation | S1 — mount/retention
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (vibration/shock retention), Demonstration
Status | Evidence-Pending (retention means and holding force to confirm)
Rationale | A device that shifts, floats, or must be wedged in changes both its thermal coupling (MER-015) and its shock signature, and can be lost in the box or damage payload; retention is the mechanical precondition for a representative record (SYS-052 rationale). Tool-free matters because placement is done gloved at the last mile with no workshop. Judgment; the retention means and holding force are Evidence-Pending, set against the shipper set (MER-001) and vibration profile (SYS-076).
MER-023 — Captive tether / clip to prevent loss and misplacement
Requirement Statement | The device shall provide a captive attachment feature (tether point, clip, or equivalent) allowing it to be secured to the shipper or an internal fixture so that it is not left loose in the payload, subject to the feature not defeating the shipper's thermal barrier or the device's ingress seal.
Parent | SYS-052, SYS-031; PRD-MECH-003, PRD-MECH-012
Allocation | S1 — mount/retention (tether/clip feature)
Type | Derived
Criticality | Medium
Safety/Security | None
Verification Method | Inspection, Test (attachment load)
Status | Preliminary
Rationale | A reusable fleet (D1) loses units when devices are left loose and swept out with packaging; a captive attachment supports retention (MER-022) and stray-recovery/TCO control (SYS-031). The feature must not become a seal or thermal-barrier defeat, so it is bounded. Judgment.
MER-024 — Anti-rattle constraint of internal assemblies
Requirement Statement | The battery, boards, display, and probe connections shall be mechanically constrained within the enclosure so that no internal element rattles, chafes, or loads its solder joints under the transport vibration and shock of SYS-076/SYS-077, preventing joint fatigue, intermittent connection, or generation of spurious shock events.
Parent | SYS-076, SYS-077, SYS-005; PRD-MECH-007
Allocation | S1 — housing (internal retention)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (vibration + function/continuity), Inspection
Status | Evidence-Pending (constraint scheme with vibration profile)
Rationale | Unconstrained internal mass is the classic fatigue and intermittent-fault driver under sustained transport vibration (SYS-076); a loose board fatigues its joints and a loose accelerometer mass can even inject false shock events (SYS-005). Internal constraint is a data-integrity and reliability requirement. Standard (IEC 60068-2-64 random vibration; revision to confirm); the constraint scheme is Evidence-Pending, set with the vibration profile.
MER-025 — Defined orientation and anti-mis-placement keying
Requirement Statement | The device and its mount shall define the intended placement orientation (e.g. display and antenna-favourable face) and, where a specific orientation is required for thermal coupling, RF, or drainage, shall provide keying or clear marking that discourages an incorrect installed orientation.
Parent | SYS-052, SYS-080, SYS-024; PRD-MECH-003, PRD-FUN-016
Allocation | S1 — mount/retention (keying) + label/marking
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Inspection, Demonstration
Status | Preliminary
Rationale | Orientation affects thermal coupling (MER-015), antenna aperture toward the shipper opening (SYS-080), and any drainage; an installer with no training can place the device upside-down or facing the metallized wall. Keying/marking is a use-error control (mechanical realisation of pre-arm placement guidance, SYS-024). Judgment (use-error prevention; IEC 62366-1 concept, revision to confirm).
MER-026 — Probe lead management and retention
Requirement Statement | The device shall provide a means to route, retain, and (where applicable) coil excess probe lead so that the lead does not create a snag, a strain-relief overload at the gland (MER-010), or a path that displaces the tip from the product zone, and so that the lead survives repeated deployment and stowage over the reuse life.
Parent | SYS-076, SYS-051; PRD-MECH-003, PRD-MECH-004
Allocation | S1 — probe (lead) + mount/retention
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Test (deploy/stow cycling), Inspection
Status | Preliminary
Rationale | Loose probe lead is a snag, a tip-displacement path, and a repeated stress on the gland (MER-010); over many trips (D1) uncontrolled lead handling drives fatigue and mis-placement. Lead management protects both retention and the probe channel. Judgment.
5. HMI mechanicals
MER-027 — Display window optically matched for cold and high-glare legibility
Requirement Statement | The display window shall transmit and protect the display such that the accept/reject/quarantine verdict remains legible to a gloved user at the −30 °C operating extreme and across the SYS-075 ambient-lighting range (dim to direct sun), including a low-reflectance/anti-glare treatment so the verdict is not lost to washout at a sunlit dock.
Parent | SYS-074, SYS-075, SYS-058; PRD-MECH-008, PRD-ENV-003, PRD-ENV-011
Allocation | S1 — window
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (cold-chamber + high-lux legibility), Demonstration
Status | Evidence-Pending (window treatment and legibility thresholds to confirm)
Rationale | The offline edge verdict (SYS-015) is the last-mile decision point (D2) and is worthless if it cannot be read where it is used — at −30 °C cold-store staging or a sun-glared dock (SYS-074/075). The window is the mechanical/optical enabler of that legibility. Judgment; the window treatment and the cold-contrast/lux thresholds are Evidence-Pending, closed by the cold-chamber and high-lux legibility tests — coupled to Display/HMI (SYS-074).
MER-028 — Anti-fog / anti-condensation display window
Requirement Statement | The display window shall resist internal and external fogging/condensation during cold-to-warm transitions (e.g. by sealed air gap, anti-fog treatment, or internal moisture control) so that the verdict remains readable immediately on retrieval from a cold environment without a wait for the window to clear.
Parent | SYS-074, SYS-073; PRD-ENV-003, PRD-ENV-004
Allocation | S1 — window
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Test (condensation/legibility)
Status | Preliminary
Rationale | A window that fogs when a cold device is brought into a warm receiving area hides the verdict at exactly the moment of the receipt decision; the user cannot wait for it to clear and may accept blind. Anti-fog is a legibility requirement tied to the same cold-to-warm cycling as SYS-073. Judgment — coupled to internal condensation management (MER-012).
MER-029 — Glove-operable button travel and actuation force
Requirement Statement | The device control(s) shall be operable by a user wearing cold-weather work gloves down to the −30 °C operating extreme, with actuation force and travel within a defined range that gives reliable single actuation without accidental activation, and target size/spacing suited to a gloved fingertip.
Parent | SYS-074, SYS-058, SYS-023; PRD-MECH-008, PRD-UX-004
Allocation | S1 — button
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (force/travel), Demonstration (gloved usability)
Status | Evidence-Pending (force/travel/size targets to confirm in usability study)
Rationale | Arming (SYS-023) and reading the verdict happen gloved in the cold; a button too small, stiff, or ambiguous to press reliably means a mis-arm (unmonitored trip) or a mis-read — both top hazards. Actuation mechanics degrade at cold as seals and domes stiffen, so the range must hold at −30 °C. Judgment; force/travel/size targets are Evidence-Pending, set in the gloved usability study (with UX) — coupled to HMI interface (SYS-058).
MER-030 — Tactile feedback on control actuation
Requirement Statement | The control shall provide positive tactile (and where specified audible/visual) feedback confirming a registered actuation, perceptible through cold-weather gloves, so that a gloved user knows an action (e.g. arm) has been registered without relying on reading the display alone.
Parent | SYS-058, SYS-023; PRD-MECH-008, PRD-UX-004
Allocation | S1 — button
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Test, Demonstration (gloved)
Status | Preliminary
Rationale | Without tactile confirmation a gloved user cannot tell a real press from a missed one and may double-press or fail to arm; positive detent feedback closes the loop on the single-actuation contract (SYS-058) and reduces mis-arm risk (SYS-023). Judgment (haptic feedback for gloved operation).
MER-031 — Window abrasion, impact, and bonding integrity
Requirement Statement | The display window shall resist the scratch, abrasion, cleaning, and impact loads of the reuse life without a loss of legibility, and shall be bonded/retained to the housing so that its seal contributes to and does not compromise the MER-007 ingress rating under drop and thermal cycling.
Parent | SYS-078, SYS-077, SYS-079; PRD-MECH-005, PRD-MECH-007
Allocation | S1 — window (+ its seal to housing)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (abrasion, drop, IP), Inspection
Status | Preliminary
Rationale | The window is both an optical surface and part of the sealed boundary; scratches/hazing from repeated cleaning progressively hide the verdict (SYS-079), and a window that debonds on a drop breaches the seal (SYS-077/078). Its durability and bonding are therefore both legibility and ingress requirements. Standard (IEC 60529 rating maintained; IEC 60068-2-27 shock; revision to confirm).
6. Materials & cleaning (D1 reuse / cleaning; −30 °C)
MER-032 — Enclosure material toughness at the −30 °C extreme
Requirement Statement | The housing, port door, and probe materials shall be selected to retain impact toughness and ductility at the −30 °C operating extreme and the −40 °C non-operating extreme, such that a drop or handling impact at cold temperature does not cause brittle fracture that breaches the seal or exposes the electronics.
Parent | SYS-069, SYS-070, SYS-077; PRD-MECH-007, PRD-ENV-001, PRD-ENV-002
Allocation | S1 — housing + port door + probe (material)
Type | Derived
Criticality | High
Safety/Security | Safety
Verification Method | Test (cold-conditioned drop), Analysis (material data)
Status | Evidence-Pending (material selection to confirm)
Rationale | Many common polymers embrittle below their ductile-brittle transition; a housing that passes a room-temperature drop can shatter when dropped at −30 °C (the exact condition the device operates in), breaching the seal and corrupting the record (SYS-077). Cold toughness is a material-selection requirement, not a general "rugged" claim. Standard (IEC 60068-2-1 cold-conditioned handling; material low-temperature impact per datasheet; revision to confirm); selection is Evidence-Pending — flag Materials engineer.
MER-033 — Cleaning and decontamination chemical compatibility
Requirement Statement | The external housing, seals, window, controls, and labels shall withstand repeated cleaning and decontamination with the agents in the service procedure — including isopropyl-alcohol wipes, quaternary-ammonium disinfectants, and dilute-bleach (hypochlorite) wipes — over the reuse-life cleaning-cycle count, without crazing, swelling, embrittlement, seal degradation, loss of ingress rating, or loss of window/label legibility.
Parent | SYS-079, SYS-073; PRD-MECH-009, PRD-ENV-008; D1
Allocation | S1 — housing + seal + window + button + label (materials)
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (cleaning-agent aging, then IP + legibility), Inspection
Status | Evidence-Pending (definitive agent list, concentration, and cycle count to confirm)
Rationale | A reusable device is cleaned between every trip (D1); chemical attack on a seal creates an ingress path and attack on the window/label creates an illegible verdict or lost attribution — both safety-relevant, not cosmetic (SYS-079 rationale). Alcohols craze some plastics, quats and bleach attack others, so material choice must be qualified against the actual agents. Judgment; the agent list, concentration, and cleaning-cycle count are Evidence-Pending, set with the service procedure (PRD-SVC-002) — flag Materials engineer.
MER-034 — UV, sunlight, and weathering durability
Requirement Statement | The external enclosure materials, window, and printed/marked surfaces shall withstand the ultraviolet, sunlight, and outdoor-staging exposure of the distribution environment over the reuse life without embrittlement, discolouration that impairs display legibility, or fading that renders identification or mandatory markings illegible.
Parent | SYS-085; PRD-MECH-010, PRD-ENV-008
Allocation | S1 — housing + window + label/marking (materials)
Type | Allocated
Criticality | Medium
Safety/Security | None
Verification Method | Test (UV / weathering exposure), Inspection
Status | Evidence-Pending (UV exposure profile to confirm)
Rationale | Outdoor staging on tarmacs and docks embrittles polymers and fades print; combined with cleaning attack (MER-033) this is a cumulative degradation path that erodes both structural integrity and the legibility on which attribution and regulatory acceptance depend (SYS-085). Standard (IEC 60068-2-5 / ISO 4892 solar-radiation and weathering concepts; revision to confirm); the exposure profile is Evidence-Pending, set in the PVL — flag Materials engineer.
MER-035 — Durable device identification and safety/transport marking
Requirement Statement | The device shall carry durable, legible device identification (serial/UDI-style) and the required safety and transport markings, applied by a method (e.g. laser-mark, sub-surface, or qualified label) that remains legible over the full reuse life through the cleaning (MER-033), UV (MER-034), and abrasion exposures, without becoming an ingress path or a soil-trapping feature.
Parent | SYS-085, SYS-079; PRD-MECH-010, PRD-REG-line; D1
Allocation | S1 — label/marking
Type | Allocated
Criticality | Medium
Safety/Security | Safety
Verification Method | Test (label/marking durability), Inspection
Status | Evidence-Pending (marking method and required marking set to confirm)
Rationale | Attribution of a record to a device, fleet management (D1), and regulatory/transport acceptance (D4) all depend on markings that stay legible; an adhesive label that peels also becomes an ingress and contamination path. The marking method must survive the same cleaning/UV attack the housing does. Standard (WHO PQS E006 marking expectations; transport marking per UN 38.3/IATA; revision to confirm); the marking set is Evidence-Pending, set in the regulatory stage (PRD-MECH-010).
MER-036 — Biocompatibility not required (patient-contact exclusion)
Requirement Statement | The device shall be designed and documented as having no patient-contact and no product-contact function; accordingly ISO 10993 biocompatibility evaluation is recorded as not applicable, while food/product-contact and cleaning-residue safety of external materials shall still be considered against the packaged-vaccine context.
Parent | SYS-079; PRD-MECH-009; concept §2/§4 (information device, non-therapeutic)
Allocation | S1 — housing + probe (material justification)
Type | Derived
Criticality | Low
Safety/Security | Safety
Verification Method | Inspection (design/labeling review), Analysis
Status | Baselined
Rationale | The SCM is an information device that monitors the shipper environment and does not contact the patient or the vaccine directly (concept §2); ISO 10993 biocompatibility therefore does not apply and asserting it would be over-scoping. The note is recorded explicitly so the exclusion is a deliberate, documented decision rather than an omission, and so that any future patient/product-contact variant re-opens it. Standard (ISO 10993 scope — not applicable; revision to confirm) — flag Regulatory engineer to confirm.
MER-037 — RoHS / REACH material compliance
Requirement Statement | All device materials, finishes, and joining media shall comply with the applicable RoHS restricted-substance limits and REACH SVHC obligations for the target markets, and the material declarations shall be recorded to support the compliance and end-of-life requirements.
Parent | SYS-079, SYS-085; PRD-MECH-009, PRD-REG-line; D4
Allocation | S1 — housing + seal + window + fastener + label (materials)
Type | Derived
Criticality | Medium
Safety/Security | None
Verification Method | Inspection (material declarations / compliance dossier), Analysis
Status | Preliminary
Rationale | Multi-market placement (D4) makes RoHS/REACH a market-access gate; restricted-substance content in a housing, seal, or solder-adjacent finish can block sale or import irrespective of function. Recording declarations at material selection avoids a late redesign. Standard (EU RoHS Directive / REACH Regulation; revision to confirm) — flag Regulatory/Materials engineer.
MER-038 — Serviceable, non-glue-only fastening for end-of-life and repair
Requirement Statement | The enclosure shall be assembled with reworkable mechanical fasteners (or an equivalent openable joint) rather than a glue-only permanent bond, so that the device can be opened for battery service, seal replacement (MER-013), and repair, and can be disassembled for end-of-life material separation, without destroying the housing.
Parent | SYS-078; PRD-MECH-009, PRD-MECH-005; D1
Allocation | S1 — fastener (assembly method)
Type | Derived
Criticality | Medium
Safety/Security | Safety
Verification Method | Inspection, Demonstration (open/re-seal/re-verify IP)
Status | Preliminary
Rationale | A multi-year reusable device (D1) needs battery replacement, replaceable seals (MER-013), and eventual EoL material separation for RoHS/REACH-driven recycling (MER-037); a glue-only enclosure cannot be opened without destruction, forcing scrap and defeating the reuse economics. Serviceable fasteners must coexist with sealed penetrations (MER-014). Judgment (design-for-service and design-for-EoL) — coupled to Power (battery service) and Sustainability.
MER-039 — Drop and mechanical-shock structural survival
Requirement Statement | The device and probe shall survive the free-fall drop and handling mechanical-shock regime set in the Verification Plan without loss of function, loss of logged or buffered data, breach of the MER-007 ingress seal, or an undetected change in measurement accuracy.
Parent | SYS-077; PRD-MECH-007, PRD-ENV-006
Allocation | S1 — housing + window + probe + seal
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (IEC 60068-2-27 shock, IEC 60068-2-31 free-fall/topple)
Status | Evidence-Pending (shock pulse and drop height to confirm)
Rationale | Drops and rough handling are routine in logistics and last-mile use; the dangerous case is a drop that cracks the seal or shifts a sensor without stopping the device, producing a plausible-looking but wrong record (SYS-077 rationale). The mechanical design must fail safe, not silent. Standard (IEC 60068-2-27 / -2-31; revision to confirm); the shock pulse and drop height are Evidence-Pending, set in the PVL — coupled to material toughness (MER-032) and window integrity (MER-031).
MER-040 — Transport vibration and fatigue structural survival
Requirement Statement | The device, probe, retention features, and internal constraints shall survive the combined air/road/rail random and sinusoidal vibration regime set in the Verification Plan without connector or solder-joint fatigue failure, loss of function or data, seal loosening, or a shift of the device or probe out of its representative position.
Parent | SYS-076; PRD-MECH-007, PRD-MECH-003, PRD-ENV-005
Allocation | S1 — housing + mount/retention + probe + fastener
Type | Allocated
Criticality | High
Safety/Security | Safety
Verification Method | Test (IEC 60068-2-64 random, IEC 60068-2-6 sinusoidal, then function + IP + accuracy)
Status | Evidence-Pending (vibration profile/severity to confirm)
Rationale | Sustained transport vibration is a cumulative fatigue driver that loosens fasteners/seals and fatigues joints, and can walk the probe out of the product zone — a shifted probe reports an unrepresentative zone and corrupts disposition (SYS-076 rationale). This is where anti-rattle constraint (MER-024) and retention (MER-022) are proven together. Standard (IEC 60068-2-64 / -2-6; revision to confirm); the profile severity is Evidence-Pending, set in the PVL against representative lane spectra.
SRS-000 — Product Overview (module SRS-OVR)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-OVR · Owner: Embedded Software Engineer (EM-ENG-007) with the Chief Systems Engineer (EM-ENG-001). The software-library front matter: it fixes the software's purpose, scope, identity, users, environments, lifecycle, assumptions, and success criteria, and states the software-scope/identity requirements that bound every other module. Decomposed from the Product Vision (PRD-VIS-001…007) and Architecture (PRD-ARCH), based on the concept (D1–D4) and product description. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Fix, for the whole Software Requirements Library, what the SCM software is for and what bounds it: the value chain it must implement (capture → detect → warn → decide → prove), its identity as information-support software (never therapeutic, never temperature-controlling), its edge-autonomous scope, its configurable-across-products constraint, its reusable multi-trip lifecycle support, its target users and use environments, and its software success criteria. The requirements here are software-scope / identity constraints — they do not restate module functions; they ensure no downstream module optimises a local sub-optimum that violates the product's reason to exist.
2. Scope
Software purpose and value-chain scope; software identity/intended-use conformance; edge-autonomous scope; configurability across products/markets without firmware change; reusable multi-trip lifecycle software support; target-user and use-environment constraints; and software success criteria. Also provides the library-wide definitions, references, abbreviations, assumptions, and constraints. Excludes the detailed functional behaviour (SRS-DEV/CONN/CLOUD/MOB/etc.), the product-level (non-software) requirements (PRD), and commercial terms; those are referenced where the software identity depends on them.
3. Applicable Standards
WHO PQS E006-line (device role; category Evidence-Pending); ISO 14971 (severity-first intent); EU GDP and 21 CFR Part 11 / ALCOA+ (evidence intent); ISO/IEC/IEEE 29148 (requirements). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Definitions
• SCM — Smart Cold Chain Monitor: a reusable, battery-operated, cellular-connected environmental data logger placed inside insulated vaccine shippers.
• Value chain (capture → detect → warn → decide → prove) — capture (sense & log), detect (excursion logic), warn (alert/annunciate), decide (offline accept/reject/quarantine verdict), prove (tamper-evident auditable record/export).
• Trip profile — the configurable per-product parameter set (thresholds, excursion logic, logging interval, upload cadence, alert routing) applied to a device before arming.
• Verdict — the on-device accept / reject / quarantine indication for a completed or in-progress trip.
• Excursion — a product-relevant temperature (freeze or heat), humidity, or shock exposure exceeding the active profile's limits per its excursion logic.
• Armed trip — the interval, started by an explicit arm action with a valid profile, over which the device is credited as monitoring.
5. Abbreviations & References
Abbreviations | ALCOA+ (attributable, legible, contemporaneous, original, accurate, +complete, consistent, enduring, available); BLE (Bluetooth Low Energy); GDP (Good Distribution Practice); HMI (human-machine interface); LTE-M (LTE Cat-M1); NVM (non-volatile memory); OTA (over-the-air update); PQS (Prequalification of Vaccine-related Devices); RBAC (role-based access control); SSO (single sign-on).
References | concept-and-decisions.md (D1–D4); product-description.md; system-block-diagram.md (I-01…I-12); PRD-VIS (000), PRD-ARCH (300), PRD-SAFE (1300), PRD-REG (1400); SYS-functional/interface/safety/regulatory; software-requirements-standard.md.
6. Assumptions & Constraints
• D1 the software supports a reusable, multi-trip device (return, clean, recharge, triage, re-provision); D2 the software is offline-first store-and-forward (log locally, upload on cadence, alert immediately when connected); D3 the software operates the device within an aggressive duty-cycled energy budget across −30 °C to +60 °C for 90 days, with charging inhibited in the cold; D4 the software serves a multi-market deployment (WHO PQS, EU GDP, US, transport, radio).
• The software is information-support software: it supports, and never automates or replaces, the human product-disposition decision, and it neither controls nor preserves product temperature.
• Numeric targets (accuracy ±0.5 °C, alert latency ≤5 min when connected, detection miss-rate, 90-day endurance, cycle-life) are held in SRS-PERF / PRD-PERF and closed by validation; where unresolved they are Evidence-Pending, not invented.
6a. Software-scope / identity requirements
SRS-OVR-001 — Implement the capture→detect→warn→decide→prove value chain
Requirement Statement | The SCM software shall implement the full value chain — capture (interval sense & gap-free log), detect (two-sided product-relevant excursion logic), warn (immediate alert when connected; local capture always), decide (offline on-device accept/reject/quarantine verdict), and prove (complete, tamper-evident, attributable, exportable trip record) — as its defining scope.
Rationale | The value chain is the product's reason to exist; fixing it as the software's defining scope keeps every module aligned to the beneficiary outcome rather than a local feature.
Priority | Must
Source | User Need
Verification Method | Test, Demonstration
Traceability | PRD-VIS-001 (intended use), SYS-001/009/011/015/026 (chain functions), product-description §2; → SRS-DEV, SRS-CONN, SRS-CLOUD, SRS-DATA.
Notes | The chain functions are detailed across the functional modules; this states the scope constraint.
SRS-OVR-002 — Information-support software identity (not therapeutic, not temperature-controlling)
Requirement Statement | The SCM software shall support — and shall neither automate nor replace — the human product-disposition decision, and shall not implement any function that controls or preserves product temperature or presents the device as a therapeutic or temperature-control device.
Rationale | Locks the software product class so no downstream requirement drifts the SCM into a temperature-control or therapeutic role, which would change its regulatory classification and safety basis.
Priority | Must
Source | User Need (Safety)
Verification Method | Inspection
Traceability | PRD-VIS-001, PRD-VIS-002 (beneficiary-outcome primacy), SYS-149 (severity anchor); → SRS-SAFE, SRS-REG. Standard: WHO PQS device role (revision to confirm).
Notes | Severity anchor STN-020: both disposition-error directions harm the beneficiary.
SRS-OVR-003 — Edge-autonomous software scope (offline-first)
Requirement Statement | The SCM device software shall perform sensing, logging, excursion detection, and the on-device verdict entirely on the device, without dependence on cloud or mobile connectivity, so that the last mile is never a black box.
Rationale | Offline-first is the core architectural commitment (D2); the least-trained last-mile user has no phone, network, or login, so the edge functions must stand alone.
Priority | Must
Source | Architecture (User Need)
Verification Method | Test
Traceability | PRD-ARCH-003 (edge-autonomous), PRD-VIS-001, SYS-013/015; D2; → SRS-DEV, SRS-CTX-001/004.
Notes | Edge verdict read stays login-free (SYS-015); connectivity is opportunistic (SRS-CONN).
SRS-OVR-004 — Configurable across products/markets without firmware change
Requirement Statement | The SCM software shall realise different vaccine/product thermal profiles and per-market operating profiles through trip-profile and regional configuration alone, without a firmware variant, wherever feasible.
Rationale | One reusable fleet must serve many products and markets; configuration-driven variance limits SKU/firmware proliferation and keeps the validated software base common.
Priority | Should
Source | Business (Architecture)
Verification Method | Demonstration
Traceability | PRD-VIS-007 (configurable across products), PRD-ARCH-010 (variant strategy), SYS-022; → SRS-DEV-005/007, SRS-CLOUD (config).
Notes | Radio/regulatory variants may still require per-market hardware/certification (PRD-REG-004).
SRS-OVR-005 — Reusable multi-trip lifecycle software support (D1)
Requirement Statement | The SCM software shall support the reusable multi-trip lifecycle — trip close, return, recharge (charge inhibited below the qualified cold limit), readiness triage (charge, calibration, firmware, profile), and re-provisioning — without loss of calibration or record integrity between trips.
Rationale | Reuse economics and record credibility depend on a defined, software-supported loop; a device cannot be re-fielded safely if readiness and record integrity are not maintained across trips.
Priority | Must
Source | Business (D1)
Verification Method | Demonstration
Traceability | PRD-VIS-005 (reusable lifecycle), SYS-030 (readiness), PRD-SVC; D1; → SRS-DEV-020/023, SRS-FLEET.
Notes | Cycle-life target Evidence-Pending (PRD-REL); readiness gating detailed in SRS-FLEET.
SRS-OVR-006 — Target-user and use-environment constraint
Requirement Statement | The SCM software shall be usable and correct for its intended users and environments — QA/dispatch provisioners, transport handlers, and minimally trained, gloved, offline last-mile users — operating unattended across the −30 °C to +60 °C envelope inside RF-attenuating shippers throughout the global distribution chain, and shall not assume a trained operator, continuous connectivity, or an ambient laboratory environment for any safety-relevant task.
Rationale | Designing to the least-trained user and the harshest environment is what makes the product safe and effective where it is actually used; assuming a benign user/environment is a use-error and reliability hazard.
Priority | Must
Source | User Need (Safety)
Verification Method | Test (summative usability), Analysis
Traceability | PRD-VIS-001, SYS-153 (use-error protection), PRD-SAFE-005; D3/D4; → SRS-SAFE-007, SRS-REG-006, SRS-CTX-009. Standard: IEC 62366-1 (revision to confirm).
Notes | Feeds the Usability Engineering File; ties SRS-CTX-009 operational-environment context.
SRS-OVR-007 — Software success criteria and goals
Requirement Statement | The SCM software shall meet its validated success criteria — no material temperature excursion (per the active profile) undetected in the trip record, and a complete, tamper-evident, ALCOA+ / Part 11-admissible record supporting every disposition — as the measurable definition of the software's headline promise.
Rationale | "Protect the dose and prove it" must be measurable; fixing the success criteria makes the headline promise verifiable rather than aspirational, and orders every trade-off toward it.
Priority | Must
Source | Business (User Need)
Verification Method | Test (validation)
Traceability | PRD-VIS-003/004 (two-sided protection, no undetected excursion), SYS-009/026, PRD-REG-003; → SRS-PERF, SRS-SAFE, SRS-DATA; → PVaL. Standard: ALCOA+ / 21 CFR Part 11 (revision to confirm).
Notes | Numeric miss-rate target Evidence-Pending; set in validation (evidence).
---
Traceability
Every requirement above traces upward to at least one Product Vision requirement (PRD-VIS-001…007) and, where applicable, Architecture (PRD-ARCH-003/010) and the Functional/Safety System Requirements, and forward to the modules that realise the scope. Principal upstream anchors: PRD-VIS-001 (intended use), PRD-VIS-002 (beneficiary-outcome primacy), PRD-VIS-003/004 (two-sided protection, no undetected excursion), PRD-VIS-005 (reusable lifecycle), PRD-VIS-007 (configurability); PRD-ARCH-003 (edge-autonomy); SYS-149 (severity anchor). Cross-module: value-chain functions → SRS-DEV / SRS-CONN / SRS-CLOUD / SRS-DATA; software risk controls → SRS-SAFE; regulatory conformance → SRS-REG; system-context constraints → SRS-CTX; numeric targets → SRS-PERF. Evidence posture: every quantitative target is Standard, Cited, or Judgment; unresolved quantities (WHO PQS category, detection miss-rate, endurance, cycle-life) are Evidence-Pending and closed by verification/validation. No standard clause number is asserted as fact.
SRS-100 — System Context (module SRS-CTX)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-CTX · Owner: Embedded Software Engineer (EM-ENG-007) with the Chief Systems Engineer (EM-ENG-001). Detailed system-context software requirements decomposed from the Architecture Product Requirements (PRD-ARCH-001…010) and the Interface System Requirements (SYS-051…068), based on the system block diagram (interfaces I-01…I-12) and the product description (subsystems S1–S9). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Fix the context constraints the SCM software must honour: the three-tier system composition, the subsystem decomposition and boundaries, the typed interface inventory (I-01…I-12), the split trust boundary, the device↔cloud and device↔mobile transports, the cloud↔enterprise integration, identity/federation, the cross-boundary data flows and schema/version contracts, and the deployment topology and operational environments. These requirements do not re-specify functions (SRS-DEV/CONN/CLOUD/etc.); they state the architectural envelope inside which every other software module operates so no module silently violates a system boundary.
2. Scope
System/logical/physical architecture context; subsystem decomposition (S1–S9) and boundary conformance; conformance to the typed interface inventory and the interface control documentation (I-01…I-12); the split trust boundary (login-free edge read vs platform-governed configuration/data); device↔cloud application transport context (MQTT / HTTPS / REST over mutually authenticated TLS, I-10); mobile/BLE app context (I-09); cloud↔enterprise integration and REST export with identity federation (I-11); cross-boundary data-flow and schema/version contracts (I-10); and deployment topology and operational environments (offline-first edge, multi-region cloud, −30 °C to +60 °C). Excludes the detailed behaviour of each subsystem (owned by SRS-DEV/CONN/CLOUD/MOB/SEC/OTA/DATA), the cryptographic mechanisms (SRS-SEC), and the physical/RF/thermal interface qualification (system/RF/mechanical stages); those are referenced where the software must conform to their contract.
3. Applicable Standards
ISO/IEC/IEEE 29148 (requirements and interface definition); IEC 62443 concepts (security zoning / trust boundaries); MQTT v5, HTTPS/REST, and TLS 1.2+/1.3 (application transport); Bluetooth Core / GATT (BLE provisioning). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Definitions
• Tier — one of the three top-level system elements: the device tier (S1–S7, in the shipper), the cloud platform tier (S8), and the mobile/BLE application tier (S9).
• Trust boundary — a partition across which identity must be proven before privileged actions; the SCM has a login-free read-only edge boundary (on-device verdict) and a platform-governed boundary (configuration, ingest, export, OTA).
• Typed interface (I-01…I-12) — a boundary in the system block diagram carrying one type: PWR, DATA, RF, THRM, or MECH. This module constrains the software-bearing DATA/RF interfaces (I-05…I-12).
• ICD — Interface Control Documentation: the authoritative definition of each interface's type, direction, and contract.
5. Assumptions
• The three-tier composition and subsystem set are fixed by the architecture (PRD-ARCH-001/002) and the block diagram; this module binds software to them, it does not redefine them.
• The typed interface inventory (I-01…I-12) and its ICD are the single authoritative interface source (SYS-066, PRD-ARCH-009); the RF-in-container interface (I-08) is Evidence-Pending (SYS-059), and the software degrades to store-and-forward when it cannot be closed.
• The cryptographic identity, mutual-authentication, and key mechanisms behind every trust boundary are owned by SRS-SEC (EM-ENG-008); this module states the boundary the software must honour, not the mechanism.
• Cloud multi-region deployment, availability, and enterprise-integration/identity-federation services are provided by SRS-CLOUD; this module states the context those services operate within.
6. Architecture Overview
The SCM is one system realised across three tiers connected by twelve typed interfaces. Software lives in the device (S3 firmware over S4–S7), the cloud (S8), and the mobile app (S9). The device is edge-autonomous — sensing, logging, detection, and the offline verdict run wholly on-device (I-05/I-06/I-07) with no dependence on the cloud or app. Opportunistic connectivity moves data and alerts out over I-08/I-10 (LTE-M/Wi-Fi to cloud) and I-09 (BLE to app), and the cloud integrates outward over I-11 (enterprise/quality systems) and inward over I-12 (signed OTA). A split trust boundary separates the login-free on-device read from the platform-governed configuration/data plane. Every software module must conform to this composition, these boundaries, and the ICD; the requirements below make those context constraints explicit and testable.
---
7. Functional Requirements (context constraints)
7.1 Composition & decomposition
SRS-CTX-001 — Three-tier system composition
Requirement Statement | The SCM software shall be structured as three cooperating tiers — device firmware (S1–S7), cloud platform (S8), and mobile/BLE application (S9) — and each tier's software shall perform its safety-relevant functions within its defined connectivity assumptions, without a tier depending on another tier being reachable to perform a function allocated to it.
Rationale | The three-tier composition and its independence assumptions are the top-level architectural invariant; a tier that silently depends on another breaks the offline-first guarantee.
Priority | Must
Source | Architecture
Verification Method | Inspection, Test
Traceability | PRD-ARCH-001 (three-tier composition), PRD-ARCH-003; SYS (architecture); → SRS-DEV, SRS-CLOUD, SRS-MOB; block diagram Fig. 1.
Notes | Edge-autonomous core function is stated in SRS-CTX-004.
SRS-CTX-002 — Subsystem decomposition and boundary conformance
Requirement Statement | The device software shall be organised over the defined device subsystems (S1 enclosure, S2 power, S3 compute/MCU+firmware, S4 sensing, S5 connectivity, S6 HMI, S7 storage) and shall interact with each subsystem only across its defined interface, such that no software component reaches across a subsystem boundary by an undocumented path.
Rationale | A clean subsystem decomposition with respected boundaries is what keeps the design analysable, testable, and free of hidden coupling that defeats energy-gating, safety, or security controls.
Priority | Must
Source | Architecture
Verification Method | Inspection, Analysis
Traceability | PRD-ARCH-002 (device subsystem set), SYS-066 (interface conformance); → SRS-DEV, SRS-CONN; product-description §3.
Notes | Interface conformance detailed in SRS-CTX-003.
SRS-CTX-003 — Conformance to the typed interface inventory / ICD (I-01…I-12)
Requirement Statement | Every software-bearing interface the SCM realises shall conform to the typed interface inventory (I-01…I-12) and the interface control documentation — matching the defined type (DATA/RF), direction, and contract — and no undocumented software interface shall exist between subsystems or across the device/cloud/mobile trust boundaries.
Rationale | Interface drift between disciplines is a leading source of integration defects, and an undocumented interface is an ungoverned attack surface; a single authoritative ICD keeps the block diagram, requirements, and design outputs in agreement.
Priority | Must
Source | Architecture (Security)
Verification Method | Inspection, Analysis
Traceability | SYS-066 (ICD conformance), SYS-051…068 (typed interfaces), PRD-ARCH-009; → all software modules; block diagram §3. Standard: ISO/IEC/IEEE 29148 (revision to confirm).
Notes | I-08 (RF-in-container) is Evidence-Pending (SYS-059); software degrades to store-and-forward.
7.2 Trust boundaries & transports
SRS-CTX-004 — Split trust boundary (login-free edge read vs platform governance)
Requirement Statement | The software shall enforce a split trust boundary in which the on-device accept/reject/quarantine verdict is readable without network, phone, or login, while all configuration, provisioning, data ingest, export, and OTA actions require authenticated, authorised access enforced at the platform-governed boundary.
Rationale | Frictionless offline edge use for the least-trained user must coexist with enterprise / Part 11 governance; the split boundary is how both are satisfied without one defeating the other.
Priority | Must
Source | Architecture (Security)
Verification Method | Analysis, Test
Traceability | PRD-ARCH-006 (split trust boundary), SYS-068 (authenticated identity), SYS-015 (login-free verdict); → SRS-SEC, SRS-DEV-030, SRS-CONN-012. Standard: IEC 62443 concepts (revision to confirm).
Notes | Mechanisms owned by SRS-SEC (EM-ENG-008); threat model Evidence-Pending (SYS-068).
SRS-CTX-005 — Device↔cloud application transport context (I-10)
Requirement Statement | The device and cloud software shall exchange telemetry, alerts, commands, and OTA metadata across the I-10 boundary using MQTT v5 and/or HTTPS over mutually authenticated TLS, with ordered, resumable, idempotent, acknowledged transfer and alert prioritisation, and shall not exchange controlled data over any unauthenticated or unencrypted channel.
Rationale | I-10 is the store-and-forward and alert boundary; the context (secure, acknowledged, idempotent, alert-prioritised transport) must be fixed so every module using it inherits the same guarantees.
Priority | Must
Source | Architecture (Security)
Verification Method | Test, Inspection
Traceability | SYS-062 (store-and-forward interface), SYS-063 (protocol/schema), PRD-ARCH-005, PRD-CONN-006/ 008; → SRS-CONN-007/011, SRS-CLOUD, SRS-SEC. Standard: MQTT v5, HTTPS/REST, TLS 1.2+/1.3 (revision to confirm).
Notes | Transport behaviour detailed in SRS-CONN; this fixes the context constraint.
SRS-CTX-006 — Mobile/BLE application context (I-09)
Requirement Statement | The device and mobile software shall interact across the I-09 boundary over a defined BLE GATT service for provisioning, configuration, arming, and local trip-verdict/record read, such that configuration and controlled read require authenticated access while the login-free local verdict read is preserved, and only defined, version-negotiated operations are accepted.
Rationale | BLE is the local provisioning and offline-read path; its context (authenticated config vs login-free read, defined versioned GATT surface) must be fixed to serve the workflow without becoming a standing exposure surface.
Priority | Must
Source | Architecture (Security)
Verification Method | Test, Inspection
Traceability | SYS-060/061 (BLE link / GATT), PRD-ARCH-005, PRD-CONN-006, PRD-MOB-002; → SRS-CONN-004/012, SRS-MOB, SRS-SEC. Standard: Bluetooth Core / GATT (revision to confirm).
Notes | GATT structure and versioning per SYS-061; mechanisms per SRS-SEC.
7.3 Enterprise integration, data flow & deployment
SRS-CTX-007 — Cloud↔enterprise integration, REST export & identity federation (I-11)
Requirement Statement | The cloud software shall expose a documented, versioned REST API and export across the I-11 boundary that delivers trip records and fleet data to authorised enterprise/quality systems in a machine-readable format preserving attribution, time, gaps, and tamper-evidence, under authenticated, role-based, audit-logged access integrated with the enterprise identity provider (federated SSO/RBAC).
Rationale | I-11 is where the compliance value is delivered to ERP/QMS for release decisions; it must preserve the record's ALCOA+ properties and be governed by federated identity and RBAC, not ad-hoc access.
Priority | Must
Source | Architecture (Security)
Verification Method | Test, Demonstration
Traceability | SYS-064 (enterprise export/API), PRD-CLOUD-005, PRD-REG-003, PRD-SEC-005; → SRS-CLOUD, SRS-SEC, SRS-DATA. Standard: ISO/IEC/IEEE 29148, EU GDP / 21 CFR Part 11 (revision to confirm).
Notes | Identity federation / RBAC mechanisms owned by SRS-CLOUD/SRS-SEC.
SRS-CTX-008 — Cross-boundary data-flow and schema/version contract
Requirement Statement | For every message crossing a software boundary (I-09, I-10, I-11, I-12), the software at both endpoints shall validate the message against the published schema, units, and negotiated schema version, and shall reject or quarantine a malformed, unknown-version, or out-of-contract message rather than misinterpreting it, so that a safety-relevant record is never silently mis-decoded across a version boundary.
Rationale | A safety-relevant record crossing a version boundary must never be silently mis-decoded; an explicit schema/version contract is the defence against fleet/cloud version skew during OTA rollouts.
Priority | Must
Source | Architecture (Security)
Verification Method | Test, Inspection
Traceability | SYS-063 (protocol/message format), SYS-065 (OTA interface), PRD-ARCH-009, PRD-CLOUD-002; → SRS-DATA, SRS-CONN, SRS-CLOUD, SRS-OTA. Standard: ISO/IEC/IEEE 29148 (revision to confirm).
Notes | Message schemas/encodings detailed in SRS-DATA; this fixes the validate-or-quarantine contract.
SRS-CTX-009 — Deployment topology and operational-environment context
Requirement Statement | The software shall operate within the defined deployment topology — an offline-first device operating unattended across the −30 °C to +60 °C envelope inside RF-attenuating shippers, a multi-region cloud platform, and mobile apps at depot/last-mile — such that no software function assumes continuous connectivity, a single region, a fixed carrier, or an ambient laboratory environment.
Rationale | The operating context (intermittent connectivity, extreme cold, RF attenuation, multi-region, multi-carrier) is the reality every module must be correct within; encoding it as a context constraint prevents modules from assuming a benign environment.
Priority | Must
Source | Architecture
Verification Method | Analysis, Test
Traceability | PRD-ARCH-004 (duty-cycled), PRD-ARCH-003 (edge-autonomous), SYS-059 (RF-in-container), SYS-116 (roaming); → SRS-DEV, SRS-CONN, SRS-CLOUD, SRS-PERF; D2/D3/D4.
Notes | RF-in-container performance Evidence-Pending (SYS-059); multi-region cloud detailed in SRS-CLOUD.
---
Traceability
Every requirement above traces upward to at least one Architecture Product Requirement (PRD-ARCH-001…010) and Interface System Requirement (SYS-051…068), and forward to the modules that realise the context. Principal upstream anchors: PRD-ARCH-001/002/003 (composition, decomposition, edge-autonomy), PRD-ARCH-005/006 (multi-bearer, split trust), PRD-ARCH-009 (interface conformance); SYS-062/063/064/065 (device↔cloud, enterprise, OTA), SYS-060/061 (BLE), SYS-066/067/068 (ICD, error handling, authenticated identity). Cross-module: device behaviour → SRS-DEV; connectivity transports → SRS-CONN; cloud/enterprise/federation → SRS-CLOUD; mobile app → SRS-MOB; cryptographic identity/trust mechanisms → SRS-SEC; message schemas/retention → SRS-DATA; OTA delivery → SRS-OTA. The RF-in-container interface (I-08) is carried Evidence-Pending (SYS-059) and closed by the RF-in-container test (PRD-VER-005).
SRS-1000 — Performance (module SRS-PERF)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-PERF · Owner: Embedded Software Engineer (EM-ENG-007) with the Principal Reliability/Power and Cloud engineers. Detailed software-performance requirements decomposed from the System Requirements (SYS-037…050, plus SYS-011/019/020) and Product Requirements (PRD-PERF 1700), quantifying the timing, capacity, throughput, responsiveness, and energy budgets that make the SCM's verdict trustworthy, its alert timely, and its 90-day endurance achievable. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify the measurable performance targets the SCM device, cloud, and mobile software shall meet: boot/wake time, compute/memory/storage utilization and headroom, flash endurance, detection and alert latency, interaction responsiveness, fault-recovery time, sustained logging throughput, concurrency, cloud response and scalability, and — the binding constraint — battery life and power consumption over the 90-day worst-case cold-duty profile. Performance is treated as a first-class, verified property, not an emergent one.
2. Scope
Device firmware startup and runtime resource budgets; on-device excursion-detection latency; end-to-end excursion-alert delivery latency; on-device verdict computation and HMI responsiveness; fault/watchdog recovery time; non-volatile log write throughput and concurrency; flash-memory endurance; cloud ingest/alert-processing response time and horizontal scalability; mobile-app responsiveness; and the software contribution to the device energy budget. Excludes the battery/BMS electrical design and cell-level cold derating (S2, Power engineer), the RF link budget and antenna (S5/S1, RF engineer — see SRS-CONN), storage media selection and wear-levelling hardware (S7), and the cloud infrastructure sizing itself (SRS-CLOUD); those are referenced where the software performance interfaces with them.
3. Applicable Standards
WHO PQS E006 (temperature-monitoring device performance/accuracy class — informs the timing and record targets); IEC 60068 (performance versus environment, for cold-extreme performance retention); ISO/IEC/IEEE 29148 (requirement verifiability); IEC 62304 (software life-cycle process). Cited by designation only; revision to confirm — no clause number or performance class is asserted as fact.
4. Assumptions
• The 90-day endurance target and the worst-case cold-duty profile are set by the system power budget (SYS-045/046, PRD-PERF-005); the achievable energy margin and the low-temperature usable-capacity figure are Evidence-Pending and close on the 90-day cold-endurance test (MR-VAL-006). No battery-life number in hours/days beyond the ≥90-day system target, and no per-function energy figure, is asserted here.
• The energy governor and duty-cycle scheduler (SRS-DEV) own the sleep/wake windows; this module specifies the performance the software delivers within those windows and the sanctioned alert exception.
• Numeric targets carried from the PRD/SYS as Judgment (±0.5 °C, 5 min alert, 3 s responsiveness, 50 % log spare) are stated as the working software target and confirmed in V&V; genuinely unresolved quantities (battery-life margin, per-class energy allocations, RF-in-container delivery latency) are marked Evidence-Pending and never invented.
• Non-volatile log capacity, atomic-write integrity, and buffer ordering are provided by SRS-DEV/SRS-DATA; this module specifies the throughput and latency of the write and read paths.
• The cloud provides idempotent ingest and elastic compute (SRS-CLOUD); the mobile app runs on a contemporary consumer smartphone (SRS-MOB).
5. Architecture Overview
Performance is budgeted across four layers. On the device, a low-power scheduler wakes the MCU from deep sleep for each sampling interval, completes acquire→timestamp→evaluate→commit within a bounded slice of the interval, and returns to sleep — the duty cycle that makes 90 days achievable. A latency-critical path carries excursion detection (on-device, immediate) and, when a bearer is present, alert delivery to the cloud inside the 5-minute budget. An interaction path computes and renders the accept/reject/quarantine verdict on the OLED within the 3-second responsiveness budget, at cold and low battery. The cloud ingests, persists, and routes alerts with a bounded processing time and scales horizontally with fleet size. The mobile app renders status and configuration within its own responsiveness budget. Every path is bounded so that no single workload — a shock-event burst, a backlog upload, a full-record verdict — starves the safety-critical sample, alert, or verdict.
---
7. Functional Requirements
7.1 Startup & resource performance
SRS-PERF-001 — Boot / wake-to-monitoring time
Requirement Statement | The device software shall complete cold boot to a fully armed, monitoring state (sensing, logging, detection, and clock active) within a bounded, documented startup time, and shall complete a scheduled wake-from-sleep to sample-committed within a bounded slice of the logging interval such that no scheduled sample is delayed into the following interval.
Rationale | A trip must not begin with an unmonitored gap, and the duty-cycle wake path must fit inside the interval with margin, or gap-free interval logging (SYS-002) fails at the low-power wake boundary.
Priority | Must
Source | Performance
Verification Method | Test
Traceability | SYS-037 (acquire-settle-stamp within interval), SYS-023/024 (arm/readiness), PRD-PERF-003; → SRS-DEV (boot sequence, scheduler), SRS-PERF-014 (energy).
Notes | Startup and wake-time budgets slow at the −30 °C extreme (MCU/oscillator warm-up); the bound is stated and verified at the cold worst case (SRS-PERF-018).
SRS-PERF-002 — Compute (CPU) utilization headroom
Requirement Statement | The device software shall keep steady-state CPU utilization, under the combined peak of sampling, excursion evaluation, event capture, logging, and a concurrent scheduled upload, at or below a documented ceiling that preserves at least 50 % processing headroom for peak transients.
Rationale | A margin of spare compute prevents priority inversion and missed deadlines when a shock-event burst coincides with an upload; a device running at the edge cannot guarantee the sample or alert deadline.
Priority | Should
Source | Performance
Verification Method | Test, Analysis
Traceability | SYS-044 (sustained write throughput), PRD-PERF-006; → SRS-PERF-011 (concurrency), SRS-DEV (task scheduling).
Notes | The 50 % spare is a Judgment target aligned with PRD-PERF (spare-capacity posture); confirmed by profiling under the worst-case concurrent load.
SRS-PERF-003 — Memory (RAM) utilization headroom
Requirement Statement | The device software shall bound peak RAM usage (stack, heap if any, and buffers) under worst-case concurrent operation to a documented ceiling that retains a defined free-memory reserve, and shall never allocate such that a safety-critical sample, event, or alert cannot be captured for lack of memory.
Rationale | Memory exhaustion on a resource-constrained MCU manifests as a missed sample or a crash; a bounded footprint with reserve is what keeps capture guaranteed under load.
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-044, SYS-002 (no silent gap), PRD-PERF-006; → SRS-DEV (fault mgmt), SRS-PERF-016.
Notes | Static allocation preferred for critical paths (implementation guidance, not mandated here); high-water-mark measured by test.
SRS-PERF-004 — Non-volatile log capacity and spare for a full 90-day offline trip
Requirement Statement | The device software shall store the complete multi-channel sample and event log for a full 90-day trip at the minimum logging interval with at least 50 % spare non-volatile capacity, with no connectivity assumed and without overwriting unsent data.
Rationale | An entirely offline trip must still yield a complete record with headroom; store-and-forward requires worst-case offline capacity (D2), and 50 % spare absorbs event bursts and sizing uncertainty.
Priority | Must
Source | Performance
Verification Method | Analysis (sizing), Test
Traceability | SYS-043 (log capacity + 50 % spare), SYS-018 (offline buffering), PRD-PERF-006; → SRS-DATA (record format), SRS-DEV (storage).
Notes | Order-of-magnitude ≈ 1.3×10^5 samples/channel at a 1-minute interval; exact sizing is Evidence-Pending on record format and channel count (SYS-043) and closes with SRS-DATA — no byte figure asserted here.
SRS-PERF-005 — Flash endurance over the reuse life
Requirement Statement | The device software shall manage non-volatile writes (wear-levelling, write coalescing, and journaling) so that cumulative program/erase cycles over the specified reuse/charge-cycle life remain within the storage medium's endurance rating with a documented margin, and shall detect and flag approaching end-of-endurance rather than fail silently.
Rationale | A reusable multi-trip device (D1) writes the log many times over life; exceeding flash endurance corrupts the record — the product's core value — so wear must be budgeted and monitored.
Priority | Must
Source | Reliability
Verification Method | Analysis, Test
Traceability | SYS-165 (endurance over reuse), SYS-168 (atomic writes), PRD-REL-002/005; → SRS-REL, SRS-DIAG (health).
Notes | Endurance rating and per-trip write budget depend on the final storage device (Evidence-Pending); requirement fixes the management obligation and margin, not the device part number.
7.2 Latency & responsiveness
SRS-PERF-006 — On-device excursion-detection latency
Requirement Statement | The device software shall evaluate each new sample against the active thresholds and product-stability logic and register a detected excursion within one logging interval of the sample that satisfies the excursion condition, independent of connectivity.
Rationale | Detection is the safety-critical first link; the end-to-end alert budget presupposes prompt on-device detection, and offline excursions must still be captured at detection time (SYS-013).
Priority | Must
Source | Safety
Verification Method | Test
Traceability | SYS-038 (detection latency), SYS-009/010/013 (two-sided detection, product logic, local capture), PRD-PERF-004; → SRS-SAFE / HAL (delayed-alert hazard, Stage 10).
Notes | Detection must complete at the sampling cadence even at cold/low-battery; reconcile with HAL.
SRS-PERF-007 — Excursion-alert delivery latency when connected (≤5 min)
Requirement Statement | When connectivity is available at the time of detection, the software shall deliver an excursion/shock alert to the cloud within 5 minutes of on-device detection, ahead of routine bulk telemetry.
Rationale | Fast alerting converts passive monitoring into prevented loss — a responder must learn of a break in time to intervene (D2); the alert value is time-bound.
Priority | Must
Source | Safety
Verification Method | Test (latency, detection→cloud notification)
Traceability | SYS-039 (5-min alert latency), SYS-011 (immediate alert), PRD-PERF-004, PRD-FUN-008; → SRS-CONN-008/022 (alert prioritisation, delivery), SRS-CLOUD (ingest/route).
Notes | 5 min is a Judgment target confirmed in validation (SYS-039); RF attenuation inside the shipper may dominate achievable latency and is Evidence-Pending on the RF-in-container test (PRD-VER-005). Offline detection is always immediate; queued-alert delivery latency is bounded by SRS-CONN / SYS-040.
SRS-PERF-008 — On-device verdict computation and display responsiveness (≤3 s)
Requirement Statement | The device software shall compute the accept/reject/quarantine verdict from the full stored trip record against the active profile and render it on the device display within 3 seconds of the user action, with no network, phone, or login, including at the −30 °C cold extreme and at low battery.
Rationale | The edge verdict is the decision point for a time-pressured last-mile user with no phone or network (SYS-015); if a full 90-day record cannot be evaluated within the responsiveness budget on the MCU, the offline value proposition fails.
Priority | Should
Source | User Need
Verification Method | Test (response time incl. cold/low-battery), Demonstration
Traceability | SYS-041 (HMI responsiveness), SYS-042 (verdict compute/display), PRD-PERF-007, PRD-FUN-010; → SRS-DEV (verdict engine), SRS-PERF-018.
Notes | 3 s is a Judgment target confirmed in the usability study; worst case is a full record evaluated at cold/low state-of-charge (display and MCU wake times slow at cold).
SRS-PERF-009 — Fault-recovery time to monitoring
Requirement Statement | Following a watchdog reset, brown-out recovery, or firmware fault, the device software shall return to an armed monitoring state (sensing, logging, detection active) within a bounded recovery time, log the recovery event, and preserve all committed records across the event.
Rationale | A prolonged post-fault outage creates an unmonitored trip believed monitored — a false-reassurance hazard; bounded recovery limits the gap and the logged event keeps monitoring honest.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-169 (watchdog recovery), SYS-167 (safe state), SYS-002/017 (gap honesty), PRD-REL-004/006; → SRS-REL, SRS-DIAG.
Notes | Any recovery gap is represented as an explicit gap marker (SYS-002) and disqualifies an unqualified accept (SYS-017).
7.3 Throughput, concurrency & scalability
SRS-PERF-010 — Sustained logging and event write throughput
Requirement Statement | The device software shall commit each scheduled sample and each detected shock event to non-volatile storage at the peak combined rate of base logging plus event capture, sustained for the full trip, without dropping, overwriting, or corrupting any record.
Rationale | Capacity is necessary but not sufficient; the write path must keep up with a burst of shock events superimposed on the base cadence, or gap-free logging is violated exactly when the data matters.
Priority | Must
Source | Performance
Verification Method | Test
Traceability | SYS-044 (sustained write throughput), SYS-005 (shock capture), PRD-PERF-006, PRD-FUN-004; → SRS-DATA, SRS-DEV, SRS-REL (atomic writes).
Notes | Peak rate set from the shock-event profile; verified under a superimposed base-plus-burst load.
SRS-PERF-011 — Concurrent workload without missed samples
Requirement Statement | The device software shall execute concurrent sampling, excursion evaluation, event capture, logging, and a scheduled connectivity/upload session such that a scheduled sample, a detected excursion, and a queued alert are never delayed beyond their deadlines by a lower-priority workload.
Rationale | The safety-critical sample/detect/alert path must have deterministic priority over discretionary upload and diagnostics; priority inversion here is a missed excursion.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-044, SYS-011/013, SYS-112 (alert priority), PRD-PERF-006; → SRS-CONN-008, SRS-DEV (scheduler priorities), SRS-PERF-002.
Notes | Alert-path prioritisation over bulk upload is realised in SRS-CONN-008; this fixes the on-device scheduling guarantee.
SRS-PERF-012 — Cloud ingest/alert response time and horizontal scalability
Requirement Statement | The cloud software shall ingest, persist, and route an excursion alert within a bounded processing time on receipt, shall serve fleet/record queries within a bounded response time, and shall scale horizontally to the target fleet size without breaching those bounds.
Rationale | The device-side 5-minute alert budget (SRS-PERF-007) includes a cloud processing/routing term; and a reusable fleet grows, so response times must hold as device count rises.
Priority | Must
Source | Performance
Verification Method | Test (load/latency), Analysis
Traceability | SYS-039 (end-to-end alert latency, S8 allocation), SYS-012 (routing/escalation), PRD-PERF-004; → SRS-CLOUD (ingest, scaling, availability), SRS-FLEET.
Notes | Absolute cloud latency and scale targets are owned/quantified by SRS-CLOUD; this binds the alert and query response budgets that the device- and user-facing latency depend on.
SRS-PERF-013 — Mobile application responsiveness (≤3 s)
Requirement Statement | The mobile application shall present requested device status, trip record summary, and configuration screens within 3 seconds of the user action under nominal network conditions, and shall remain responsive (non-blocking) while a background sync or BLE offload is in progress.
Rationale | A time-pressured operator configuring or reading a device must not be left waiting; a blocking UI during sync is a usability failure that pushes users to skip configuration steps.
Priority | Should
Source | User Need
Verification Method | Test (response time)
Traceability | PRD-PERF-007, SYS-041 (responsiveness posture), SYS-022 (config over mobile); → SRS-MOB.
Notes | 3 s is a Judgment target aligned with PRD-PERF-007, confirmed in the usability study; excludes first-time large record download, which shows progress.
---
8. Non-functional Requirements
SRS-PERF-014 — Battery life and power consumption over the 90-day cold-duty profile (Evidence-Pending)
Requirement Statement | The device software shall constrain its cumulative energy consumption — across wake/compute, sensing, logging, HMI, connectivity, and location — so that the device achieves at least 90 days of continuous armed operation from a single full charge under the defined worst-case cold-duty profile spanning −30 °C, retaining sensing, logging, on-device verdict, and connected-alert capability throughout, with a documented energy margin.
Rationale | The 90-day/−30 °C endurance is the headline program target and the binding constraint (D3); software duty-cycling is the primary lever, and a device that dies mid-trip produces an incomplete record and a false-reassurance hazard.
Priority | Must
Source | Performance
Verification Method | Test (90-day cold-endurance, MR-VAL-006), Analysis (energy budget)
Traceability | SYS-045 (90-day energy performance), SYS-046 (−30 °C retention), SYS-090 (energy-budget allocation), PRD-PERF-005; → SRS-REL, SRS-CONN-021 (radio energy), SRS-DEV (governor).
Notes | Evidence-Pending — the energy margin, the low-temperature usable-capacity figure, and the per-class (radio/GNSS/compute) energy allocations are unresolved and close on the 90-day cold-endurance test (MR-VAL-006) and the power budget (SYS-090). No battery-life figure beyond the ≥90-day system target is asserted (evidence). The program's top risk (MR-RISK-001), validated by test.
SRS-PERF-015 — Software energy conformance to the per-function allocation
Requirement Statement | Each software function that draws energy (sampling cadence, upload cadence, location cadence, HMI wake, diagnostics) shall operate within its allocated share of the documented energy budget and shall be individually configurable so its energy cost can be traded against the endurance target.
Rationale | The 90-day budget is only defensible if each energy-consuming function is bounded and tunable; an un-budgeted function silently erodes the margin.
Priority | Must
Source | Performance
Verification Method | Analysis, Test
Traceability | SYS-048 (location cadence within budget), SYS-090, SYS-171 (graceful degradation), PRD-PERF-005/010; → SRS-REL, SRS-CONN-021.
Notes | Per-function allocations are Evidence-Pending (SYS-090) and close on the power budget; this binds the software to whatever the allocation resolves to.
---
Error Handling
SRS-PERF-016 — Performance-degradation handling without missed samples
Requirement Statement | Under resource pressure (low memory, storage-write backpressure, or CPU saturation), the device software shall shed discretionary work (diagnostics, non-urgent upload, location) before ever delaying or dropping a scheduled sample, a detected excursion, or a queued alert, and shall log the degradation event.
Rationale | When performance headroom is exhausted, the safety-critical capture path must survive by design; degradation must be graceful and visible, never a silent missed sample.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-171 (graceful degradation order), SYS-021 (buffer-full behaviour), SYS-002, PRD-REL-008; → SRS-REL, SRS-PERF-011.
Notes | Shed order aligns with the system load-shedding order (SYS-171), which is Evidence-Pending on the depletion test.
---
Diagnostics
SRS-PERF-017 — Performance metrics collection
Requirement Statement | The device software shall record performance metrics — boot/wake time, per-cycle compute time, RAM high-water mark, write throughput and latency, missed-deadline count, and per-function energy estimates — to diagnostics for fleet triage and for closing the Evidence-Pending energy budget.
Rationale | Performance can only be managed and its Evidence-Pending targets closed if it is measured in the field, not just on the bench; missed-deadline counters surface latent timing risk before it corrupts a record.
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-113 (diagnostics), SYS-090, PRD-PERF-005; → SRS-DIAG, SRS-FLEET.
Notes | Metrics must themselves be energy-cheap (SRS-PERF-015) and are subordinate to the alert path.
---
Security Considerations
SRS-PERF-018 — Performance retention under security processing and at environmental extremes
Requirement Statement | The device software shall meet its latency, responsiveness, and endurance targets while performing the required cryptographic and integrity operations (TLS session, record signing, secure boot verification) and while operating at the −30 °C cold extreme and at low battery.
Rationale | Security processing and cold operation both add time and energy cost; a target met only without crypto or only at room temperature is not a real target for this device.
Priority | Must
Source | Performance
Verification Method | Test (incl. cold/low-battery), Analysis
Traceability | SYS-046 (cold retention), SYS-041 (responsiveness at cold/low-battery), PRD-PERF-005/007; → SRS-SEC (crypto mechanisms, EM-ENG-008), SRS-CONN-011.
Notes | Cryptographic mechanism performance is owned by SRS-SEC; this states the end-to-end retention obligation.
---
Verification Requirements
SRS-PERF-019 — Performance verification under worst-case and cold/low-battery conditions
Requirement Statement | The software performance requirements shall be verified by boot/wake-timing tests, detection- and alert-latency tests, verdict/HMI responsiveness tests at −30 °C and low battery, sustained write-throughput and concurrency tests, flash-endurance analysis/test, cloud load/latency tests, mobile-responsiveness tests, and the 90-day cold-endurance energy test (MR-VAL-006), with results traced to the requirements above.
Rationale | The timing, throughput, and endurance claims are only credible if verified at the worst-case environmental and load conditions the device actually meets in service, not extrapolated from room-temp bench runs.
Priority | Must
Source | Performance
Verification Method | Test, Analysis
Traceability | SYS-045/046, SYS-037/039/041/044, PRD-PERF-005/007, PRD-VER-004 (D3 validate-by-test); → SRS-VER.
Notes | The battery-life/energy targets are Evidence-Pending; this fixes the verification obligation (MR-VAL-006) that closes them.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-037/038/039 (timing/latency), SYS-041/042 (responsiveness/verdict), SYS-043/044 (capacity/throughput), SYS-045/046 (90-day/−30 °C endurance, D3), SYS-048/090 (location/energy allocation), SYS-169/171 (recovery/degradation); PRD-PERF-003…012, PRD-FUN-008/010, PRD-REL-004/006/008. Cross-module: cryptographic/TLS performance → SRS-SEC (EM-ENG-008); radio/GNSS energy and alert delivery → SRS-CONN; storage/scheduler/verdict engine → SRS-DEV; record format/sizing → SRS-DATA; cloud ingest/scaling/response → SRS-CLOUD; mobile responsiveness → SRS-MOB; degradation/recovery reliability → SRS-REL; delayed-alert hazard reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (battery-life margin, low-temperature usable capacity, per-class energy allocations, exact log sizing, RF-in-container alert latency, flash endurance rating) are carried as gaps, not invented, and close on their associated tests — chiefly the 90-day cold-endurance test (MR-VAL-006) and the RF-in-container test (PRD-VER-005).
SRS-1100 — Reliability (module SRS-REL)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-REL · Owner: Embedded Software Engineer (EM-ENG-007) with the Principal Reliability/Materials and Power engineers. Detailed software-reliability requirements decomposed from the System Requirements (SYS-163…176, plus SYS-002/007/017/167) and Product Requirements (PRD-REL 1200), realising fault-tolerant, fail-safe, self-healing behaviour that preserves the record and honest monitoring across power, link, and firmware faults — including the D3 −30 °C / 90-day cold-endurance tension. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM software survives power interruption, brown-out, crash, communication failure, and resource exhaustion without corrupting or losing the record or silently ceasing to monitor: power-fail-safe writes, crash and watchdog recovery, bounded retry policies, availability self-check and silent-outage prevention, graceful degradation under low energy, health monitoring and self-healing, and the software contribution to cold-endurance reliability over the reuse life.
2. Scope
Power-interruption and brown-out handling; crash/watchdog recovery; power-fail-safe atomic persistence; flash-endurance-aware writing over reuse; communication-failure tolerance (retry/back-off, no loss, no duplication); availability self-check and silent-outage prevention; graceful degradation and load-shedding under low energy; software fault tolerance, redundancy, and self-healing; device↔cloud reconciliation; health monitoring; and cold-endurance reliability (D3). Excludes the battery/BMS electrical reliability and cold derating (S2, Power engineer), storage-media endurance rating (S7), RF link margin (S5, SRS-CONN), cloud platform DR/SLA mechanisms (SRS-CLOUD), cryptographic integrity mechanisms (SRS-SEC), and the DFMEA/reliability prediction program itself (ERFKL, Stage 11); those are referenced where the software reliability interfaces with them.
3. Applicable Standards
IEC 62304 (software life-cycle, incl. software risk management); ISO 14971 (fail-safe/risk-control concepts); IEC 60068 (environmental reliability, cold endurance); IEC 62368-1 / IEC 62133 (single-fault and cell-level charge-limit concepts, referenced for the software charge-inhibit interface); handbook reliability-prediction methods (MIL-HDBK-217 / Telcordia SR-332 concepts) referenced by the reliability program. Cited by designation only; revision/method to confirm — no clause number or failure-rate figure is asserted as fact.
4. Assumptions
• The D3 −30 °C / 90-day cold-endurance set is the headline reliability promise and the top open program risk (MR-RISK-001); endurance margin, cold usable-capacity, reuse-cycle fade, and the graceful-degradation thresholds are Evidence-Pending and validated by test (MR-VAL-006, depletion test). No MTBF/failure-rate, cycle-count, availability-fraction, or reserve figure is invented here.
• Non-volatile media, atomic-write primitives, and buffer ordering are provided by SRS-DEV/SRS-DATA; this module specifies the reliability behaviour of the software that uses them.
• The battery/BMS provides charge-inhibit-below-limit and state-of-health signals (S2); this module specifies the software behaviour that consumes those signals (readiness gating, degradation, self-flagging).
• The cloud provides idempotent ingest, durable storage, and a reconciliation endpoint (SRS-CLOUD); this module specifies the device-side robust, resumable, de-duplicated behaviour and the reconciliation client.
• Safe-state, watchdog, and availability requirements are reconciled with the Hazard Analysis (HAL, Stage 10); several carry a reconcile-with-HAL status accordingly.
5. Architecture Overview
Reliability is built in three concentric rings. The persistence ring commits every sample, event, and audit record atomically so no power loss, brown-out, reset, or update can produce a torn or corrupt record. The recovery ring — brown-out detection, watchdog supervision, safe-state entry, and autonomous restart — returns a faulted device to armed monitoring without user intervention and logs every recovery so a recovered device is never assumed to have monitored continuously. The honesty ring — availability self-check, sensor-fault flagging, and graceful low-energy degradation — guarantees that an armed device is at all times either verifiably monitoring or visibly flagged as not, and that a depleting device protects the record to the last available energy. Above the device, device↔cloud reconciliation and bounded, idempotent, resumable upload compose the links into a guaranteed end-to-end delivery, with redundancy or self-healing applied by exception where the DFMEA identifies a single point of failure.
---
7. Functional Requirements
7.1 Power interruption, brown-out & persistence
SRS-REL-001 — Power-fail-safe atomic log writes
Requirement Statement | The device software shall commit each sample, event, and audit record to non-volatile storage atomically such that any committed record survives power loss, brown-out, reset, and firmware update intact, with no corruption of previously committed data and no partial/torn record surviving as valid.
Rationale | Record completeness/integrity is the product's core value; low-temperature brown-out under voltage sag (D3) makes power-fail-safe writes essential — a torn write must fail the record, not silently corrupt it.
Priority | Must
Source | Reliability
Verification Method | Test (power-cycle/brown-out interruption)
Traceability | SYS-168 (atomic writes), SYS-002 (gap-free), PRD-REL-005, PRD-FUN-017; → SRS-DATA, SRS-DEV (storage/journal).
Notes | Journaled/atomic-write design; brown-out detection threshold set with the Power engineer (PRD-ELEC-011).
SRS-REL-002 — Brown-out detection and orderly commit-then-park
Requirement Statement | On detecting supply voltage falling below the defined brown-out threshold, the device software shall complete or safely abort the in-flight write, flush the safe state, and park in a known state that preserves committed data, rather than continuing operations into an uncontrolled reset.
Rationale | Cold-induced voltage sag (D3) makes brown-out a routine event; an orderly commit-then-park converts a would-be corruption into a clean, recoverable stop.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-168, SYS-167 (safe state), PRD-REL-004/005, PRD-ELEC-011; → SRS-DEV, SRS-REL-001.
Notes | Brown-out threshold is Evidence-Pending on the Power engineer's confirmation; requirement fixes the behaviour at whatever threshold resolves.
SRS-REL-003 — Crash recovery to armed monitoring with preserved record
Requirement Statement | On a firmware crash, exception, or unexpected reset, the device software shall recover to an armed monitoring state, resume logging at the next scheduled interval, preserve all previously committed records, and record the crash and recovery as an attributable, time-stamped event.
Rationale | A crash must degrade to a logged, bounded gap and automatic resumption — never to a corrupted record or a silent unmonitored trip (false reassurance).
Priority | Must
Source | Reliability
Verification Method | Test (fault injection)
Traceability | SYS-167 (safe state/recovery), SYS-169 (recovery event), SYS-002/017, PRD-REL-004; → SRS-DIAG (crash dump), SRS-PERF-009 (recovery time).
Notes | The recovery gap is an explicit gap marker (SYS-002) that disqualifies an unqualified accept (SYS-017); reconcile with HAL.
SRS-REL-004 — Flash-endurance-aware writing over the reuse life
Requirement Statement | The device software shall bound cumulative program/erase cycles over the specified reuse/charge-cycle life through wear-levelling, write coalescing, and journaling, and shall detect and flag approaching end-of-endurance rather than allow silent write failure or record corruption.
Rationale | A reusable multi-trip device (D1) rewrites the log many times over life; exceeding flash endurance corrupts the record, so wear must be budgeted and monitored, and end-of-life must self-flag.
Priority | Must
Source | Reliability
Verification Method | Analysis, Test
Traceability | SYS-165 (endurance over reuse), SYS-168, PRD-REL-002/005; → SRS-PERF-005, SRS-DIAG.
Notes | Endurance rating and per-trip write budget are Evidence-Pending on the final storage device; requirement fixes the management and self-flag obligation.
7.2 Recovery, watchdog & safe state
SRS-REL-005 — Defined safe state on any fault
Requirement Statement | On any power, connectivity, firmware, sensor, or reset fault, the device software shall enter a defined known safe state that preserves committed logged data and an honest monitoring status, attempt automatic recovery to normal monitoring, and never enter a state that corrupts the record or falsely reads as monitoring.
Rationale | A fault must never corrupt the authoritative record nor leave the device silently non-monitoring; the safe state is the fail-safe anchor of the reliability design.
Priority | Must
Source | Reliability
Verification Method | Test (fault injection)
Traceability | SYS-167 (safe state), PRD-REL-004; → SRS-SAFE / HAL (Stage 10), SRS-REL-003.
Notes | Safe-state definition is Judgment pending HAL reconciliation (Stage 10); closed by fault-injection test. Cross-flag Safety engineer.
SRS-REL-006 — Watchdog supervision and autonomous recovery with logged recovery event
Requirement Statement | The device firmware shall be supervised by a watchdog that returns a hung or unresponsive device to an armed monitoring state without user intervention, and shall record every watchdog recovery as an attributable, time-stamped event so a recovered device is never assumed to have monitored continuously.
Rationale | A silent hang creates an unmonitored trip believed monitored — a false-reassurance hazard; the logged recovery event feeds the honest-monitoring indication (SYS-017).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-169 (watchdog recovery), SYS-017 (honest indication), PRD-REL-006; → SRS-DIAG, SRS-PERF-009.
Notes | Watchdog must also cover the low-power sleep/wake path so a wake that fails to complete is caught.
SRS-REL-007 — Bounded, energy-aware retry policy on communication failure
Requirement Statement | On a failed connection, session, or transfer, the device software shall retry using a bounded exponential back-off with a configured maximum interval and per-window attempt ceiling, such that repeated failures never drain the energy budget through unbounded retries and never cause loss of a committed record.
Rationale | Naïve immediate retry in poor coverage (e.g. inside the shipper) is a primary energy leak; bounded back-off recovers when coverage returns without exhausting the budget.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-174 (retry/back-off, no data loss), SYS-089 (energy governor), PRD-REL-011, PRD-CONN-008; → SRS-CONN-009 (transport back-off), SRS-PERF-015.
Notes | Transport-layer realisation is in SRS-CONN-009; alerts retain priority within the ceiling (SRS-CONN-008). Ties to the D3 energy budget.
7.3 Communication robustness & end-to-end delivery
SRS-REL-008 — Loss-free, duplicate-free upload across intermittent links
Requirement Statement | The device software shall preserve all buffered records across any number of connectivity interruptions and, on reconnection, resume upload in intact chronological order with idempotent (exactly-once) cloud ingest, such that transient link conditions never cause loss of a committed record and never produce a duplicate record in the cloud store.
Rationale | Global lanes have unreliable connectivity; ordered, idempotent, resumable upload protects the ALCOA+ completeness and consistency of the record (SYS-020, D2).
Priority | Must
Source | Reliability
Verification Method | Test (lossy/intermittent link)
Traceability | SYS-174 (communication robustness), SYS-020 (ordered/resumable/de-duplicated), PRD-REL-011, PRD-CONN-008; → SRS-CONN-007, SRS-CLOUD (idempotent ingest).
Notes | Transport realisation in SRS-CONN-007; buffer ordering in SRS-DEV. Cross-flag Connectivity engineer on RF-in-container link margin (SYS-104-line).
SRS-REL-009 — End-to-end delivery reliability with device↔cloud reconciliation and self-healing
Requirement Statement | The software shall ensure that every committed sample, event, and excursion is eventually and correctly delivered to the authoritative cloud record over a trip, reconciling device and cloud stores so any undelivered or divergent record is detected, re-delivered, or flagged — with redundancy or self-healing applied where a single point of failure would otherwise threaten the record.
Rationale | Reliability of the whole exceeds the sum of links: offline-first capture, robust upload, and a durable platform must compose into guaranteed end-to-end delivery, with reconciliation catching any silent gap.
Priority | Must
Source | Reliability
Verification Method | Test (end-to-end trip reliability)
Traceability | SYS-176 (end-to-end delivery/reconciliation), SYS-018/174/175, PRD-REL-005/011/012; → SRS-CLOUD, SRS-CONN, SRS-DATA.
Notes | Redundancy/self-healing applied by exception where the DFMEA (ERFKL, Stage 11) identifies an SPOF; reconcile with HAL.
7.4 Availability, degradation & self-healing
SRS-REL-010 — Availability self-check and silent-outage prevention
Requirement Statement | While armed, the device software shall periodically self-check its sensing→evaluation→logging chain and, on detecting an outage, shall flag the device as not monitoring (locally on the HMI and to the cloud) such that an armed device is at all times either verifiably monitoring or visibly flagged as not, never silently outaged.
Rationale | "Armed" must reliably mean "monitoring"; a device that appears alive while not monitoring is a false-reassurance hazard, so any gap must surface in the verdict and to the fleet (SYS-017).
Priority | Must
Source | Reliability
Verification Method | Test (availability/self-check)
Traceability | SYS-170 (availability self-check), SYS-047 (availability, no silent outage), SYS-017, PRD-REL-007, PRD-PERF-009; → SRS-DIAG, SRS-SAFE / HAL.
Notes | Self-check cadence and availability fraction are Evidence-Pending pending HAL; closed by the self-check test. Cross-flag Safety and HMI engineers.
SRS-REL-011 — Sensor plausibility and fault flagging without discarding raw data
Requirement Statement | The device software shall evaluate each sensor reading against physically plausible bounds and record a data-quality flag when a reading is implausible, out-of-range, or from a faulted channel, without discarding the raw value, and shall not compute an unqualified accept verdict from flagged data.
Rationale | A silently-wrong reading is more dangerous than a flagged one (no false reassurance); the raw value is retained for forensic reconstruction.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-007 (plausibility/fault flag), SYS-017 (honest verdict), PRD-REL-004, PRD-SAFE-003; → SRS-DEV, SRS-DIAG.
Notes | Feeds the honest-monitoring indication and readiness gating; reconcile with HAL.
SRS-REL-012 — Graceful degradation under low energy (defined load-shedding order)
Requirement Statement | As available energy depletes, the device software shall shed discretionary functions in a defined priority order — location and non-essential uploads first, then upload cadence, then non-essential HMI — while preserving core logging, excursion detection/alerting, and the last-committed authoritative record to the last available energy, retaining a reserve sufficient for an orderly graceful-degradation sequence.
Rationale | A depleting device (heightened at −30 °C, D3) must protect the record over convenience; core logging/alerting is the last thing shed, and end-of-trip must be a controlled landing, not an uncontrolled shutdown.
Priority | Must
Source | Reliability
Verification Method | Test (depletion)
Traceability | SYS-171 (degradation order), SYS-164 (end-of-trip reserve), SYS-021 (buffer-full), PRD-REL-008; → SRS-PERF-016, SRS-CONN.
Notes | Shed order, energy thresholds, and reserve fraction are Evidence-Pending and close on the depletion test (SYS-171/164); Power engineer confirms per-load budget.
7.5 Health monitoring & cold-endurance reliability (D3)
SRS-REL-013 — Continuous health monitoring and self-healing
Requirement Statement | The device software shall continuously monitor its own health indicators (task liveness, memory high-water mark, storage-write success, sensor liveness, clock discipline, battery state-of-health) and shall take defined self-healing actions (task restart, buffer re-init, resource reclaim, safe-state entry) on a detected anomaly, logging each action as an attributable event.
Rationale | Early detection and bounded self-healing of latent faults prevent a small anomaly from becoming a corrupted record or an unmonitored trip; logging keeps self-healing honest.
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-170 (self-check), SYS-169 (recovery), SYS-007, PRD-REL-006/007; → SRS-DIAG, SRS-FLEET.
Notes | Self-healing must never mask a fault from the record or the verdict (no false reassurance).
SRS-REL-014 — Software contribution to cold-endurance reliability and readiness self-flagging (D3)
Requirement Statement | The device software shall sustain sensing, logging, on-device verdict, and scheduled connectivity attempts across the 90-day worst-case cold-duty profile with zero loss of committed data, and shall reliably detect when battery state-of-health can no longer meet that endurance and block the device from being armed/redeployed until serviced.
Rationale | The D3 headline promise: a device that dies in the cold loses the record when it matters, and capacity fade over the reuse loop can silently cause mid-life dead-in-cold; a reusable fleet must fail safe by self-flagging, not by an unmonitored trip (MR-RISK-001).
Priority | Must
Source | Reliability
Verification Method | Test (90-day cold-endurance MR-VAL-006; cycle-life)
Traceability | SYS-163 (90-day cold endurance), SYS-164/165 (reserve, cycle-life), SYS-030 (readiness gating), PRD-REL-001/002; → SRS-PERF-014, SRS-FLEET, SRS-DIAG.
Notes | Evidence-Pending — endurance margin, cold usable-capacity, and per-cycle fade budget are unresolved and validated by test (MR-VAL-006, cycle-life test). No cycle count or margin figure asserted here (evidence); the program's top risk. Charge-inhibit reliability (SYS-166) is a Power/BMS interface, cross-flagged to Safety.
---
8. Non-functional Requirements
SRS-REL-015 — Software reliability program participation and failure-rate allocation
Requirement Statement | The device and cloud software shall be developed and analysed under the documented reliability program (software DFMEA/SWFMEA, fault-injection coverage, and failure-rate allocation), providing the evidence and the software failure-mode inputs the program requires to demonstrate the endurance, availability, and delivery targets.
Rationale | Reliability must be engineered and demonstrated, not assumed; the software must contribute its failure modes and coverage evidence to the system reliability prediction (ERFKL, Stage 11).
Priority | Should
Source | Reliability
Verification Method | Analysis
Traceability | SYS-173 (reliability program), PRD-REL-010; → ERFKL SWFMEA (Stage 11), SRS-VER.
Notes | MTBF/software-failure-rate targets and prediction method are Evidence-Pending (SYS-173) and set by the program; not invented here.
---
Error Handling
SRS-REL-016 — No-silent-failure invariant across all software faults
Requirement Statement | The software shall handle every detected fault (power, storage, sensor, link, firmware, resource) by preserving committed data, logging an attributable failure event, and either recovering to monitoring or visibly flagging not-monitoring, and shall never respond to a fault by silently discarding data or silently continuing in a degraded-but-unflagged state.
Rationale | The unifying reliability invariant: every failure is either recovered-and-logged or visibly-flagged; a silent failure is the one outcome the design forbids (false reassurance).
Priority | Must
Source | Reliability
Verification Method | Test (fault injection across fault classes)
Traceability | SYS-167/169/170, SYS-017/021, PRD-REL-004/006/007; → SRS-DEV, SRS-DIAG, SRS-CONN-023.
Notes | This is the reliability counterpart of the connectivity no-silent-failure requirement (SRS-CONN-023); reconcile with HAL.
---
Diagnostics
SRS-REL-017 — Reliability event and health telemetry to the fleet
Requirement Statement | The software shall record and make available to fleet tooling all reliability events — resets, watchdog recoveries, brown-outs, crashes, self-heal actions, availability outages, sensor faults, degradation transitions, and battery state-of-health — with attributable timestamps for triage and reliability-growth tracking.
Rationale | A reusable fleet's reliability can only be managed and improved if failures are observable in the field; this telemetry also feeds the reliability program and the DFMEA (ERFKL).
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-170/173, SYS-030 (readiness), PRD-REL-007/010; → SRS-DIAG, SRS-FLEET.
Notes | Reliability telemetry is energy-cheap and subordinate to the alert path (SRS-PERF-015).
---
Security Considerations
SRS-REL-018 — Reliability mechanisms shall not weaken record integrity or tamper-evidence
Requirement Statement | Recovery, self-healing, retry, and degradation mechanisms shall preserve the tamper-evidence, attribution, and append-only integrity of the record and audit trail, and shall not create a path by which a record can be altered, deleted, or fabricated under the guise of a recovery action.
Rationale | Fault-recovery is a common weak point where integrity guarantees are accidentally bypassed; a recovery must never become an integrity or non-repudiation gap.
Priority | Must
Source | Security
Verification Method | Test, Analysis
Traceability | SYS-026/028 (tamper-evident, append-only audit), SYS-167, PRD-SEC-004; → SRS-SEC (mechanisms, EM-ENG-008), SRS-DATA.
Notes | Cryptographic/tamper-evidence mechanisms owned by SRS-SEC; this states the reliability-side obligation not to undermine them.
---
Verification Requirements
SRS-REL-019 — Reliability verification by fault injection, power-cycle, lossy-link, depletion, and endurance test
Requirement Statement | The software reliability requirements shall be verified by power-cycle/brown-out storage-integrity tests, fault-injection and crash-recovery tests, watchdog and availability self-check tests, lossy/intermittent-link and end-to-end delivery tests, low-energy depletion/degradation tests, and the 90-day cold-endurance and cycle-life tests (MR-VAL-006), with results traced to the requirements above.
Rationale | The fail-safe, no-loss, no-silent-outage, and cold-endurance claims are only credible if verified under the actual fault, link, energy, and cold conditions the device meets in service.
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-163…176, PRD-REL-001…012, PRD-VER-004 (D3 validate-by-test); → SRS-VER.
Notes | The D3 cold-endurance targets and degradation thresholds are Evidence-Pending; this fixes the verification obligation (MR-VAL-006) that closes them.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-163/164/165 (cold endurance, reserve, cycle-life, D3), SYS-166 (charge-safety interface), SYS-167/168/169 (safe state, atomic writes, watchdog), SYS-170/171 (availability, degradation), SYS-172/173 (robustness over reuse, reliability program), SYS-174/175/176 (communication robustness, platform, end-to-end delivery), SYS-002/007/017/021 (gap honesty, plausibility, honest verdict, buffer-full); PRD-REL-001…012, PRD-FUN-017, PRD-SAFE-003, PRD-ELEC-011. Cross-module: atomic-write/storage/scheduler → SRS-DEV; record format/reconciliation → SRS-DATA; transport robustness/retry → SRS-CONN; idempotent ingest/durable storage/DR → SRS-CLOUD; energy budget and degradation → SRS-PERF; tamper-evidence/integrity → SRS-SEC (EM-ENG-008); software DFMEA → ERFKL (Stage 11); safe-state/availability hazard reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (cold-endurance margin, cold usable-capacity, reuse-cycle fade, brown-out threshold, availability fraction, degradation thresholds/reserve, flash endurance rating, software failure-rate) are carried as gaps, not invented, and close on their associated tests — chiefly the 90-day cold-endurance and cycle-life tests (MR-VAL-006), the depletion test, and fault-injection/power-cycle tests.
SRS-1200 — Safety (module SRS-SAFE)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-SAFE · Owner: Embedded Software Engineer (EM-ENG-007) with the Safety/Risk Engineer and, for the battery-thermal controls, the Battery Engineer. Detailed software safety / risk-control requirements decomposed from the Safety System Requirements (SYS-149…162) and the safety-relevant Functional and Interface System Requirements (SYS-011/013/017/023/032, SYS-057/058/067), realising the preliminary Product safety requirements (PRD-SAFE-001…012). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
MODULE STATUS — PRELIMINARY (reconcile with the Hazard Analysis Library, Stage 10, ORCH-017).
Per the org risk process (CLAUDE.md rule 3; ISO 14971), software risk-control requirements originate in
the Hazard Analysis Library (HAL). Every requirement in this module is a preliminary software
risk-control requirement: it is decomposed from a parent SYS-/PRD-SAFE requirement and the concept, and
each shall be traced, at HAL reconciliation, to one or more HAL hazard records (HA-###) and Risk-Control
Requirements (RCR-###) — in particular the software-controlled-hazard set (HAL-1000 line) — and then
re-baselined here from Preliminary to Baselined. Action Priority is severity-first: severity is fixed by
outcome (stated per requirement); probability and detectability targets are Evidence-Pending until HAL.
All requirements in this module carry Priority: Must. Standards are cited by designation, revision to
confirm — no clause number is asserted as fact.
The two program safety tensions are carried explicitly in software: the D3 dead-in-cold / charge-at-low-
temperature hazard (SRS-SAFE-010, SRS-SAFE-011) and the severity anchor that both disposition-error
directions harm the beneficiary (SRS-SAFE-001, SRS-SAFE-004, SRS-SAFE-005).
1. Purpose
Specify the software behaviours on the SCM device (and, where the device software cooperates with it, the cloud/mobile tier) that act as risk controls — the software mechanisms that keep the device from contributing to a wrong product-disposition decision or to a physical (battery/thermal) hazard. The module defines how software enters and holds a safe state, keeps alerts and the verdict honest, annunciates safety-relevant conditions, degrades and fails safe, controls charging in the cold, detects a dead-in-cold record gap, and logs and self-checks its own safety functions.
2. Scope
Software risk-control framework and its HAL traceability; software-commanded safe-state entry and maintenance; honest-gap / no-false-reassurance verdict behaviour; verdict integrity; alarm management (no fatigue-induced miss); operator notification of safety-relevant states; emergency/degraded operation; fail-safe behaviour on fault/reset/OTA; software charge-inhibit control in the cold; dead-in-cold detection and annunciation; independent safety monitoring; diagnostic coverage of safety functions; and safety logging. Excludes the hazard analysis itself and the residual-risk acceptance (HAL / ISO 14971 risk file, Stage 10, SYS-161/162); the BMS/cell protection hardware and the physical charge-inhibit gate (S2, PRD-ELEC — this module owns the software command and annunciation); the cryptographic mechanisms behind tamper-evidence and signing (SRS-SEC); and the numeric detection/alarm/energy targets (SRS-PERF, closed by validation). Those are referenced where the safety software interfaces with them.
3. Applicable Standards
ISO 14971 (risk management, severity-dominant); IEC 62304 (medical-device software life-cycle and software safety classification); IEC 62366-1 (use-related risk / use-error safety); IEC 62368-1 and Li-ion charge-temperature limits (battery/thermal safety, for the software charge-inhibit control); alarm-management principles (alert balance); ALCOA+ / EU GDP and 21 CFR Part 11 (record completeness underpinning no-false-reassurance). Cited by designation only; revision to confirm — no clause number is asserted as fact. IEC 62304 software safety class is Evidence-Pending on the SYS-178 regulatory classification.
4. Assumptions
• The HAL (Stage 10) will confirm each HA-###/RCR-###, the software safety class per IEC 62304, and the severity-first Action Priority; until then every requirement here is Preliminary and its probability/detectability targets are Evidence-Pending (SYS-162).
• The physical protections — BMS over-charge/-discharge/-current/-temperature/short protection and the hardware charge-inhibit gate below the low-temperature limit — are provided by the power subsystem (S2, SYS-155/156, PRD-ELEC); this module specifies the software command, supervision, and annunciation of those states, not the hardware protection itself.
• Non-volatile record durability, tamper-evidence, atomic/power-fail-safe writes, gap marking, and the trusted clock are provided by the device-software module (SRS-DEV-008/009/010/011/021/027); the safe-state and honest-gap behaviours here build on those guarantees.
• The offline on-device accept/reject/quarantine verdict and its display path exist (SRS-DEV / SYS-015, SYS-058, I-07); this module constrains their safety behaviour (honesty, integrity, failed-element detection).
• The final low-temperature charge limit (nominally ≥0 °C) is Evidence-Pending on the cell datasheet (SYS-155); the software applies the configured qualified limit, whatever its final value.
5. Architecture Overview
A safety supervisor within the device firmware sits above the normal capture→detect→decide path. It continuously evaluates a small set of safety invariants — record integrity, monitoring continuity, verdict consistency, charge-temperature state, energy/thermal state — and can command the device into a defined safe state independently of the normal application flow. The supervisor is the single authority that (a) blocks or qualifies an "accept" verdict when monitoring was incomplete, (b) commands and annunciates the cold charge-inhibit, (c) annunciates a dead-in-cold power-loss gap on restoration, (d) drives operator notification of any safety-relevant state, and (e) writes a safety log distinct from routine diagnostics. Alarm balance (hysteresis/debounce/consolidation) is applied in the detection path but governed here as a risk control. All safety functions are covered by self-diagnostics whose failure itself forces the safe state (no-silent-failure, SRS-DEV-028). Every control in this module is a preliminary risk control pending HAL.
---
7. Functional Requirements (software risk controls)
7.1 Risk-control framework & safe state
SRS-SAFE-001 — Software risk controls implemented and HAL-traceable
Requirement Statement | Each software behaviour in this module designated a risk control shall be implemented as an identifiable, verifiable software function, and shall carry a bidirectional trace from its parent safety requirement to its HAL hazard record(s) (HA-###) and Risk-Control Requirement(s) (RCR-###), the software design output implementing it, and the evidence verifying its effectiveness.
Rationale | A risk control that cannot be pointed to in the software and traced to a hazard and its verification cannot be shown to reduce risk; traceability is what makes the software safety case auditable and reconcilable with the HAL (severity anchor: both disposition-error directions harm the beneficiary).
Priority | Must
Source | Regulatory (Safety)
Verification Method | Inspection, Analysis
Traceability | SYS-162 (hazard traceability), SYS-149 (severity anchor), PRD-SAFE-001/011; → HAL HA-###/RCR-### (Stage 10), SRS-VER, ORCH-018/EM-ENG-012. Standard: ISO 14971, IEC 62304 (revision to confirm).
Notes | PRELIMINARY. Software safety class per IEC 62304 is Evidence-Pending on SYS-178.
SRS-SAFE-002 — Software-commanded safe-state entry
Requirement Statement | On detection of any fault condition from which continued normal operation would be unsafe or misleading, the device software shall command the device into the defined safe state — preserving stored trip data intact, holding the charge subsystem in its non-hazardous condition, and not presenting a misleading monitoring or verdict indication — within a bounded, verifiable transition.
Rationale | Fail-safe behaviour requires an explicit, testable transition into a defined state the software fails into; an undefined or unbounded transition cannot be verified as a risk control.
Priority | Must
Source | Security (Safety)
Verification Method | Test (fault injection)
Traceability | SYS-159 (defined safe state), SYS-158 (fail-safe), PRD-SAFE-010, PRD-ARCH-008; → SRS-DEV-024 (safe mode), SRS-DEV-019 (fault mgmt); → HAL HA-###/RCR-###. Standard: ISO 14971 safe-state (revision to confirm).
Notes | PRELIMINARY. Severity: unsafe/misleading operation. Safe-state definition owned jointly with SYS-159; software owns entry.
SRS-SAFE-003 — Safe-state maintenance and honest indication while faulted
Requirement Statement | While in the safe state, the device software shall maintain data integrity and the non-hazardous charge/energy condition, shall continue to represent the device's true state to the operator, and shall not resume normal monitoring or present an unqualified "accept" verdict until the fault is cleared and readiness is re-established.
Rationale | A device that silently self-recovers from a fault into a normal-looking state can mask a lapse; the safe state must be honest and sticky until the condition is genuinely resolved.
Priority | Must
Source | Security (Safety)
Verification Method | Test (fault injection), Demonstration
Traceability | SYS-159, SYS-151 (no false reassurance), PRD-SAFE-010/003; → SRS-DEV-024/028, SRS-SAFE-004; → HAL HA-###/RCR-###.
Notes | PRELIMINARY. Ties the no-silent-failure invariant (SRS-DEV-028).
7.2 Verdict & record integrity (informational safety)
SRS-SAFE-004 — Honest-gap / no-false-reassurance verdict behaviour
Requirement Statement | When the trip record contains a monitoring gap, sensor fault, power-loss period, interrupted armed interval, or unresolved data-quality flag, the device software shall not present an unqualified clean-pass / "accept" verdict for that trip and shall make the incompleteness visible on-device and in the record.
Rationale | A silent gap that reads as a clean pass produces a wrong disposition on an incomplete record; an honest gap is safer than false confidence — the dominant informational hazard direction (severity anchor).
Priority | Must
Source | Security (Safety)
Verification Method | Test
Traceability | SYS-151 (no false reassurance), SYS-017, SYS-002 (gap marking), PRD-SAFE-003, PRD-FUN-010; → SRS-DEV-011 (gap marking), SRS-DEV-007; → HAL HA-###/RCR-###. Standard: ALCOA+ completeness (revision to confirm).
Notes | PRELIMINARY. Severity: wrong disposition on incomplete record.
SRS-SAFE-005 — Verdict integrity and device/cloud consistency
Requirement Statement | The device software shall derive the on-device accept/reject/quarantine verdict from the same excursion evaluation and active profile as the authoritative record, shall protect the verdict and its inputs against undetected corruption, and shall detect and flag any divergence between the on-device verdict and the cloud verdict for the same trip and profile rather than presenting either as authoritative in silence.
Rationale | Two sources of truth that silently disagree, or a corrupted verdict shown as valid, destroy trust in disposition and can direct a wrong accept/reject; consistency and integrity are the informational risk control.
Priority | Must
Source | Security (Safety)
Verification Method | Test, Analysis
Traceability | SYS-016 (verdict consistency), SYS-149, SYS-058 (no false HMI state), PRD-FUN-011, PRD-SAFE-001; → SRS-DEV-009 (tamper-evidence), SRS-SEC; → HAL HA-###/RCR-###.
Notes | PRELIMINARY. A failed display element must not turn "reject" into "accept" (SYS-058).
7.3 Alarm management & operator notification
SRS-SAFE-006 — Alarm management to avoid fatigue-induced miss
Requirement Statement | The software shall apply the configured hysteresis, debounce, and alert- consolidation parameters to excursion alerting so that the false-alarm rate and the missed-excursion rate both meet their validated targets, and shall not suppress or consolidate a genuine, product-relevant excursion in a way that could cause it to be missed.
Rationale | Over-sensitivity causes fatigue and muted real alerts; under-sensitivity causes missed excursions — both directions are safety-relevant. Alarm balance is a two-sided risk control.
Priority | Must
Source | Security (Safety)
Verification Method | Test (validation)
Traceability | SYS-152 (alarm balance), SYS-014, SYS-010, PRD-SAFE-004, PRD-FUN-007; → SRS-PERF (numeric false-alarm/miss targets), SRS-DEV-013; → HAL HA-###/RCR-###. Standard: alarm-management principles (revision to confirm).
Notes | PRELIMINARY. Numeric targets held in SRS-PERF/validation (evidence).
SRS-SAFE-007 — Operator notification of safety-relevant states
Requirement Statement | The device software shall annunciate each safety-relevant state — void/mis-armed trip, incomplete-monitoring/gap condition, charge-inhibited, low-battery-degradation, sensor/verdict fault, and safe-state entry — unmistakably on the on-device HMI (and, where connectivity allows, to the responder), such that the state cannot be mistaken for normal healthy operation.
Rationale | A safety-relevant state that is not annunciated to the accountable human is not a risk control; the minimally-trained, gloved, offline edge user must be able to see that something is wrong.
Priority | Must
Source | User Need (Safety)
Verification Method | Test (summative usability), Demonstration
Traceability | SYS-154 (mis-arm unmistakable), SYS-153 (use-error protection), SYS-058, PRD-SAFE-005/006, PRD-UX-004; → SRS-DEV-024, SRS-SAFE-002; → HAL HA-###/RCR-###. Standard: IEC 62366-1 (revision to confirm).
Notes | PRELIMINARY. Feeds the Usability Engineering File; login-free edge read preserved.
7.4 Emergency / degraded operation & fail-safe
SRS-SAFE-008 — Emergency / degraded operation preserving monitoring and record
Requirement Statement | Under resource-constrained conditions (low battery, storage near capacity, connectivity loss, or a partial subsystem fault), the device software shall degrade in a defined order that preserves — for as long as physically possible — excursion detection, local capture of every excursion, and the integrity of the trip record and audit trail, shall shed only non-safety-relevant functions first, and shall never silently cease monitoring.
Rationale | Graceful degradation must protect the most safety-relevant capability last; the last mile is never a black box, and a silent stop-monitoring is a false-negative record hazard.
Priority | Must
Source | Reliability (Safety)
Verification Method | Test (fault injection)
Traceability | SYS-158, SYS-021 (buffer-full graceful behaviour), SYS-013 (local capture), PRD-SAFE-010, PRD-REL-004; → SRS-DEV-017 (low-battery degradation ordering), SRS-DEV-008; → HAL HA-###/RCR-###.
Notes | PRELIMINARY. Degradation ordering owned with SRS-DEV-017.
SRS-SAFE-009 — Fail-safe behaviour on fault, reset, and OTA
Requirement Statement | After any power, connectivity, firmware, watchdog-reset, or OTA event, the device software shall fail safe — preserving the logged trip data intact, returning to a defined known safe or known-good state, and neither corrupting the record nor silently ceasing to monitor — and shall never leave a fielded device bricked or unmonitored-without-notice as a result of an update.
Rationale | A fault or update must never corrupt the evidentiary record or leave the device unsafe or silently unmonitored; fail-safety preserves both safety and the record.
Priority | Must
Source | Security (Safety)
Verification Method | Test (fault injection, interrupted-OTA)
Traceability | SYS-158 (fail-safe), SYS-032 (fail-safe OTA), SYS-065 (OTA interface), PRD-SAFE-010, PRD-ARCH-008; → SRS-DEV-010/018/024, SRS-DEV-031, SRS-OTA; → HAL HA-###/RCR-###. Standard: IEC 62304 (revision to confirm).
Notes | PRELIMINARY. OTA integrity/rollback mechanism in SRS-OTA; this states the fail-safe outcome.
7.5 Battery / thermal software safety (D3)
SRS-SAFE-010 — Software charge-inhibit control and annunciation in the cold
Requirement Statement | The device software shall command inhibition of battery charging whenever cell temperature is below the configured qualified low-temperature charge limit (nominally ≥0 °C, final limit per cell datasheet), shall verify that the charge subsystem has entered the inhibited state, and shall clearly annunciate the charge-inhibited condition to the user rather than allowing charging to fail silently.
Rationale | Charging a cold Li-ion cell drives lithium plating — an internal-short and thermal-event precursor — a physical hazard created directly by the D3 low-temperature operating envelope; the software command, verification, and annunciation are the software layer of that control.
Priority | Must
Source | Security (Safety)
Verification Method | Test (charge-inhibit)
Traceability | SYS-155 (charge-inhibit below safe cell temp), SYS-053 (charge-path inhibit), PRD-SAFE-007, PRD-ELEC-005/006; → SRS-DEV-015 (charging state), S2/BMS (physical gate); → HAL HA-###/RCR-###. Standard: IEC 62368-1, Li-ion charge-temperature limits (final limit Evidence-Pending on cell datasheet; revision to confirm).
Notes | PRELIMINARY. Physical protection owned by S2/BMS; software owns command + verify + annunciate. Severity: thermal event.
SRS-SAFE-011 — Dead-in-cold detection and annunciation (D3 informational failure)
Requirement Statement | On power restoration after a deep-cold power-loss or brown-out, the device software shall detect that an armed-trip interval was not monitored, shall represent that interval as an explicit power-loss gap in the record rather than resuming as if monitoring were continuous, and shall annunciate the dead-in-cold gap on-device so the trip cannot read as uneventful across the missing interval.
Rationale | A device that dies silently in the cold produces a false-negative record — the missing interval reads as uneventful — which can force a wrong disposition; reframing the D3 endurance tension as a disposition-safety control means the software must never let a cold-death gap masquerade as clean data.
Priority | Must
Source | Reliability (Safety)
Verification Method | Test (90-day cold-endurance, power-loss injection)
Traceability | SYS-160 (dead-in-cold prevention), SYS-002/017, PRD-SAFE-012, PRD-REL-001; → SRS-DEV-011 (gap marking), SRS-DEV-027 (retention), SRS-SAFE-004; → HAL HA-###/RCR-###. Standard: ALCOA+ (revision to confirm).
Notes | PRELIMINARY. Endurance numbers Evidence-Pending, closed by cold-endurance test.
7.6 Safety monitoring, diagnostic coverage & logging
SRS-SAFE-012 — Independent safety monitoring / supervision
Requirement Statement | The device software shall run a safety supervisor, independent of the normal application flow, that continuously checks the safety invariants (record-write integrity, monitoring continuity, verdict consistency, charge-temperature state, energy/thermal state), and that commands the safe state on any invariant violation it detects.
Rationale | Safety-relevant functions must be watched by a mechanism that does not itself depend on the path it supervises; independent monitoring is what catches a hung or misbehaving application before it produces an unsafe outcome.
Priority | Must
Source | Security (Safety)
Verification Method | Test (fault injection), Analysis
Traceability | SYS-158/159, SYS-067 (interface fallback to safe state), PRD-SAFE-010, PRD-ARCH-008; → SRS-DEV-018 (watchdog), SRS-DEV-019, SRS-SAFE-002; → HAL HA-###/RCR-###. Standard: IEC 62304 (revision to confirm).
Notes | PRELIMINARY. Cooperates with the hardware watchdog (SRS-DEV-018).
SRS-SAFE-013 — Diagnostic coverage of safety functions
Requirement Statement | The device software shall provide diagnostic coverage of each safety function and its sensing/actuation path — including sensor plausibility/fault detection, storage-write integrity, HMI element integrity, charge-inhibit path, and clock validity — such that a latent failure of a safety function is detected and forced to a fail-safe or annunciated state rather than remaining silently ineffective.
Rationale | A risk control that has silently failed provides no protection; diagnostic coverage converts a latent safety-function failure into a detected, annunciated, or fail-safe condition.
Priority | Must
Source | Reliability (Safety)
Verification Method | Test (fault injection)
Traceability | SYS-007 (sensor plausibility), SYS-057 (write integrity), SYS-058 (HMI integrity), SYS-006 (clock), PRD-SAFE-003, PRD-REL-004; → SRS-DEV-003/028/029, SRS-DIAG; → HAL HA-###/RCR-###.
Notes | PRELIMINARY. Diagnostic-coverage target Evidence-Pending until HAL.
SRS-SAFE-014 — Safety event logging
Requirement Statement | The device software shall record, to durable append-only storage distinct from routine diagnostics, every safety-relevant event — safe-state entry/exit and cause, charge-inhibit activation, dead-in-cold gap, verdict divergence, alarm-suppression decisions, and safety-function diagnostic failures — each attributable, time-stamped against the trusted clock, and tamper-evident.
Rationale | The effectiveness and residual risk of a control cannot be evaluated post-market without an attributable, tamper-evident record of when it acted; safety logging is the evidence base for the HAL and post-market surveillance.
Priority | Must
Source | Regulatory (Safety)
Verification Method | Test, Inspection
Traceability | SYS-028 (audit trail), SYS-006 (trusted timestamp), SYS-162, PRD-SAFE-011, PRD-SEC-004; → SRS-DEV-009/012/021, SRS-DATA; → HAL HA-###/RCR-###. Standard: 21 CFR Part 11 audit trail, ALCOA+ (revision to confirm).
Notes | PRELIMINARY. Feeds residual-risk evaluation (SYS-161) and post-market (Stage 17).
---
Verification Requirements
SRS-SAFE-015 — Safety software verification and HAL reconciliation
Requirement Statement | Each software risk control in this module shall be verified by the method stated in its record — fault-injection, charge-inhibit, cold-endurance/power-loss, summative-usability, and record-integrity tests as applicable — and, on completion of the HAL (Stage 10), each requirement shall be reconciled to its resolved HA-###/RCR-###, its severity-first Action Priority, and its probability/detectability targets, and migrated from Preliminary to Baselined only where the hazard analysis confirms the control.
Rationale | The safety claims of this module are only credible when each control is verified under the conditions it must operate in and each preliminary requirement is closed against a confirmed hazard record; this fixes the obligation that closes the Preliminary status.
Priority | Must
Source | Regulatory (Safety)
Verification Method | Test, Analysis, Inspection
Traceability | SYS-162, SYS-161, PRD-SAFE-011; → HAL (Stage 10), SRS-VER, PVL/PVaL, ORCH-018/EM-ENG-012. Standard: ISO 14971, IEC 62304 (revision to confirm).
Notes | PRELIMINARY. The reconciliation gate for the whole module.
---
Traceability
Every requirement above traces upward to at least one Safety System Requirement (SYS-149…162) and Product safety/architecture requirement (PRD-SAFE-…, PRD-ARCH-008), and forward — at HAL reconciliation — to a HAL hazard record (HA-###) and Risk-Control Requirement (RCR-###), the software design output, and the verification/validation evidence. Principal upstream anchors: SYS-149 (severity anchor), SYS-151/152/153/154 (honesty, alarm balance, use-error, mis-arm), SYS-155/156/160 (charge/thermal, dead-in-cold), SYS-158/159 (fail-safe / safe state), SYS-161/162 (residual risk, traceability); PRD-SAFE-001…012. Cross-module: safe mode / fault / watchdog / gap-marking / durable writes / trusted clock → SRS-DEV; OTA integrity/rollback → SRS-OTA; tamper-evidence/signing cryptography → SRS-SEC; numeric detection/alarm/energy targets → SRS-PERF; audit-trail record → SRS-DATA; delayed-alert connectivity hazard → SRS-CONN-008. All requirements are PRELIMINARY and reconcile with the HAL (Stage 10, ORCH-017); Action Priority is severity-first, probability/detectability Evidence-Pending until HAL.
SRS-1300 — Regulatory (module SRS-REG)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-REG · Owner: Quality / Regulatory Engineer (EM-ENG-006) with the Embedded Software Engineer (EM-ENG-007). Detailed software regulatory / compliance requirements decomposed from the Regulatory System Requirements (SYS-177…190) and the safety-relevant record/interface System Requirements (SYS-006/026/027/028, SYS-063/064), realising the Product regulatory requirements (PRD-REG-001…012). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md). Research-driven per the standard (identify → search KB → ingest missing → cite).
---
EVIDENCE & STATUS POSTURE. Consistent with the org evidence posture (CLAUDE.md rule 1; index.md
"Applicability & evidence posture"), every standard below is cited by designation only, revision to
confirm, and no standard clause number is asserted as fact. Two named market-access gaps carry
forward from concept §7 and remain open at this baseline: the exact WHO PQS temperature-monitoring
device/logger category and specification revision (SRS-REG-001, WHO PQS category Evidence-Pending),
and the US FDA regulatory classification of the SCM as a distribution-monitoring information accessory
(Evidence-Pending). Requirements that depend on either — notably the CSV scope (SRS-REG-013) and the
IEC 62304 software safety class (SRS-REG-004) — inherit that status and are conditional on their
resolution. These requirements state what the SCM software shall be/do to make the system saleable,
deployable, flyable, and trusted; they do not restate the system-level or physical means of compliance
(radio/EMC/transport hardware, labeling, QMS), which are owned at the system/regulatory stage and referenced
here.
1. Purpose
Specify the software behaviours across the device firmware, cloud platform, and mobile application that are required to achieve and hold the SCM's market-access and data-integrity obligations: WHO PQS conformance, EU GDP data integrity, 21 CFR Part 11 electronic records and signatures / ALCOA+, the IEC 62304 software life-cycle, the ISO 14971 risk-management contribution, IEC 62366-1 usability, software's role in radio/EMC/ electrical-safety and transport conformance, hazardous-substance/end-of-life support, GDPR/privacy data handling, computerized-system validation (CSV/GAMP 5), audit, and maintenance of conformity across OTA.
2. Scope
Software conformance to: WHO PQS E006-line (device category Evidence-Pending); EU GDP data-integrity expectations; 21 CFR Part 11 electronic records/signatures and ALCOA+; IEC 62304 software life-cycle and safety classification; ISO 14971 risk-management (software contribution); IEC 62366-1 use-related risk (UI); FCC Part 15 / CE RED radio conformance (software radio governance); IEC 62368-1 electrical/thermal safety (software's supporting role); EMC per CISPR 32/35 and IEC 61000 series (software immunity behaviour); UN 38.3 / IATA DGR transport (shipping-mode software); RoHS/REACH/WEEE (software support for material declarations and end-of-life decommission); GDPR and regional privacy law (data minimisation in software); CSV per GAMP 5 (validated software state); audit requirements (cross-tier audit trail); and maintenance-of-conformity across OTA change. Excludes the physical/system means of compliance — radio/EMC/electrical-safety hardware qualification, UN 38.3 cell testing, labeling/marking, and the QMS itself (SYS-180…188, PRD-REG, PRD-ELEC, PRD-MECH) — which are referenced where the software interfaces with them.
3. Applicable Standards
WHO PQS E006-line temperature-monitoring device/logger category (category & revision Evidence-Pending); EU GDP (Good Distribution Practice, data-integrity expectations); 21 CFR Part 11 (electronic records and signatures) and ALCOA+; IEC 62304 (medical-device software life-cycle; safety class Evidence-Pending on SYS-178); ISO 14971 (risk management); IEC 62366-1 (usability / use-related risk); FCC Part 15 and CE RED (radio); IEC 62368-1 (electrical/thermal safety); CISPR 32 / CISPR 35 and IEC 61000 series (EMC); UN 38.3 and IATA DGR (lithium-battery transport); RoHS, REACH, WEEE and applicable battery directives (hazardous-substance / end-of-life); GDPR and applicable regional data-protection law (privacy); GAMP 5 (computerized-system validation). US FDA classification Evidence-Pending. All cited by designation, revision to confirm; no clause number asserted as fact.
4. Assumptions
• The system-level regulatory obligations (radio type approval, EMC/safety/transport qualification, labeling, and the QMS/ISO 13485 applicability) are owned at Stage 06 regulatory (SYS-177…190) and the regulatory stage; this module specifies only the software contribution to each.
• The regulated software tiers are the device firmware (S3), the cloud platform (S8), and the mobile app (S9); each is subject to CSV commensurate with the regulated data it creates, stores, transforms, or presents (SYS-189).
• The electronic record substrate — tamper-evident append-only storage, atomic writes, trusted clock, gap marking, and store-and-forward de-duplication — is provided by SRS-DEV / SRS-CONN / SRS-DATA; this module states the compliance properties those substrates must satisfy.
• The WHO PQS category/revision and US FDA classification are Evidence-Pending (concept §7, SYS-177/178); requirements fix the software behaviour, not the unresolved category/classification.
• Cryptographic mechanisms for signatures, tamper-evidence, and mutual authentication are owned by SRS-SEC (EM-ENG-008); this module states the regulatory obligations those mechanisms must meet.
5. Architecture Overview
Regulatory conformance is realised in software as a set of cross-tier properties on the evidentiary trip record and on the regulated software life-cycle. On the device, firmware guarantees attributable, contemporaneous, gap-honest, tamper-evident capture; the cloud assembles and stores the ALCOA+ record, enforces electronic-signature and access controls, serves the audited enterprise export, applies retention and privacy data-minimisation, and holds the fleet in a validated state across signed OTA; the mobile app performs authenticated configuration with attributable actions. Above these, the software life-cycle (IEC 62304), risk-management contribution (ISO 14971), usability (IEC 62366-1), and CSV (GAMP 5) processes keep the regulated software developed, validated, and maintained. Radio/EMC behaviours in software (region band governance, immunity-safe behaviour) support the hardware conformance, and a shipping-mode supports UN 38.3/IATA transport. Maintenance-of-conformity gating ensures an OTA change that touches a conformance basis is re-assessed before it is fielded.
---
7. Functional Requirements (software compliance obligations)
7.1 Market access & records
SRS-REG-001 — WHO PQS temperature-monitoring conformance (software contribution)
Requirement Statement | The software shall implement the logging, detection, indication, and record functions to the degree required by the applicable WHO PQS temperature-monitoring device/logger category such that the system can achieve and hold WHO PQS listing for the target donor-funded/LMIC procurement markets.
Rationale | WHO PQS listing is a hard market-access gate for donor-funded and LMIC procurement (D4); the accuracy, detection, logging-interval, and record capabilities are largely software-realised and must meet the category performance specification.
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-177 (WHO PQS eligibility), SYS-001 (interval logging), PRD-REG-001, PRD-FUN-006; → SRS-DEV-013/020, SRS-PERF. Standard: WHO PQS E006-line (category & revision Evidence-Pending; revision to confirm).
Notes | Evidence-Pending — exact WHO PQS category/code and spec revision are a named concept gap (§7). No clause asserted.
SRS-REG-002 — EU GDP data-integrity conformance
Requirement Statement | The software shall handle the trip record across the device, cloud, and mobile tiers in conformance with EU GDP data-integrity expectations — preserving completeness, contemporaneity, and traceability of temperature-history data — so that a trip record is defensible for distribution release/acceptance and GDP audit.
Rationale | EU GDP governs distribution of medicinal products in the EU market (D4); the record's data-integrity is the product's compliance value and must survive a GDP inspection.
Priority | Must
Source | Regulatory
Verification Method | Analysis, Test
Traceability | SYS-179 (GDP/Part 11 records), SYS-026 (complete record), PRD-REG-003, PRD-FUN-017; → SRS-DATA, SRS-DEV-008/011. Standard: EU GDP (revision to confirm).
Notes | Realised jointly with SRS-REG-003 (ALCOA+) and SRS-REG-014 (audit).
SRS-REG-003 — 21 CFR Part 11 electronic records & signatures / ALCOA+
Requirement Statement | The software shall create, store, and present electronic records and, where used, electronic signatures that are attributable, legible, contemporaneous, original, accurate, and additionally complete, consistent, enduring, and available (ALCOA+), with time-stamped attributable authorship and tamper-evidence, such that records and signatures conform to 21 CFR Part 11 for admissibility in release/acceptance and audit.
Rationale | Part 11 / ALCOA+ conformance is what makes the record admissible electronic evidence for the US and data-integrity-regulated markets; a record that cannot be attributed, timed, or shown unaltered is not defensible.
Priority | Must
Source | Regulatory
Verification Method | Analysis, Test
Traceability | SYS-179, SYS-006 (trusted timestamp), SYS-026/028 (record/audit), PRD-REG-003, PRD-SEC-001/002/004; → SRS-DEV-009/021, SRS-SEC, SRS-DATA. Standard: 21 CFR Part 11, ALCOA+ (revision to confirm).
Notes | Electronic-signature use scope confirmed with SYS-178 classification.
7.2 Software life-cycle, risk & usability
SRS-REG-004 — IEC 62304 software life-cycle process
Requirement Statement | The regulated software (device firmware, cloud platform, mobile application) shall be developed and maintained under an IEC 62304-consistent software life-cycle process — including software safety classification, architecture and detailed design, verification, configuration and problem management, and maintenance — commensurate with the safety class assigned to each software item.
Rationale | A recognized software life-cycle is required to make safety-relevant software defensible to regulated customers and to underpin the ISO 14971 risk file; the process rigour scales with the software safety class.
Priority | Must
Source | Regulatory
Verification Method | Analysis, Inspection
Traceability | SYS-189 (CSV), SYS-188 (quality-system basis), PRD-REG-011/012; → SRS-VER, HAL (Stage 10). Standard: IEC 62304 (software safety class Evidence-Pending on SYS-178; revision to confirm).
Notes | Evidence-Pending — safety class conditional on the SYS-178 classification decision.
SRS-REG-005 — ISO 14971 risk-management contribution of software
Requirement Statement | The software shall implement the software risk controls identified by the ISO 14971 risk-management process (SRS-SAFE / HAL) and shall contribute the software-side evidence — control implementation, diagnostic coverage, and safety logging — required for the risk-management file and the overall residual-risk evaluation.
Rationale | ISO 14971 requires that identified risk controls be implemented and their effectiveness evidenced; the software contribution to the risk file is what shows software-controlled hazards are actually mitigated.
Priority | Must
Source | Regulatory (Safety)
Verification Method | Analysis, Test
Traceability | SYS-161/162 (residual risk / traceability), PRD-REG-011, PRD-SAFE-011; → SRS-SAFE (all), HAL HA-###/RCR-### (Stage 10). Standard: ISO 14971 (revision to confirm).
Notes | Bridges SRS-SAFE to the regulatory risk file; PRELIMINARY controls reconcile at HAL.
SRS-REG-006 — IEC 62366-1 usability / use-related risk (software UI)
Requirement Statement | The software user interfaces (on-device HMI, mobile app, and cloud disposition views) shall be designed and evaluated under an IEC 62366-1-consistent usability engineering process so that foreseeable use errors by the intended users — including minimally trained, gloved, offline last-mile users — do not lead to a wrong disposition, meeting the summative usability acceptance criteria for safety-critical tasks.
Rationale | Use-related risk in the UI is a leading disposition-error source; usability engineering controls it by design rather than by training, and its evidence feeds the Usability Engineering File.
Priority | Must
Source | Regulatory (Safety)
Verification Method | Test (summative usability)
Traceability | SYS-153 (use-error protection), SYS-058, PRD-REG-011, PRD-SAFE-005, PRD-UX-001/007; → SRS-SAFE-007, SRS-MOB. Standard: IEC 62366-1 (revision to confirm).
Notes | Feeds the Usability Engineering File; login-free edge read preserved.
7.3 Radio, EMC, safety & transport (software's role)
SRS-REG-007 — Regional radio conformance governance in software (FCC / CE RED)
Requirement Statement | The software shall operate the radios (BLE, Wi-Fi, LTE-M, GNSS) only within the bands, power, and regulatory profile configured for the device's region, and shall not initiate transmission outside the permitted band plan/limits for that configured region, supporting the FCC Part 15 / CE RED and regional radio type approvals held for that variant.
Rationale | Radio approval is a legal-to-sell/legal-to-operate gate per geography; software that could transmit outside the permitted band plan is a compliance and legal-operation failure.
Priority | Must
Source | Regulatory
Verification Method | Inspection, Test
Traceability | SYS-180 (radio type approval), SYS-115 (regional radio config), PRD-REG-004; → SRS-CONN-020, SRS-DEV-005. Standard: FCC Part 15 / CE RED, 3GPP band plans (revision to confirm).
Notes | Region config integrity per SRS-DEV-005; per-market hardware variants may still be required.
SRS-REG-008 — Software's role in electrical/thermal safety (IEC 62368-1)
Requirement Statement | The software shall command and supervise the safety-relevant electrical/thermal behaviours within its control — the low-temperature charge-inhibit and the fail-safe response to over- temperature/fault signals from the power subsystem — such that the software does not defeat, delay, or mask the IEC 62368-1 electrical/thermal safety protections of the charge path and battery subsystem.
Rationale | IEC 62368-1 safety is primarily a hardware property, but software that mismanages charging or ignores a fault signal can create the very hazard the standard guards against; software must reinforce, not undermine, the protections.
Priority | Must
Source | Regulatory (Safety)
Verification Method | Test
Traceability | SYS-182 (electrical/thermal safety), SYS-155/156, PRD-REG-005, PRD-SAFE-007/008; → SRS-SAFE-010, SRS-DEV-015. Standard: IEC 62368-1 (revision to confirm).
Notes | Physical protection owned by S2/BMS; software owns the non-defeat obligation. Reconcile with HAL.
SRS-REG-009 — EMC-immunity-safe software behaviour (CISPR / IEC 61000)
Requirement Statement | The software shall behave safely under the electromagnetic disturbances defined for the system's use and transport environments — detecting corrupted sensor reads, bus errors, or resets induced by disturbance and responding with plausibility flagging, safe-state entry, or recovery — such that an immunity event does not silently suppress a genuine excursion alert or corrupt the record undetected.
Rationale | An immunity-induced malfunction that mutes an alert or corrupts a record is a safety and data-integrity failure; the software must fail safe and honest under the disturbances the system is required to withstand (CISPR/IEC 61000).
Priority | Must
Source | Regulatory (Safety)
Verification Method | Test (EMC immunity, fault injection)
Traceability | SYS-181 (EMC), SYS-007 (plausibility), SYS-067, PRD-REG-005; → SRS-SAFE-009/013, SRS-DEV-018/028. Standard: CISPR 32/35, IEC 61000 series (revision to confirm).
Notes | Immunity-induced malfunction is safety-relevant; ties SRS-SAFE fail-safe.
SRS-REG-010 — Shipping-mode software support for UN 38.3 / IATA transport
Requirement Statement | The software shall provide a low-power shipping/transport mode that holds the device and battery in the state assumed by the UN 38.3 / IATA DGR transport configuration (radios quiescent as required, charge inhibited as required, record preserved), and shall not autonomously exit that state or initiate transmission in a way inconsistent with the device's lawful transport configuration.
Rationale | Lawful, safe air/ground transport of the Li battery is a market-access gate; software that wakes radios or manages the battery inconsistently with the transport configuration undermines the UN 38.3 / IATA basis.
Priority | Should
Source | Regulatory (Safety)
Verification Method | Test, Inspection
Traceability | SYS-183 (transport), SYS-157, PRD-REG-006, PRD-SAFE-009; → SRS-DEV-024 (shipping mode), SRS-CONN-020. Standard: UN 38.3, IATA DGR (test summary Evidence-Pending; revision to confirm).
Notes | Evidence-Pending — UN 38.3 test summary / IATA guidance to be produced (SYS-183).
7.4 Environmental, privacy, validation & audit
SRS-REG-011 — Software support for hazardous-substance / end-of-life compliance
Requirement Statement | The software shall support the RoHS/REACH/WEEE and battery end-of-life obligations by providing a secure decommission/data-erasure function that renders the trip record and credentials unrecoverable at retirement, and by maintaining the firmware/hardware and material-declaration identifiers needed to associate a device with its compliance documentation, without impeding the defined take-back/recycling path for the reusable fleet.
Rationale | A reusable, rechargeable, globally distributed device must be retired lawfully and its data disposed of responsibly; software must enable clean decommission and traceable material identification, not block it.
Priority | Should
Source | Regulatory (Sustainability)
Verification Method | Inspection, Demonstration
Traceability | SYS-185 (hazardous-substance/EoL), PRD-REG-008, PRD-VIS-006; → SRS-DEV-023 (factory reset), SRS-FLEET (decommission), SRS-SEC (erase). Standard: RoHS/REACH/WEEE, battery directives (per-market EoL path Evidence-Pending; revision to confirm).
Notes | Evidence-Pending — per-market EoL path and declarations to confirm (SYS-185).
SRS-REG-012 — GDPR / regional privacy data handling (data minimisation)
Requirement Statement | The software shall handle any personal or location-derived data (device GNSS position, operator/actor identifiers in the audit trail) in conformance with the applicable regional data-protection law of each market of deployment, applying data-minimisation appropriate to a cargo-monitoring (not people-monitoring) device — collecting, retaining, and exporting only what the monitoring and audit purpose requires.
Rationale | Although the device monitors cargo rather than people, location and actor-attribution data can be personal data; the platform must be lawful to deploy per region, with minimisation as the default.
Priority | Should
Source | Regulatory (Security)
Verification Method | Inspection, Analysis
Traceability | SYS-186 (data protection/privacy), SYS-029 (location), PRD-REG-009, PRD-SEC-010; → SRS-DATA (retention/privacy), SRS-CLOUD, SRS-SEC. Standard: GDPR and regional data-protection law (applicable-law mapping Evidence-Pending; revision to confirm).
Notes | Evidence-Pending — per-market applicable-law mapping to confirm (SYS-186).
SRS-REG-013 — Computerized-system validation (CSV / GAMP 5)
Requirement Statement | The regulated software across the device firmware, cloud platform, and mobile application shall be validated and maintained in a validated state under a GAMP 5-consistent computerized- system-validation approach, risk-scaled to the regulated data each component creates, stores, transforms, or presents, and shall remain in a validated state across configuration and OTA changes.
Rationale | Records and dispositions produced by a computerized system are only defensible to regulated customers if that system is validated to a recognized framework and kept validated across change.
Priority | Should
Source | Regulatory
Verification Method | Analysis
Traceability | SYS-189 (CSV), SYS-032 (validated change), PRD-REG-012, PRD-REG-003; → SRS-VER, SRS-OTA, SRS-REG-015. Standard: GAMP 5 (revision to confirm).
Notes | Risk-based scope conditional on SYS-178 classification (Evidence-Pending).
SRS-REG-014 — Cross-tier audit trail and record export
Requirement Statement | The software shall maintain an append-only, attributable, time-stamped, tamper-evident audit trail of configuration, arming, calibration, firmware, and record actions across the device and cloud tiers, and shall export the complete trip record and its audit trail on demand in a documented, versioned, human- and machine-readable format that preserves attribution, time, gaps, and tamper-evidence, under authenticated, access-controlled, audit-logged access.
Rationale | Attributability, non-repudiation, and an admissible export are core to a regulated record and to passing GDP/Part 11 audit at every gate — customs, procurement, and clinical intake.
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-028 (audit trail), SYS-027 (export), SYS-064 (enterprise export), PRD-REG-003/010, PRD-CLOUD-005/007; → SRS-DATA, SRS-CLOUD, SRS-SEC. Standard: 21 CFR Part 11 audit trail, EU GDP (revision to confirm).
Notes | Export/API governance (RBAC/SSO) owned with SRS-CLOUD/SRS-SEC.
---
Verification Requirements
SRS-REG-015 — Regulatory verification and maintenance of conformity across change
Requirement Statement | The software's regulatory conformance shall be verified by the method stated in each requirement — CSV, records/audit-simulation, radio/EMC/immunity, usability, and transport-mode tests as applicable — and the software shall support maintenance of conformity such that any firmware/configuration change affecting a conformance basis (WHO PQS, radio approval, EMC/safety, transport, Part 11/GDP records, or CSV validated state) is identified and re-assessed before it is fielded by OTA.
Rationale | In a multi-market fleet, a device fielded without current conformance, or changed by OTA in a way that invalidates a conformance basis, is both a legal and a data-integrity exposure; conformance must be gated at deployment and maintained across change.
Priority | Must
Source | Regulatory
Verification Method | Analysis, Inspection, Test
Traceability | SYS-190 (conformance gate / maintenance of conformity), SYS-189/032, PRD-REG-001…012; → SRS-VER, SRS-OTA, SRS-FLEET, ORCH-018/EM-ENG-012. Standard: WHO PQS, FCC/CE RED, GAMP 5 (revision to confirm).
Notes | Evidence-Pending — per-market approval status and change-control process to confirm (SYS-190).
---
Traceability
Every requirement above traces upward to at least one Regulatory System Requirement (SYS-177…190) and Product regulatory requirement (PRD-REG-001…012), and forward to the certifications/listings and regulatory reviews that close them. Principal upstream anchors: SYS-177/178 (WHO PQS / classification, Evidence-Pending), SYS-179/189 (records / CSV), SYS-180/181/182 (radio / EMC / safety), SYS-183/184 (transport / calibration), SYS-185/186 (EoL / privacy), SYS-187/188/190 (labeling / QMS / conformance gate). Cross-module: tamper-evident record substrate, atomic writes, trusted clock, gap marking → SRS-DEV; store-and-forward de-duplication, region radio governance, time sync → SRS-CONN; record schema, retention, privacy minimisation, export → SRS-DATA / SRS-CLOUD; cryptographic signatures/tamper-evidence/mutual auth → SRS-SEC; software risk controls and their risk file contribution → SRS-SAFE / HAL (Stage 10); OTA validated change → SRS-OTA. Evidence-Pending items — WHO PQS category/revision, US FDA classification, IEC 62304 software safety class, UN 38.3 test summary, per-market EoL path and privacy-law mapping, per-market approval status — are carried as gaps, not invented, and close in the Regulatory stage. All standards cited by designation, revision to confirm; no clause number asserted as fact.
SRS-1400 — Diagnostics (module SRS-DIAG)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-DIAG · Owner: Embedded Software Engineer (EM-ENG-007) with the Reliability and Manufacturing/Test engineers. Detailed software-diagnostics requirements decomposed from the System Requirements (SYS-024/030, SYS-113, SYS-169/170/173, SYS-007/017) and Product Requirements (PRD-FUN-016/020, PRD-REL-006/007/010), realising the self-test, health-monitoring, metrics, logging, and remote/manufacturing/service diagnostics that make faults observable, keep monitoring honest, and gate deployment. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM software proves it is healthy and diagnoses itself: power-on and built-in self-test, continuous health monitoring, performance and reliability metrics, structured logging and crash dumps, remote diagnostics for the fleet, manufacturing test diagnostics, service/depot diagnostics, failure reporting, debug interfaces, and the pre-arm readiness self-check that blocks deployment of an un-ready device. Diagnostics exist so that no fault is silent and every deployed device is provably ready.
2. Scope
Power-on self-test (POST); built-in self-test (BIST) of sensing, storage, clock, radios, and power path; continuous health monitoring; performance and reliability metrics; structured, attributable logging; crash capture/dumps; remote diagnostics and log retrieval over any bearer; manufacturing test diagnostics and end-of-line coverage; service/depot diagnostics and triage; failure reporting to the fleet; secured debug interfaces; and the pre-arm readiness self-check. Excludes the diagnostics transport mechanics (bearer selection, SRS-CONN), the cloud-side fleet dashboards and analytics (SRS-CLOUD/SRS-FLEET), the mobile diagnostics UI (SRS-MOB), cryptographic protection of the debug port and logs (SRS-SEC), and OTA update diagnostics (SRS-OTA); those are referenced where the diagnostics software interfaces with them.
3. Applicable Standards
IEC 62304 (software life-cycle, incl. anomaly/problem records); WHO PQS E006 (device readiness/performance verification concepts); ISO/IEC/IEEE 29148 (verifiability of diagnostic requirements); IEC 62368-1 / JTAG-protection concepts (secure debug). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• Diagnostics must be energy-cheap: the diagnostics subsystem operates within its share of the D3 energy budget (SYS-090, Evidence-Pending) and is always subordinate to the excursion-alert path (SRS-PERF-015).
• Battery state-of-charge and state-of-health, charge-path status, and charge-inhibit signals are provided by the BMS (S2); this module specifies the software that reads, evaluates, and reports them.
• Structured logs and crash dumps are stored using the atomic, tamper-evident persistence of SRS-DEV/SRS-DATA; this module specifies their content, capture, and retrieval, not the storage primitives.
• Remote diagnostics ride the connectivity bearers and are authenticated per SRS-CONN/SRS-SEC; this module specifies the diagnostic content and access model, not the transport or the crypto.
• Readiness thresholds (adequate charge, valid calibration, valid firmware/profile, probe present) are set by the system readiness model (SYS-024/030); several depend on Evidence-Pending values (calibration interval, cold state-of-health) resolved by test.
5. Architecture Overview
A self-test layer runs at power-on (POST) and on demand (BIST), exercising sensing, storage, clock, radios, HMI, and the power/charge path and reporting a structured pass/fail per subsystem. A health-monitor layer runs continuously while armed, watching task liveness, memory, storage-write success, sensor liveness, clock discipline, and battery health, feeding the availability self-check and the reliability event stream. A metrics/logging layer records performance and reliability metrics and structured, attributable log entries, and captures a crash dump on an unexpected fault. A diagnostics-access layer exposes this evidence three ways — remotely to the fleet (authenticated, energy-budgeted), at manufacturing end-of-line (full coverage), and at the service depot (triage self-test) — plus a secured developer debug interface. The pre-arm readiness self-check composes the self-tests and health state into a single go/no-go that blocks a device from being armed while any readiness pre-condition fails.
---
7. Functional Requirements
7.1 Self-test
SRS-DIAG-001 — Power-on self-test (POST)
Requirement Statement | On every power-on and reset, the device software shall execute a power-on self-test of the MCU core, non-volatile storage, clock/RTC, sensing channels, radios, HMI, and power/charge path, and shall record a structured pass/fail result per subsystem before declaring the device ready.
Rationale | A device must not present as ready with a silently failed subsystem; POST establishes a known, attested starting state for every trip and every recovery.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-024 (pre-arm checks), SYS-169 (recovery), SYS-007 (fault flag), PRD-REL-004, PRD-FUN-016; → SRS-DEV (boot), SRS-DIAG-011 (readiness).
Notes | A POST failure feeds the readiness self-check (SRS-DIAG-011) and must be visible on the HMI and to the fleet; POST must fit the cold-boot timing budget (SRS-PERF-001).
SRS-DIAG-002 — Built-in self-test (BIST) on demand
Requirement Statement | The device software shall provide an authenticated, on-demand built-in self-test that exercises each sensing channel (against internal references or plausibility), the storage write/read path, clock discipline, each radio bearer, the HMI, and the power/charge path, and reports a structured per-subsystem result with a diagnostic code for each failure.
Rationale | Deploying or triaging a reusable device requires confirming, on demand, that every subsystem actually works — not just that it powered on; diagnostic codes make failures actionable at the depot.
Priority | Must
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness), SYS-113 (diagnostics), SYS-024, PRD-FUN-020, PRD-REL-007; → SRS-CONN-024 (connectivity self-test), SRS-DIAG-009 (service).
Notes | Access authenticated per the SRS-DEV mode model; connectivity BIST is realised in SRS-CONN-024.
SRS-DIAG-003 — Pre-arm readiness self-check as an arming gate
Requirement Statement | Before confirming a device as armed, the software shall verify and report the readiness pre-conditions — adequate state of charge for the intended trip, valid calibration, valid firmware, valid trip profile, and probe present — and shall block arming and guide the user to correct any failed pre-condition.
Rationale | A device armed while un-ready produces a worthless or misleading record; the readiness gate is the last defence against a silently-doomed trip (mis-arm is a top hazard).
Priority | Must
Source | Safety
Verification Method | Test
Traceability | SYS-024 (readiness/pre-arm), SYS-023 (arm confirmation), SYS-030, PRD-FUN-016/020, PRD-SVC-004; → SRS-DIAG-001/002, SRS-SAFE / HAL.
Notes | State-of-charge-for-trip and calibration-validity thresholds depend on Evidence-Pending values (cold state-of-health SYS-165; calibration interval SYS-034); reconcile with HAL. This is the diagnostics counterpart of the "pre-arm readiness" topic and gates deployment (SYS-030).
7.2 Health monitoring & metrics
SRS-DIAG-004 — Continuous health monitoring while armed
Requirement Statement | While armed, the device software shall continuously monitor task liveness, memory high-water mark, storage-write success, sensor liveness/plausibility, clock discipline, and battery state-of-charge/health, and shall raise an attributable health event on any indicator crossing its defined bound.
Rationale | Faults that appear between self-tests must still be caught; continuous health monitoring is what turns a latent anomaly into a logged event and an honest not-monitoring flag (SYS-170).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-170 (self-check), SYS-007 (plausibility), SYS-169, PRD-REL-006/007; → SRS-REL-010/013, SRS-DIAG-006.
Notes | Health monitoring feeds the availability self-check (SRS-REL-010) and self-healing (SRS-REL-013); it must be energy-cheap.
SRS-DIAG-005 — Performance and reliability metrics collection
Requirement Statement | The device software shall collect and persist performance and reliability metrics — boot/wake time, per-cycle compute time, RAM high-water mark, write throughput/latency, missed-deadline count, reset/watchdog/brown-out/crash counts, per-function energy estimates, and battery state-of-health — and shall make them available to remote diagnostics.
Rationale | Performance and reliability can only be managed, and their Evidence-Pending targets closed, if they are measured in the field; these metrics also feed the reliability program and the energy budget.
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-113 (diagnostics/metrics), SYS-090 (energy), SYS-173 (reliability program), PRD-PERF-005, PRD-REL-010; → SRS-PERF-017, SRS-REL-017, SRS-DIAG-007.
Notes | Metrics support closing the D3 energy budget (MR-VAL-006) and the reliability prediction (SYS-173, Evidence-Pending).
7.3 Logging, crash capture & failure reporting
SRS-DIAG-006 — Structured, attributable diagnostic logging
Requirement Statement | The device software shall record diagnostic events in a structured, timestamped, attributable log with a defined severity and diagnostic code per entry, retained in non-volatile storage under the same integrity guarantees as the record and without displacing safety-critical trip data.
Rationale | Triage and audit require a consistent, machine-readable diagnostic history; unstructured logs are not actionable at fleet scale, and diagnostics must never evict excursion or audit data.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-113, SYS-021 (retention priority), SYS-006 (trusted timestamp), PRD-REL-007; → SRS-DATA, SRS-DEV (storage), SRS-DIAG-007.
Notes | Diagnostic logs are subordinate to excursion/audit retention (SYS-021); integrity per SRS-SEC.
SRS-DIAG-007 — Crash and fault dump capture
Requirement Statement | On a firmware crash, exception, watchdog reset, or brown-out, the device software shall capture a bounded diagnostic dump (fault cause/exception context, relevant state, recent event history) to non-volatile storage for later retrieval, without corrupting committed trip data.
Rationale | A recovered fault whose cause is not captured recurs undiagnosed; the crash dump is the primary evidence for root-cause and reliability growth (feeds the DFMEA/ERFKL).
Priority | Should
Source | Reliability
Verification Method | Test (fault injection)
Traceability | SYS-169 (recovery event), SYS-167, PRD-REL-004/006; → SRS-REL-003, SRS-DIAG-006/008.
Notes | Dump size is bounded to protect the energy/storage budget; retrieval via remote/service diagnostics.
SRS-DIAG-008 — Failure reporting to the fleet
Requirement Statement | The software shall report detected faults, self-test failures, crashes, availability outages, and readiness failures to the fleet platform as attributable, coded failure events, such that a fleet operator can identify and locate a faulted or not-ready device without physical access.
Rationale | A reusable fleet cannot be kept healthy if faults are only visible on the device; remote failure reporting is what enables recall, service, and stray-recovery decisions.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-030 (readiness reporting), SYS-031 (fleet visibility), SYS-113, PRD-FUN-020, PRD-REL-007; → SRS-FLEET, SRS-CLOUD, SRS-CONN.
Notes | Failure reports ride the energy-budgeted diagnostics path and never compete with the alert path.
7.4 Remote, manufacturing & service diagnostics
SRS-DIAG-009 — Remote diagnostics and log retrieval
Requirement Statement | The software shall support authenticated remote retrieval of diagnostic logs, metrics, crash dumps, and self-test results over the available bearer, and authenticated remote invocation of a non-intrusive self-test, without loss of buffered trip data and within the energy budget.
Rationale | Fleet-scale triage of geographically dispersed devices requires remote access to diagnostics; it must be authenticated, non-destructive, and energy-aware.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-113 (diagnostics), SYS-031 (fleet), PRD-REL-007, PRD-CONN-010; → SRS-CONN (transport), SRS-SEC (auth), SRS-FLEET.
Notes | Access authenticated per SRS-SEC; a remote self-test must never disrupt active monitoring or the alert path.
SRS-DIAG-010 — Manufacturing (end-of-line) test diagnostics
Requirement Statement | The software shall provide a manufacturing diagnostics mode that exercises and reports coverage of every sensor channel, storage, clock, each radio, HMI, and the power/charge path with pass/fail and measured values, and shall record the end-of-line result and calibration/provisioning state to the device for traceability.
Rationale | End-of-line test coverage is what keeps latent manufacturing defects out of the field; the recorded result establishes the device's birth record for the reuse-life traceability.
Priority | Must
Source | Architecture
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness), SYS-113, SYS-173 (reliability program), PRD-REL-010, PRD-SVC-004; → SRS-DEV (manufacturing mode), SRS-VER (manufacturing verification).
Notes | Manufacturing mode is access-controlled and disabled/locked before field deployment (SRS-SEC); calibration provisioning ties to SYS-034.
SRS-DIAG-011 — Service / depot diagnostics and triage
Requirement Statement | The software shall provide an authenticated service diagnostics function for the depot that runs the BIST, reports battery state-of-health and readiness status, surfaces stored crash dumps and failure history, and gives a clear serviceable/return/retire recommendation for the reuse loop.
Rationale | Reuse economics (D1) depend on fast, correct depot triage; a technician must be able to decide in minutes whether a returned device is fit to redeploy.
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-030, SYS-165 (state-of-health/redeploy block), SYS-113, PRD-REL-002, PRD-SVC-004; → SRS-DIAG-002/003, SRS-REL-014, SRS-MOB.
Notes | State-of-health redeploy threshold depends on Evidence-Pending cold cycle-life data (SYS-165); requirement fixes the triage behaviour.
7.5 Debug interface
SRS-DIAG-012 — Secured developer debug interface
Requirement Statement | The software shall provide a developer debug interface (trace/console/JTAG-class) that is authenticated and access-controlled, disabled or locked in production/field configuration, and auditable, such that it exposes diagnostics for development without becoming an unauthenticated attack or data-exfiltration surface in the field.
Rationale | A debug interface is essential for development but a serious attack surface if left open in the field; it must be present-but-locked, not present-and-exposed.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-113, PRD-SEC-005/008 (secure debug/JTAG protection concepts); → SRS-SEC (JTAG/debug protection, EM-ENG-008), SRS-DEV (developer mode).
Notes | Debug-port protection mechanism owned by SRS-SEC (EM-ENG-008); this states the diagnostics-side access and audit obligation.
---
8. Non-functional Requirements
SRS-DIAG-013 — Diagnostics energy and priority conformance
Requirement Statement | The diagnostics subsystem shall operate within its allocated share of the device energy budget over the 90-day cold-duty profile and shall always yield the radio, compute, and storage to the excursion-alert and sample-capture paths, never delaying a sample, detection, or alert to run diagnostics.
Rationale | Diagnostics that erode the energy budget or preempt the alert path would defeat the very reliability they exist to prove; they must be strictly subordinate.
Priority | Must
Source | Performance
Verification Method | Analysis, Test
Traceability | SYS-090 (energy allocation), SYS-112 (alert priority), PRD-PERF-005; → SRS-PERF-015, SRS-CONN-014.
Notes | Diagnostics energy allocation is Evidence-Pending (SYS-090) and closes on the power budget.
---
Error Handling
SRS-DIAG-014 — Robust diagnostics that never mask or evict the record
Requirement Statement | A failure within the diagnostics subsystem itself (log-write failure, metric overflow, dump-capture failure) shall be contained and flagged without corrupting or evicting trip records and without falsely reporting the device as healthy.
Rationale | The diagnostics subsystem must not become a new failure mode; a broken logger must not mask a real fault or displace safety-critical data.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-021 (retention priority), SYS-170 (honest status), PRD-REL-007; → SRS-REL-016, SRS-DATA.
Notes | Ties to the no-silent-failure invariant (SRS-REL-016).
---
Security Considerations
SRS-DIAG-015 — Authenticated, integrity-protected diagnostic access
Requirement Statement | All diagnostic access — remote retrieval, self-test invocation, manufacturing, service, and debug — shall require authentication appropriate to its mode, shall protect diagnostic data in transit and at rest, and shall not expose trip-record contents or credentials beyond the accessing role's authorisation.
Rationale | Diagnostics can leak sensitive record and device data or provide an unauthenticated control path; access must be gated and diagnostic data protected.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-110 (mutual auth), PRD-SEC-005/006; → SRS-SEC (mechanisms, EM-ENG-008), SRS-CONN-012, SRS-DIAG-012.
Notes | Cryptographic and role mechanisms owned by SRS-SEC; this states the diagnostics-side posture.
---
Verification Requirements
SRS-DIAG-016 — Diagnostics verification and coverage demonstration
Requirement Statement | The diagnostics requirements shall be verified by POST/BIST coverage tests, fault-injection tests confirming each seeded fault is detected/reported, readiness-gate tests, remote and service diagnostics tests, manufacturing end-of-line coverage demonstration, and debug-interface lock-down inspection, with results traced to the requirements above.
Rationale | Diagnostic coverage claims (that faults are actually caught and reported) are only credible if demonstrated by seeding faults and confirming detection — not by inspection of the code alone.
Priority | Must
Source | Reliability
Verification Method | Test, Inspection
Traceability | SYS-113, SYS-024/030/170, PRD-REL-007/010, PRD-VER-001; → SRS-VER, ERFKL (Stage 11).
Notes | Diagnostic-coverage evidence feeds the reliability program (SYS-173) and the software DFMEA (ERFKL).
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-113 (diagnostics & signal logging), SYS-024/030 (pre-arm readiness, readiness reporting), SYS-169/170 (recovery event, availability self-check), SYS-007/017 (plausibility, honest status), SYS-173 (reliability program), SYS-031 (fleet visibility); PRD-FUN-016/020, PRD-REL-004/006/007/010, PRD-SVC-004, PRD-SEC-005/006. Cross-module: diagnostics transport and connectivity self-test → SRS-CONN; fleet dashboards/failure aggregation → SRS-FLEET; cloud storage/analytics → SRS-CLOUD; mobile/service UI → SRS-MOB; log/dump storage & integrity → SRS-DEV / SRS-DATA; debug-port and diagnostic-data protection → SRS-SEC (EM-ENG-008); energy allocation → SRS-PERF; self-check/ self-healing → SRS-REL; software DFMEA and coverage → ERFKL (Stage 11); readiness/mis-arm hazard → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (diagnostics energy allocation, calibration interval and cold state-of-health thresholds feeding the readiness gate) are carried as gaps, not invented, and close on their associated tests (power budget MR-VAL-006, calibration/cycle-life tests).
SRS-1500 — Verification (module SRS-VER)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-VER · Owner: Quality/Regulatory Engineer (EM-ENG-006) with the Embedded Software Engineer (EM-ENG-007) and the Verification lead (ORCH-014). Detailed software-verification requirements decomposed from the Product Requirements (PRD-VER 1900) and the System Requirements (SYS-045/046, SYS-104/173), governing how the SCM software requirements are traced, verified across the test hierarchy, and accepted — including the headline D3 validate-by-test and the RF-in-container qualification. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify the software verification strategy: bidirectional requirement traceability and objective acceptance criteria for every SRS requirement; the verification matrix and method allocation; the test hierarchy (unit → integration → system → performance → reliability → stress → environmental); manufacturing verification; field validation; regression; security testing; compliance verification (WHO PQS / EU GDP / 21 CFR Part 11); documentation verification; and the D3 endurance and RF-in-container test hooks. This module is the software source from which the Verification Plan, System Test Plan, and Verification/Traceability Matrices are derived.
2. Scope
Requirement verifiability and acceptance criteria; bidirectional traceability; method allocation by risk; the verification/test-level hierarchy for firmware, cloud, and mobile software; performance, reliability, stress, and environmental (cold/endurance) testing; the RF-in-container test hook; manufacturing (end-of-line) software verification; field validation; regression and CI verification; security testing (penetration, fuzz, crypto/interface); compliance verification against the WHO PQS / EU GDP / 21 CFR Part 11 regime; and documentation/records verification. Excludes the detailed test procedures and pass/fail limits (Verification Plan, ORCH-014), the validation scenarios (Validation, PVaL, ORCH-015), the safety-case content (HAL, Stage 10), and the hardware/environmental qualification test execution itself; those are referenced where the software verification interfaces with them.
3. Applicable Standards
ISO/IEC/IEEE 29148 (requirements verification); ISO/IEC/IEEE 29119 (software testing — concepts, levels, and techniques); IEC 62304 (software life-cycle verification and problem resolution); ISO 14971 (risk-control verification); WHO PQS E006 (device performance verification class); EU GDP and 21 CFR Part 11 / GAMP 5 (data-integrity and computerised-system validation, electronic records); IEC 60068 (environmental/endurance methods); FCC Part 15 / CE RED (radio conformance, referenced for the RF path). Cited by designation only; revision to confirm — no clause number, class, or method is asserted as fact.
4. Assumptions
• Every SRS requirement across the library carries a Verification Method and shall carry objective acceptance criteria; this module makes that obligation explicit for the software modules and drives the matrix.
• Safety-related software requirements are verified jointly with the Hazard Analysis (HAL, Stage 10), each tracing to an HA-/RCR- record; several requirements here carry a reconcile-with-HAL status.
• The D3 endurance targets (energy margin, cold usable-capacity, cycle-life fade) and the RF-in-container delivery-success/latency are Evidence-Pending; this module fixes the verification obligations (MR-VAL-006, PRD-VER-005) that close them — it does not assert the results.
• Detailed procedures, limits, sample sizes, and test-house selection live in the Verification Plan (ORCH-014); test houses are Evidence-Pending (concept §7).
• The exact WHO PQS device category/revision and the US regulatory classification are open items (concept §7, D4) and are confirmed during the Regulatory stage; compliance verification here references the regime by designation.
5. Architecture Overview
Verification is layered to mirror the software build. Static verification (requirement review, code review, static analysis) confirms each requirement is atomic, testable, and has acceptance criteria and a method. Dynamic verification runs the test hierarchy: unit tests per module, integration tests across module boundaries (device↔cloud↔mobile), and system tests against the SYS/SRS requirements end to end. Overlaid on this are the specialised campaigns — performance, reliability/fault-injection, stress, and environmental (cold-endurance) — plus the two program-defining qualifications: the 90-day cold-endurance test (MR-VAL-006, closing the D3 tension) and the RF-in-container test (PRD-VER-005, closing the principal open technical risk). Manufacturing verification confirms end-of-line software coverage; field validation confirms behaviour in representative use; regression guards every change through CI; security testing adversarially probes the attack surface; and compliance and documentation verification close the regulated record. A bidirectional traceability matrix binds every requirement to its verification result and back to its user need, and is assembled by ORCH-018 / EM-ENG-012.
---
7. Functional Requirements
7.1 Traceability, acceptance & method allocation
SRS-VER-001 — Every software requirement is verifiable with objective acceptance criteria
Requirement Statement | Every SCM software requirement shall carry a verification method (Test/Analysis/Inspection/Demonstration) and objective, measurable acceptance criteria, and no requirement shall be accepted as met without evidence satisfying those criteria.
Rationale | An unverifiable requirement cannot be accepted or defended; objective acceptance criteria are what make "shall" enforceable.
Priority | Must
Source | Verification
Verification Method | Inspection
Traceability | PRD-VER-001, ISO/IEC/IEEE 29148; → all SRS modules, Verification Matrix (ORCH-014).
Notes | Enforced by the library's field set; Evidence-Pending numeric targets carry the acceptance criterion "confirmed by named test" until closed.
SRS-VER-002 — Bidirectional requirement traceability
Requirement Statement | The software verification program shall maintain bidirectional traceability from each SRS requirement back to its System/Product requirement and user need, and forward to design output(s) and verification result(s), with no orphan requirements and no untraced verification.
Rationale | Traceability proves coverage, enables impact analysis when a requirement or design changes, and is a regulated-record expectation (GDP/Part 11).
Priority | Must
Source | Regulatory
Verification Method | Inspection
Traceability | PRD-VER-002, ISO 14971 / 29148; → Traceability Matrix (ORCH-018 / EM-ENG-012), SRS-DIAG, all SRS modules.
Notes | Matrix assembled by ORCH-018 / EM-ENG-012; each SRS module's Traceability section feeds it.
SRS-VER-003 — Verification matrix and risk-appropriate method allocation
Requirement Statement | The software verification program shall produce a verification matrix that allocates the most rigorous appropriate method to each requirement, preferring Test for safety- and function-critical requirements, and shall justify any use of Analysis, Inspection, or Demonstration in place of Test for a critical requirement.
Rationale | Severity-first (CLAUDE.md rule 3): critical requirements — detection, alerting, record integrity, endurance — warrant the strongest evidence, and any weaker method must be explicitly justified.
Priority | Must
Source | Safety
Verification Method | Analysis
Traceability | PRD-VER-003, ISO 14971; → SRS-SAFE, SRS-REL, SRS-PERF; Verification Plan (ORCH-014).
Notes | Reconcile with HAL (Stage 10); safety-critical software verification is Test-preferred.
7.2 Test hierarchy
SRS-VER-004 — Unit and static verification
Requirement Statement | The device, cloud, and mobile software shall be verified at unit level by automated unit tests with defined coverage targets and by static analysis and code review, such that each software unit is shown to meet its detailed design and coding standards before integration.
Rationale | Defects are cheapest and safest to catch at unit level; unit/static verification is the foundation of the IEC 62304 verification hierarchy.
Priority | Must
Source | Verification
Verification Method | Test, Inspection
Traceability | PRD-VER-001, IEC 62304, ISO/IEC/IEEE 29119; → SRS-DEV, SRS-CLOUD, SRS-MOB; Verification Plan.
Notes | Coverage targets set in the Verification Plan; safety-critical units carry higher coverage expectations (reconcile with HAL).
SRS-VER-005 — Integration verification across module and system boundaries
Requirement Statement | The software shall be verified at integration level across module boundaries and across the device↔connectivity↔cloud↔mobile boundaries, confirming interface contracts, data ordering, idempotent ingest, and end-to-end record consistency between integrated components.
Rationale | Many SCM failure modes live at the seams — store-and-forward resume, de-duplication, device↔cloud verdict consistency; these are only exposed by integration testing.
Priority | Must
Source | Verification
Verification Method | Test
Traceability | PRD-VER-001, SYS-016/020/176, ISO/IEC/IEEE 29119; → SRS-CONN, SRS-CLOUD, SRS-REL-009.
Notes | Includes device↔cloud verdict-consistency (SYS-016) and reconciliation (SYS-176).
SRS-VER-006 — System verification against SYS/SRS requirements
Requirement Statement | The integrated system software shall be verified end to end against the System and Software requirements — capture, detect, warn, decide, prove — under representative trip scenarios, including offline-first operation, reconnection, and the on-device verdict, with results traced to each requirement.
Rationale | System verification demonstrates that the composed software delivers the value chain the requirements promise, not just that each part works in isolation.
Priority | Must
Source | Verification
Verification Method | Test, Demonstration
Traceability | PRD-VER-001/002, SYS-001…032, ISO/IEC/IEEE 29119; → all SRS modules; System Test Plan (ORCH-014).
Notes | Offline-verdict and reconnection scenarios are headline system tests (D2).
7.3 Specialised campaigns (performance, reliability, stress, environmental)
SRS-VER-007 — Performance and reliability verification
Requirement Statement | The software shall be subjected to performance verification (latency, throughput, responsiveness, resource utilization) and reliability verification (fault injection, power-cycle/brown-out, watchdog, availability self-check, lossy-link, and end-to-end delivery) with results traced to the SRS-PERF and SRS-REL requirements.
Rationale | The performance and fail-safe claims are only credible if exercised by dedicated campaigns under worst-case load and fault conditions, not inferred from functional tests.
Priority | Must
Source | Verification
Verification Method | Test
Traceability | PRD-VER-001, SYS-037…047, SYS-163…176; → SRS-PERF-019, SRS-REL-019; Verification Plan.
Notes | Executes the verification obligations defined in SRS-PERF-019 and SRS-REL-019.
SRS-VER-008 — Stress and environmental (cold-endurance) software verification with the D3 test hook
Requirement Statement | The software shall be verified under stress (resource exhaustion, buffer-full, sustained event-burst) and under environmental extremes, and the software-relevant behaviour shall be exercised within the 90-day cold-endurance and cycle-life tests (MR-VAL-006) that close the D3 tension — confirming sensing, logging, verdict, alerting, graceful degradation, and zero data loss across the worst-case cold-duty profile.
Rationale | The D3 −30 °C / 90-day endurance is the top program risk and must be closed by test, not analysis; software degradation and data-integrity behaviour are integral to that endurance test.
Priority | Must
Source | Verification
Verification Method | Test
Traceability | PRD-VER-004 (D3 validate-by-test), SYS-045/046, SYS-163/164/165/171, IEC 60068, MR-VAL-006; → SRS-PERF-014, SRS-REL-012/014.
Notes | Evidence-Pending targets (energy margin, cold capacity, cycle-life fade, degradation thresholds) are closed by this test; no result asserted here. Headline validate-by-test.
SRS-VER-009 — RF-in-container qualification test hook for the connectivity software
Requirement Statement | The connectivity and store-and-forward software shall be verified by the RF-in-container test (PRD-VER-005), qualifying data-egress and excursion-alert delivery from inside representative metallized/insulated shippers, and confirming buffered records and queued alerts are delivered loss-free and in order on the first successful link.
Rationale | RF attenuation inside the shipper is the principal open technical risk; the store-and-forward, retry/back-off, and alert-priority behaviour must be qualified under the actual attenuated conditions.
Priority | Must
Source | Verification
Verification Method | Test, Analysis (link-budget model)
Traceability | PRD-VER-005, SYS-104/109/049, PRD-CONN-007; → SRS-CONN-019/026, SRS-REL-008.
Notes | Evidence-Pending — attenuation, achievable margin, delivery-success fraction, and alert latency are unresolved and close on this test; the antenna-strategy decision (SYS-109) is informed by it.
7.4 Manufacturing, field, regression & security
SRS-VER-010 — Manufacturing (end-of-line) software verification
Requirement Statement | The software shall be verified at manufacturing by an end-of-line test that confirms correct firmware version and integrity, exercises the manufacturing diagnostics coverage (SRS-DIAG-010), verifies provisioning/calibration state, and records an attributable end-of-line result to the device and the record system.
Rationale | End-of-line software verification keeps latent build/provisioning defects out of the field and establishes the device's birth record for reuse-life traceability.
Priority | Must
Source | Verification
Verification Method | Test, Inspection
Traceability | PRD-VER-001, SYS-030, SYS-032 (firmware integrity); → SRS-DIAG-010, SRS-DEV, SRS-OTA.
Notes | Firmware authenticity/integrity verification ties to SRS-SEC/SRS-OTA.
SRS-VER-011 — Field validation of intended use
Requirement Statement | The software shall be validated in representative field/use conditions — including the offline-verdict usability study, excursion-detection/false-alarm validation over real trip profiles, and audit-record/export simulation — confirming the software meets user needs in actual use, not only bench conditions.
Rationale | Verification proves requirements met; validation proves the right product was built — the offline verdict and alarm-balance behaviour especially must be confirmed with real users and real lanes.
Priority | Must
Source | Validation
Verification Method | Test (validation)
Traceability | PRD-VER-008, SYS-014/015/017 (alarm balance, offline verdict, honest status), IEC 62366-1 (usability); → Validation (PVaL, ORCH-015), SRS-MOB.
Notes | Usability study is a critical validation (IEC 62366-1); executed under PVaL (ORCH-015).
SRS-VER-012 — Regression and continuous-integration verification
Requirement Statement | Every software change (firmware, cloud, mobile) shall be subject to automated regression verification in continuous integration, re-running the affected unit/integration/system and security tests before release, such that a change cannot regress a previously verified safety- or function-critical requirement undetected.
Rationale | A reusable, OTA-updatable fleet (D1) evolves continuously; without automated regression a fix can silently break a verified behaviour across the field.
Priority | Must
Source | Verification
Verification Method | Test
Traceability | PRD-VER-001, SYS-032 (fail-safe update), IEC 62304; → SRS-OTA, SRS-DEV; Verification Plan.
Notes | OTA rollout is gated on regression pass; ties to SRS-OTA staged rollout and rollback.
SRS-VER-013 — Security testing of the software attack surface
Requirement Statement | The software shall be subjected to security testing — interface/penetration testing, fuzzing of external inputs, authentication/authorization and session testing, TLS/certificate validation testing, and debug-interface lock-down verification — with findings triaged and closed before release.
Rationale | The connectivity, provisioning, diagnostics, and OTA surfaces are the primary remote attack vectors; security must be tested adversarially, not assumed from design review.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | PRD-VER-001, PRD-SEC-005/006, SYS-110; → SRS-SEC (owner EM-ENG-008), SRS-CONN-025, SRS-DIAG-012/015.
Notes | Security-test scope and methods owned jointly with the Product Security Engineer (EM-ENG-008).
7.5 Compliance & documentation
SRS-VER-014 — Compliance verification against the WHO PQS / GDP / Part 11 regime
Requirement Statement | The software shall be verified for compliance with the applicable regulated-record and device-performance regime — WHO PQS device performance, EU GDP data-integrity (ALCOA+), and 21 CFR Part 11 electronic-records/audit-trail expectations — demonstrating attributable, complete, tamper-evident, exportable records and a compliant audit trail.
Rationale | The auditable cold-chain record is the product's compliance value; its data-integrity and electronic-records behaviour must be verified against the target regime (D4), not merely implemented.
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | PRD-VER-006, SYS-026/027/028 (record, export, audit trail), SYS-006 (trusted timestamp); → SRS-REG, SRS-DATA, SRS-CLOUD; GAMP 5 (CSV).
Notes | Exact WHO PQS category/revision and US classification are open items (concept §7, D4) confirmed in the Regulatory stage; cited by designation, revision to confirm.
SRS-VER-015 — Documentation and records verification
Requirement Statement | The software verification program shall verify that the required life-cycle documentation and records exist, are consistent with the as-built software, and are complete — including the requirements, design, traceability matrix, verification results, anomaly/problem records, and release records — such that the software's conformity is demonstrable and auditable.
Rationale | In a regulated device, undocumented or inconsistent verification is unverified; the record of verification is itself an object of verification (IEC 62304 / GAMP 5).
Priority | Must
Source | Regulatory
Verification Method | Inspection
Traceability | PRD-VER-001/002, IEC 62304, GAMP 5; → Traceability Matrix (ORCH-018 / EM-ENG-012), all SRS modules.
Notes | Documentation completeness is a release gate; anomaly/problem records per IEC 62304 problem resolution.
---
8. Non-functional Requirements
SRS-VER-016 — Verification environment representativeness and independence
Requirement Statement | Software verification shall be executed in environments (hardware-in-the-loop, cold chamber, RF chamber/shipper, network emulation, cloud staging) representative of the field, and safety-critical verification shall be reviewed with appropriate independence from the implementing engineer.
Rationale | Verification results are only trustworthy if the environment reflects real conditions (cold, RF-attenuated, intermittent) and if critical evidence is independently reviewed.
Priority | Should
Source | Verification
Verification Method | Inspection
Traceability | PRD-VER-004/005, SYS-046/104; → Verification Plan (ORCH-014), SRS-VER-008/009.
Notes | Test houses/chambers are Evidence-Pending (concept §7); representativeness is the binding criterion.
---
Error Handling
SRS-VER-017 — Verification of error, fault, and degraded-mode behaviour
Requirement Statement | The verification program shall explicitly verify the software's error-handling, fault-recovery, and graceful-degradation behaviour by fault and error injection, confirming the no-silent-failure invariant — that every fault is recovered-and-logged or visibly-flagged, never silently dropped.
Rationale | Error paths are the least-exercised and highest-risk code; the fail-safe and no-silent-failure claims must be positively demonstrated by injecting the faults, not assumed.
Priority | Must
Source | Reliability
Verification Method | Test (fault/error injection)
Traceability | SYS-167/169/170, PRD-REL-004; → SRS-REL-016, SRS-DIAG-014, SRS-CONN-023.
Notes | Executes the verification side of the no-silent-failure invariant (SRS-REL-016).
---
Diagnostics
SRS-VER-018 — Use of built-in diagnostics as verification evidence
Requirement Statement | The verification program shall use the device's built-in diagnostics (POST/BIST results, health metrics, fault/crash logs, diagnostic-coverage reports) as corroborating verification evidence, and shall verify that those diagnostics themselves report faults accurately (diagnostic-coverage verification).
Rationale | Built-in diagnostics both provide verification evidence and are themselves a requirement to be verified; a diagnostic that mis-reports health is worse than none.
Priority | Should
Source | Verification
Verification Method | Test
Traceability | SYS-113/170, PRD-REL-007/010; → SRS-DIAG-016, ERFKL (Stage 11).
Notes | Ties to diagnostics-coverage verification (SRS-DIAG-016) and the reliability program (SYS-173).
---
Security Considerations
SRS-VER-019 — Integrity and attributability of verification records
Requirement Statement | Verification records, test evidence, and the traceability matrix shall be attributable, retained under version control, and protected against undetected alteration, such that the software's verification history is itself a trustworthy, auditable record.
Rationale | In a regulated context, the verification evidence must be as trustworthy as the product record it supports; tampered or unattributable evidence invalidates the conformity claim.
Priority | Must
Source | Regulatory
Verification Method | Inspection
Traceability | PRD-VER-002, SYS-028 (audit trail concepts), GAMP 5 / 21 CFR Part 11; → SRS-SEC, SRS-REG.
Notes | Evidence-integrity mechanisms coordinate with SRS-SEC (EM-ENG-008).
---
Verification Requirements
SRS-VER-020 — Verification completeness and release gate
Requirement Statement | No software shall be released for field deployment or OTA rollout until every applicable SRS requirement has a passing verification result (or an approved, traced deviation) recorded in the verification matrix, safety-critical verification is reconciled with the HAL, and the D3 and RF-in-container qualifications relevant to the release are closed or explicitly deferred with rationale.
Rationale | The release gate is what makes verification consequential; shipping unverified critical behaviour into a public-health cold-chain device is the outcome this module exists to prevent.
Priority | Must
Source | Verification
Verification Method | Inspection
Traceability | PRD-VER-001/002/004/005/007, SYS-032 (fail-safe update); → all SRS modules, HAL (Stage 10), Verification Plan / Validation (ORCH-014/015).
Notes | Deferrals of Evidence-Pending items (D3 endurance, RF-in-container) must be explicit, traced, and risk-assessed; reconcile with HAL.
---
Traceability
Every requirement above traces upward to at least one Product Requirement (PRD-VER-…) and/or System Requirement (SYS-…), and forward to the Verification Plan, Validation, and the Traceability Matrix. Principal upstream anchors: PRD-VER-001 (verifiability/acceptance), PRD-VER-002 (traceability), PRD-VER-003 (method allocation), PRD-VER-004 (D3 validate-by-test), PRD-VER-005 (RF-in-container), PRD-VER-006 (environmental/regulatory suite), PRD-VER-007 (safety/HAL reconciliation), PRD-VER-008 (validation); SYS-045/046 (endurance), SYS-104 (RF), SYS-016/020/176 (integration/consistency/reconciliation), SYS-026/027/028 (record/export/audit), SYS-032 (fail-safe update), SYS-173 (reliability program). Cross-module: this module governs the verification obligations discharged in SRS-PERF-019, SRS-REL-019, SRS-DIAG-016, and SRS-CONN-026; security testing → SRS-SEC (EM-ENG-008); compliance/records → SRS-REG and SRS-DATA/SRS-CLOUD; validation → PVaL (ORCH-015); safety verification → SRS-SAFE / HAL (Stage 10); matrix assembly → ORCH-018 / EM-ENG-012. Evidence-Pending items (D3 endurance results, RF-in-container delivery/latency, WHO PQS category/revision, US classification, test houses) are carried as verification obligations that close on their named tests (MR-VAL-006, PRD-VER-005) — the results are not asserted here (evidence).
SRS-200 — Device Software (module SRS-DEV)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-DEV · Owner: Embedded Software Engineer (EM-ENG-007). Detailed device-firmware software requirements decomposed from the System Requirements (SYS-…) and Product Requirements (PRD-…), realising the on-device half of the value chain capture → detect → warn → decide → prove. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify, at the firmware level, how the SCM device software boots, initialises, senses, schedules, manages power, persists an evidentiary record, recovers from faults, and is provisioned/serviced — such that the device delivers a gap-free, tamper-evident cold-chain record and an honest on-device verdict for a 90-day trip while living inside the −30 °C to +60 °C energy budget (D3) and behaving offline-first (D2).
2. Scope
Boot and secure boot; startup diagnostics/POST; subsystem initialisation; configuration management; factory and field provisioning; persistent storage, the trip log, tamper-evident append-only and power-fail-safe writes; logging/diagnostics services; the sensor framework and drivers; task scheduling; the power-state machine and energy governor (sleep / logging / alert / charging / OTA / fault / shipping); watchdog; fault management and recovery; calibration coefficient management; manufacturing/field/developer operating modes; time synchronisation and the trusted clock; configuration backup and factory reset; and safe mode. Excludes connectivity/transport behaviour (see SRS-300 / SRS-CONN), cloud (SRS-CLOUD), mobile app (SRS-MOB), cryptographic mechanism detail (SRS-SEC / EM-ENG-008), and OTA campaign/packaging (SRS-OTA); those are referenced where the device firmware interfaces with them.
3. Applicable Standards
IEC 62304 (medical-device software life-cycle processes — safety classification and process rigour, to be confirmed with EM-ENG-006); ISO 14971 (risk-control software requirements); IEC 62443 concepts (secure device software); 21 CFR Part 11 / EU GDP / ALCOA+ (attributable, legible, contemporaneous, original, accurate — enduring/available electronic records); ISO/IEC 17025 concepts (calibration coefficient handling); JEDEC/eMMC or serial-flash endurance guidance (wear management). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The MCU provides a hardware root of trust / secure-boot primitive and a secure key store sufficient to anchor image authentication and provisioned credentials (mechanism owned by SRS-SEC / EM-ENG-008; confidence Medium — silicon selection pending).
• Non-volatile media is serial NOR/NAND flash or eMMC with finite program/erase endurance; exact part and endurance figure are Evidence-Pending and drive wear-levelling sizing (SRS-DEV-008).
• A low-power RTC with an independent time domain and battery/super-cap backing keeps time across resets and deep sleep (SRS-DEV-020); its free-running drift is bounded but not yet measured (Evidence-Pending).
• Final battery chemistry, usable cold capacity, and per-load energy allocations are Evidence-Pending (SYS-086/087/090, D3); the energy governor here is written against the allocation, not against invented capacity numbers.
• An RTOS or equivalent cooperative scheduler is available on the target MCU (SRS-DEV-013).
5. Architecture Overview
Firmware is layered: (a) an immutable boot ROM / first-stage loader that verifies the second-stage image; (b) a second-stage bootloader supporting A/B image slots, rollback, and safe mode entry; (c) a hardware abstraction layer (HAL) and driver framework for sensors, storage, radios, HMI, and power; (d) a device-services middleware layer (config store, log/record engine, calibration, time service, fault manager, diagnostics); and (e) the application implementing sensing, excursion evaluation, the on-device verdict, and the power-state machine. A single power-state / energy governor (SYS-089) owns transitions between sleep, logging, alert, charging, OTA, fault, and shipping states and hard-gates every high-drain load. A watchdog and fault manager provide the last line of defence, driving the device to safe mode without ever silently abandoning an armed trip.
---
7. Functional Requirements
7.1 Boot, secure boot & startup
SRS-DEV-001 — Deterministic staged boot sequence
Requirement Statement | The device firmware shall execute a deterministic staged boot sequence — first-stage loader → image verification → HAL/clock init → device-services init → application start — and shall reach an armed-trip monitoring state within the boot-time budget defined in SRS-PERF, or else enter safe mode (SRS-DEV-023).
Rationale | A bounded, ordered boot guarantees the device resumes monitoring predictably after any reset or power event without indeterminate states that could drop samples.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-099 (clean deterministic state), PRD-REL-005; → SRS-PERF (boot budget).
Notes | Boot-time numeric target is set in SRS-PERF; not asserted here (evidence).
SRS-DEV-002 — Secure boot / image authentication
Requirement Statement | The device firmware shall verify the authenticity and integrity of each firmware image against the hardware root of trust before transferring execution to it, and shall refuse to execute an image that fails verification.
Rationale | Only genuine, unmodified firmware may run; an unverified image could disarm monitoring, corrupt the record, or exfiltrate credentials.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-032 (signed, fail-safe firmware), PRD-SEC-008, PRD-FUN-022; → SRS-SEC (key/RoT mechanism, EM-ENG-008), SRS-OTA-… (image acceptance).
Notes | Cryptographic mechanism (algorithm, key hierarchy, revocation) owned by SRS-SEC; this requirement mandates the verification gate. Standard: IEC 62443 concepts (revision to confirm).
SRS-DEV-003 — Power-on self-test and startup diagnostics (POST)
Requirement Statement | On each boot the device firmware shall execute a power-on self-test that checks MCU core state, RAM, non-volatile storage accessibility, RTC validity, each sensor channel's presence/health, and battery/fuel-gauge communication, and shall record the pass/fail result of each check to the diagnostic log.
Rationale | A device that boots with a faulted subsystem must detect it before it is credited as monitoring, so a defect never masquerades as a clean trip.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-007 (sensor plausibility/fault flag), SYS-024 (pre-arm readiness), PRD-REL-004; → SRS-DIAG-…, SRS-DEV-024 (pre-arm gate).
Notes | POST outcome feeds the pre-arm readiness check (SRS-DEV-019 mode logic / SYS-024).
SRS-DEV-004 — Ordered subsystem initialisation with fault isolation
Requirement Statement | The device firmware shall initialise subsystems in a defined dependency order (clock/power → storage → time service → sensors → HMI → radios) and, if a non-critical subsystem fails to initialise, shall continue booting with that subsystem marked unavailable rather than aborting the boot.
Rationale | Sensing and logging must survive the loss of a non-essential subsystem (e.g. a radio); only a failure that prevents safe monitoring should block the boot.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-098 (function shedding preserves core sensing/logging), PRD-REL-008; → SRS-DEV-017.
Notes | "Critical" set (storage, clock, temperature sensing) versus "non-critical" (radios, GNSS, Wi-Fi) is defined in design; a critical-subsystem failure routes to safe mode (SRS-DEV-023).
7.2 Configuration & provisioning
SRS-DEV-005 — Versioned configuration management
Requirement Statement | The device firmware shall maintain device configuration and the active trip profile in a versioned, integrity-checked configuration store, applying a configuration change only after it validates against defined constraints, and shall retain the previously active configuration until the new one is committed.
Rationale | Configuration is part of the evidentiary record; a corrupt or half-applied config must never silently take effect or leave the device un-configured.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-022 (profile config), SYS-025 (config validation/error prevention), PRD-FUN-014; → SRS-DEV-021 (backup/restore).
Notes | Constraint validation (low < high threshold, interval in range) mirrors SYS-025; invalid config is rejected before it can be armed.
SRS-DEV-006 — Factory provisioning of identity and credentials
Requirement Statement | The device firmware shall support a controlled factory-provisioning operation that writes the device's unique identity, cryptographic credentials, and immutable factory parameters to secure storage exactly once, and shall thereafter prevent overwrite of those immutable parameters in field operation.
Rationale | A unique, tamper-resistant identity is the anchor for authenticated connectivity, attributable records, and fleet enrolment; it must be set once under controlled conditions.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-110 (authenticated provisioning), SYS-030 (readiness/identity), PRD-CONN-006, PRD-SEC-005; → SRS-SEC (key provisioning, EM-ENG-008).
Notes | Key generation/injection mechanism owned by SRS-SEC; this mandates the one-time, protected device-side write.
SRS-DEV-007 — Field provisioning of the trip profile
Requirement Statement | The device firmware shall accept a field-provisioned trip profile over an authenticated bearer or on-device, validate it (SRS-DEV-005), record which profile is active with its version, and require this profile to be present and valid before the device can be armed.
Rationale | Monitoring must match what harms the specific product, and the active profile must be part of the attributable record.
Priority | Must
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-022, SYS-024, PRD-FUN-014; → SRS-CONN-006 (secure provisioning transport).
Notes | Authenticated pairing/transport is specified in SRS-CONN; the profile is applied before arming (SYS-023).
7.3 Persistent storage, the trip log & data integrity
SRS-DEV-008 — Non-volatile trip log with wear management and full-trip capacity
Requirement Statement | The device firmware shall store all samples, events, excursions, gap markers, configuration, and lifecycle actions in a non-volatile log sized to hold a complete 90-day trip at the minimum logging interval with defined margin, applying wear-levelling so that no single storage region is worn out within the device's specified reuse life.
Rationale | Offline-first store-and-forward requires capacity for the worst-case offline duration, and a reusable device must not wear its flash out mid-life and start losing records.
Priority | Must
Source | Reliability
Verification Method | Analysis, Test
Traceability | SYS-018 (offline buffering of a full trip), SYS-002 (gap-free logging), PRD-FUN-012; → SRS-PERF (flash endurance), SRS-REL.
Notes | Storage sizing and flash endurance figure are Evidence-Pending (SYS-018 status; part TBD); requirement is stated against the trip duration, not an invented byte count.
SRS-DEV-009 — Tamper-evident, append-only log writes
Requirement Statement | The device firmware shall write log records append-only and shall protect the log with a tamper-evidence mechanism (e.g. a per-record chained integrity value) such that any alteration, deletion, or re-ordering of committed records is detectable on later verification.
Rationale | A regulated cold-chain record must be attributable and defensible; silent alteration of the record would destroy its evidentiary value.
Priority | Must
Source | Regulatory
Verification Method | Test, Analysis
Traceability | SYS-026 (tamper-evident record), SYS-028 (append-only audit trail), PRD-FUN-017, PRD-SEC-002; → SRS-SEC (integrity primitive), SRS-DATA. Standard: 21 CFR Part 11 / ALCOA+ (revision to confirm).
Notes | Cryptographic strength of the integrity value is governed by SRS-SEC; this mandates append-only + detectability at the firmware layer.
SRS-DEV-010 — Power-fail-safe atomic writes
Requirement Statement | The device firmware shall commit each log or configuration write atomically, such that a power loss, brown-out, or reset during a write leaves the store with either the fully written new record or the last valid prior state, and never a partially written or corrupt record.
Rationale | A brown-out mid-write must never corrupt the evidentiary record; the storage integrity guarantee underpins every other data claim.
Priority | Must
Source | Reliability
Verification Method | Test (power-interruption / brown-out injection)
Traceability | SYS-099 (brown-out clean safe-state, no corruption), SYS-002, PRD-REL-005; → SRS-REL.
Notes | Verified by repeated power-cut-during-write cycling; ties to the regulated-rail brown-out behaviour in SYS-099.
SRS-DEV-011 — Explicit gap marking on any missed interval
Requirement Statement | The device firmware shall write an explicit gap marker into the log whenever a scheduled sample cannot be taken or committed, so that no armed-trip interval is silently absent from the record.
Rationale | A missing sample must never masquerade as in-range data (no false reassurance); ALCOA+ completeness requires gaps be visible, not concealed.
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-002 (gap-free logging with explicit gap marker), SYS-017 (honest incomplete monitoring), PRD-FUN-017; → SRS-DEV-023 (safe mode records the gap).
Notes | Gap markers feed the honest-verdict logic (SYS-017): a gapped record cannot read as an unqualified "accept".
7.4 Logging, diagnostics & the sensor framework
SRS-DEV-012 — Diagnostic and operational logging service
Requirement Statement | The device firmware shall provide a leveled logging/diagnostics service that records boot results, faults, resets (with cause), state transitions, radio attempts, and battery events to non-volatile diagnostic storage, retained independently of the trip log and available to fleet tooling.
Rationale | Silent-comms or dead-in-cold devices can only be triaged if the device kept a durable, attributable diagnostic trail separate from the evidentiary trip record.
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-113 (connectivity diagnostics), SYS-030 (readiness reporting), PRD-CONN-010; → SRS-DIAG-…, SRS-CONN-013.
Notes | Diagnostic logging is energy-cheap (SYS-113); it must not contend with the trip log for integrity-critical storage.
SRS-DEV-013 — Sensor framework and channel drivers
Requirement Statement | The device firmware shall provide a sensor framework that acquires the internal temperature, external-probe temperature, relative humidity, and 3-axis acceleration channels through a uniform driver interface, timestamps each acquisition, and reports per-channel health so a faulted channel is flagged without discarding acquired raw values.
Rationale | Dual-channel, multi-sensor acquisition must be uniform, timestamped, and self-checking; a silently-wrong reading is more dangerous than a flagged one.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-001 (interval sampling), SYS-003 (dual-channel temperature), SYS-004 (humidity), SYS-005 (shock), SYS-007 (plausibility/fault flag), PRD-FUN-001/002/003/004.
Notes | Plausibility bounds and data-quality flags per SYS-007; raw value is always retained.
SRS-DEV-014 — Deterministic task scheduling
Requirement Statement | The device firmware shall schedule sampling, storage commit, excursion evaluation, and connectivity tasks under a deterministic scheduler that guarantees the sampling task meets its configured interval regardless of the load of lower-priority tasks such as uploads.
Rationale | Gap-free logging is only guaranteed if the sampling task cannot be starved by a busy or blocked upload; sampling is the highest-priority periodic function.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-001, SYS-002, PRD-FUN-001; → SRS-DEV-016 (watchdog on task liveness).
Notes | Task priorities and worst-case timing are established in design; scheduler starvation is a tested failure mode.
7.5 Power management & the energy-governor state machine (D3)
SRS-DEV-015 — Power-state machine (sleep / logging / alert / charging / OTA / fault / shipping)
Requirement Statement | The device firmware shall implement a power-state machine with defined states — deep-sleep, logging, alert, charging, OTA, fault, and shipping — and defined, deterministic transitions between them, holding the device in the lowest viable state consistent with its current obligations (armed monitoring, queued alert, active charge, or update).
Rationale | A single, explicit state machine is the enabler that makes 90-day cold endurance achievable and makes device behaviour analysable against the energy budget (D3).
Priority | Must
Source | Architecture
Verification Method | Test, Analysis
Traceability | SYS-089 (duty-cycled energy governor), SYS-086 (90-day operation), PRD-ELEC-008, PRD-ARCH-004; D3. → SRS-PERF, SRS-REL.
Notes | State transitions and residency are the basis of the energy-budget analysis (SYS-090, Evidence-Pending); shipping state is the pre-deployment ultra-low-power hold (SRS-DEV-024).
SRS-DEV-016 — Hard energy-gating of high-drain loads
Requirement Statement | The energy governor shall gate every high-drain load — cellular, Wi-Fi, BLE, and GNSS in particular — so that no radio or GNSS activity occurs outside its scheduled, energy-budgeted window, with the sole exception of a queued excursion/shock alert which may pre-empt to deliver within its latency obligation.
Rationale | Radios dominate the energy budget; continuous live radio for 90 days is prohibitive, so loads must be bounded hard while never blocking a rescue-enabling alert.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-089 (hard radio gating), SYS-112 (alert-path priority exception), SYS-011, PRD-ELEC-008; D2/D3. → SRS-CONN-016 (alert prioritisation).
Notes | The alert exception is the single sanctioned bypass; all other radio use is scheduled.
SRS-DEV-017 — Low-battery graceful degradation ordering
Requirement Statement | As battery energy depletes, the device firmware shall shed discretionary functions in a defined order (location/GNSS, then uploads, then non-essential HMI) before core sensing and logging, warn of the low-battery state, and shall never silently stop monitoring an armed device.
Rationale | A depleting device must protect the evidentiary record over convenience — losing the log is worse than losing an upload — and must remain honest about its status to the end.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-098 (low-battery graceful degradation), SYS-021 (retain excursion/audit first), PRD-ELEC-010, PRD-REL-008.
Notes | Degradation order and warning thresholds are Evidence-Pending (SYS-098); requirement fixes the ordering principle, not the numeric thresholds.
7.6 Watchdog, fault management & recovery
SRS-DEV-018 — Watchdog supervision and recovery
Requirement Statement | The device firmware shall service a hardware watchdog only while its critical tasks (sampling, storage commit, and the state machine) are demonstrably alive, and on watchdog timeout shall reset into a recovery path that resumes monitoring and records the reset with its cause.
Rationale | A hung firmware must be recovered automatically without losing the armed trip or the reason it hung.
Priority | Must
Source | Reliability
Verification Method | Test (fault injection)
Traceability | SYS-099 (clean deterministic recovery), PRD-REL-005; → SRS-DEV-012 (reset-cause log), SRS-REL.
Notes | Watchdog kick must be conditional on task liveness, not unconditional, or it defeats its purpose.
SRS-DEV-019 — Fault management and classification
Requirement Statement | The device firmware shall maintain a fault manager that detects, classifies (recoverable vs. critical), records, and acts on faults — retrying or isolating recoverable faults and transitioning to safe mode (SRS-DEV-023) on a critical fault — without ever leaving an armed device silently unmonitored.
Rationale | Faults must be handled by policy, not ad hoc; the safety-relevant invariant is that monitoring is either continued or its loss is made visible.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-007, SYS-017, SYS-098, PRD-REL-004/008; → SRS-DEV-023 (safe mode), SRS-SAFE.
Notes | Critical-fault set and actions to be reconciled with the Hazard Analysis (HAL, Stage 10).
7.7 Calibration, time & modes
SRS-DEV-020 — Calibration coefficient management
Requirement Statement | The device firmware shall store per-sensor calibration coefficients with their calibration date and validity period in protected storage, apply them to every raw reading before it is logged, and flag the device as out-of-calibration (blocking arming) when the validity period has expired.
Rationale | A defensible temperature record depends on applied, in-date calibration; an out-of-calibration device must not be deployed unnoticed.
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-030 (readiness: calibration validity), SYS-024 (pre-arm), PRD-FUN-020; → SRS-DATA. Standard: ISO/IEC 17025 concepts (revision to confirm).
Notes | Calibration coefficients are attributable record data; changes are audit-logged (SRS-DEV-009).
SRS-DEV-021 — Trusted clock and time discipline
Requirement Statement | The device firmware shall maintain a monotonic real-time clock referenced to UTC, discipline it to a trusted network time reference whenever a link is available, timestamp every sample and event with the time source and synchronisation status, and continue free-running with disclosed drift when no reference is available.
Rationale | A defensible record and correct excursion timing require a trusted clock; offline-first means the clock must free-run with bounded, disclosed drift and re-discipline on reconnection.
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-006 (trusted timestamp), SYS-118 (network time sync), PRD-FUN-019, PRD-ARCH-007; → SRS-CONN-017 (network-time transport). Standard: 21 CFR Part 11 attributable/accurate (revision to confirm).
Notes | Free-running drift bound is Evidence-Pending (RTC part TBD); the disclosure of drift is mandatory regardless.
SRS-DEV-022 — Manufacturing, field, and developer operating modes
Requirement Statement | The device firmware shall support distinct manufacturing, field, and developer operating modes with mode-appropriate capabilities and access controls, shall record every mode entry/exit in the audit trail, and shall disable manufacturing/developer diagnostic and debug capabilities in field mode.
Rationale | Test and debug capabilities that are legitimate in manufacturing become an attack surface and a data-integrity risk in the field; mode must be explicit, controlled, and logged.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-028 (audit trail of lifecycle actions), PRD-SEC-004; → SRS-SEC (debug/JTAG lock, EM-ENG-008), SRS-DEV-024.
Notes | Secure-debug / JTAG protection mechanism owned by SRS-SEC; this mandates the mode model and its field lockdown.
7.8 Backup, reset, shipping & safe mode
SRS-DEV-023 — Configuration backup and restore; factory reset
Requirement Statement | The device firmware shall keep a recoverable backup of committed device configuration and calibration, restore from it if the primary configuration store is found corrupt, and provide an authenticated factory-reset operation that clears field configuration and trip data while preserving factory identity, credentials, and calibration.
Rationale | Reusability requires safe re-provisioning between trips without destroying the device's immutable identity and calibration; a corrupt config must be recoverable rather than bricking the device.
Priority | Must
Source | Reliability
Verification Method | Test, Demonstration
Traceability | SYS-030 (reuse-loop readiness), SYS-022, PRD-FUN-020, PRD-SVC-004; → SRS-DEV-005/006.
Notes | Factory reset is authenticated (SRS-DEV-022 mode / SYS-110); it must not erase calibration or identity.
SRS-DEV-024 — Safe mode and shipping (long-term storage) mode
Requirement Statement | The device firmware shall provide (a) a safe mode entered on critical fault or failed image verification that preserves committed data, indicates the fault, and supports recovery/OTA without executing suspect application code; and (b) a shipping/storage mode that holds the un-armed device in an ultra-low-power state preserving battery and calibration until it is next provisioned and armed.
Rationale | The device must fail into a defined, recoverable, honest state rather than bricking, and must survive shelf/transit storage without draining the budget or losing readiness.
Priority | Must
Source | Reliability
Verification Method | Test, Demonstration
Traceability | SYS-032 (never bricked/unmonitored-without-notice), SYS-024 (pre-arm readiness), SYS-092 (quiescent budget), PRD-SAFE-010, PRD-REL-008; → SRS-DEV-015 (states), SRS-OTA (recovery).
Notes | Safe mode never presents an unqualified "accept"; shipping mode wakes to full readiness on provisioning/arming.
---
8. Non-functional Requirements
SRS-DEV-025 — Boot and wake timing
Requirement Statement | The device firmware shall complete cold boot to armed-monitoring, and wake from deep sleep to a completed sample commit, each within the timing budgets defined in SRS-PERF.
Rationale | Wake time is both a data-integrity concern (a slow wake can miss an interval) and an energy concern (wake overhead is charged against every interval).
Priority | Must
Source | Performance
Verification Method | Test
Traceability | SYS-001, SYS-089/090; → SRS-PERF (numeric budgets).
Notes | Numeric targets held in SRS-PERF (evidence); this requirement binds firmware to them.
SRS-DEV-026 — Memory and storage headroom
Requirement Statement | The device firmware shall operate with defined RAM, stack, and non-volatile storage headroom margins under worst-case load, and shall detect and log exhaustion of any of these resources as a fault.
Rationale | Silent resource exhaustion is a classic cause of mid-trip crashes and lost samples; headroom must be designed in and monitored.
Priority | Should
Source | Reliability
Verification Method | Analysis, Test
Traceability | SYS-021 (buffer-full graceful behaviour), PRD-REL-008; → SRS-PERF, SRS-DEV-019.
Notes | Margin figures set in SRS-PERF; exhaustion routes through the fault manager.
SRS-DEV-027 — Data-retention durability across power loss and reset
Requirement Statement | The device firmware shall guarantee that all committed samples, events, configuration, calibration, and audit entries survive any number of power losses, brown-outs, and resets without loss or corruption for the device's specified reuse life.
Rationale | Enduring, available records (ALCOA+) are the product's compliance value; durability across power events is non-negotiable.
Priority | Must
Source | Reliability
Verification Method | Test (power-cycle endurance)
Traceability | SYS-018, SYS-099, SYS-026, PRD-REL-005; → SRS-DEV-010, SRS-REL.
Notes | Verified alongside the power-fail-safe write test (SRS-DEV-010).
---
Error Handling
SRS-DEV-028 — No-silent-failure invariant
Requirement Statement | For every detected error affecting sensing, timekeeping, storage, or the armed state, the device firmware shall record the error in the log/diagnostic trail and reflect it in device status, such that no error affecting the record's completeness or the monitoring state is handled silently.
Rationale | The overarching safety principle is honesty: any condition that could compromise the record must be visible, never suppressed.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-007, SYS-017, SYS-021, PRD-SAFE-003; → SRS-DEV-011/019/023.
Notes | Ties directly to the honest-incomplete-monitoring requirement (SYS-017).
---
Diagnostics
SRS-DEV-029 — On-demand self-test and health readout
Requirement Statement | The device firmware shall provide an on-demand built-in self-test and a structured health readout (POST results, sensor health, storage state, battery SoC/SoH, calibration validity, firmware version, last reset cause) accessible to service tooling over an authenticated local interface.
Rationale | Depot triage and readiness decisions require a complete, trustworthy device health picture on demand.
Priority | Should
Source | Reliability
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness reporting), SYS-113, PRD-FUN-020, PRD-SVC-004; → SRS-DIAG-…, SRS-CONN-013.
Notes | Access is authenticated per SRS-DEV-022 mode model.
---
Security Considerations
SRS-DEV-030 — Least-privilege field runtime and protected secrets
Requirement Statement | In field mode the device firmware shall expose no debug, memory-dump, or unauthenticated configuration interface, and shall keep provisioned credentials and calibration in protected storage inaccessible to application-level read-out.
Rationale | The fielded device is physically accessible to many parties; its runtime attack surface and secret exposure must be minimised.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | PRD-SEC-004/005, SYS-110, SYS-028; → SRS-SEC (mechanisms, EM-ENG-008), SRS-DEV-022.
Notes | Detailed cryptographic and secure-storage mechanisms owned by SRS-SEC (EM-ENG-008); this states the device-software posture.
---
Verification Requirements
SRS-DEV-031 — Power-interruption and cold-environment firmware verification
Requirement Statement | The device software shall be verified by power-interruption/brown-out injection during writes, watchdog and fault-injection testing, resource-exhaustion testing, and functional regression across the −30 °C to +60 °C operating range, with results traced to the requirements above.
Rationale | The firmware's reliability and integrity claims are only credible if verified under the power and thermal stresses the device actually meets in service.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-086, SYS-099, PRD-REL-005; → SRS-VER, SRS-DEV-010/018/019/027.
Notes | Cold-environment endurance is Evidence-Pending (D3); this fixes the verification obligation.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-001/002/003/006/007 (sense & log), SYS-017/018/021/026/028 (record integrity), SYS-030/032 (readiness & OTA), SYS-086/089/090/092/098/099 (power/energy, D3). Cross-module: cryptographic and secure-boot mechanisms → SRS-SEC (EM-ENG-008); transport, provisioning, and network-time bearers → SRS-CONN; OTA acceptance/recovery → SRS-OTA; numeric timing/endurance budgets → SRS-PERF; safety risk-control reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (storage sizing, RTC drift bound, energy allocations, degradation thresholds) are carried as gaps, not invented, and close on the associated tests.
SRS-300 — Connectivity (module SRS-CONN)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-CONN · Owner: Embedded Software Engineer (EM-ENG-007) with the RF/Connectivity Engineer. Detailed connectivity-software requirements decomposed from the System Requirements (SYS-104…118, plus SYS-011/019/
020) and Product Requirements (PRD-CONN 800, PRD-FUN), realising the warn → prove data path offline-first (D2) and within the 90-day energy budget (D3). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM device software establishes, selects, prioritises, and manages its communication bearers (LTE-M, Wi-Fi, BLE, USB), moves buffered telemetry and excursion alerts to the cloud, is provisioned locally, handles offline operation and reconnection without data loss, and diagnoses its own connectivity — all under an energy budget that forbids continuous live radio and inside an RF-hostile metallized/insulated shipper.
2. Scope
LTE-M (PSM/eDRX, APN); Wi-Fi opportunistic offload; BLE provisioning and short-range offload; USB device transport; opportunistic multi-bearer selection and fallback; MQTT and HTTPS application transport; offline store-and-forward and reconnection; connection retry and back-off; bandwidth/latency management; network diagnostics and signal-quality logging; certificate handling; connection-quality assessment; roaming; network time synchronisation; alert-path prioritisation; and RF-in-container degradation handling. Excludes the radio hardware/antenna design and link-budget (S5/S1, RF engineer), cloud-side ingest and routing (SRS-CLOUD), OTA campaign/packaging (SRS-OTA), on-device storage integrity (SRS-DEV), and cryptographic mechanisms (SRS-SEC); those are referenced where the connectivity software interfaces with them.
3. Applicable Standards
3GPP LTE-M / Cat-M1 (LPWAN, PSM/eDRX power-saving); IEEE 802.11 (Wi-Fi); Bluetooth Core Specification (BLE, pairing/security); USB (device/PD); MQTT v5 (application messaging); TLS 1.2+/1.3 (secure transport); FCC Part 15 / CE RED (radio conformance, see SRS-REG / PRD-REG-004); IEC 62304 (software life-cycle process). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The LTE-M modem supports PSM and eDRX power-saving and roaming-capable SIM/eSIM provisioning; per-market carrier coverage and the SIM strategy are Evidence-Pending (SYS-116, PRD-CONN-001).
• RF attenuation inside a representative metallized/insulated shipper, the achievable link margin, and the per-trip delivery-success target are Evidence-Pending and close on the RF-in-container test (PRD-VER-005 / SYS-104); no attenuation or success-rate figure is asserted here.
• The energy governor (SRS-DEV-015/016) owns the scheduled windows in which radios may operate; this module specifies behaviour within those windows and the sanctioned alert exception.
• On-device non-volatile buffering, ordering, and integrity are provided by SRS-DEV (storage/log); this module specifies the transport that drains and resumes that buffer.
• Cloud endpoints provide idempotent (exactly-once) ingest and a trusted network-time reference (SRS-CLOUD).
5. Architecture Overview
A bearer-selection layer ranks available bearers — LTE-M (primary WAN), then Wi-Fi, then BLE, plus USB when tethered — and picks the lowest-energy path that meets the applicable cadence/latency target, falling back automatically and never requiring any single bearer to be present. Above it, an application-transport layer carries excursion/shock alerts (latency-prioritised) and bulk telemetry (cadence-driven) over MQTT v5 and/or HTTPS on mutually authenticated TLS. A store-and-forward controller cooperates with the on-device buffer (SRS-DEV) to resume interrupted uploads in intact chronological order with exactly-once ingest. A retry/back-off governor bounds reconnection energy. A diagnostics collector logs bearer, signal quality, attempts/failures, and backlog depth for fleet triage and to feed RF-in-container evidence. All radio activity is hard-gated by the energy governor except a queued alert.
---
7. Functional Requirements
7.1 Bearers
SRS-CONN-001 — LTE-M primary wide-area bearer with power-saving
Requirement Statement | The connectivity software shall use LTE-M (3GPP Cat-M1) as the primary wide-area bearer for store-and-forward upload and excursion-alert delivery, operating the modem in a power-saving mode (PSM and/or eDRX) between scheduled windows consistent with the 90-day energy budget.
Rationale | LTE-M is a low-power WAN suited to duty-cycled operation across global lanes; PSM/eDRX preserves the energy budget while keeping the primary alert path available (D2/D3).
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-105 (LTE-M primary), SYS-089 (energy governor), PRD-CONN-001, PRD-ELEC-008; D2/D3. Standard: 3GPP LTE-M/Cat-M1 (revision to confirm).
Notes | Coverage/roaming per market is Evidence-Pending (SRS-CONN-015). Radio use only within governor windows (SRS-DEV-016), alert exception aside.
SRS-CONN-002 — APN and carrier-session configuration
Requirement Statement | The connectivity software shall establish the LTE-M packet-data session using a configurable APN and carrier/credential profile, validate session establishment before attempting data transfer, and record session establishment success/failure to connectivity diagnostics.
Rationale | Multi-market, multi-carrier operation requires configurable APN/credentials, and a failed session must be diagnosable rather than presenting as generic "no data".
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-105, SYS-116 (roaming), SYS-113 (diagnostics), PRD-CONN-001/012; → SRS-CONN-011.
Notes | APN/credential profile is part of provisioned configuration (SRS-DEV-005/007).
SRS-CONN-003 — Wi-Fi opportunistic bulk offload
Requirement Statement | The connectivity software shall support Wi-Fi (IEEE 802.11) for opportunistic bulk-data offload to a provisioned network where available and shall prefer Wi-Fi for backlog upload when it is the lowest-energy path satisfying the cadence target.
Rationale | Higher-throughput, lower-cost offload at depots and facilities drains the buffered backlog while preserving the energy budget.
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-106 (Wi-Fi offload), SYS-108 (bearer selection), PRD-CONN-002; → SRS-CONN-006. Standard: IEEE 802.11 (revision to confirm).
Notes | Enterprise-Wi-Fi authentication (802.1X/EAP) is Evidence-Pending (PRD-CONN-002).
SRS-CONN-004 — BLE provisioning and short-range offload
Requirement Statement | The connectivity software shall support Bluetooth Low Energy for local provisioning/configuration, local record read, and short-range data offload to an authenticated, paired app or gateway (pairing per SRS-CONN-012).
Rationale | BLE enables frictionless local setup and offload where WAN is absent, keeping the login-free edge read available while configuration stays access-controlled (D2/AD-5).
Priority | Must
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-107 (BLE provisioning/offload), SYS-110 (authenticated pairing), PRD-CONN-003, PRD-MOB-001; → SRS-DEV-007 (field provisioning). Standard: Bluetooth Core / BLE (revision to confirm).
Notes | Edge verdict read stays login-free (SRS-DEV / SYS-015); configuration requires authentication.
SRS-CONN-005 — USB device transport for offload and update
Requirement Statement | The connectivity software shall support a USB-C data interface for local record offload and firmware-update transfer at the depot, available without a wide-area link and without compromising the tamper-evidence or ordering of the record.
Rationale | A wired path guarantees offload and update at the depot regardless of RF conditions and provides a fallback when radios are unavailable.
Priority | Should
Source | Architecture
Verification Method | Test
Traceability | SYS-114 (USB-C update transport), PRD-CONN-011; → SRS-DEV-009 (integrity), SRS-OTA.
Notes | USB charge input and its protection are covered in the power domain (SYS-093/102/103); this covers the data function only.
7.2 Bearer selection & offload behaviour
SRS-CONN-006 — Opportunistic multi-bearer selection and fallback
Requirement Statement | The connectivity software shall select among available bearers — LTE-M (primary), then Wi-Fi, then BLE, plus USB when tethered — preferring the lowest-energy path that satisfies the applicable cadence/latency target, and shall fall back automatically without requiring any single bearer to be present for the device to eventually deliver its record.
Rationale | Energy-aware fail-over bearer choice keeps data flowing under intermittent coverage without exhausting the budget; offline-first means loss of all bearers degrades to buffered capture, never data loss (D2/AD-1).
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-108 (multi-bearer selection), SYS-019 (multi-bearer offload), PRD-CONN-004, PRD-FUN-013; → SRS-CONN-007.
Notes | Selection policy is configurable; ranking and thresholds set in configuration (SRS-DEV-005).
SRS-CONN-007 — Store-and-forward on connectivity loss (resume, no loss, no duplication)
Requirement Statement | The connectivity software shall detect connectivity loss, cooperate with the on-device buffer to preserve samples/events/alerts, and on reconnection resume upload in intact chronological order with idempotent (exactly-once) ingest, such that no record is duplicated or lost across any number of interruptions.
Rationale | Intermittent links must never corrupt or lose the evidentiary record; ordering, resume, and de-duplication preserve ALCOA+ completeness and consistency (D2).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-111 (store-and-forward), SYS-020 (ordered/resumable/de-duplicated), SYS-018, PRD-CONN-008, PRD-FUN-012; → SRS-DEV-008 (buffer), SRS-CLOUD (idempotent ingest). Standard: ALCOA+ / EU GDP (revision to confirm).
Notes | Buffer sizing and ordering guarantees are provided by SRS-DEV; this specifies the resumable, idempotent transport.
SRS-CONN-008 — Alert-path prioritisation
Requirement Statement | When connectivity or energy is constrained, the connectivity software shall transmit excursion/shock alerts ahead of routine bulk telemetry over the selected bearer, subject to the energy budget, so that a rescue-enabling alert is never queued behind bulk data.
Rationale | Alert latency is connectivity-bound; the alert is captured offline the instant it is detected and must be delivered as soon as a bearer is available, ahead of backlog — a late alert is a recognised hazard.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-112 (alert-path prioritisation), SYS-011 (immediate alert), PRD-CONN-009, PRD-FUN-008; → SRS-DEV-016 (governor alert exception), HAL delayed-alert hazard (Stage 10).
Notes | Source classed Security/Safety; forward to HAL for the delayed-alert hazard reconciliation.
SRS-CONN-009 — Connection retry with bounded, energy-aware back-off
Requirement Statement | On a failed connection or transfer, the connectivity software shall retry using a bounded exponential back-off with a configured maximum interval and attempt ceiling per window, so that repeated failures cannot drain the energy budget through unbounded retries.
Rationale | Naïve immediate retry in poor coverage (e.g. inside the shipper) is a primary energy leak; back-off bounds the cost while still recovering when coverage returns.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-108, SYS-089 (energy governor), PRD-CONN-004, PRD-ELEC-008; → SRS-CONN-018.
Notes | Back-off parameters configurable; alerts retain priority within the ceiling (SRS-CONN-008).
SRS-CONN-010 — Cadence-driven scheduled upload
Requirement Statement | The connectivity software shall upload buffered telemetry on a configurable cadence (default hourly, range 15 min–24 h) within the energy-budgeted radio window, batching records to minimise per-upload connection overhead.
Rationale | Store-and-forward on a cadence, not continuous live upload, is what makes the 90-day budget achievable; batching amortises the fixed energy cost of establishing a session.
Priority | Must
Source | Performance
Verification Method | Test
Traceability | SYS-019 (scheduled cadence upload), PRD-FUN-013, PRD-CONN-004; → SRS-DEV-015 (windows).
Notes | Cadence default hourly (D2); the energy budget is the binding constraint.
7.3 Application transport & security
SRS-CONN-011 — MQTT / HTTPS application transport over mutually authenticated TLS
Requirement Statement | The connectivity software shall exchange telemetry, alerts, and commands with the cloud using MQTT v5 and/or HTTPS over mutually authenticated TLS (device and server both authenticated), and shall not transmit controlled data over an unauthenticated or unencrypted channel.
Rationale | The evidentiary data path must be confidential, integrity-protected, and mutually authenticated so neither the device nor the endpoint can be impersonated.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-110 (mutual authentication), SYS-020 (idempotent ingest), PRD-CONN-006, PRD-SEC-006; → SRS-SEC (TLS/key mechanism, EM-ENG-008), SRS-CLOUD. Standard: MQTT v5, TLS 1.2+/1.3 (revision to confirm).
Notes | Cipher suites, key hierarchy, and session policy owned by SRS-SEC; this mandates the transport and the mutual-auth gate.
SRS-CONN-012 — Authenticated pairing and provisioning over any bearer
Requirement Statement | The connectivity software shall permit configuration, provisioning, and controlled-data read only after mutual authentication of the requesting party over the bearer in use (BLE, Wi-Fi, or LTE-M), rejecting any configuration or controlled read from an unauthenticated party.
Rationale | Provisioning and pairing are attack surfaces; unauthorised configuration could disarm monitoring or corrupt a profile.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-110 (secure provisioning/pairing), PRD-CONN-006, PRD-SEC-005; → SRS-SEC, SRS-DEV-006/ 007/022. Standard: Bluetooth Core pairing / TLS (revision to confirm).
Notes | Login-free edge verdict read is preserved (AD-5); only configuration/controlled read is gated.
SRS-CONN-013 — Certificate and credential handling
Requirement Statement | The connectivity software shall store its device certificate/credentials in protected storage, validate the server certificate chain and validity on each secure session, reject an expired or untrusted certificate, and support in-field credential/certificate renewal without loss of buffered data.
Rationale | A trustworthy authenticated channel depends on valid, renewable credentials; an expired or spoofed certificate must break the session, not be accepted.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-110, PRD-SEC-006, PRD-CONN-006; → SRS-SEC (certificate lifecycle, EM-ENG-008), SRS-DEV-006.
Notes | Certificate lifecycle/rotation policy owned by SRS-SEC; this states the device connectivity behaviour and the no-data-loss constraint during renewal.
7.4 Diagnostics, quality & bandwidth
SRS-CONN-014 — Connectivity diagnostics and signal-quality logging
Requirement Statement | The connectivity software shall record, per connection attempt, the bearer used, signal quality/RSSI(-class), carrier/network identity, attempt and failure counts, session outcomes, and buffer-backlog depth, and shall make these diagnostics available to fleet tooling.
Rationale | Fleet operators must be able to diagnose silent-comms devices and coverage gaps; the signal-quality history also feeds the RF-in-container evidence base.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-113 (diagnostics & signal logging), SYS-104 (RF evidence), PRD-CONN-010; → SRS-DEV-012, SRS-DIAG.
Notes | Diagnostics are energy-cheap (SYS-113); they must not compete with alert delivery for the radio window.
SRS-CONN-015 — Roaming across carriers and regions
Requirement Statement | The connectivity software shall maintain wide-area connectivity across carriers and regions encountered on a trip via roaming-capable SIM/eSIM provisioning, automatically selecting an available permitted network without manual intervention, and shall log the network/carrier used.
Rationale | Global vaccine lanes cross carrier and national boundaries mid-trip; without automatic roaming the alert/upload path breaks at each border.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-116 (roaming), SYS-115 (regional radio config), PRD-CONN-001/012; → SRS-CONN-014, SRS-REG.
Notes | Per-market carrier coverage and the SIM/eSIM strategy are Evidence-Pending (SYS-116); requirement fixes the behaviour, not the carrier list.
SRS-CONN-016 — Bandwidth, latency, and payload management
Requirement Statement | The connectivity software shall compress and/or batch telemetry payloads and size transfers to the constrained LPWAN bandwidth so that a scheduled upload completes within its energy and latency window, and shall segment large transfers to remain resumable.
Rationale | LTE-M throughput is low and each on-air second costs energy; payload discipline is what keeps uploads inside both the latency window and the energy budget.
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-019, SYS-104, PRD-CONN-004, PRD-PERF-004; → SRS-CONN-007 (resume), SRS-DATA (serialization/compression).
Notes | Serialization/compression scheme detailed in SRS-DATA; this binds it to the connectivity window.
7.5 Time, connection quality & RF-in-container
SRS-CONN-017 — Network time synchronisation of the device clock
Requirement Statement | The connectivity software shall obtain a trusted network time reference whenever a link is available, deliver it to the device time service for RTC discipline, and record the synchronisation event and residual error with connectivity diagnostics.
Rationale | A defensible record and correct excursion timing require a trusted, UTC-referenced clock; the network is the re-discipline source for the free-running RTC (offline-first).
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-118 (network time sync), SYS-006 (trusted timestamp), PRD-FUN-019, PRD-ARCH-007; → SRS-DEV-021 (trusted clock/discipline). Standard: 21 CFR Part 11 attributable/accurate (revision to confirm).
Notes | RTC free-running and drift disclosure are owned by SRS-DEV-021; this provides the network reference.
SRS-CONN-018 — Connection-quality assessment for energy-aware decisions
Requirement Statement | Before committing a bulk upload, the connectivity software shall assess current connection quality on the candidate bearer and defer or down-scope a non-urgent bulk transfer when quality is too poor to complete within the energy window, while never deferring a queued alert on those grounds.
Rationale | Attempting a large upload over a marginal link (typical inside the shipper) wastes energy on a transfer that will not complete; quality-gating bulk transfers protects the budget without penalising alerts.
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-104, SYS-108, SYS-112, PRD-CONN-004/007; → SRS-CONN-008 (alert exemption), SRS-CONN-009.
Notes | Alerts are exempt from deferral (SRS-CONN-008); quality thresholds configurable and informed by the RF-in-container test.
SRS-CONN-019 — RF-in-container degradation handling (Evidence-Pending)
Requirement Statement | When operating inside a representative metallized/insulated shipper where the wide-area link cannot be closed, the connectivity software shall continue buffering all records locally, keep retrying within the energy-bounded back-off, and deliver all buffered data and queued alerts on the first successful link, such that RF failure inside the container defers but never loses the record or an alert.
Rationale | RF attenuation inside the shipper can exceed the fade margin of a low-power WAN radio; if egress fails, offline capture must still preserve the record while alerts are delivered as soon as egress is possible (D2).
Priority | Must
Source | Reliability
Verification Method | Test (RF-in-container, PRD-VER-005), Analysis (link-budget model)
Traceability | SYS-104 (RF link budget/delivery, Evidence-Pending), SYS-109 (antenna strategy), SYS-111, PRD-CONN-007, PRD-ENV-009; → SRS-CONN-007/009/014.
Notes | Evidence-Pending — the attenuation figure, achievable margin, and delivery-success target are unresolved and close on PRD-VER-005. No numeric RF value is asserted here (evidence). The external-antenna decision (SYS-109) is a live architectural branch.
SRS-CONN-020 — Regional radio compliance enforcement in software
Requirement Statement | The connectivity software shall operate the radios only within the bands, power, and regulatory profile configured for the device's region, and shall not initiate transmission outside the permitted band plan/limits for that configured region.
Rationale | Multi-market operation requires per-region radio conformance; transmitting outside the permitted band plan is a compliance and legal-operation failure.
Priority | Must
Source | Regulatory
Verification Method | Inspection, Test
Traceability | SYS-115 (regional radio configurability), PRD-CONN-012, PRD-REG-004; → SRS-REG, SRS-DEV-005 (region config). Standard: FCC Part 15 / CE RED; 3GPP band plans (revision to confirm).
Notes | Some regions may require per-market hardware variants (PRD-REG-004); flag Regulatory (EM-ENG-006). Region config integrity per SRS-DEV-005.
---
8. Non-functional Requirements
SRS-CONN-021 — Connectivity energy conformance
Requirement Statement | The connectivity software shall keep the cumulative energy consumed by all bearer activity (sessions, retries, transfers, roaming scans) within the radio/GNSS energy allocations of the documented energy budget over the worst-case 90-day cold-duty profile.
Rationale | Radios are the dominant load; connectivity behaviour is only acceptable if it lives inside its budget allocation, or 90-day endurance fails (D3).
Priority | Must
Source | Performance
Verification Method | Analysis, Test
Traceability | SYS-090 (energy-budget allocation), SYS-089, PRD-ELEC-008; → SRS-PERF.
Notes | Per-class radio/GNSS allocations are Evidence-Pending and close on the RF-in-container test (SYS-090); requirement binds the software to the allocation.
SRS-CONN-022 — Alert-delivery latency when a bearer is available
Requirement Statement | When a usable bearer is present at detection time, the connectivity software shall deliver an excursion/shock alert to the cloud within the alert-latency target defined in SRS-PERF.
Rationale | The value of an alert is time-bound; when a link exists, delivery must be prompt enough for a responder to act.
Priority | Must
Source | Performance
Verification Method | Test
Traceability | SYS-011, SYS-112, PRD-FUN-008, PRD-PERF-004; → SRS-CONN-008, SRS-PERF.
Notes | Numeric latency target held in SRS-PERF (evidence); offline detection is always immediate (SRS-DEV / SYS-013).
---
Error Handling
SRS-CONN-023 — Graceful, logged handling of transport failures
Requirement Statement | The connectivity software shall handle bearer establishment, TLS, and transfer failures without data loss — preserving buffered records, logging the failure cause to connectivity diagnostics, and returning to the retry/back-off governor — and shall never discard telemetry or an alert because a transport attempt failed.
Rationale | A failed send must degrade to buffered-and-retried, never to lost data; the failure cause must be visible for fleet triage.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-111, SYS-113, SYS-021, PRD-CONN-008/010; → SRS-CONN-007/009/014.
Notes | Ties to the no-silent-failure invariant in SRS-DEV-028.
---
Diagnostics
SRS-CONN-024 — Field connectivity self-test
Requirement Statement | The connectivity software shall provide an authenticated, on-demand connectivity self-test reporting reachable bearers, last signal quality per bearer, last successful session per bearer, certificate validity, and current backlog depth, for depot triage.
Rationale | Deploying or triaging a device requires confirming it can actually reach the cloud on at least one bearer before it goes back into service.
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness), SYS-113, PRD-FUN-020, PRD-CONN-010; → SRS-DEV-029, SRS-CONN-014.
Notes | Access authenticated per SRS-DEV-022 mode model.
---
Security Considerations
SRS-CONN-025 — Confidentiality and integrity of the connectivity path
Requirement Statement | The connectivity software shall protect all controlled telemetry, alerts, and configuration in transit for confidentiality and integrity via mutually authenticated TLS, and shall not expose an unauthenticated remote command, configuration, or debug interface over any bearer.
Rationale | The connectivity path is the primary remote attack surface; controlled data and control must be authenticated, encrypted, and integrity-protected end to end.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-110, PRD-SEC-006, PRD-CONN-006; → SRS-SEC (mechanisms, EM-ENG-008), SRS-CONN-011/012/ 013.
Notes | Cryptographic mechanisms owned by SRS-SEC (EM-ENG-008); this states the connectivity-software security posture.
---
Verification Requirements
SRS-CONN-026 — Connectivity verification under intermittent and RF-hostile conditions
Requirement Statement | The connectivity software shall be verified by bearer tests (LTE-M/Wi-Fi/BLE/ USB), reconnection/store-and-forward integrity tests, alert-priority and latency tests, retry/back-off energy tests, roaming tests, and the RF-in-container test (PRD-VER-005), with results traced to the requirements above.
Rationale | The offline-first, energy-bounded, RF-hostile claims are only credible if verified under the intermittent and attenuated conditions the device actually meets in service.
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-104, SYS-111, SYS-112, SYS-116, PRD-VER-005, PRD-CONN-007/008/009; → SRS-VER.
Notes | RF-in-container performance is Evidence-Pending; this fixes the verification obligation that closes it.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-105/106/107 (bearers), SYS-108/111/112 (selection, store-and-forward, alert priority), SYS-113 (diagnostics), SYS-115/116 (regional radio & roaming), SYS-118 (network time), SYS-104/109/117 (RF-in-container, Evidence-Pending); PRD-CONN-001…012 and PRD-FUN-008/012/013. Cross-module: cryptographic/TLS/certificate mechanisms → SRS-SEC (EM-ENG-008); on-device buffer/storage/clock → SRS-DEV; idempotent ingest and endpoints → SRS-CLOUD; OTA transport → SRS-OTA; radio conformance → SRS-REG / PRD-REG-004; numeric latency/energy budgets → SRS-PERF; delayed-alert hazard reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (RF attenuation and delivery-success target, per-market carrier coverage/SIM strategy, radio/GNSS energy allocations, enterprise-Wi-Fi auth) are carried as gaps, not invented, and close on their associated tests.
SRS-400 — Cloud Software (module SRS-CLOUD)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-CLOUD · Owner: Principal Cloud/Platform Engineer with the Product Security Engineer (EM-ENG-008) and the Quality/Regulatory Engineer (EM-ENG-006). Detailed cloud-platform software requirements decomposed from the System Requirements (SYS-011/012/019/020, SYS-026/027/028, SYS-030/031, SYS-119…131, SYS-132…148, SYS-174/175/176) and the Product Requirements (PRD-CLOUD 900, with PRD-FUN, PRD-SEC), realising the warn → decide → prove value chain on the platform side of the split trust boundary (AD-5) and the store-and-forward evidence model (D2) for a reusable fleet (D1). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM cloud platform registers and identifies devices; maintains each device's desired/reported state (shadow/twin); ingests buffered store-and-forward telemetry idempotently and integrity-preserving; processes commands, routes messages, evaluates rules and events, and drives alert notification and escalation; assembles, stores, searches, analyses, reports and exports the tamper-evident trip record (21 CFR Part 11-style); governs users, roles, federated identity, and audit; and delivers this as a horizontally scalable, observable, available, recoverable, multi-tenant service. It specifies platform software behaviour and obligations, not the specific cloud provider, product, or infrastructure choice.
2. Scope
Device registry and identity; device shadow/digital twin; fleet-management platform services (device state, grouping — deep fleet workflow in SRS-FLEET); store-and-forward telemetry ingestion; command processing and message routing; rules engine, event processing, and the alert notification/escalation engine; record assembly, durable storage, analytics, search, reporting and export; user management, role-based access control, identity federation/SSO, and audit logging; API gateway and secured integration APIs; microservice/container deployment, horizontal scaling, availability, backup, disaster recovery, multi-region operation, and monitoring/observability. Excludes the device-side capture/verdict/buffer (SRS-DEV), the connectivity transport that delivers to the platform (SRS-CONN), cryptographic mechanisms and key lifecycle (SRS-SEC, EM-ENG-008), OTA packaging/campaign internals (SRS-OTA — the platform orchestration surface is referenced), the mobile app (SRS-MOB), data schema/retention detail (SRS-DATA), and numeric performance targets (SRS-PERF); each is referenced where the cloud software interfaces with it.
3. Applicable Standards
21 CFR Part 11 / ALCOA+ (electronic records, signatures, audit trail); EU GDP (data governance, data-integrity expectations); GAMP 5 (computerised-system validation, CSV); ISO/IEC 27001 concepts and SOC 2 Trust Services Criteria concepts (platform security, availability, confidentiality); IEC 62443 concepts (system security, secure device onboarding); OWASP ASVS (application security verification); OAuth 2.0 / OpenID Connect and SAML 2.0 (identity federation/SSO); TLS 1.2+/1.3 (transport). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• Devices deliver ordered, resumable, de-duplicable buffered payloads over mutually authenticated TLS via the connectivity layer (SRS-CONN-007/011), each carrying a versioned schema and integrity artefact (SYS-119, SYS-132); the platform is the point at which the record becomes the authoritative evidence.
• Cryptographic tamper-evidence is applied at capture on the device and must be verified and preserved by the platform, not created by it (SYS-132); key/credential lifecycle and PKI are owned by SRS-SEC (SYS-140).
• Concrete record export formats (e.g. PDF + CSV/JSON), retention periods per market/customer, availability SLA, and RPO/RTO targets are Evidence-Pending (SYS-124/126/128, SYS-175, PRD-CLOUD-005/011/012) and are not invented here; requirements fix the behaviour, not the unresolved number.
• The platform enforces the governed side of the split trust boundary (AD-5): all configuration, data access, export, administration and record modification are authenticated, role-authorised and audited, while the login-free on-device edge read (SYS-147/148) is entirely device-side and is never gated by the platform.
• Multi-tenancy: each customer owns and is isolated to its own records and configuration (SYS-125).
5. Architecture Overview
The platform is a set of independently deployable microservices on a container/orchestration substrate behind an API gateway that terminates mutually authenticated TLS, authenticates callers, and enforces rate limits. A device registry holds durable per-device identity, configuration, calibration, and firmware state and admits data only from registered, authenticated devices; a device shadow/twin service reconciles desired vs reported state. An ingestion pipeline accepts store-and-forward payloads, validates and de-duplicates them idempotently, verifies and preserves tamper-evidence, and writes to a durable, multi-tenant record store; a rules/event engine evaluates ingested telemetry and device-fault events and feeds a notification/escalation engine; a command service issues authenticated commands to devices via routing. A record-assembly service composes complete, attributable trip records for search, analytics, reporting, and Part 11-style export. An identity/access layer (user management, RBAC, OIDC/SAML federation) and an append-only audit service govern all human access. Cross-cutting observability (metrics, logs, traces, health) and backup/DR/multi-region infrastructure make the service scalable, available, and recoverable. Every governed action is authenticated, authorised, and audited.
---
7. Functional Requirements
7.1 Device registry, identity & shadow
SRS-CLOUD-001 — Authoritative device registry and identity
Requirement Statement | The cloud platform shall maintain an authoritative registry entry for every device recording its unique cryptographic identity, durable serial/asset identity, current configuration and active profile version, calibration status and validity date, firmware version, ownership/tenant, and lifecycle state, and shall accept telemetry, events, and audit data only from a registered device whose identity is authenticated.
Rationale | Attribution, tenant isolation, and fleet control all depend on an authoritative device identity; an unregistered or unauthenticated source must never be able to inject data into the evidentiary record.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-133 (attribution), SYS-134 (device identity), SYS-121 (provenance), SYS-030 (readiness), PRD-CLOUD-001, PRD-SEC-006; → SRS-SEC (identity/keys, EM-ENG-008), SRS-FLEET-001. Standard: IEC 62443 secure onboarding concepts; 21 CFR Part 11 attributable (revision to confirm).
Notes | Device durable identity ties PRD-MECH-010; registry enrolment/provisioning shared with SRS-SEC (SYS-141) and SRS-FLEET.
SRS-CLOUD-002 — Device shadow / digital twin state reconciliation
Requirement Statement | The cloud platform shall maintain, per device, a shadow/twin comprising the last-reported device state (firmware, profile version, readiness, connectivity health, last-seen) and the desired state (target profile, target firmware, pending commands), and shall reconcile the two so that a desired change is applied and confirmed when the device next connects, without loss or duplication of the desired intent across intermittent connectivity.
Rationale | Store-and-forward devices are frequently offline (D2); a twin lets operators author intent and read last-known state asynchronously, and guarantees a pending change is neither lost nor applied twice when a duty-cycled device reappears.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-019 (cadence upload), SYS-030 (readiness), SYS-176 (device↔cloud reconciliation), PRD-CLOUD-001/006, PRD-FUN-020; → SRS-CLOUD-006 (commands), SRS-FLEET-005.
Notes | Desired-state changes are audited (SRS-CLOUD-020) and role-gated (SRS-CLOUD-018).
7.2 Ingestion, command & routing
SRS-CLOUD-003 — Store-and-forward telemetry ingestion, ordered and idempotent
Requirement Statement | The cloud platform shall ingest buffered device telemetry, events, excursions, gap markers, and audit records in intact chronological order, shall store each record exactly once (idempotent de-duplication keyed on device identity and record identity/sequence), and shall correctly reassemble a record set delivered across any number of interrupted, resumed, or out-of-session uploads without loss or duplication.
Rationale | Opportunistic multi-bearer offload (D2) can deliver the same buffered records more than once and across many partial sessions; exactly-once, order-preserving ingest is what keeps the reassembled cloud record complete and consistent (ALCOA+).
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-020 (ordered/resumable/de-duplicated), SYS-011 (alert), SYS-176, SYS-122 (serialisation), PRD-CLOUD-002, PRD-CONN-008; → SRS-CONN-007, SRS-DATA. Standard: ALCOA+ enduring/consistent, EU GDP (revision to confirm).
Notes | Idempotency key and sequencing scheme detailed in SRS-DATA; this binds the platform behaviour.
SRS-CLOUD-004 — Integrity-preserving ingest with validation, quality and consistency checks
Requirement Statement | On ingest the cloud platform shall verify each record's cryptographic tamper-evidence and attribution, validate it against its declared schema version, check chronological consistency and de-duplication, and shall flag — never silently drop, repair, or overwrite — any record that is malformed, out-of-schema, unattributable, integrity-failing, duplicated, or inconsistent, recording the anomaly in the trip record and the audit trail.
Rationale | Integrity preservation across an intermittent link requires the ingest boundary to detect corruption and gaps rather than paper over them; silent repair would fabricate evidence, breaching ALCOA+ original/accurate.
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-129 (ingest validation), SYS-132 (tamper-evidence), SYS-002 (gaps), SYS-007 (quality flags), PRD-CLOUD-002, PRD-SEC-001; → SRS-CLOUD-011, SRS-SEC. Standard: 21 CFR Part 11 / ALCOA+ (revision to confirm).
Notes | Flags, not fixes (SYS-129). Integrity mechanism verified here; created on device (SRS-SEC).
SRS-CLOUD-005 — Command processing and confirmed delivery to devices
Requirement Statement | The cloud platform shall accept authenticated, role-authorised commands (e.g. profile update, firmware-update target, reboot, diagnostic/log-pull, decommission), queue each command against the device shadow, deliver it to the authenticated device over mutually authenticated transport, and record command issuance, delivery, and device acknowledgement/result in the audit trail, applying a command at most once.
Rationale | Remote control of a duty-cycled fleet must be authorised, attributable, and confirmed; an unconfirmed or duplicated command (e.g. a double reboot or double profile change) could disrupt a live trip or corrupt configuration.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-022 (profile config), SYS-032/139 (firmware), SYS-028 (audit), SYS-176, PRD-CLOUD-008, PRD-FUN-014; → SRS-CLOUD-002/020, SRS-FLEET-004/006, SRS-CONN-012. Standard: IEC 62443 concepts (revision to confirm).
Notes | Command authenticity/keys owned by SRS-SEC; unauthenticated remote command is prohibited (SYS-137, SRS-CONN-025).
SRS-CLOUD-006 — Message routing across platform services
Requirement Statement | The cloud platform shall route each ingested message to the applicable downstream services (record store, rules/event engine, shadow, analytics) based on device identity, tenant, message type, and content, without cross-tenant leakage and without loss of a message under normal and surge load.
Rationale | A decoupled ingest→process→store→notify path requires reliable internal routing; a dropped or mis-routed message would either lose evidence or leak it across a tenant boundary.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-125 (tenant isolation), SYS-129, SYS-020, PRD-CLOUD-002/009; → SRS-CLOUD-008/011, SRS-CLOUD-023.
Notes | Delivery guarantees compose with idempotent ingest (SRS-CLOUD-003) to give end-to-end exactly-once semantics.
7.3 Rules engine, event processing & alerting
SRS-CLOUD-007 — Configurable rules engine for alert and event policy
Requirement Statement | The cloud platform shall provide a rules engine allowing authorised users to configure alert and event rules (excursion and device-fault conditions, thresholds, recipients, channels, escalation timing) per product, shipment, tenant, and role within governance controls, and shall version and audit every rule change.
Rationale | Alerts must match product stability and each organisation's response process; a governed, versioned, audited rule set makes alerting adaptable without becoming an untraceable configuration surface.
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-012 (routing/escalation), SYS-014 (alarm balance), PRD-CLOUD-004, PRD-FUN-009, PRD-SEC-005; → SRS-CLOUD-010/018/020.
Notes | Alarm-fatigue balance (hysteresis/consolidation) reconciles with HAL delayed/missed-alert hazard.
SRS-CLOUD-008 — Event processing of telemetry and device-fault streams
Requirement Statement | The cloud platform shall process ingested telemetry and device-status streams to detect excursion events, device-fault/readiness events, and connectivity/liveness events, evaluate them against the active rules, and generate the corresponding alert and dashboard events with attribution and trusted time, without dropping a safety-relevant event under load.
Rationale | Turning raw ingested data into actionable events (excursion, fault, overdue) is the platform's core detect→warn function; a lost excursion or fault event is a lost intervention.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-011 (excursion alert), SYS-016 (verdict consistency), SYS-031 (stray/overdue), PRD-CLOUD-003/006, PRD-FUN-008; → SRS-CLOUD-010, SRS-FLEET-011.
Notes | Cloud excursion evaluation must be consistent with the on-device verdict (SYS-016); divergence is detected and flagged.
SRS-CLOUD-009 — Alert notification and escalation engine
Requirement Statement | The cloud platform shall deliver each generated alert to the responder(s) configured for the device's trip/owner over the configured channels (e.g. push, email, SMS/webhook), shall track acknowledgement, and shall escalate an unacknowledged alert to the next responder/tier per the configurable escalation policy within its defined time, recording every notification, acknowledgement, and escalation with time and actor.
Rationale | An alert that reaches no accountable human is not a risk control; acknowledgement tracking and time-bounded escalation ensure a warning is actually received and acted on.
Priority | Must
Source | User Need
Verification Method | Test
Traceability | SYS-012 (routing/escalation), SYS-011, PRD-CLOUD-003, PRD-MOB-005, PRD-FUN-009; → SRS-MOB-008 (push), HAL delayed/unacknowledged-alert hazard (Stage 10). Standard: alarm-management escalation principles (revision to confirm).
Notes | Safety-classed; reconcile with HAL. Delivery latency target held in SRS-PERF.
7.4 Record, storage, analytics, search, reporting & export
SRS-CLOUD-010 — Trip record assembly (Part 11-style)
Requirement Statement | The cloud platform shall assemble, per trip, a complete record comprising all samples, events, excursions, gap markers, quality flags, the active profile version, provenance metadata (device, firmware, calibration, actor), the trusted time base and sync/drift disclosure, and the device/cloud audit trail, each attributable to its source and time and protected by verified tamper-evidence.
Rationale | The assembled, attributable, tamper-evident trip record is the product's compliance value; it must reconstruct exactly what was measured, under what configuration, on what device, in what state.
Priority | Must
Source | Regulatory
Verification Method | Test, Analysis
Traceability | SYS-026 (complete record), SYS-121 (provenance), SYS-131 (time base), SYS-028/135 (audit), PRD-CLOUD-005, PRD-FUN-017; → SRS-CLOUD-004/012/020, SRS-DATA. Standard: 21 CFR Part 11 records, EU GDP, ALCOA+ (revision to confirm).
Notes | Profile version binding per SYS-120; a verdict is meaningful only against the profile that produced it.
SRS-CLOUD-011 — On-demand record export in an admissible, integrity-preserving format
Requirement Statement | The cloud platform shall export a complete trip record on demand in documented, human- and machine-readable formats that preserve samples, events, gap markers, quality flags, provenance, time base, and tamper-evidence, and shall record each export as an attributable, time-stamped audit event.
Rationale | Evidence must move to where release/acceptance decisions and disputes live (ERP/QMS/partner systems) without losing its ALCOA+ properties; every export is itself a governed record action.
Priority | Must
Source | Regulatory
Verification Method | Test, Demonstration
Traceability | SYS-027 (export), SYS-126 (import/integration format), PRD-CLOUD-005/010, PRD-REG-003; → SRS-CLOUD-020/021, SRS-DATA. Standard: 21 CFR Part 11 records, ALCOA+ (revision to confirm).
Notes | Concrete formats (e.g. PDF + CSV/JSON) are Evidence-Pending (SYS-126); the integrity obligation is fixed, the format list is not invented here.
SRS-CLOUD-012 — Durable multi-tenant record storage with retention and isolation
Requirement Statement | The cloud platform shall store all trip records, events, and audit data in durable storage owned by and isolated to the originating tenant, shall retain them for the configured retention period, and shall not delete, expire, or alter a retained record except through an audited, authorised action, recording the applicable retention policy with the data.
Rationale | Regulated customers require ownership, provable isolation, and controlled retention of their cold-chain evidence; cross-tenant leakage or silent deletion would be both a security breach and a data-integrity failure.
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-124 (retention), SYS-125 (ownership/isolation), SYS-130 (historical), PRD-CLOUD-011, PRD-SEC-005/007; → SRS-CLOUD-018/020, SRS-SEC (at-rest encryption). Standard: EU GDP data governance, ISO/IEC 27001 concepts (revision to confirm).
Notes | Retention periods per market/customer are Evidence-Pending (SYS-124); at-rest encryption keys per SRS-SEC (SYS-138/140).
SRS-CLOUD-013 — Search and query over records and fleet
Requirement Statement | The cloud platform shall allow authorised users to search and query records and device state by device, trip, tenant, time range, excursion/event, verdict, and readiness/lifecycle state, returning only data within the caller's authorised tenant and role scope.
Rationale | Retained records are only useful if they remain findable for release, audit, trend, and dispute; scoped query enforces isolation while making evidence retrievable years later.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-130 (queryable historical), SYS-125 (isolation), SYS-031 (fleet visibility), PRD-CLOUD-005/006; → SRS-CLOUD-012/018, SRS-FLEET-009.
Notes | Query results are read-only against governed state (AD-5); no query path may modify a record.
SRS-CLOUD-014 — Analytics and reporting over the fleet and record base
Requirement Statement | The cloud platform shall compute and present analytics and reports (e.g. excursion rates, alert response times, fleet readiness and utilisation, trip outcomes) scoped to the caller's tenant and role, derived from the authoritative records without modifying them.
Rationale | Operational and compliance oversight (trend, TCO, response performance) needs aggregate views of the reusable fleet (D1) and the evidence base, without ever mutating the source of truth.
Priority | Should
Source | User Need
Verification Method | Demonstration
Traceability | SYS-031 (fleet status), SYS-130, PRD-CLOUD-006/010, PRD-FUN-021; → SRS-FLEET-010, SRS-CLOUD-013.
Notes | Analytics are derived/read-only; the authoritative record is immutable except via audited action.
7.5 Identity, access, audit & APIs
SRS-CLOUD-015 — User management
Requirement Statement | The cloud platform shall provide user lifecycle management (create, modify, disable, remove) with unique user identity per tenant, and shall bind each governed action to an authenticated user, recording user-lifecycle changes in the audit trail.
Rationale | Attribution of human actions (Part 11) requires uniquely identified, managed user accounts; orphaned or shared accounts destroy non-repudiation.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-133 (attribution), SYS-136 (access control), SYS-146 (sessions), PRD-CLOUD-009, PRD-SEC-005; → SRS-CLOUD-016/017/020. Standard: 21 CFR Part 11 attributable (revision to confirm).
Notes | Session/authentication mechanisms per SRS-SEC (SYS-146).
SRS-CLOUD-016 — Least-privilege role-based access control
Requirement Statement | The cloud platform shall enforce least-privilege role-based access control on every configuration, data-access, export, and administrative function, denying by default and permitting only actions granted to the caller's role within the caller's tenant, and shall never gate the login-free on-device edge read.
Rationale | Only authorised roles may configure devices or read/modify governed data; RBAC on the platform side of the trust boundary (AD-5) is what keeps the frictionless edge read from becoming an ungoverned data path.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (RBAC), SYS-147 (split trust boundary), SYS-125 (isolation), PRD-CLOUD-009, PRD-SEC-005; → SRS-MOB-011, SRS-CLOUD-021. Standard: OWASP ASVS access control, 21 CFR Part 11 access limitation (revision to confirm).
Notes | Edge read is device-side and out of scope of platform RBAC (SYS-147/148).
SRS-CLOUD-017 — Identity federation and enterprise single sign-on
Requirement Statement | The cloud platform shall support federation to an enterprise identity provider via standard single sign-on (OIDC and/or SAML 2.0), mapping federated identities and group claims to platform roles, so that a customer can govern access from its own IdP without a parallel credential store.
Rationale | Enterprise/regulated deployment is gated by IT security and CSV; SSO avoids a duplicate credential store and lets the customer's IdP enforce its own password, MFA, and de-provisioning policy.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (SSO federation), PRD-CLOUD-009, PRD-SEC-005, PRD-MOB-007; → SRS-CLOUD-016, SRS-MOB-001. Standard: OAuth 2.0 / OpenID Connect, SAML 2.0 (revision to confirm).
Notes | Token/session protection per SRS-SEC (SYS-146); role mapping is audited (SRS-CLOUD-020).
SRS-CLOUD-018 — Append-only, non-editable audit trail
Requirement Statement | The cloud platform shall maintain a secure, append-only, time-stamped, attributable audit trail of all configuration changes, authentication/access events, data actions (view/export/modify), command issuance, and administrative actions, such that no entry can be altered or deleted without detection, and shall reconcile the cloud audit trail with the uploaded device audit trail.
Rationale | The governance of the record must itself be inspectable and defensible (Part 11); a cloud-only trail misses offline device actions, so the two are reconciled after upload.
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-028 (audit trail), SYS-135 (append-only device+cloud), SYS-142 (security events), PRD-CLOUD-007, PRD-SEC-004; → SRS-CLOUD-010, SRS-SEC. Standard: 21 CFR Part 11 audit trail, SOC 2 concepts (revision to confirm).
Notes | Retained per policy (SRS-CLOUD-012); security-event logging detailed in SRS-CLOUD-024/SRS-SEC.
SRS-CLOUD-019 — Secured API gateway and integration APIs
Requirement Statement | The cloud platform shall expose all external access through an API gateway that terminates mutually authenticated TLS, authenticates and authorises every request (device identity for device APIs; user/role or service credential for integration APIs), enforces per-tenant rate limiting and input validation, and shall expose secured APIs enabling authorised enterprise/OEM systems to retrieve records, alerts, and fleet status within their tenant scope.
Rationale | The API surface is the primary remote attack surface and the integration path into ERP/QMS/partner systems; every request must be authenticated, authorised, tenant-scoped, and rate-limited.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-137 (mutual TLS), SYS-136 (RBAC), SYS-125 (isolation), PRD-CLOUD-010, PRD-SEC-006; → SRS-CLOUD-016/018, SRS-SEC, SRS-CONN-011. Standard: OWASP ASVS, OAuth 2.0, TLS 1.2+/1.3 (revision to confirm).
Notes | Cryptographic mechanisms and token issuance owned by SRS-SEC (EM-ENG-008).
---
8. Non-functional Requirements
SRS-CLOUD-020 — Microservice/container deployment and horizontal scalability
Requirement Statement | The cloud platform shall be deployed as independently deployable, horizontally scalable microservices on a container orchestration substrate, such that ingest, processing, alerting, query, and API capacity can be scaled out to meet fleet growth and surge load without downtime and without loss or reordering of in-flight records.
Rationale | A growing reusable fleet (D1) and bursty store-and-forward offload (many devices draining backlog at a depot) require elastic capacity; scaling must never compromise the exactly-once, ordered ingest guarantee.
Priority | Should
Source | Performance
Verification Method | Test, Analysis
Traceability | SYS-175 (platform reliability), SYS-020, PRD-CLOUD-012; → SRS-CLOUD-003/006, SRS-PERF. Standard: SOC 2 availability concepts (revision to confirm).
Notes | Numeric throughput/latency targets held in SRS-PERF (evidence).
SRS-CLOUD-021 — Platform availability for ingest and alerting
Requirement Statement | The cloud platform shall meet its defined availability target for the ingest and alerting paths and shall not lose or silently corrupt any ingested record under any single-component failure, degrading gracefully (e.g. accepting and queuing ingest) rather than rejecting or dropping data during partial outages.
Rationale | Alerts and records depend on a durable, available platform; a single-component failure must neither drop an alert nor lose an ingested record.
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-175 (availability/durability), SYS-176 (end-to-end delivery), PRD-CLOUD-012, PRD-REL-012; → SRS-CLOUD-022, SRS-PERF. Standard: SOC 2 availability concepts (revision to confirm).
Notes | Availability SLA is Evidence-Pending (SYS-175); the no-loss/no-silent-corruption obligation is fixed here.
SRS-CLOUD-022 — Backup, disaster recovery, and multi-region durability (RPO/RTO)
Requirement Statement | The cloud platform shall maintain durable, geographically redundant backups of all ingested records, audit trails, and configuration, and shall be able to restore them to a consistent, integrity-verified, attribution-preserving state within defined recovery-point (RPO) and recovery-time (RTO) objectives following a component or site failure, without silent loss or corruption, and shall support multi-region operation where data-residency obligations require it.
Rationale | The cloud is where the record becomes authoritative evidence; a site failure must not lose or corrupt it, and restore must re-establish tamper-evidence and attribution, not merely availability.
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-128 (backup/restore/DR), SYS-175 (RPO/RTO), SYS-148 (residency/privacy), PRD-CLOUD-012, PRD-REL-006/012; → SRS-CLOUD-012, SRS-DATA. Standard: SOC 2 availability, EU GDP residency, ISO/IEC 27001 concepts (revision to confirm).
Notes | RPO/RTO targets and residency mapping are Evidence-Pending (SYS-128/175, D4); the recovery obligation is fixed, the numbers are not invented here.
---
Error Handling
SRS-CLOUD-023 — Graceful, logged handling of ingest and processing failures
Requirement Statement | On any ingest, validation, routing, storage, or processing failure the cloud platform shall preserve the affected data without loss (e.g. dead-letter/quarantine with retry), record the failure cause, and surface it to operators, and shall never silently discard telemetry, an event, an alert, or an audit record because a processing step failed.
Rationale | A failed processing step must degrade to preserved-and-retried, never to lost evidence; the failure cause must be visible for platform operations and for the no-silent-failure invariant.
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-129 (flag, never drop), SYS-021 (never discard), SYS-176, PRD-CLOUD-002/012; → SRS-CLOUD-003/004/024.
Notes | Quarantined records are themselves attributable and auditable (SRS-CLOUD-018).
---
Diagnostics
SRS-CLOUD-024 — Monitoring, logging, and observability
Requirement Statement | The cloud platform shall emit metrics, structured logs, and distributed traces for the ingest, processing, alerting, storage, and API paths, shall monitor service health and security-relevant events, and shall alert operators on degradation, backlog growth, integrity/signature failures, and availability breaches, retaining operational and security event records per policy.
Rationale | A connected, OTA-updatable, multi-tenant fleet platform can only be operated safely if its own health, backlog, and security posture are observable; undetected degradation or compromise silently corrupts the evidentiary product.
Priority | Should
Source | Reliability
Verification Method | Demonstration, Test
Traceability | SYS-142 (security monitoring/incident logging), SYS-175, SYS-031 (liveness), PRD-CLOUD-007/012, PRD-SEC-011; → SRS-CLOUD-021/023, SRS-SEC. Standard: NIST CSF detect/respond, SOC 2 concepts (revision to confirm).
Notes | Incident-response runbook is Evidence-Pending (SYS-142); security-event detection shared with SRS-SEC (EM-ENG-008).
---
Security Considerations
SRS-CLOUD-025 — Platform security posture and split-trust enforcement
Requirement Statement | The cloud platform shall authenticate and authorise every governed request, isolate every tenant's data, encrypt controlled data at rest and in transit, protect authenticated sessions and tokens against theft/fixation/replay, and shall enforce that all configuration, data access, export, administration, and record modification occur only through authenticated, role-authorised, audited paths — while never imposing authentication on, or accepting modification from, the login-free on-device edge read.
Rationale | The platform is the governed side of the split trust boundary (AD-5); its security posture is what makes the record trustworthy and controlled without defeating the frictionless offline edge read.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-147 (split trust boundary), SYS-136/137/138/146 (RBAC/TLS/at-rest/session), SYS-125 (isolation), PRD-CLOUD-009, PRD-SEC-005/006/007; → SRS-SEC (mechanisms, EM-ENG-008), SRS-CLOUD-016/018/019. Standard: OWASP ASVS, IEC 62443 concepts, SOC 2 concepts, TLS 1.2+/1.3 (revision to confirm).
Notes | Cryptographic mechanisms and key lifecycle owned by SRS-SEC; this states the platform-software posture. Formal threat model is Evidence-Pending (SYS-143).
---
Verification Requirements
SRS-CLOUD-026 — Cloud verification for integrity, availability, access, and recovery
Requirement Statement | The cloud platform shall be verified by ingest-integrity and idempotency tests, order/resume/de-duplication tests, ingest-validation (flag-not-fix) tests, alert routing/escalation and latency tests, record-assembly and export-integrity tests, RBAC/SSO and tenant-isolation tests, audit-trail-immutability tests, availability/scaling tests, and backup/restore/DR (RPO/RTO) tests, with results traced to the requirements above and to CSV/validation review (GAMP 5) for regulated customers.
Rationale | The store-and-forward integrity, split-trust governance, availability, and recoverability claims are only credible if verified under the intermittent, multi-tenant, surge, and failure conditions the platform actually meets in service.
Priority | Must
Source | Regulatory
Verification Method | Test, Analysis, Demonstration
Traceability | SYS-020, SYS-125, SYS-128, SYS-129, SYS-136, SYS-175/176, PRD-CLOUD-002/005/009/012; → SRS-VER. Standard: GAMP 5 CSV, 21 CFR Part 11, SOC 2 concepts (revision to confirm).
Notes | Evidence-Pending items (export formats, retention periods, availability SLA, RPO/RTO) are carried as gaps and close on their associated tests/validation.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-011/012 (alert/escalation), SYS-019/020/176 (store-and-forward, ordered/idempotent, end-to-end delivery), SYS-026/027/028 (record/export/ audit), SYS-030/031 (readiness, fleet visibility), SYS-119…131 (data/schema/retention/time/DR), SYS-132…148 (integrity, identity, RBAC/SSO, session, split trust boundary), SYS-175 (platform reliability); PRD-CLOUD-001…012, PRD-SEC-005/006/007, PRD-FUN-008/009/017/020, PRD-REG-003. Cross-module: cryptographic/ TLS/key/at-rest mechanisms → SRS-SEC (EM-ENG-008); device capture/verdict/buffer → SRS-DEV; transport that delivers to the platform → SRS-CONN; OTA campaign/packaging → SRS-OTA; deep fleet workflows → SRS-FLEET; mobile app → SRS-MOB; schema/serialisation/retention detail → SRS-DATA; numeric availability/latency/RPO/RTO → SRS-PERF; safety-relevant alert/escalation reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items (concrete export formats, per-market/customer retention periods, availability SLA, RPO/RTO targets, residency mapping, formal threat model, IR runbook) are carried as gaps, not invented, and close on their associated tests and CSV/validation activities.
SRS-500 — Mobile Application (module SRS-MOB)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-MOB · Owner: Principal Mobile/App Engineer with Human-Factors, the Product Security Engineer (EM-ENG-008), and EM-ENG-001. Detailed mobile-application software requirements decomposed from the System Requirements (SYS-022…025, SYS-030, SYS-136/146/147, SYS-006/016, SYS-032) and the Product Requirements (PRD-MOB 1000, with PRD-CLOUD 900, PRD-SEC, PRD-UX), realising the provisioning/configuration/notification aid on the platform-governed side of the split trust boundary (AD-5) — the app augments, never gates, the login-free on-device verdict (C9). Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM mobile application authenticates and authorises its users; pairs with a device over authenticated BLE; onboards/provisions and re-provisions devices; authors and applies trip profiles with error-prevention; assists arming with positive confirmation without becoming an arming dependency; reads local device status/history/diagnostics; triggers governed firmware updates; presents push alert notifications; operates in the background and offline with later synchronisation; enforces role-based permissions and OS-level permissions; and meets accessibility and localisation expectations across markets. It specifies app behaviour and obligations; it is an aid for dispatch, service, and fleet users and is never a dependency for the safety-critical edge use.
2. Scope
User authentication/authorisation (via enterprise SSO/RBAC federation); authenticated BLE pairing; device onboarding/provisioning and field re-provisioning; trip-profile configuration with validation/poka-yoke; arming aid with confirmation; local device read over BLE; firmware-update trigger; push alert notifications; background execution; offline mode and synchronisation with conflict handling; role-based permissions and OS permissions; device sharing; accessibility; localisation; and crash reporting/telemetry. Excludes the device-side capture/verdict/buffer/arming fallback (SRS-DEV), the login-free on-device verdict itself (SRS-DEV / SYS-015/147), cloud services (SRS-CLOUD), the BLE/transport stack on the device (SRS-CONN), cryptographic mechanisms and identity keys (SRS-SEC, EM-ENG-008), and OTA packaging (SRS-OTA); each is referenced where the app interfaces with it.
3. Applicable Standards
IEC 62366-1 (usability engineering, use-error prevention); platform accessibility guidelines (iOS/Android, e.g. WCAG-aligned platform APIs); OAuth 2.0 / OpenID Connect and SAML 2.0 (federated authentication/SSO); OWASP ASVS and OWASP MASVS (application/mobile application security verification); Bluetooth Core Specification (BLE pairing/security); TLS 1.2+/1.3 (app-to-cloud transport). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The login-free on-device accept/reject/quarantine verdict (SYS-015/147) is entirely device-side and is never provided, gated, or required by the app; the app is optional to the safety-critical edge read.
• The device performs its own on-device arming fallback (PRD-FUN-015 / SYS-023); the app assists arming but is not required for a device to be armed.
• Enterprise identity, RBAC roles, and SSO are provided by the cloud platform (SRS-CLOUD-016/017); the app authenticates against and enforces those roles rather than maintaining a parallel credential store.
• Cryptographic pairing, device identity, and app-to-cloud TLS mechanisms are owned by SRS-SEC (SYS-134/137) and the connectivity/BLE stack by SRS-CONN; the app consumes them.
• The target language set and the concrete accessibility conformance level are Evidence-Pending (PRD-MOB-008, D4) and are not invented here.
5. Architecture Overview
The app is a role-aware client with three cooperating surfaces: an identity layer that authenticates the user through enterprise SSO (OIDC/SAML) and enforces the resulting role-based permissions on every governed action; a device-proximity layer that pairs over authenticated BLE and performs onboarding/provisioning, profile configuration, arming assistance, local read, and firmware-update trigger against a nearby device; and a cloud-sync layer that exchanges configuration, records, notifications, and audit with the platform over TLS and reconciles state when connectivity returns. A local store lets the app provision and read devices while the phone is offline and queues changes for later sync with audited conflict resolution. Push notifications and background execution keep alert delivery and sync current without the app in the foreground. Throughout, the app treats the device edge read as out-of-band and never substitutes for it.
---
7. Functional Requirements
7.1 Identity & access
SRS-MOB-001 — User authentication via enterprise SSO
Requirement Statement | The mobile app shall authenticate each user against the enterprise identity provider via standard single sign-on (OIDC and/or SAML 2.0) before permitting any governed action, shall protect authentication tokens against theft/replay in platform-provided secure storage, and shall re-authenticate the user when the session expires or before a sensitive action.
Rationale | App actions on configuration and data are governed (Part 11 access); authenticating through the customer's IdP avoids a parallel credential store and lets the customer enforce its own MFA and de-provisioning.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (SSO/RBAC), SYS-146 (session management), PRD-MOB-007, PRD-CLOUD-009, PRD-SEC-005; → SRS-CLOUD-017, SRS-SEC. Standard: OAuth 2.0 / OpenID Connect, SAML 2.0, OWASP MASVS (revision to confirm).
Notes | Does not gate the login-free on-device edge read (AD-5). Token/session mechanisms per SRS-SEC.
SRS-MOB-002 — Role-based permission enforcement
Requirement Statement | The mobile app shall enforce the authenticated user's role-based permissions (via SSO/RBAC) on every configuration, arming, firmware, data-read, export, and device-sharing action, denying by default and permitting only actions granted to the user's role and tenant scope.
Rationale | App actions must obey the same enterprise governance as the platform; a client that ignored RBAC would be an ungoverned back door to configuration and data.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (least-privilege RBAC), SYS-147 (split trust boundary), PRD-MOB-007, PRD-CLOUD-009, PRD-SEC-005; → SRS-CLOUD-016, SRS-MOB-001. Standard: OWASP ASVS access control (revision to confirm).
Notes | Consistent with cloud RBAC (SRS-CLOUD-016); enforcement is defence-in-depth, authoritative check remains server-side.
7.2 Pairing, provisioning & sharing
SRS-MOB-003 — Secure authenticated BLE pairing
Requirement Statement | The mobile app shall pair with a device over authenticated BLE such that only an authorised user can configure the device or read controlled data, shall reject an unauthenticated or spoofed device, and shall establish the paired channel with confidentiality and integrity protection.
Rationale | Provisioning and controlled read are attack surfaces; unauthorised pairing could disarm monitoring or corrupt a profile, so pairing must be authenticated without breaking edge simplicity.
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-110-line / SYS-133 (authenticated actor), SYS-147, PRD-MOB-001, PRD-CONN-003/006, PRD-SEC-005; → SRS-CONN-004/012, SRS-SEC. Standard: Bluetooth Core pairing, OWASP MASVS (revision to confirm).
Notes | Edge verdict read remains login-free and does not require pairing (AD-5). BLE mechanisms per SRS-CONN/SRS-SEC.
SRS-MOB-004 — Device onboarding and field provisioning
Requirement Statement | The mobile app shall onboard a device by enrolling it (or confirming its enrolment) in the cloud registry under the correct tenant/owner, applying initial configuration and identity association, and shall support authenticated field re-provisioning of a returned/refurbished device between trips, recording each onboarding/provisioning action for the audit trail.
Rationale | The reusable/refurbish loop (D1) requires devices to be enrolled, associated to an owner, and re-provisioned between trips; onboarding must be authenticated and attributable so the registry and the record's provenance stay trustworthy.
Priority | Must
Source | Architecture
Verification Method | Test, Demonstration
Traceability | SYS-141 (field re-provisioning), SYS-030 (readiness), SYS-121 (provenance), PRD-CLOUD-001, PRD-MOB-001, PRD-SEC-009; → SRS-CLOUD-001, SRS-FLEET-002, SRS-SEC. Standard: IEC 62443 secure onboarding concepts (revision to confirm).
Notes | Private key material is never exposed to the app in cleartext (SYS-141); provisioning is audited (SRS-CLOUD-018).
SRS-MOB-005 — Device sharing with scoped, revocable access
Requirement Statement | The mobile app shall allow an authorised user to share access to a device or shipment with another authorised user within the tenant at a defined role/scope, and shall support revoking that access, recording each grant and revocation in the audit trail.
Rationale | Hand-offs across dispatch, transport, and receiving require controlled sharing of visibility and action; sharing must be scoped, revocable, and audited so it does not become an ungoverned access path.
Priority | Could
Source | User Need
Verification Method | Test
Traceability | SYS-136 (RBAC), SYS-125 (tenant isolation), PRD-MOB-007, PRD-CLOUD-009; → SRS-CLOUD-016/ 018, SRS-MOB-002.
Notes | Sharing never extends beyond tenant scope; grants are role-bounded (SRS-CLOUD-016).
7.3 Configuration, arming & local read
SRS-MOB-006 — Trip-profile configuration with validation and poka-yoke
Requirement Statement | The mobile app shall let an authorised user select or define a trip profile (thresholds, excursion logic, logging interval, upload cadence, shock threshold, alert routing), shall present a valid product-appropriate default, shall validate the profile against defined constraints (e.g. low < high threshold, interval within bounds) and prevent applying an invalid or internally inconsistent configuration, and shall apply the validated profile to the device before arming.
Rationale | Configuration must match the product and be error-proof; use-error protection prevents a misconfiguration from becoming a missed excursion.
Priority | Must
Source | Safety
Verification Method | Test
Traceability | SYS-022 (profile configuration), SYS-025 (config validation), SYS-120 (profile schema), PRD-MOB-002, PRD-FUN-014, PRD-UX-007; → SRS-CLOUD-005, SRS-DEV (apply/arm). Standard: IEC 62366-1 use-error (revision to confirm).
Notes | Poka-yoke selection; reconcile with HAL (misconfiguration hazard). Profile version bound to the record (SYS-120).
SRS-MOB-007 — Arming aid with confirmation (not an arming dependency)
Requirement Statement | The mobile app shall assist arming by guiding the user through pre-arm readiness checks and shall display positive confirmation of the armed + profile-valid + placement state, without being required for arming — the device shall remain able to arm without the app.
Rationale | The app speeds and de-risks arming, but mis-arm (an unmonitored trip) is a top hazard and the device must still arm offline without the app (offline resilience, AD-6).
Priority | Should
Source | User Need
Verification Method | Demonstration
Traceability | SYS-023 (arm with confirmation), SYS-024 (readiness/pre-arm checks), PRD-MOB-003, PRD-FUN-015, PRD-UX-004; → SRS-DEV (on-device arm fallback), HAL mis-arm hazard (Stage 10).
Notes | On-device arming fallback is mandatory (PRD-FUN-015); the app must never be a single point of failure for arming.
SRS-MOB-008 — Local device read for field triage
Requirement Statement | The mobile app shall read a device's current status, recent history, diagnostics, readiness, and connectivity self-test over the authenticated BLE channel for field/service triage, presenting detail beyond the on-device summary without altering any governed record.
Rationale | Field and service users need more detail than the on-device verdict; local read complements the edge verdict and supports triage where WAN is absent, while remaining read-only against the record.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-030 (readiness), SYS-147 (read-only edge/governed split), PRD-MOB-004, PRD-CONN-003, PRD-FUN-020; → SRS-CONN-024, SRS-DEV.
Notes | Complements, does not substitute for, the on-device verdict; local read cannot modify a tamper-evident record (SYS-132/147).
7.4 Firmware, notifications & background
SRS-MOB-009 — Firmware-update trigger
Requirement Statement | The mobile app shall allow an authorised user to trigger a governed firmware update for a targeted device (initiating or confirming a cloud-orchestrated, signed update), shall present update progress and outcome, and shall not itself deliver unsigned firmware or bypass the platform's OTA orchestration and signature verification.
Rationale | Field service may need to trigger an update at the device, but firmware authenticity, staged rollout, and fail-safety are the platform's and device's responsibility; the app triggers, it does not become an unsigned side-channel.
Priority | Should
Source | Security
Verification Method | Test
Traceability | SYS-032/139 (signed fail-safe OTA), PRD-MOB-001, PRD-CLOUD-008, PRD-SEC-008; → SRS-CLOUD-005, SRS-OTA, SRS-SEC. Standard: secure-update principles, IEC 62443 concepts (revision to confirm).
Notes | Signing/verification/rollback owned by device+OTA+SEC (SYS-139); the app triggers and observes.
SRS-MOB-010 — Push alert notifications with acknowledgement
Requirement Statement | The mobile app shall present excursion and device-fault notifications routed to the user by the cloud via platform push, shall let the user acknowledge an alert, and shall report the acknowledgement to the cloud so escalation state is updated.
Rationale | Mobile is a primary responder channel; acknowledgement feeds the cloud escalation engine so a received alert stops escalating and an unreceived one escalates.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-012 (routing/escalation), SYS-011 (alert), PRD-MOB-005, PRD-CLOUD-003, PRD-FUN-009; → SRS-CLOUD-009, HAL delayed/unacknowledged-alert hazard (Stage 10).
Notes | One channel among several (SRS-CLOUD-009); the device's offline alert capture is independent of the app.
SRS-MOB-011 — Background execution for sync and notification delivery
Requirement Statement | The mobile app shall use platform background-execution facilities to receive push notifications and complete queued synchronisation with the cloud when not in the foreground, within the constraints imposed by the mobile OS, without loss of a queued change or a delivered notification.
Rationale | Field users rarely keep the app foreevidence-backed; background delivery keeps alerts timely and keeps offline-authored changes syncing without loss.
Priority | Could
Source | Performance
Verification Method | Test
Traceability | SYS-176 (reconciliation), PRD-MOB-005/006; → SRS-MOB-010/012.
Notes | Behaviour bounded by OS background policy; no queued change is dropped when backevidence-backed (SRS-MOB-012).
7.5 Offline, sync & permissions
SRS-MOB-012 — Offline operation and synchronisation with audited conflict resolution
Requirement Statement | The mobile app shall support local (BLE) provisioning, configuration, and read when the phone has no network, shall queue changes locally, and shall synchronise with the cloud when connectivity returns, resolving any sync conflict deterministically and recording the resolution in the audit trail without silent overwrite of governed data.
Rationale | Field users often lack phone connectivity at the point of work; offline authoring with later audited sync preserves both usability and record governance.
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-176 (device↔cloud reconciliation), SYS-028 (audit), PRD-MOB-006, PRD-CONN-003, PRD-CLOUD-010; → SRS-CLOUD-018, SRS-MOB-011.
Notes | Conflicts resolved with audit (PRD-MOB-006); no governed record is silently overwritten.
SRS-MOB-013 — Runtime OS permission handling
Requirement Statement | The mobile app shall request only the OS-level permissions it needs for its active function (e.g. Bluetooth, notifications, and — only where required — location), shall degrade gracefully when a permission is denied, and shall explain to the user why a permission is needed before requesting it.
Rationale | Least-privilege permission use limits privacy exposure and respects platform policy; the app must remain usable (with reduced function) when a permission is withheld.
Priority | Should
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-148 (data minimisation/privacy), PRD-MOB-008, PRD-SEC-010; → SRS-MOB-016. Standard: OWASP MASVS, platform permission guidelines (revision to confirm).
Notes | Location minimisation aligns with data-minimisation (SYS-148); regional privacy mapping is Evidence-Pending (D4).
---
8. Non-functional Requirements
SRS-MOB-014 — Accessibility and localisation
Requirement Statement | The mobile app shall conform to the target-platform accessibility guidelines (e.g. dynamic type, screen-reader labelling, sufficient contrast, and touch-target sizing) and shall support localisation of user-facing text, dates/times (UTC-referenced), units, and number formats for the target markets.
Rationale | Diverse users across global markets (D4) and inclusive design require accessible, localised interfaces; ambiguous units or untranslated safety text is a use-error risk.
Priority | Could
Source | User Need
Verification Method | Inspection
Traceability | PRD-MOB-008, PRD-UX-010; D4; → accessibility/localisation review. Standard: iOS/Android accessibility guidelines, IEC 62366-1 (revision to confirm).
Notes | Target language set and conformance level are Evidence-Pending (PRD-MOB-008); the obligation is fixed, the list is not invented here.
---
Error Handling
SRS-MOB-015 — Graceful error handling and crash reporting/telemetry
Requirement Statement | The mobile app shall handle connectivity, pairing, sync, and permission errors without loss of a queued governed change, presenting an actionable message, and shall capture crash and diagnostic telemetry, transmitting it only with the applicable consent and with personal/tenant data minimised, to support defect triage.
Rationale | A robust field tool must fail visibly and recoverably, and its own defects must be diagnosable; crash telemetry must not itself become a privacy or tenant-leakage exposure.
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-148 (data minimisation), SYS-176, PRD-MOB-006, PRD-SEC-010; → SRS-MOB-012/013, SRS-CLOUD-024. Standard: OWASP MASVS logging/privacy (revision to confirm).
Notes | Telemetry consent and minimisation align with SYS-148; no governed change is lost on error (SRS-MOB-012).
---
Security Considerations
SRS-MOB-016 — Mobile application security posture
Requirement Statement | The mobile app shall protect data in transit to the cloud with TLS 1.2+/1.3, store tokens and any cached controlled data in platform-provided secure storage, enforce role-based permissions locally as defence-in-depth, and shall not expose an unauthenticated configuration, control, or debug path — while never imposing authentication on, or providing, the login-free on-device edge read.
Rationale | The app is a governed-side client (AD-5); its security posture protects credentials, controlled data, and configuration on an untrusted mobile platform without defeating the frictionless offline edge read (which is device-side and out of the app's scope).
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-137 (transport), SYS-146 (session), SYS-138 (at-rest), SYS-147 (split trust boundary), PRD-SEC-005/006/007, PRD-MOB-007; → SRS-SEC (mechanisms, EM-ENG-008), SRS-MOB-001/002. Standard: OWASP MASVS, TLS 1.2+/1.3 (revision to confirm).
Notes | Cryptographic mechanisms and key handling owned by SRS-SEC; this states the app-software posture.
---
Verification Requirements
SRS-MOB-017 — Mobile verification across pairing, configuration, offline, security and accessibility
Requirement Statement | The mobile app shall be verified by SSO/RBAC and permission-enforcement tests, authenticated-pairing and onboarding tests, configuration-validation and arming-confirmation demonstrations, local-read and firmware-trigger tests, push-notification/acknowledgement and background-execution tests, offline-authoring and audited-sync-conflict tests, mobile-security (MASVS) tests, and accessibility/ localisation review, with results traced to the requirements above.
Rationale | The augment-not-gate posture, the governed-side security, and the offline field usability are only credible if verified under the connectivity, permission, and platform conditions the app actually meets in service.
Priority | Must
Source | Security
Verification Method | Test, Demonstration, Inspection
Traceability | SYS-022/023/024/025, SYS-136/146/147, SYS-176, PRD-MOB-001…008, PRD-CLOUD-003/009; → SRS-VER. Standard: OWASP MASVS, IEC 62366-1 (revision to confirm).
Notes | Evidence-Pending items (target language set, accessibility conformance level, regional privacy mapping) are carried as gaps and close on their associated reviews/tests.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-022…025 (configure/arm/validate), SYS-030 (readiness), SYS-136/146/147 (RBAC/SSO, session, split trust boundary), SYS-141 (field re-provisioning), SYS-032/139 (firmware), SYS-011/012 (alert/escalation), SYS-176 (reconciliation), SYS-148 (data minimisation); PRD-MOB-001…008, PRD-CLOUD-003/008/009, PRD-SEC-005/006/007/009, PRD-FUN-009/014/015/020. Cross-module: cloud identity/RBAC/SSO, alerting, onboarding registry → SRS-CLOUD; cryptographic/TLS/pairing/key mechanisms → SRS-SEC (EM-ENG-008); BLE/transport stack → SRS-CONN; device capture/verdict/arming fallback/buffer → SRS-DEV; OTA packaging/campaign → SRS-OTA; deep fleet workflows → SRS-FLEET; usability/HMI → SRS-SAFE / HAL for the mis-arm and delayed-alert hazards (Stage 10). Evidence-Pending items (target languages, accessibility conformance level, regional privacy-law mapping) are carried as gaps, not invented, and close on their associated reviews.
SRS-600 — Cybersecurity (module SRS-SEC)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-SEC · Owner: Product Security Engineer (EM-ENG-008) with the Embedded Software Engineer (EM-ENG-007) and Cloud/Platform Engineer. Detailed cybersecurity-software requirements decomposed from the Security & Data-Integrity System Requirements (SYS-132…148) and Product Requirements (PRD-SEC 1100, security-relevant parts of PRD-CLOUD 900), realising the trustworthy-record posture (D2) and the split trust boundary (AD-5: login-free edge read vs platform-governed configuration/data). This module owns the cryptographic and identity mechanisms that SRS-CONN, SRS-DEV, SRS-CLOUD, SRS-OTA and SRS-DATA depend on. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM software establishes and defends a chain of trust from a hardware root of trust through secure boot, signed firmware, protected key storage, mutually authenticated encrypted transport, and encryption at rest; how it manages device and user identity, authentication, authorization, and sessions; how cryptographic keys and certificates are generated, provisioned, rotated, and revoked over the product's fielded life; and how the software detects, records, and responds to tampering, intrusion, and disclosed vulnerabilities — without ever gating the frictionless login-free on-device edge read (AD-5).
2. Scope
Hardware root of trust and secure element usage; secure boot and firmware authenticity verification; firmware signing/verification; secure storage of keys, credentials, and records; certificate lifecycle and device/field provisioning; TLS and mutual authentication mechanisms; device and user identity management; authentication and password/credential policy; role-based authorization; session management; the split trust boundary; cryptographic key generation, rotation, and revocation; encryption at rest and in transit; cryptographic tamper-evidence of records; physical and logical tamper detection; secure debug / JTAG-SWD lockdown; supply-chain security and SBOM; threat monitoring, incident logging, and the security audit trail; security updates; and vulnerability management. Excludes the transport bearers and connection management (SRS-CONN, which consumes the TLS/certificate mechanisms here), OTA campaign packaging and orchestration (SRS-OTA, which consumes the signing/verification mechanisms here), on-device buffer/log storage mechanics (SRS-DEV), cloud application/ingest functions (SRS-CLOUD), data schema/serialisation (SRS-DATA), and safety framing of security-relevant hazards (SRS-SAFE / HAL). Those are referenced where the security software interfaces with them.
3. Applicable Standards
IEC 62443 concepts (device/system security, secure development, defence-in-depth); OWASP ASVS (application security verification — authentication, session, access control, cryptography); NIST CSF (identify/protect/detect/respond/recover) and NIST 800-series concepts (SP 800-57 key management, SP 800-63 digital identity, SP 800-53/-207 access control good practice); 21 CFR Part 11 / ALCOA+ (electronic records, signatures, access limitation, audit trail); EU GDP (data governance); GDPR (personal-data protection, where applicable); TLS 1.2+/1.3 (secure transport). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The compute subsystem (S3) provides a hardware root of trust — a secure element or an MCU with an equivalent protected key store and immutable boot ROM; the specific part and its certified assurance level are Evidence-Pending (SYS-134, design).
• The formal threat model is Evidence-Pending and is authored/maintained with the security persona (SYS-143); requirements whose specific mechanism, algorithm, key length, rotation interval, or regional mapping is unresolved are marked Evidence-Pending and are not invented here (evidence).
• The key-lifecycle process and the factory/field provisioning process are Evidence-Pending (SYS-140/141) and close in design/CSV; this module fixes the software behaviour, not the ceremony detail.
• The cloud platform (S8) and mobile app (S9) implement the platform-governed trust domain (identity, RBAC, SSO federation, session, audit); this module specifies the security-software requirements allocated to each tier and the boundary contract between them.
• Cryptographic primitives are drawn from an approved, maintained algorithm set; the concrete cipher suite, curve/key sizes, and hash are selected in design against the threat model (Evidence-Pending).
5. Architecture Overview
Trust is anchored in a hardware root of trust (HRoT) on the device (S3) that holds the device's unique private identity and boot-verification keys and performs secure boot, refusing to execute unverified or altered firmware. Above it, a crypto services layer provides signing/verification, key storage, key generation/rotation, at-rest encryption, and the record tamper-evidence primitive consumed by SRS-DEV and SRS-DATA. A secure transport layer provides mutually authenticated TLS and certificate validation consumed by SRS-CONN, using the HRoT-held device identity. On the platform side (S8/S9) an identity and access layer provides user identity, authentication, RBAC, SSO federation, and session management; a PKI/registry issues, rotates, and revokes device and transport credentials; and a security operations layer collects the append-only security audit trail, monitors for security events, and drives incident response and vulnerability management. Cutting across all of it is the split trust boundary (AD-5): the on-device accept/reject/quarantine edge read is strictly read-only and requires no credential, while every configuration, data-access, administrative, and record-modifying action is authenticated, role-authorised, and audited. Production debug/programming ports are locked. A supply-chain/SBOM discipline governs what is built and shipped, feeding vulnerability management for the fielded life.
---
7. Functional Requirements
7.1 Root of trust, secure boot & firmware authenticity
SRS-SEC-001 — Hardware root of trust and protected key store
Requirement Statement | The device security software shall anchor all device-side cryptographic operations (identity, signing/verification, key storage, at-rest key protection) in a hardware root of trust — a secure element or equivalent protected key store — and shall not hold any long-term private key or device credential in general-purpose flash or RAM in cleartext.
Rationale | A software-only key can be extracted or cloned from a device that travels in uncontrolled hands; anchoring in hardware is what makes identity, attribution, mutual TLS, and signed OTA trustworthy (SYS-134). Every higher security control collapses to key protection.
Priority | Must
Source | Security
Verification Method | Analysis, Test
Traceability | SYS-134 (HRoT/secure boot), SYS-140 (key store), PRD-SEC-003; → SRS-DEV (key store use), SRS-SEC-005/017. Standard: IEC 62443 concepts; NIST SP 800-57 (revision to confirm).
Notes | Specific secure element / MCU secure enclave and its assurance level are Evidence-Pending (design). Threat model Evidence-Pending (SRS-SEC-023).
SRS-SEC-002 — Unique, non-cloneable device identity
Requirement Statement | The device security software shall provision and use a unique, non-cloneable cryptographic device identity, whose private key is generated in or injected into the hardware root of trust and never leaves it in cleartext, as the single source of device identity for mutual TLS (SRS-SEC-010), record/action attribution (SRS-SEC-018), and firmware trust.
Rationale | Attribution and non-repudiation (SYS-133) and registry-gated ingest (PRD-CLOUD-001) require each device to be uniquely and unforgeably identifiable; a cloneable identity would let an impostor inject or repudiate records.
Priority | Must
Source | Security
Verification Method | Test, Analysis
Traceability | SYS-134, SYS-133 (attribution), PRD-SEC-003, PRD-CLOUD-001; → SRS-SEC-009/010/018, SRS-CONN-013. Standard: IEC 62443 concepts (revision to confirm).
Notes | Provisioning ceremony per SRS-SEC-009; identity feeds cloud device-registry enrolment (S8).
SRS-SEC-003 — Secure boot with firmware authenticity and integrity verification
Requirement Statement | The device security software shall verify the authenticity and integrity of each firmware image against the hardware root of trust at every boot before transferring execution to it, and shall refuse to execute — entering a defined safe/recovery state — any image that is unsigned, signed by an untrusted key, altered, or below the anti-rollback minimum version.
Rationale | A tampered image that runs could silently corrupt the evidentiary record, exfiltrate keys, or disable monitoring; secure boot ensures only genuine, unaltered firmware executes (SYS-134).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-134 (secure boot), SYS-139 (anti-rollback), PRD-SEC-003; → SRS-SEC-004, SRS-OTA-004/006, SRS-DEV (boot/safe mode). Standard: IEC 62443; secure-boot good practice (rev. to confirm).
Notes | Anti-rollback minimum enforced jointly with SRS-OTA-006. Safe/recovery-state behaviour owned by SRS-DEV; this fixes the refuse-to-run gate.
SRS-SEC-004 — Firmware signing and verification chain
Requirement Statement | The security software shall accept and apply only firmware whose cryptographic signature is verified against a trusted verification key rooted in the hardware root of trust, shall verify the signature over the complete image (including version and compatibility metadata), and shall protect the signing keys such that no fielded device holds a firmware signing (private) key.
Rationale | Firmware is the highest-privilege code on the device; only images signed by the controlled release authority may run, and the signing key must never be exposed on a fielded unit (SYS-139).
Priority | Must
Source | Security
Verification Method | Test, Analysis
Traceability | SYS-139 (signed update), SYS-140 (key mgmt), PRD-SEC-008; → SRS-OTA-004/005, SRS-SEC-017. Standard: IEC 62443; secure-update principles (revision to confirm).
Notes | OTA packaging/transport in SRS-OTA; this owns the signature verification mechanism and key custody. Signing infrastructure is a platform/PKI responsibility (Evidence-Pending, SRS-SEC-017).
7.2 Secure storage & encryption
SRS-SEC-005 — Protected storage of keys and credentials
Requirement Statement | The security software shall store all device private keys, transport credentials, and at-rest encryption keys inside the hardware root of trust or under keys sealed to it, and shall expose only cryptographic operations (sign, verify, encrypt, decrypt, derive) — never the raw private key material — to application firmware.
Rationale | Confining key use to the HRoT means a firmware compromise or memory dump cannot yield usable keys, preserving identity and confidentiality even under partial compromise (SYS-138/140/144).
Priority | Must
Source | Security
Verification Method | Test, Analysis
Traceability | SYS-138 (at-rest), SYS-140 (key store), SYS-144 (debug lockdown), PRD-SEC-007/009; → SRS-SEC-001/006/020. Standard: NIST SP 800-57; IEC 62443 (revision to confirm).
Notes | Complements JTAG/SWD lockdown (SRS-SEC-020); key material never in cleartext at rest or on any bus.
SRS-SEC-006 — Encryption of controlled data at rest on the device
Requirement Statement | The security software shall encrypt the device non-volatile record log and credential/key store at rest using keys managed under SRS-SEC-017, such that physical recovery of a lost or stolen device does not expose record contents, personal/location data, or credentials in cleartext.
Rationale | A reusable device passes through uncontrolled hands and can be lost or stolen; at-rest encryption protects confidentiality, complementing the integrity/tamper-evidence of SRS-SEC-018 (SYS-138).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-138 (at-rest encryption), PRD-SEC-007; → SRS-DEV (NVM/log), SRS-SEC-005/017, SRS-DATA-007. Standard: ISO/IEC 27001 concepts; NIST crypto good practice (revision to confirm).
Notes | Algorithm/key length Evidence-Pending (design). At-rest encryption addresses confidentiality; integrity is addressed by SRS-SEC-018.
SRS-SEC-007 — Encryption of controlled data in transit
Requirement Statement | The security software shall protect all controlled data in transit — device-to-cloud, app-to-cloud, and any BLE/USB configuration or controlled read — for confidentiality and integrity using authenticated encryption over TLS 1.2 or later (or an equivalent authenticated, integrity-protected channel), and shall reject any attempt to transmit controlled data over an unauthenticated, unencrypted, or downgraded channel.
Rationale | The connectivity path is the primary remote attack surface; records, alerts, commands, and configuration must not be intercepted, altered, or spoofed in transit (SYS-137).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-137 (mutual TLS), PRD-SEC-006, PRD-CONN-006; → SRS-CONN-011/025, SRS-SEC-010. Standard: TLS 1.2+/1.3; OWASP ASVS (revision to confirm).
Notes | Bearers and connection management owned by SRS-CONN; this owns the transport-security mechanism. Login-free edge read is not a data-in-transit path (AD-5).
SRS-SEC-008 — Encryption of controlled data at rest in the cloud
Requirement Statement | The platform security software shall encrypt controlled data at rest in the cloud store (records, audit trails, credentials, configuration, personal/location data) using platform-managed keys under SRS-SEC-017, such that compromise of a storage medium or backup does not expose cleartext record or credential contents.
Rationale | The cloud holds the authoritative record and cross-tenant data; a compromised disk, snapshot, or backup must not disclose customer cold-chain evidence or credentials (SYS-138, SYS-125).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-138 (at-rest, cloud), SYS-125 (tenant isolation), PRD-SEC-007, PRD-CLOUD-011; → SRS-DATA-007/016, SRS-SEC-017. Standard: ISO/IEC 27001 concepts (revision to confirm).
Notes | Backup/restore integrity per SRS-DATA-014/015; tenant key separation reconciled with SRS-DATA-009.
7.3 Certificates, TLS & mutual authentication
SRS-SEC-009 — Certificate and credential lifecycle with secure provisioning
Requirement Statement | The security software shall support the full lifecycle of the device certificate and transport credentials — controlled, authenticated, audit-logged provisioning at manufacture; enrolment into the cloud device registry; authenticated field re-provisioning or renewal on refurbishment/redeploy; and revocation — such that private key material is never exposed in cleartext at any lifecycle stage and no buffered record is lost across a renewal.
Rationale | Identity and attribution are only as trustworthy as the provisioning that created them; the reuse/refurbish loop (D1) requires field re-provisioning, and a leaky ceremony would undermine the whole trust chain (SYS-141).
Priority | Must
Source | Security
Verification Method | Analysis, Inspection, Test
Traceability | SYS-141 (provisioning), SYS-140 (lifecycle), PRD-SEC-009/003, PRD-CLOUD-001; → SRS-CONN-013, SRS-SEC-002/017. Standard: IEC 62443 secure provisioning (revision to confirm).
Notes | Evidence-Pending — factory/field provisioning ceremony to confirm (SYS-141). No-data-loss on renewal reconciled with SRS-CONN-013.
SRS-SEC-010 — Mutually authenticated TLS
Requirement Statement | The security software shall establish device-to-cloud and app-to-cloud sessions using mutually authenticated TLS in which the device authenticates with its hardware-rooted identity (SRS-SEC-002) and validates the server certificate chain, and shall reject any session that is unauthenticated on either side, downgraded, or presents an untrusted, expired, or revoked certificate.
Rationale | Mutual authentication ensures only registered, genuine devices deliver data and that the device only ever talks to the genuine platform, closing impersonation on both directions (SYS-137).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-137 (mutual auth), PRD-SEC-006, PRD-CONN-006, PRD-CLOUD-001; → SRS-CONN-011/025, SRS-SEC-007/011. Standard: TLS 1.2+/1.3; IEC 62443 (revision to confirm).
Notes | Cipher suite selection Evidence-Pending (design against threat model). Transport orchestration in SRS-CONN.
SRS-SEC-011 — Certificate validation and revocation checking
Requirement Statement | On each secure session the security software shall validate the peer certificate's chain of trust, validity period, and revocation status against the current trust configuration, shall reject an expired, untrusted, or revoked certificate, and shall support in-field update of trust anchors and the revocation set through an authenticated, integrity-protected mechanism.
Rationale | A trustworthy channel depends on live validity; an expired or revoked certificate must break the session rather than be accepted, and trust anchors must be renewable over a multi-year fielded life (SYS-137/140).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-137, SYS-140, PRD-SEC-006/009; → SRS-CONN-013, SRS-SEC-009/017. Standard: TLS; OWASP ASVS (revision to confirm).
Notes | Revocation distribution mechanism (CRL/OCSP-class vs registry push) Evidence-Pending given the intermittent/offline link (D2).
7.4 Identity, authentication, authorization & sessions (platform-governed domain)
SRS-SEC-012 — User identity management and authentication
Requirement Statement | The platform and mobile security software shall authenticate every human user of the governed domain against a managed identity, shall support federation to an enterprise identity provider via single sign-on (SAML/OIDC-class), and shall bind the authenticated user identity to every governed action for attribution (SRS-SEC-018).
Rationale | Enterprise/regulated deployment is gated by IT security and CSV; only authenticated, identifiable users may configure devices or access data, and SSO avoids a parallel credential store (SYS-136/146).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (RBAC/SSO), SYS-146 (auth), PRD-SEC-005, PRD-CLOUD-009; → SRS-SEC-014/015/016, SRS-CLOUD (user mgmt). Standard: NIST SP 800-63; OWASP ASVS (revision to confirm).
Notes | Multi-factor authentication for privileged roles Should where the IdP supports it; reconciled with customer IT policy (Evidence-Pending).
SRS-SEC-013 — Authentication and credential (password) policy
Requirement Statement | Where the platform authenticates users with locally managed credentials rather than federated SSO, the security software shall enforce a configurable credential policy (minimum strength, screening against known-breached credentials, no forced-periodic-rotation without cause, secure recovery), shall store credentials only as salted, computationally hard one-way hashes, and shall apply account lockout/rate-limiting against credential-guessing attacks.
Rationale | Weak or reused passwords and unlimited guessing are a primary account-takeover path; modern credential policy and rate-limiting protect the governed domain without imposing counter-productive rotation (SYS-146).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-146 (auth/lockout), PRD-SEC-005, PRD-CLOUD-009; → SRS-SEC-012/015. Standard: NIST SP 800-63B; OWASP ASVS (revision to confirm).
Notes | Federated SSO (SRS-SEC-012) is preferred; this governs the local-credential fallback path.
SRS-SEC-014 — Least-privilege role-based authorization
Requirement Statement | The platform and mobile security software shall enforce least-privilege role-based access control on every configuration, data-access, export, administrative, and record-modifying function, shall deny by default any action not explicitly granted to the actor's role, and shall not gate the login-free on-device edge read with any access-control mechanism.
Rationale | Only authorised roles may configure devices, read, export, or modify data; deny-by-default and least privilege bound the blast radius of a compromised account — while AD-5's frictionless edge read must remain ungated (SYS-136/147).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-136 (RBAC), SYS-147 (trust split), PRD-SEC-005, PRD-CLOUD-009; → SRS-SEC-012/016, SRS-DATA-009 (tenant scope). Standard: 21 CFR Part 11 access limitation; OWASP ASVS access control (revision to confirm).
Notes | Authorization is enforced server-side within the owning tenant scope (SRS-DATA-009); the mobile app is not the sole enforcement point.
SRS-SEC-015 — Session and token management
Requirement Statement | The platform and mobile security software shall protect authenticated sessions and tokens against theft, fixation, and replay (integrity-protected tokens, secure transport-only cookies where applicable, rotation on privilege change), shall enforce configurable idle and absolute session expiry, and shall require re-authentication for defined sensitive actions.
Rationale | RBAC/SSO is only as strong as the session that carries the authenticated identity; a stolen or replayed session bypasses access control and enables record tampering or misconfiguration (SYS-146).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-146 (session mgmt), PRD-SEC-005/006; → SRS-SEC-012/014. Standard: OWASP ASVS session management (revision to confirm).
Notes | Sensitive actions (e.g. profile change, export, deletion, key/credential action) require step-up/re-auth; the sensitive-action set is finalised with the threat model (Evidence-Pending).
SRS-SEC-016 — Split trust boundary enforcement (edge read vs governed domain)
Requirement Statement | The security software shall enforce two distinct trust domains: (a) an on-device edge domain in which the accept/reject/quarantine verdict and trip status are readable on the device display with no network, phone, login, or credential, strictly read-only with respect to governed state; and (b) a platform-governed domain in which all configuration, data access, export, administration, and record modification are authenticated, role-authorised, and audited — such that no edge-read interaction can alter, delete, or reconfigure any governed data or record.
Rationale | AD-5 / concept C9: the last-mile user has no phone, network, or training, so the verdict must be readable with zero friction, yet the authoritative record must stay governed — the boundary is a security requirement, not a UX convenience (SYS-147).
Priority | Must
Source | Security
Verification Method | Test, Analysis, Demonstration
Traceability | SYS-147 (split trust boundary), SYS-148 (privacy), PRD-SEC-005/001, PRD-CLOUD-009; → SRS-DEV (edge read/HMI), SRS-SEC-014, SRS-CONN-004/012. Standard: IEC 62443 zoning concepts (rev. to confirm).
Notes | The edge read exposes only the verdict/status, never governed data export or configuration; this is the security invariant behind SRS-CONN-004/012.
7.5 Key lifecycle, tamper-evidence & audit
SRS-SEC-017 — Cryptographic key generation, rotation, and revocation
Requirement Statement | The security software shall generate cryptographic keys from an approved random source of sufficient entropy, shall support rotation of device identity, transport, signing/ verification, and at-rest encryption keys on a defined schedule and on demand, and shall support fleet-wide revocation of a compromised or expired key or credential — all without loss of the historical record's integrity or attribution.
Rationale | Integrity (SRS-SEC-018), identity (SRS-SEC-002), transport (SRS-SEC-010), signing (SRS-SEC-004), and at-rest encryption (SRS-SEC-006/008) all collapse to key management; without defined rotation/revocation a single compromise is unrecoverable (SYS-140).
Priority | Must
Source | Security
Verification Method | Analysis, Test
Traceability | SYS-140 (key lifecycle), PRD-SEC-009; → SRS-SEC-002/004/006/008/009/011. Standard: NIST SP 800-57 key management; IEC 62443 (revision to confirm).
Notes | Evidence-Pending — key hierarchy, rotation intervals, and revocation distribution to confirm in design/CSV (SYS-140). Rotation must preserve verifiability of records signed under prior keys.
SRS-SEC-018 — Cryptographic tamper-evidence of records and actions
Requirement Statement | The security software shall apply a cryptographic integrity mechanism (digital signature and/or hash-chain) to every logged record and lifecycle action at the point of capture on the device, bind it to the authenticated device (and, where a human is the actor, the authenticated user/role) and the trusted UTC timestamp, and preserve it unbroken through transport and cloud storage, such that any addition, deletion, reordering, or alteration of a stored record or action is computationally detectable and attributable.
Rationale | A release/acceptance disposition and any audit rest on the record; a silent edit must be impossible to hide, and integrity must begin at capture — not at ingest — so the last mile is never a trust gap (SYS-132/133, D2).
Priority | Must
Source | Security
Verification Method | Test, Analysis
Traceability | SYS-132 (tamper-evidence), SYS-133 (attribution/non-repudiation), SYS-131 (trusted time), PRD-SEC-001/002, PRD-CLOUD-002; → SRS-DATA-003/013, SRS-DEV (capture), SRS-CLOUD (verify/preserve). Standard: 21 CFR Part 11 / ALCOA+; NIST crypto good practice (revision to confirm).
Notes | Ingest-side verification per SRS-DATA-013 / SRS-CLOUD; this owns the device-side capture mechanism. Signing key custody per SRS-SEC-005/017.
SRS-SEC-019 — Append-only, non-editable security audit trail
Requirement Statement | The security software shall maintain an append-only audit trail on both device and cloud recording security-relevant events — authentication attempts, access-control decisions, configuration/arming/calibration/firmware/key/data actions, integrity/signature-check outcomes — where each entry is attributable, UTC-time-stamped, and cannot be modified or deleted without detection, and shall reconcile the device and cloud trails after upload.
Rationale | The governance of the record must itself be inspectable and defensible; a device-only trail is lost with a lost unit and a cloud-only trail misses offline actions (SYS-135, D2).
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-135 (audit trail), SYS-028, PRD-SEC-004, PRD-CLOUD-007; → SRS-DATA-003/013, SRS-SEC-022. Standard: 21 CFR Part 11 audit trail; ALCOA+ (revision to confirm).
Notes | Audit-trail integrity is itself protected under SRS-SEC-018. Retention per SRS-DATA-010 and security-event policy (SRS-SEC-022).
SRS-SEC-020 — Production debug and programming-port lockdown
Requirement Statement | The device security software/configuration shall disable, lock, or cryptographically authenticate all hardware debug and programming interfaces (e.g. JTAG/SWD) on production devices, such that a fielded device cannot expose keys, memory, or firmware, or be re-flashed, through a debug port without cryptographic authorisation.
Rationale | An open debug port is the most direct path to key extraction, memory dump, and firmware replacement, defeating secure boot, at-rest encryption, and provisioning; physical access is realistic for a device that travels in uncontrolled hands (SYS-144).
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-144 (debug lockdown), PRD-SEC-003/007/009; → SRS-SEC-001/005, SRS-DEV (manufacturing/ developer modes). Standard: IEC 62443 device hardening (revision to confirm).
Notes | Any authenticated debug re-enable path is itself a governed, audited action (SRS-SEC-019). Reconcile with threat model (SRS-SEC-023).
7.6 Supply chain, monitoring, updates & vulnerability management
SRS-SEC-021 — Supply-chain security and software bill of materials (SBOM)
Requirement Statement | The security software process shall maintain a software bill of materials for device firmware, cloud, and mobile components, shall verify the authenticity and integrity of third-party and build artefacts in the build/release pipeline, and shall enable rapid identification of affected releases and fielded devices when a component vulnerability is disclosed.
Rationale | A signed image (SRS-SEC-004) guarantees only what shipped, not that its components are free of known vulnerabilities or upstream tampering; an SBOM makes vulnerability management actionable when a dependency CVE is disclosed (SYS-145).
Priority | Should
Source | Security
Verification Method | Inspection, Analysis
Traceability | SYS-145 (SBOM/supply chain), PRD-SEC-008/012/011; → SRS-SEC-023/024, SRS-OTA-013. Standard: NIST CSF supply chain; SBOM good practice; IEC 62443 (revision to confirm).
Notes | SBOM format and pipeline signing tooling Evidence-Pending (design/DevOps).
SRS-SEC-022 — Threat monitoring, security-event detection and incident logging
Requirement Statement | The platform security software shall detect, log, and alert on security-relevant events (authentication failures, unauthorised-device data attempts, anomalous access, failed integrity or signature checks, revoked-credential use), shall record them in the security audit trail with retention per policy, and shall support a defined incident-response workflow.
Rationale | A connected, OTA-updatable fleet needs detection and response, not prevention alone; an undetected compromise silently corrupts the evidentiary product (SYS-142).
Priority | Should
Source | Security
Verification Method | Demonstration, Test
Traceability | SYS-142 (monitoring/IR), PRD-SEC-011, PRD-CLOUD-007/012; → SRS-SEC-019/023, SRS-CLOUD (observability). Standard: NIST CSF detect/respond; IEC 62443 (revision to confirm).
Notes | Evidence-Pending — incident-response runbook to confirm (SYS-142). Feeds vulnerability management (SRS-SEC-023).
SRS-SEC-023 — Security-by-design threat model and vulnerability management
Requirement Statement | The security software shall be developed against a documented, maintained threat model whose identified risks trace to the security requirements and controls in this module, and shall be governed by an ongoing vulnerability-management process (monitoring of disclosures against the SBOM, triage by severity, and remediation) across device, cloud, and mobile for the product's fielded life.
Rationale | Security requirements must be driven by an explicit, reviewed threat model rather than assumed; a fielded connected device faces vulnerabilities discovered after release that must be tracked and patched (SYS-143).
Priority | Must
Source | Security
Verification Method | Analysis
Traceability | SYS-143 (threat model/vuln mgmt), PRD-SEC-012/008/011; → all SRS-SEC, SRS-OTA-011/013, SRS-SAFE / HAL. Standard: IEC 62443; NIST CSF identify; OWASP ASVS (revision to confirm).
Notes | Evidence-Pending — formal threat model to confirm (SYS-143); reconcile safety-relevant security with HAL (Stage 10). Severity-dominant Action Priority (CLAUDE.md rule 3).
SRS-SEC-024 — Timely security updates via the trustworthy update path
Requirement Statement | The security software shall enable timely remediation of a triaged vulnerability by delivering a signed security update through the OTA path (SRS-OTA), prioritising security-critical fixes in campaign scheduling, and enforcing the anti-rollback minimum so a device cannot be regressed to a known vulnerable version after patching.
Rationale | Detection and an SBOM are only useful if a fix can be shipped and cannot be undone; secure, anti-rollback-protected updates close the disclosed-vulnerability loop for the fielded fleet (SYS-139/143).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-139 (signed/anti-rollback), SYS-143 (vuln mgmt), PRD-SEC-008; → SRS-OTA-004/006/010, SRS-SEC-004/023. Standard: IEC 62443; secure-update principles (revision to confirm).
Notes | Update mechanism/packaging owned by SRS-OTA; this fixes the security-update obligation and anti-rollback tie.
---
8. Non-functional Requirements
SRS-SEC-025 — Cryptographic operations within the energy and latency budget
Requirement Statement | The device security software shall perform its cryptographic operations (secure boot verification, per-record signing/hashing, TLS handshakes, at-rest encryption) within the compute-time and energy allocations of the documented energy budget over the worst-case 90-day cold-duty profile, without causing a sample, event, or alert to miss its timing target.
Rationale | Security must not defeat endurance (D3) or delay a rescue-enabling alert; per-record and per-session crypto cost must live inside the budget, or 90-day operation and alert latency fail.
Priority | Must
Source | Performance
Verification Method | Analysis, Test
Traceability | SYS-089/090 (energy budget), SYS-011 (alert latency), PRD-ELEC-008; → SRS-PERF, SRS-CONN-021. Standard: — (evidence: no numeric budget asserted here).
Notes | Hardware-accelerated crypto in the secure element is preferred to meet this; per-operation budget Evidence-Pending (SRS-PERF/power budget).
SRS-SEC-026 — Cryptographic agility over the fielded life
Requirement Statement | The security software shall implement cryptographic algorithms, key sizes, and certificate parameters as configurable, updatable choices (via SRS-OTA and trust-anchor update) rather than hard-wired constants, so that a deprecated or broken primitive can be migrated fleet-wide without hardware replacement.
Rationale | A device with a multi-year fielded life will outlast the security margin of today's primitives; crypto-agility is what lets the fleet migrate before a primitive becomes a liability (SYS-143).
Priority | Should
Source | Security
Verification Method | Analysis
Traceability | SYS-143, SYS-140, PRD-SEC-012/009; → SRS-SEC-017/024, SRS-OTA. Standard: NIST SP 800-57 crypto transition; IEC 62443 (revision to confirm).
Notes | Bounded by the fixed HRoT primitives; the migratable set is defined with the threat model (Evidence-Pending).
---
Error Handling
SRS-SEC-027 — Fail-secure handling of security-control failures
Requirement Statement | On failure of a security control — failed boot verification, failed signature or integrity check, failed authentication, expired/revoked certificate, or key-store fault — the security software shall fail secure (deny the action, break or refuse the session, or enter the defined safe/recovery state), shall never fail open to an unauthenticated or unverified path, and shall record the failure to the security audit trail without exposing secret material in any diagnostic.
Rationale | A security control that fails open is worse than none; a failed check must deny and be visible, and no error path may leak keys or record contents (SYS-134/135/142).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-134, SYS-135, SYS-142, PRD-SEC-003/008; → SRS-SEC-003/010/019/022, SRS-DEV (safe mode). Standard: OWASP ASVS error handling; IEC 62443 (revision to confirm).
Notes | Ties the no-silent-failure invariant; fail-secure must still preserve buffered data (no record loss) per SRS-CONN-023 / SRS-DEV.
---
Diagnostics
SRS-SEC-028 — Security posture self-check and reporting
Requirement Statement | The security software shall provide an authenticated on-device and platform security self-check reporting secure-boot status, firmware/signature validity, certificate validity and expiry, key-store health, debug-port lock state, and last security-relevant events, for depot triage and fleet security monitoring, without disclosing secret key material.
Rationale | Deploying or triaging a device, and monitoring the fleet's security posture, requires confirming the trust chain is intact before the device returns to service (SYS-142, SYS-030).
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-142, SYS-030 (readiness), PRD-SEC-011, PRD-FUN-020; → SRS-DEV-029 (self-test), SRS-DIAG, SRS-SEC-022. Standard: IEC 62443; NIST CSF (revision to confirm).
Notes | Access authenticated per the governed-domain model (SRS-SEC-014); posture feeds readiness reporting (SYS-030).
---
Security Considerations
SRS-SEC-029 — Defence-in-depth and no single point of security failure
Requirement Statement | The security architecture shall apply defence-in-depth such that no single control failure exposes record integrity, device identity, or credential confidentiality — combining hardware-rooted keys, secure boot, signed firmware, mutual TLS, at-rest encryption, tamper-evidence, least-privilege access, debug lockdown, and monitoring — and each control shall be independently verifiable.
Rationale | The evidentiary and safety value of the device rests on layered controls; a single defeated layer must not collapse the whole trust chain, consistent with severity-first risk treatment (CLAUDE.md rule 3; SYS-143).
Priority | Must
Source | Security
Verification Method | Analysis
Traceability | SYS-143, SYS-132/134/137/138/144, PRD-SEC-012; → all SRS-SEC, SRS-SAFE / HAL. Standard: IEC 62443 defence-in-depth; NIST CSF (revision to confirm).
Notes | This is the module's governing posture; the threat model (SRS-SEC-023) sizes each layer's contribution (Evidence-Pending).
---
Verification Requirements
SRS-SEC-030 — Security verification against the threat model
Requirement Statement | The security software shall be verified by secure-boot and firmware-signature tests, tamper-evidence and attribution tests, mutual-TLS and certificate-validation/revocation tests, at-rest encryption tests, access-control/session/authentication tests, debug-lockdown inspection, key-lifecycle review, penetration/security assessment against the threat model, and CSV/validation review, with all results traced to the requirements above.
Rationale | The trustworthy-record and controlled-fleet claims are only credible if verified against an explicit threat model under realistic attack conditions, and the security posture is a regulated (Part 11 / CSV) obligation (SYS-143, PRD-SEC).
Priority | Must
Source | Security
Verification Method | Test, Analysis, Inspection
Traceability | SYS-143, all SYS-132…148, PRD-SEC-001…012, PRD-CLOUD-002/007/009; → SRS-VER, SRS-SAFE / HAL. Standard: OWASP ASVS; IEC 62443; GAMP 5 CSV (revision to confirm).
Notes | Threat model, key-lifecycle process, and IR runbook are Evidence-Pending; this fixes the verification obligation that closes them.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-132…148, plus SYS-006/028/030/ 089/090/125/131/133) and Product Requirement (PRD-SEC-001…012, PRD-CLOUD-001/002/007/009/011). Principal upstream anchors: SYS-134 (HRoT/identity/secure boot), SYS-132/133 (tamper-evidence/attribution), SYS-135 (audit trail), SYS-137/138 (transport/at-rest encryption), SYS-139 (signed fail-safe update), SYS-140/141 (key & provisioning lifecycle), SYS-136/146 (RBAC/SSO/session), SYS-142/143/144/145 (monitoring, threat model/vuln mgmt, debug lockdown, SBOM), SYS-147 (split trust boundary), SYS-148 (privacy). Cross-module: transport/certificate consumption → SRS-CONN; firmware signing/anti-rollback consumption → SRS-OTA; on-device key store, log capture, safe mode → SRS-DEV; ingest verification, user mgmt, observability, tenant isolation → SRS-CLOUD / SRS-DATA; numeric energy/latency budgets → SRS-PERF; safety-relevant security hazards → SRS-SAFE / HAL (Stage 10); CSV/regulatory posture → SRS-REG. Evidence-Pending items — the formal threat model, key-lifecycle process, factory/field provisioning ceremony, cipher-suite and key-size selection, revocation distribution, SBOM tooling, incident-response runbook, secure-element part/assurance level — are carried as gaps, not invented, and close on their associated design/CSV activities and SRS-SEC-030.
SRS-700 — OTA Update (module SRS-OTA)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-OTA · Owner: Embedded Software Engineer (EM-ENG-007) with the Product Security Engineer (EM-ENG-008) and Cloud/Platform Engineer. Detailed over-the-air-update software requirements decomposed from the System Requirements (SYS-032, SYS-139, plus SYS-020/030/113) and Product Requirements (PRD-FUN-022, PRD-SEC-008, PRD-CLOUD-008), realising signed, resumable, fail-safe fleet updates that never brick a fielded device and never leave it unmonitored without notice (AD-7, D2). This module consumes the firmware-signing/ verification and anti-rollback mechanisms owned by SRS-SEC and the resumable transport owned by SRS-CONN. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how SCM firmware updates are packaged, downloaded, verified, applied, and — if necessary — rolled back, across an energy-constrained, intermittently connected, RF-hostile fleet, such that an update is always authentic and integral, survives power interruption without bricking or data loss, can be resumed after a dropped link, is compatible with the target device, and is orchestrated as observable, reversible, staged fleet campaigns from the cloud.
2. Scope
Update package format and metadata; firmware download over the selected bearer; resumable/differential download; integrity and signature verification; compatibility and version verification; anti-rollback; apply/activate with fail-safe (A/B or equivalent) and power-interruption handling; rollback and recovery; data preservation across update; version and dependency management; cloud-orchestrated campaign management, per-device targeting, staged/fleet rollout, and campaign rollback; failure reporting; and update metrics. Excludes the cryptographic signing/verification and key mechanisms (SRS-SEC, which OTA consumes), the bearer/transport and store-and-forward resume mechanics (SRS-CONN), on-device storage/partitioning and safe mode (SRS-DEV), the record data preserved across an update (SRS-DATA), and the safety framing of a bricked/unmonitored hazard (SRS-SAFE / HAL). Those are referenced where the OTA software interfaces with them.
3. Applicable Standards
IEC 62443 concepts (secure update, integrity); secure-update / firmware-update good practice (e.g. IETF SUIT-class manifest concepts, A/B fail-safe update patterns); NIST 800-series concepts (firmware protection, resilience); IEC 62304 (software life-cycle process); TLS 1.2+/1.3 and MQTT v5 (transport, via SRS-CONN). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The device holds a hardware root of trust and firmware verification key (SRS-SEC-001/004); OTA verifies against it and never holds a signing (private) key on a fielded unit.
• The device storage layout supports fail-safe activation (dual-bank A/B or an equivalent staging-plus-verified-swap scheme); the concrete partitioning is Evidence-Pending and owned by SRS-DEV.
• The bearer selection, resumable transfer, and back-off are provided by SRS-CONN (LTE-M primary, Wi-Fi/BLE/ USB); OTA specifies update behaviour over whatever bearer SRS-CONN selects, within the energy budget.
• The cloud provides OTA campaign orchestration, per-device targeting, staged rollout, and status tracking (PRD-CLOUD-008); this module specifies the device-side and orchestration-contract requirements.
• Updates run only within the energy governor's sanctioned windows and must not compromise the armed-trip record or the alert path.
5. Architecture Overview
A campaign orchestrator in the cloud (S8) defines an update, targets a device cohort, and drives a staged rollout (canary → expanding rings) with per-device status, pause, and rollback. On the device, an update agent downloads the signed update package over the SRS-CONN-selected bearer — resumably and, where supported, differentially — into a staging area without touching the running image. A verifier checks the package's signature, integrity, target compatibility, and anti-rollback minimum against the hardware root of trust (SRS-SEC) before any activation. An applier activates the verified image using a fail-safe (A/B) scheme so that an interruption at any point leaves a bootable image; a boot confirmer validates the new image on first boot and automatically rolls back to the known-good image if it fails to confirm. Trip data, configuration, calibration, identity, and keys are preserved across the update. The agent reports progress, success, and failure with cause to the campaign, feeding update metrics and fleet readiness. Every stage is energy-gated and non-destructive to the record.
---
7. Functional Requirements
7.1 Packaging, download & resume
SRS-OTA-001 — Signed update package with self-describing manifest
Requirement Statement | The OTA software shall accept firmware only as a self-describing update package containing the image(s), a cryptographic signature, and a manifest declaring firmware version, target hardware/device compatibility, dependency/precondition metadata, anti-rollback minimum, and integrity digests, such that the device can verify authenticity, integrity, and applicability from the package alone.
Rationale | A self-describing, signed package is the unit of trust and decidability; the device must be able to decide — offline if needed — whether a package is genuine, intact, and applicable before acting on it (SYS-139).
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-139 (signed update), PRD-SEC-008, PRD-FUN-022; → SRS-OTA-004/005/007, SRS-SEC-004. Standard: SUIT-class manifest concepts; IEC 62443 (revision to confirm).
Notes | Signature/verification mechanism owned by SRS-SEC-004; this fixes the package/manifest contract.
SRS-OTA-002 — Firmware download over the selected bearer within the energy budget
Requirement Statement | The OTA software shall download the update package over the bearer selected by the connectivity software (LTE-M primary, Wi-Fi/BLE/USB where available), only within the energy governor's sanctioned windows, and shall not begin or continue a download that would breach the energy budget or displace excursion-alert delivery.
Rationale | A multi-hundred-kilobyte image over LPWAN is energy-expensive; download must respect the 90-day budget and never starve the safety-relevant alert path (D3, SYS-090).
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-139, SYS-090 (energy budget), PRD-CLOUD-008, PRD-CONN-011; → SRS-CONN-006/008/016, SRS-OTA-003. Standard: — (evidence: no energy figure asserted here).
Notes | Bearer selection and back-off owned by SRS-CONN; alerts retain priority (SRS-CONN-008).
SRS-OTA-003 — Resumable download after interruption
Requirement Statement | The OTA software shall resume an interrupted package download from the last verified offset after a dropped link, power cycle, or window closure, without re-downloading already received, verified segments, and shall validate the integrity of each segment and of the reassembled package before use.
Rationale | LTE-M sessions inside a shipper are short and intermittent; a non-resumable download would rarely complete and would waste the energy budget on repeated restarts (D2/D3).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-020 (resumable transfer), PRD-CONN-008, PRD-CLOUD-008; → SRS-CONN-007/016, SRS-OTA-004. Standard: secure-update resilience (revision to confirm).
Notes | Uses the resumable transport of SRS-CONN-007; segment integrity ties SRS-OTA-004.
SRS-OTA-004 — Differential / partial update
Requirement Statement | The OTA software shall support differential (delta) update packages that transfer only the change from a known base version, shall verify the reconstructed full image's signature and integrity against the hardware root of trust before activation, and shall fall back to a full-image update when the device's base version is not a valid delta source.
Rationale | Delta updates cut on-air bytes and therefore energy and time on a constrained link, but the reconstructed image must still be fully verified so a delta can never introduce an unverified image (D3, SYS-139).
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-139, SYS-090, PRD-CONN-004, PRD-CLOUD-008; → SRS-OTA-001/005, SRS-SEC-004. Standard: delta-update good practice (revision to confirm).
Notes | Verification is always over the full reconstructed image, never the delta alone.
7.2 Verification before activation
SRS-OTA-005 — Integrity and signature verification before activation
Requirement Statement | The OTA software shall verify the update package's cryptographic signature and integrity digests against the hardware root of trust and shall refuse to activate any image that is unsigned, signed by an untrusted key, or integrity-failing, recording the rejection to the audit trail and leaving the running image untouched.
Rationale | Activation is the highest-consequence step; only an authentic, intact image may ever be made bootable, and a rejection must never disturb the currently running, known-good firmware (SYS-139).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-139 (verify against RoT), PRD-SEC-008; → SRS-SEC-003/004/018, SRS-OTA-001/006. Standard: IEC 62443; secure-boot/update (revision to confirm).
Notes | Verification mechanism owned by SRS-SEC-004; secure boot (SRS-SEC-003) re-verifies at next boot.
SRS-OTA-006 — Compatibility, version and anti-rollback verification
Requirement Statement | Before activation the OTA software shall verify that the package's declared target hardware/device compatibility matches the device, that declared dependencies/preconditions are met, and that the firmware version is at or above the anti-rollback minimum, and shall reject a mismatched, unsatisfied, or below-minimum package.
Rationale | Applying firmware for the wrong hardware variant, an unmet dependency, or a downgrade to a known-vulnerable version can brick a device or reopen a closed vulnerability; compatibility and anti-rollback gates prevent both (SYS-139).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-139 (anti-rollback), SYS-115 (regional variants), PRD-SEC-008, PRD-CLOUD-008; → SRS-OTA-001/012, SRS-SEC-003/024. Standard: IEC 62443; secure-update (revision to confirm).
Notes | Anti-rollback minimum enforced jointly with secure boot (SRS-SEC-003); regional-variant compatibility ties SRS-CONN-020.
7.3 Apply, fail-safe & recovery
SRS-OTA-007 — Fail-safe (A/B) activation that never bricks the device
Requirement Statement | The OTA software shall apply a verified image using a fail-safe activation scheme (dual-bank A/B or equivalent staged swap) in which the currently running image remains intact and bootable until the new image is verified and confirmed, such that no fielded device can be left non-bootable by an update.
Rationale | A bricked device on a shipment is unrecoverable in the field and loses the record; fail-safe activation guarantees a bootable image at every instant (AD-7, SYS-139/032).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-032 (never bricked), PRD-SEC-008, PRD-SAFE-010; → SRS-DEV (partitioning/boot), SRS-OTA-008/009. Standard: A/B fail-safe update pattern; IEC 62443 (revision to confirm).
Notes | Storage partitioning owned by SRS-DEV; concrete A/B vs staged scheme Evidence-Pending (design).
SRS-OTA-008 — Power-interruption and window-closure safety during apply
Requirement Statement | The OTA software shall be resilient to power loss, brownout, or governor-window closure at any point during download, staging, verification, activation, or first-boot confirmation, such that after any interruption the device boots either the prior known-good image or the confirmed new image — never a partially written or unverified image — and resumes monitoring.
Rationale | Cold-chain power is duty-cycled and low-temperature battery behaviour is uncertain (D3); an update interrupted by power loss must be atomic with respect to the bootable image, or the device is bricked or silently corrupted (SYS-139).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-032, PRD-SEC-008, PRD-REL-005; → SRS-OTA-003/007/009, SRS-DEV (safe boot). Standard: secure-update resilience; power-fail-safe design (revision to confirm).
Notes | Activation must be a single atomic commit; interrupted staging discards to the prior image.
SRS-OTA-009 — First-boot confirmation with automatic rollback
Requirement Statement | After activating a new image the OTA software shall require the image to self-confirm healthy operation on first boot within a defined confirmation window, and shall automatically roll back to the prior known-good image if confirmation fails, recording the failed update and its cause, so that a bad image cannot leave the device stuck or unmonitored.
Rationale | A signed, compatible image can still fail at runtime; automatic first-boot rollback bounds the blast radius of a bad release to a single reboot without human field access (AD-7, SYS-139).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-032, PRD-SEC-008, PRD-CLOUD-008; → SRS-OTA-007/011, SRS-DEV (watchdog/boot), SRS-SEC-019. Standard: A/B confirm-or-rollback pattern (revision to confirm).
Notes | Confirmation criteria (boot, self-test pass, connectivity) defined with SRS-DEV; watchdog-backed.
SRS-OTA-010 — Data, configuration, identity and key preservation across update
Requirement Statement | The OTA software shall preserve, across any update or rollback, all buffered and historical trip data, the active trip profile/configuration, calibration data, device identity, and cryptographic keys, and shall migrate persisted data across schema changes without loss or corruption, such that no update erases or invalidates the evidentiary record or the device's trust anchors.
Rationale | An update must never destroy the record or force re-provisioning; store-and-forward (D2) can hold a full offline trip, and losing it or the identity on update would be a record and trust failure (SYS-139/032).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-032, SYS-018 (buffer), PRD-FUN-022, PRD-SEC-008; → SRS-DATA-002/016, SRS-DEV (storage), SRS-SEC-005. Standard: — (evidence).
Notes | Schema migration reconciled with SRS-DATA-002/016; identity/key custody per SRS-SEC-005.
7.4 Campaign management & orchestration
SRS-OTA-011 — Cloud campaign management with staged/fleet rollout and rollback
Requirement Statement | The OTA orchestration software shall manage updates as campaigns with per-device targeting, staged rollout (canary then expanding rings), pause/resume, per-device and per-cohort status tracking, and campaign rollback, and shall halt or pause a campaign automatically when the observed failure rate exceeds a configured threshold, to prevent mass bricking of the fleet.
Rationale | A fleet-wide simultaneous update of a bad image is a mass-failure event; staged rollout with automatic halt and rollback contains a bad release to a small canary cohort (AD-7, PRD-CLOUD-008).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-139, SYS-032, PRD-CLOUD-008, PRD-SEC-008; → SRS-OTA-009/012, SRS-CLOUD (orchestration), SRS-SEC-022. Standard: staged-rollout good practice; IEC 62443 (revision to confirm).
Notes | Cloud campaign functions owned by SRS-CLOUD; this fixes the OTA orchestration contract and the auto-halt safety behaviour. Failure-rate threshold configurable (Evidence-Pending).
SRS-OTA-012 — Version and dependency management across the fleet
Requirement Statement | The OTA software shall track each device's current firmware version and dependency state in the fleet registry, shall target campaigns by version/hardware/region eligibility, and shall prevent applying an update to a device whose declared dependencies or hardware/region variant do not satisfy the package's manifest.
Rationale | A mixed fleet spans firmware versions, hardware variants, and regional radio profiles; version/dependency management ensures each device gets an applicable update and none gets an incompatible one (SYS-030/115, SYS-139).
Priority | Should
Source | Architecture
Verification Method | Test
Traceability | SYS-030 (readiness/version), SYS-115 (regional), SYS-139, PRD-CLOUD-008, PRD-FUN-020; → SRS-OTA-001/006/011, SRS-CONN-020. Standard: — (evidence).
Notes | Ties fleet readiness/version reporting (SYS-030); regional eligibility ties SRS-CONN-020.
7.5 Reporting & metrics
SRS-OTA-013 — Update failure reporting with cause
Requirement Statement | The OTA software shall report the outcome of each update attempt — success, or failure with a specific cause (download incomplete, signature/integrity failure, incompatibility, anti-rollback rejection, activation failure, first-boot-confirmation failure, rolled back) — to the campaign orchestrator and the audit trail, and shall never fail an update silently.
Rationale | Fleet operators must know why a device did not update to triage a stalled campaign, a bad image, or a coverage gap; a silent OTA failure hides both security-patch gaps and reliability problems (PRD-CLOUD-008, SYS-142).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-139, SYS-113 (diagnostics), SYS-142 (security events), PRD-CLOUD-008, PRD-SEC-011; → SRS-CONN-014, SRS-SEC-019/022, SRS-DIAG. Standard: NIST CSF detect (revision to confirm).
Notes | Ties connectivity diagnostics (SRS-CONN-014) and the security audit trail (SRS-SEC-019).
SRS-OTA-014 — Update metrics and campaign observability
Requirement Statement | The OTA software shall collect and expose update metrics — per-device and per-cohort download progress, verification and activation outcomes, rollback counts, time-to-complete, and energy consumed — to fleet tooling so that campaign health and cost can be observed and a decision to pause, continue, or roll back can be made on evidence.
Rationale | A staged rollout decision (SRS-OTA-011) is only as good as the metrics behind it; observable campaign health turns rollout gating from guesswork into an evidence-based decision (PRD-CLOUD-008).
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-113, SYS-139, PRD-CLOUD-008/006; → SRS-OTA-011/013, SRS-CLOUD (observability), SRS-DIAG. Standard: — (evidence).
Notes | Energy-per-update metric feeds the connectivity/energy budget evidence (SRS-CONN-021).
---
8. Non-functional Requirements
SRS-OTA-015 — OTA energy conformance
Requirement Statement | The OTA software shall keep the cumulative energy consumed by update download, verification, and activation within the OTA allocation of the documented energy budget over the worst-case 90-day cold-duty profile, deferring or fragmenting a non-urgent update across windows rather than breaching the budget.
Rationale | An update that drains the battery defeats endurance and can strand a device mid-trip; OTA must live inside its energy allocation like every other radio consumer (D3, SYS-090).
Priority | Must
Source | Performance
Verification Method | Analysis, Test
Traceability | SYS-090 (energy allocation), SYS-089, PRD-ELEC-008; → SRS-CONN-021, SRS-PERF. Standard: — (evidence: no energy figure asserted here).
Notes | OTA energy allocation Evidence-Pending and closes on the power budget (SYS-090). Security- critical updates (SRS-SEC-024) may be prioritised within the allocation.
---
Error Handling
SRS-OTA-016 — Non-destructive handling of any update failure
Requirement Statement | On any download, verification, activation, or confirmation failure the OTA software shall leave the device running a known-good, monitoring-capable image with its record and trust anchors intact, shall log the failure cause, and shall return the device to normal operation and, where appropriate, retry under campaign control — never leaving the device bricked, unmonitored-without-notice, or in an unverified state.
Rationale | The controlling invariant of OTA is that a failed update is a non-event for the record and for monitoring; the fallback is always the known-good image, and the failure is always visible (AD-7, SYS-032/139).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-032, SYS-139, PRD-SEC-008, PRD-SAFE-010; → SRS-OTA-007/008/009/013, SRS-SEC-027. Standard: fail-safe update (revision to confirm).
Notes | "Unmonitored-without-notice" is a HAL-reconciled hazard; forward to SRS-SAFE / HAL (Stage 10).
---
Diagnostics
SRS-OTA-017 — Update-readiness and status self-report
Requirement Statement | The OTA software shall provide, on demand to authorised fleet tooling and depot triage, the device's current firmware version, active/backup bank status, last update outcome and cause, pending-campaign eligibility, and update-readiness (charge, storage headroom, connectivity), without disturbing an armed trip.
Rationale | Deciding whether to include a device in a campaign, and diagnosing one that failed to update, requires visibility of its firmware and readiness state before acting (SYS-030, PRD-CLOUD-008).
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness), SYS-113, PRD-FUN-020, PRD-CLOUD-008; → SRS-OTA-012/013/014, SRS-DEV-029, SRS-DIAG. Standard: — (evidence).
Notes | Access authenticated per the governed-domain model (SRS-SEC-014).
---
Security Considerations
SRS-OTA-018 — OTA as a governed, authenticated, auditable action
Requirement Statement | The OTA software shall treat every update as a security-relevant action — authenticated to the platform, authorised for the target cohort, verified against the hardware root of trust, and recorded in the audit trail — and shall not accept an update instruction, package, or campaign command from an unauthenticated or unauthorised source over any bearer.
Rationale | The update path is a high-value attack surface (a malicious update could compromise the whole fleet); it must be authenticated, verified, and audited end to end, consistent with the trustworthy-update posture (SYS-139/142, SRS-SEC).
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-139, SYS-142, PRD-SEC-008, PRD-CLOUD-008; → SRS-SEC-004/010/018/019/022, SRS-CONN-011/012/025. Standard: IEC 62443; secure-update (revision to confirm).
Notes | Cryptographic mechanisms owned by SRS-SEC; transport authentication by SRS-CONN; this states the OTA security posture.
---
Verification Requirements
SRS-OTA-019 — OTA verification under interruption, RF-hostile, and adversarial conditions
Requirement Statement | The OTA software shall be verified by signed/tampered-package acceptance tests, compatibility and anti-rollback tests, resumable and differential download tests, power-interruption-during- apply tests, first-boot-confirmation and automatic-rollback tests, data-preservation-across-update tests, staged-rollout/auto-halt/campaign-rollback tests, and RF-hostile intermittent-link tests, with results traced to the requirements above.
Rationale | The never-brick, never-lose-the-record, never-accept-a-bad-image claims are only credible if verified under power loss, dropped links, and rejected/tampered packages — the conditions the fleet actually meets (AD-7, SYS-139).
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-032, SYS-139, PRD-SEC-008, PRD-CLOUD-008, PRD-FUN-022; → SRS-VER, SRS-SAFE / HAL, SRS-SEC-030. Standard: IEC 62304; secure-update verification (revision to confirm).
Notes | Power-interruption and rollback tests are the load-bearing evidence for the never-brick invariant; reconcile the bricked/unmonitored hazard with HAL (Stage 10).
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-032, SYS-139, plus SYS-018/020/ 030/090/113/115/142) and Product Requirement (PRD-FUN-022, PRD-SEC-008, PRD-CLOUD-008, plus PRD-CONN-004/011, PRD-SAFE-010, PRD-ELEC-008). Principal upstream anchors: SYS-139 (signed, resumable, fail-safe, anti-rollback update), SYS-032 (never bricked / never unmonitored-without-notice), SYS-090 (energy allocation), SYS-030/113 (readiness/diagnostics), SYS-115 (regional variants). Cross-module: firmware signing/verification, anti- rollback, and audit → SRS-SEC; bearer selection, resumable transport, and back-off → SRS-CONN; storage partitioning, safe boot, watchdog, and schema migration → SRS-DEV / SRS-DATA; campaign orchestration, targeting, and observability → SRS-CLOUD; numeric energy/latency budgets → SRS-PERF; bricked/unmonitored-without-notice hazard → SRS-SAFE / HAL (Stage 10). Evidence-Pending items — the A/B vs staged partitioning scheme, delta-update tooling, the auto-halt failure-rate threshold, the OTA energy allocation, and confirmation criteria — are carried as gaps, not invented, and close on their associated design activities and SRS-OTA-019.
SRS-800 — Fleet Management (module SRS-FLEET)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-FLEET · Owner: Principal Cloud/Platform Engineer with the Service/Reliability Engineer, the Product Security Engineer (EM-ENG-008), and EM-ENG-001. Detailed fleet-management software requirements decomposed from the System Requirements (SYS-030/031/032, SYS-133/134/135/139/141, SYS-165, SYS-172, SYS-176) and the Product Requirements (PRD-CLOUD 900, with PRD-FUN, PRD-SEC, PRD-SVC, PRD-REL), realising the reusable/ rechargeable multi-trip fleet economics (D1) and the readiness-gated deploy loop. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM platform manages a reusable fleet across its service life: enrolling and provisioning devices; transferring ownership; grouping and tagging; remotely configuring, diagnosing, rebooting, and pulling logs; monitoring health and heartbeat; maintaining a firmware inventory; reporting calibration and readiness compliance; capturing usage statistics and fleet analytics; managing fleet-level alerts; governing device lifecycle and decommissioning; and detecting stray, overdue, and unreported devices for recovery. It specifies the fleet-management behaviour and obligations that sit on top of, and depend on, the core cloud platform (SRS-CLOUD); it does not re-specify the underlying registry, ingest, RBAC, audit, or availability mechanisms owned there.
2. Scope
Device enrollment; fleet provisioning and field re-provisioning; ownership/tenant transfer; grouping and tagging; remote configuration, diagnostics, reboot, and log-pull; health monitoring and heartbeat; firmware inventory; calibration and readiness compliance reporting and deployment gating; usage statistics; fleet analytics; fleet alert management; lifecycle management and decommissioning; and stray/overdue/unreported detection. Excludes the core registry/identity, shadow/twin, ingest, RBAC/SSO, audit, storage, and availability/DR mechanisms (SRS-CLOUD — referenced as the substrate); OTA packaging/campaign internals (SRS-OTA); device-side readiness computation, calibration, and safe firmware application (SRS-DEV); cryptographic identity/keys and provisioning mechanisms (SRS-SEC, EM-ENG-008); the mobile onboarding client (SRS-MOB); and connectivity transport (SRS-CONN); each is referenced where fleet management interfaces with it.
3. Applicable Standards
21 CFR Part 11 / ALCOA+ (attributable, auditable fleet/lifecycle actions); EU GDP (data governance, device qualification/calibration expectations for monitoring equipment); GAMP 5 (CSV of the managed system); IEC 62443 concepts (secure device onboarding, ownership transfer, decommissioning); SOC 2 concepts (availability and change governance); ISO/IEC 27001 concepts (asset management). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The authoritative device registry, identity, shadow/twin, RBAC/SSO, audit trail, durable storage, and availability/DR are provided by SRS-CLOUD; fleet management is a governed set of workflows over that substrate and never bypasses its access control or audit.
• Devices report readiness (state of charge, calibration validity/date, firmware version, last-seen) and battery state-of-health, and compute their own readiness/verdict on-device (SYS-030/165); the platform consumes and aggregates these, and gates deployment on them.
• All remote fleet actions (configure, reboot, diagnostic, log-pull, firmware target, ownership transfer, decommission) are authenticated device commands delivered per SRS-CLOUD-005 and are audited per SRS-CLOUD-018.
• Firmware update packaging, signature verification, staged rollout, and fail-safe application are owned by SRS-OTA and the device (SYS-032/139); fleet management provides the inventory and targeting view.
• Cryptographic identity, key lifecycle, and secure provisioning/decommissioning mechanisms are owned by SRS-SEC (SYS-134/140/141).
5. Architecture Overview
Fleet management is a set of platform services layered on the core cloud (SRS-CLOUD). An enrollment/ provisioning service admits devices to the registry under an owning tenant and supports authenticated field re-provisioning across the reuse loop (D1). An ownership/organisation service transfers devices between tenants/sites and groups/tags them for bulk operation. A remote-operations service issues authenticated configure/reboot/diagnostic/log-pull commands and firmware targeting via the device shadow. A health/ heartbeat service tracks liveness, connectivity health, and readiness, feeding a compliance/readiness service that reports calibration and readiness status and gates deployment of unfit devices. A usage/ analytics service aggregates trip counts, cycle counts, and utilisation. A lifecycle service governs state transitions from enrollment through active use to decommissioning, and a recovery service flags stray, overdue, and unreported devices. Every fleet action is authenticated, role-authorised, and audited via the core platform.
---
7. Functional Requirements
7.1 Enrollment, provisioning & ownership
SRS-FLEET-001 — Device enrollment into the managed fleet
Requirement Statement | The fleet-management platform shall enroll each device into the managed fleet by creating or confirming its authoritative registry entry under an owning tenant, binding its cryptographic and durable identity, initial configuration, calibration status, and firmware version, and shall record the enrollment as an attributable, time-stamped audit event.
Rationale | A reusable fleet (D1) can only be managed, tracked, and made accountable if every device is enrolled under a known owner with a bound identity; an unenrolled device is invisible to fleet control and cannot contribute a trustworthy record.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-133 (attribution), SYS-134 (identity), SYS-031 (fleet visibility), PRD-CLOUD-001/006, PRD-SEC-006; → SRS-CLOUD-001/018, SRS-SEC. Standard: IEC 62443 secure onboarding, ISO/IEC 27001 asset management concepts (revision to confirm).
Notes | Registry/identity substrate owned by SRS-CLOUD-001; enrollment is the fleet workflow over it.
SRS-FLEET-002 — Fleet provisioning and field re-provisioning across the reuse loop
Requirement Statement | The fleet-management platform shall provision a device's configuration and operational parameters at enrollment and shall support authenticated re-provisioning of a returned or refurbished device between trips (updating configuration, ownership, and readiness state), without exposing private key material and recording each provisioning action for audit.
Rationale | The reusable/refurbish loop (D1) requires devices to be reconfigured and re-associated between trips; provisioning must be authenticated and attributable so provenance and readiness stay trustworthy across many trips.
Priority | Must
Source | Architecture
Verification Method | Test
Traceability | SYS-141 (field re-provisioning), SYS-121 (provenance), SYS-030 (readiness), PRD-CLOUD-001, PRD-SVC-004, PRD-SEC-009; → SRS-CLOUD-002/005, SRS-MOB-004, SRS-SEC. Standard: IEC 62443 provisioning concepts (revision to confirm).
Notes | Key provisioning mechanisms owned by SRS-SEC (SYS-141); mobile onboarding client per SRS-MOB-004.
SRS-FLEET-003 — Ownership and tenancy transfer
Requirement Statement | The fleet-management platform shall support transferring a device's ownership between tenants/sites through an authenticated, role-authorised, audited action that reassigns the device's registry ownership and future data scope while preserving the integrity and attribution of the historical record produced under prior ownership.
Rationale | Reusable devices move between customers, sites, and owners over their life; transfer must reassign control without cross-tenant leakage and without rewriting or orphaning the historical evidence.
Priority | Should
Source | Architecture
Verification Method | Test
Traceability | SYS-125 (ownership/isolation), SYS-135 (audit), SYS-133, PRD-CLOUD-009/011, PRD-SEC-005; → SRS-CLOUD-012/016/018. Standard: IEC 62443 ownership-transfer concepts, EU GDP data governance (revision to confirm).
Notes | Historical records remain owned/isolated per their originating tenant (SYS-125); transfer affects future scope and control.
SRS-FLEET-004 — Grouping and tagging for bulk fleet operations
Requirement Statement | The fleet-management platform shall allow authorised users to organise devices into groups and apply tags (e.g. site, lane, customer, hardware revision, firmware cohort) within their tenant scope, and shall support targeting fleet operations (configuration, firmware, reporting, queries) by group or tag.
Rationale | Managing a large reusable fleet requires bulk operation and filtering; groups/tags let operators target cohorts (e.g. a firmware campaign to one hardware revision) rather than device-by-device.
Priority | Should
Source | User Need
Verification Method | Demonstration
Traceability | SYS-031 (fleet visibility), SYS-125 (isolation), PRD-CLOUD-006, PRD-FUN-021; → SRS-CLOUD-013, SRS-FLEET-005/007.
Notes | Group/tag scoping never crosses tenant boundaries (SYS-125).
7.2 Remote operations & health
SRS-FLEET-005 — Remote configuration, diagnostics, reboot, and log-pull
Requirement Statement | The fleet-management platform shall allow an authorised user to remotely apply configuration, request diagnostics, trigger a reboot, and pull device logs from a targeted device or cohort, delivering each as an authenticated, at-most-once command via the device shadow and recording issuance, delivery, and result in the audit trail, without disrupting a live trip's committed record.
Rationale | A duty-cycled fleet must be diagnosable and correctable remotely; remote operations must be authenticated, attributable, confirmed, and non-destructive to the in-progress record.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-176 (reconciliation), SYS-028 (audit), SYS-113-line (diagnostics), PRD-CLOUD-006/008, PRD-CONN-010; → SRS-CLOUD-002/005/018, SRS-CONN-014/024. Standard: IEC 62443 concepts (revision to confirm).
Notes | Commands delivered per SRS-CLOUD-005; a reboot/config must never discard committed telemetry or an alert (SYS-021, SRS-CONN-007).
SRS-FLEET-006 — Health monitoring and heartbeat
Requirement Statement | The fleet-management platform shall monitor each device's health and liveness from its reported heartbeat, last-seen time, connectivity health, state of charge, and battery state-of-health, and shall present per-device and fleet-wide health status to authorised users, flagging degraded or unreported devices.
Rationale | Fleet operators must know at a glance which devices are healthy, which are degrading (e.g. battery fade over the reuse loop), and which have gone quiet; liveness is connectivity-bound and must be presented honestly.
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-030 (readiness/last-seen), SYS-031 (fleet status), SYS-165 (battery SoH over reuse), PRD-CLOUD-006, PRD-FUN-021; → SRS-CLOUD-002/008, SRS-FLEET-008/012.
Notes | Liveness is connectivity-bound (C7); an unreported device is flagged, not assumed healthy (SRS-FLEET-012).
7.3 Inventory, compliance & analytics
SRS-FLEET-007 — Firmware inventory
Requirement Statement | The fleet-management platform shall maintain an inventory of the firmware version in force on each device and across the fleet, present it by group/tag/cohort, and identify devices eligible for or requiring update, feeding the OTA orchestration and targeting.
Rationale | Controlled, observable fleet updates and vulnerability response (when a component CVE is disclosed) require an accurate view of what firmware is running where; without inventory, a campaign is blind.
Priority | Should
Source | Security
Verification Method | Test
Traceability | SYS-032/139 (firmware/OTA), SYS-145 (SBOM/affected releases), SYS-030, PRD-CLOUD-008, PRD-SEC-008; → SRS-CLOUD-005, SRS-OTA. Standard: NIST CSF supply-chain, IEC 62443 concepts (revision to confirm).
Notes | OTA campaign packaging/rollout owned by SRS-OTA; this provides the inventory/targeting view.
SRS-FLEET-008 — Calibration and readiness compliance reporting with deployment gating
Requirement Statement | The fleet-management platform shall report each device's calibration validity and date, firmware validity, charge/energy state, and overall readiness, shall produce fleet-wide readiness/ calibration compliance reports, and shall block deployment (arming/redeployment) of a device that fails readiness — including one whose battery state-of-health can no longer meet the endurance requirement — until it is serviced.
Rationale | An out-of-calibration, under-charged, or fade-limited device must not be deployed (D1 reuse loop, D3 cold-endurance risk); a worthless or misleading record is worse than none, so readiness gating is a safety-relevant control.
Priority | Must
Source | Safety
Verification Method | Test
Traceability | SYS-030 (readiness reporting/gating), SYS-165 (endurance over reuse / block-if-unfit), SYS-172 (drift/robustness over reuse), PRD-CLOUD-006, PRD-SVC-004, PRD-REL-002; → SRS-CLOUD-002/013, SRS-DEV (readiness computation), HAL deploy-unfit hazard (Stage 10). Standard: EU GDP equipment qualification concepts (revision to confirm).
Notes | Readiness/SoH computed on-device (SYS-030/165); the platform aggregates, reports, and gates. Safety-classed; reconcile with HAL.
SRS-FLEET-009 — Usage statistics and fleet analytics
Requirement Statement | The fleet-management platform shall capture per-device and fleet usage statistics (trip counts, charge/reuse cycle counts, active/idle time, utilisation, excursion and alert rates) and shall present fleet analytics scoped to the caller's tenant and role, derived from the authoritative records without modifying them.
Rationale | Reuse economics and total-cost-of-ownership management (D1), plus reliability tracking of cycle-driven battery fade (D3), depend on aggregate usage and analytics across the fleet.
Priority | Could
Source | User Need
Verification Method | Demonstration
Traceability | SYS-031 (fleet status), SYS-165 (cycle life), SYS-130 (historical), PRD-CLOUD-006/010, PRD-FUN-021; → SRS-CLOUD-014, SRS-FLEET-006.
Notes | Analytics are derived/read-only; the authoritative record is immutable except via audited action (SRS-CLOUD-014).
7.4 Alerts, lifecycle & recovery
SRS-FLEET-010 — Fleet alert management
Requirement Statement | The fleet-management platform shall present, filter, and manage device-fault, readiness, connectivity, and stray/overdue alerts at the fleet level (by group/tag/site), track their acknowledgement and resolution, and integrate with the platform notification/escalation engine, so that a fleet operator can triage fleet-health alerts distinctly from per-shipment excursion alerts.
Rationale | A fleet operator managing hundreds of reusable devices needs fleet-health alert triage (dead/stray/unfit devices) separate from, but coordinated with, the per-trip excursion alerting so neither overwhelms the other.
Priority | Should
Source | User Need
Verification Method | Test
Traceability | SYS-031 (fleet flags), SYS-012 (routing/escalation), SYS-030, PRD-CLOUD-003/006, PRD-FUN-021; → SRS-CLOUD-009, SRS-FLEET-012.
Notes | Excursion (safety) alerts route per SRS-CLOUD-009; this manages fleet-operational alerts.
SRS-FLEET-011 — Device lifecycle management and secure decommissioning
Requirement Statement | The fleet-management platform shall govern each device's lifecycle state (enrolled, provisioned, deployed/armed, in-service, returned, maintenance, retired/decommissioned) through authenticated, audited transitions, and shall support secure decommissioning that revokes the device's credentials and removes it from active service while preserving the integrity and attribution of its historical records.
Rationale | A reusable fleet (D1) has a full-life state model; controlled decommissioning must revoke a retired or lost device's ability to act while never destroying or invalidating the evidence it produced.
Priority | Should
Source | Security
Verification Method | Test
Traceability | SYS-140 (key revocation), SYS-135 (audit), SYS-125 (record retention/isolation), PRD-CLOUD-011, PRD-SEC-009; → SRS-CLOUD-012/018, SRS-SEC. Standard: IEC 62443 decommissioning concepts, ALCOA+ enduring (revision to confirm).
Notes | Credential revocation mechanisms owned by SRS-SEC (SYS-140); historical records retained per SYS-124/125.
SRS-FLEET-012 — Stray, overdue, and unreported device detection for recovery
Requirement Statement | The fleet-management platform shall detect and flag devices that are stray (location/behaviour inconsistent with an assigned trip), overdue (past expected return), or unreported (silent beyond an expected heartbeat window), and shall present them for recovery, distinguishing a silent device from a healthy one so that a lost or unreturned reusable asset is never assumed in-service.
Rationale | A reusable fleet must be trackable to control loss and total cost of ownership (D1); a device that goes silent or fails to return must be surfaced for recovery, not silently written off or assumed fine.
Priority | Should
Source | User Need
Verification Method | Demonstration
Traceability | SYS-031 (stray/overdue/unreported flagging), SYS-030 (last-seen), SYS-029 (location), PRD-CLOUD-006, PRD-FUN-021; → SRS-CLOUD-008, SRS-FLEET-006/010.
Notes | Location cadence is energy-constrained and RF-attenuated in the shipper (SYS-029, Evidence-Pending); detection uses best-available last-seen/location.
---
8. Non-functional Requirements
SRS-FLEET-013 — Fleet-scale performance and availability
Requirement Statement | The fleet-management services shall scale to the managed fleet size and shall present fleet status, health, and readiness within the platform's defined responsiveness and availability targets, degrading gracefully under partial platform outage rather than presenting stale data as current.
Rationale | Fleet oversight is only useful if it remains responsive and honest as the reusable fleet grows; stale-as-current status would mislead deploy and recovery decisions.
Priority | Should
Source | Performance
Verification Method | Test, Analysis
Traceability | SYS-175 (platform reliability), SYS-031, PRD-CLOUD-006/012; → SRS-CLOUD-020/021, SRS-PERF. Standard: SOC 2 availability concepts (revision to confirm).
Notes | Scaling/availability substrate owned by SRS-CLOUD; numeric targets held in SRS-PERF (evidence).
---
Error Handling
SRS-FLEET-014 — Safe handling of failed fleet operations
Requirement Statement | On a failed fleet operation (provisioning, remote command, transfer, firmware target, or decommission) the fleet-management platform shall leave the device and registry in a well-defined, recoverable state, record the failure with cause, and surface it to the operator, and shall never leave a device in an ambiguous ownership, credential, or lifecycle state or silently lose the intent of the failed operation.
Rationale | A half-completed transfer, provisioning, or decommission is a governance and safety risk; failures must resolve to a known, recoverable, auditable state, never an ambiguous one.
Priority | Should
Source | Reliability
Verification Method | Test
Traceability | SYS-176 (reconciliation), SYS-135 (audit), PRD-CLOUD-002/012; → SRS-CLOUD-005/018/023, SRS-FLEET-002/003/011.
Notes | Command semantics (at-most-once, confirmed) per SRS-CLOUD-005; failures are auditable.
---
Diagnostics
SRS-FLEET-015 — Fleet-wide diagnostic visibility
Requirement Statement | The fleet-management platform shall aggregate and present per-device and fleet-wide diagnostics (connectivity health, signal quality history, readiness failures, watchdog/recovery events, integrity/anomaly flags) to authorised users for triage, drawing on device- and connectivity-side diagnostics without altering any governed record.
Rationale | Triaging a large fleet requires aggregated diagnostics so operators can find and act on patterns (e.g. a coverage gap on one lane, a cohort with recurring recoveries) rather than inspecting devices one by one.
Priority | Should
Source | User Need
Verification Method | Demonstration
Traceability | SYS-113-line (device/connectivity diagnostics), SYS-169 (watchdog recovery events), SYS-129 (anomaly flags), PRD-CLOUD-006, PRD-CONN-010; → SRS-CONN-014, SRS-CLOUD-024, SRS-DIAG.
Notes | Underlying diagnostics owned by SRS-DEV/SRS-CONN; this is the fleet aggregation/presentation.
---
Security Considerations
SRS-FLEET-016 — Governed, authenticated, audited fleet actions
Requirement Statement | The fleet-management platform shall require authenticated, role-authorised access for every fleet action (enrollment, provisioning, transfer, grouping, remote operation, firmware targeting, lifecycle transition, decommissioning), shall confine every action to the caller's tenant scope, and shall record every action in the append-only audit trail, without ever creating an unauthenticated fleet-control path.
Rationale | Fleet management is a high-privilege control surface over a connected fleet; unauthorised remote configuration, transfer, or decommissioning could disarm monitoring, hijack, or brick devices, so every action must be governed and attributable.
Priority | Must
Source | Security
Verification Method | Test, Inspection
Traceability | SYS-136 (RBAC), SYS-135 (audit), SYS-125 (isolation), SYS-140 (revocation), PRD-CLOUD-009, PRD-SEC-005; → SRS-CLOUD-016/018/019, SRS-SEC. Standard: IEC 62443 concepts, OWASP ASVS, 21 CFR Part 11 access/audit (revision to confirm).
Notes | RBAC/SSO/audit mechanisms owned by SRS-CLOUD/SRS-SEC; this states the fleet-action posture.
---
Verification Requirements
SRS-FLEET-017 — Fleet-management verification
Requirement Statement | The fleet-management platform shall be verified by enrollment/provisioning and field re-provisioning tests, ownership-transfer and tenant-isolation tests, grouping/tagging and bulk-targeting tests, remote-operation (config/reboot/diagnostic/log-pull) tests, health/heartbeat and stray/overdue detection tests, firmware-inventory and readiness/calibration compliance-gating tests, lifecycle/decommissioning (credential-revocation with record-preservation) tests, and fleet-action access-control/audit tests, with results traced to the requirements above.
Rationale | The reuse-loop economics, readiness gating, and governed remote control claims are only credible if verified across the enrollment→deploy→return→refurbish→decommission lifecycle the fleet actually follows.
Priority | Must
Source | Security
Verification Method | Test, Demonstration
Traceability | SYS-030/031/032, SYS-135/136/140/141, SYS-165/172/176, PRD-CLOUD-001/006/008/009/011, PRD-SVC-004, PRD-REL-002; → SRS-VER. Standard: GAMP 5 CSV, IEC 62443 concepts (revision to confirm).
Notes | Readiness-gating verification reconciles with HAL (deploy-unfit hazard, Stage 10).
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-…) and Product Requirement (PRD-…), and forward to verification. Principal upstream anchors: SYS-030 (readiness reporting/gating), SYS-031 (fleet visibility, stray/overdue), SYS-032 (firmware), SYS-133/134/135 (identity/attribution/audit), SYS-140/141 (key revocation, provisioning), SYS-165 (endurance over reuse), SYS-172 (robustness over reuse), SYS-176 (device↔cloud reconciliation); PRD-CLOUD-001/006/008/009/011, PRD-FUN-020/021, PRD-SVC-004, PRD-SEC-005/006/008/009, PRD-REL-002. Cross-module: registry/identity/shadow/ingest/RBAC/SSO/audit/storage/ availability substrate → SRS-CLOUD; cryptographic identity/key lifecycle/provisioning/revocation → SRS-SEC (EM-ENG-008); OTA packaging/campaign/rollout → SRS-OTA; device readiness/calibration/safe firmware/buffer → SRS-DEV; connectivity/diagnostics transport → SRS-CONN; mobile onboarding client → SRS-MOB; numeric responsiveness/availability → SRS-PERF; deploy-unfit and delayed-alert reconciliation → SRS-SAFE / HAL (Stage 10). Evidence-Pending items inherited from the substrate (availability SLA, RPO/RTO, retention periods, location cadence in-container) are carried as gaps, not invented, and close on their associated tests.
SRS-900 — Data (module SRS-DATA)
Smart Cold Chain Monitor (SCM) · Stage 07 — Subsystem / Detailed Software Requirements · Module code SRS-DATA · Owner: Embedded Software Engineer (EM-ENG-007) with the Cloud/Platform Engineer, Product Security Engineer (EM-ENG-008), and Quality/Regulatory Engineer (EM-ENG-006). Detailed data-software requirements decomposed from the Data/Records/Time/Storage System Requirements (SYS-119…131) and Product Requirements (PRD-FUN records/export/time, PRD-CLOUD ingest/retention/ownership/export/DR, PRD-SEC integrity/ at-rest), realising a structured, serialisable, sized, retained, protected, time-referenced, and provable evidentiary record across device → transport → cloud (D2, ALCOA+ / 21 CFR Part 11 / EU GDP). This module defines the data structures and handling; it consumes the cryptographic mechanisms owned by SRS-SEC and the transport owned by SRS-CONN. Field set per the Software Requirements Library Standard (00-governance/software-requirements-standard.md).
---
1. Purpose
Specify how the SCM's evidentiary data is schematised (telemetry, configuration, metadata), time-referenced, serialised, compressed, encrypted, minimised for privacy, owned and isolated per tenant, retained, backed up, restored, exported, imported, validated for quality and consistency, and stored for the long term — such that a record written by any tier is deterministically and losslessly reconstructable by another and its ALCOA+ properties (attributable, legible, contemporaneous, original, accurate, complete, consistent, enduring, available) survive every boundary crossing.
2. Scope
Versioned telemetry sample schema; versioned configuration / trip-profile schema; record metadata and provenance; the trusted time base recorded with data; defined serialisation for on-the-wire and at-rest; lossless compression for the constrained uplink; data encryption at rest and in transit (via SRS-SEC); privacy and data-minimisation; data ownership and tenant isolation; configurable retention; cloud backup; restore and disaster recovery; on-demand export; documented import/integration format; ingest-side data validation, quality and consistency checks; and historical/long-term storage and queryability. Storage sizing for a full 90-day trip is traced to SYS-127. Excludes the storage-medium/NVM mechanics and buffer management (SRS-DEV), the cryptographic primitives themselves (SRS-SEC), the bearers/transport (SRS-CONN), OTA packaging (SRS-OTA), and cloud microservice/infra design (SRS-CLOUD). Those are referenced where the data software interfaces with them.
3. Applicable Standards
21 CFR Part 11 / ALCOA+ (electronic records, attributable/original/accurate/enduring); EU GDP (data governance, retention); GDPR (personal-data minimisation/retention, where applicable); ISO 8601 (UTC time representation); ISO/IEC 27001 concepts (data protection); NIST 800-series concepts (data at rest/in transit). Cited by designation only; revision to confirm — no clause number is asserted as fact.
4. Assumptions
• The device provides non-volatile storage and buffer management (SRS-DEV); this module defines what is written and how it is structured, not the medium. Storage sizing for a full 90-day trip is Evidence-Pending and owned by SYS-127 (computed by analysis/test, not invented here).
• Cryptographic integrity (tamper-evidence), signing, and encryption keys are provided by SRS-SEC; this module binds those artefacts into the data structures and specifies data-side handling.
• The trusted UTC time base and its drift/sync disclosure are provided by SRS-DEV (RTC) and SRS-CONN (network time); this module specifies how time is carried in the record.
• The cloud provides idempotent, ordered ingest and durable storage (SRS-CLOUD); this module specifies the data validation, retention, isolation, export/import, and historical-storage behaviour.
• Concrete export/import formats (e.g. PDF + CSV/JSON), retention periods, and RPO/RTO targets are Evidence-Pending and close in the Regulatory/SLA stages (SYS-124/126/128).
5. Architecture Overview
Every measurement enters a versioned telemetry schema — value bound to channel, unit, resolution, a per-value quality flag, and its trusted timestamp — and the governing configuration/trip-profile schema is version-bound to the trip it governs. Provenance metadata stamps each record with device identity, firmware, calibration, active profile, and the authenticated actor/time of each lifecycle action. A single defined serialisation carries data losslessly on the wire and at rest, self-describing via its schema version, and lossless compression shrinks it for the constrained uplink without altering any value, flag, timestamp, or integrity artefact. Encryption (in transit via SRS-CONN/SRS-SEC, at rest via SRS-SEC) protects confidentiality; data-minimisation bounds location/personal-data exposure. In the cloud, data is owned per tenant and isolated, retained under configurable policy, backed up and restorable to an integrity- and attribution-preserving state, exportable on demand and importable via a documented format, all while ingest-side validation flags (never silently drops or repairs) malformed, unattributable, integrity-failing, or inconsistent records. Historical storage keeps records decodable, integrity-verifiable, and queryable across schema evolution for the full retention period.
---
7. Functional Requirements
7.1 Schemas & structure
SRS-DATA-001 — Versioned telemetry sample schema
Requirement Statement | The data software shall represent every telemetry sample using a defined, versioned schema in which each measurement is bound to an explicit channel identifier (internal temperature, external-probe temperature, relative humidity, 3-axis acceleration), a per-value data-quality flag, an engineering unit and resolution, and the sample's trusted timestamp, and shall record the schema version with the data.
Rationale | A record is defensible only if each value is unambiguously typed, attributed to a channel, and self-describing; an explicit schema version prevents silent mis-parsing as fields evolve (SYS-119).
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-119 (telemetry schema), SYS-007 (quality flag), PRD-FUN-001/017, PRD-CLOUD-002; → SRS-DATA-005/013, SRS-DEV (capture). Standard: ALCOA+ legible/accurate; EU GDP (revision to confirm).
Notes | Carries the per-value quality flag from SYS-007 into the record; channel set matches the sensing subsystem.
SRS-DATA-002 — Versioned configuration / trip-profile schema
Requirement Statement | The data software shall represent the trip profile and device configuration (thresholds, excursion logic, logging interval, upload cadence, shock threshold, alert routing) using a defined, versioned schema, shall bind the exact profile version in force to every trip record it governs, and shall migrate persisted configuration across schema-version changes without loss.
Rationale | A verdict is meaningful only against the profile that produced it; the active configuration is part of the record, so it must be schematised and version-bound so a reviewer can reconstruct exactly what limits applied (SYS-120).
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-120 (config schema), PRD-FUN-014, PRD-CLOUD-004/002; → SRS-OTA-010 (preserve across update), SRS-DEV (config store). Standard: ALCOA+ consistent (revision to confirm).
Notes | Config migration across updates reconciled with SRS-OTA-010.
SRS-DATA-003 — Record metadata and provenance
Requirement Statement | The data software shall stamp every trip record with provenance metadata — device serial / durable identity, firmware version, calibration identity and validity date, active profile version, and the authenticated actor and time for each lifecycle action (configure, arm, close, export) — such that the record is attributable to its originating device, software, calibration state, and responsible party.
Rationale | ALCOA+ attributable/original: a disposition must be traceable to a specific device, firmware, calibration, and actor; missing provenance renders the record inadmissible (SYS-121).
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-121 (metadata/provenance), PRD-FUN-017, PRD-SEC-002, PRD-CLOUD-001/007; → SRS-SEC-018 (bind/attribution), SRS-DATA-013. Standard: 21 CFR Part 11 attributable; ALCOA+ (rev. to confirm).
Notes | Actor/action binding uses the authenticated identity from SRS-SEC-012/018.
SRS-DATA-004 — Trusted time base recorded with data
Requirement Statement | The data software shall reference all timestamps to UTC via the trusted, monotonic time base, and shall record, with each record set, the time source, the synchronisation status at the time of logging, and the disclosed bound on clock drift, so that the temporal accuracy of every sample and event is knowable to a reviewer.
Rationale | Correct excursion timing and defensible attribution require a trusted clock whose drift and sync state travel with the data; an undisclosed clock offset would misplace an excursion in time (SYS-131, SYS-006).
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-131 (trusted time in data), SYS-006, PRD-FUN-019, PRD-ARCH-007, PRD-SEC-002; → SRS-DEV-021 (RTC/discipline), SRS-CONN-017 (network time). Standard: 21 CFR Part 11 accurate; ISO 8601; ALCOA+ contemporaneous (revision to confirm).
Notes | Time discipline owned by SRS-DEV/SRS-CONN; this fixes how time metadata is carried in the record.
7.2 Serialisation, format & compression
SRS-DATA-005 — Defined lossless serialisation for on-the-wire and at-rest data
Requirement Statement | The data software shall serialise telemetry, events, configuration, and metadata in a defined, documented format for both on-the-wire transfer and at-rest storage on device and in the cloud, such that a record written by one tier is deterministically and losslessly reconstructable by another, and the format shall carry the schema version needed to decode it.
Rationale | Store-and-forward (D2) crosses device NVM, LTE-M/Wi-Fi/BLE bearers, and cloud storage; without a single defined lossless format, integrity and tamper-evidence cannot survive the boundary crossings (SYS-122).
Priority | Must
Source | Architecture
Verification Method | Test, Analysis
Traceability | SYS-122 (serialisation), PRD-CLOUD-002, PRD-FUN-012, PRD-SEC-001; → SRS-DATA-006/013, SRS-CONN-016, SRS-SEC-018. Standard: ALCOA+ enduring/consistent (revision to confirm).
Notes | Underpins idempotent ordered ingest (SYS-020) and the integrity artefact carried by SRS-SEC-018.
SRS-DATA-006 — Lossless compression for the constrained uplink
Requirement Statement | The data software shall support lossless compression of buffered data for uplink such that no logged value, event, gap marker, quality flag, timestamp, or integrity artefact is altered or lost by the compress/decompress round-trip, and shall record sufficient information to decompress each payload independently.
Rationale | The energy and LTE-M-in-container link budgets (D2/D3) reward smaller payloads, but compression must be lossless — a lossy transform of an evidentiary record would breach ALCOA+ accuracy; and independent decompressibility protects resumable/idempotent upload (SYS-123).
Priority | Should
Source | Performance
Verification Method | Test
Traceability | SYS-123 (compression), PRD-FUN-013, PRD-CONN-004, PRD-PERF-006; → SRS-CONN-016 (bandwidth), SRS-CONN-007 (resume). Standard: — (evidence: no ratio asserted here).
Notes | Compression is bound to the connectivity window by SRS-CONN-016; must be lossless without exception.
SRS-DATA-007 — Data encryption at rest and in transit
Requirement Statement | The data software shall ensure all controlled data is encrypted in transit (via the mutually authenticated TLS channel) and at rest on the device and in the cloud using the mechanisms and keys owned by SRS-SEC, such that no controlled record, configuration, or personal/location datum is written to persistent media or transmitted in cleartext.
Rationale | A reusable device and a multi-tenant cloud both hold data that must not be exposed if a medium, backup, or channel is compromised; encryption complements the integrity/tamper-evidence of SRS-DATA-013 (SYS-138/137).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-138 (at-rest), SYS-137 (in transit), PRD-SEC-007/006, PRD-CLOUD-011; → SRS-SEC-006/007/ 008/017. Standard: ISO/IEC 27001 concepts; TLS (revision to confirm).
Notes | Cryptographic mechanisms and keys owned by SRS-SEC; this binds them to the data-handling path. Encryption addresses confidentiality; integrity is SRS-DATA-013.
7.3 Privacy, ownership & isolation
SRS-DATA-008 — Privacy and data-minimisation
Requirement Statement | The data software shall limit collection, transmission, and retention of location and any personal data to the minimum needed for the operational purpose (asset recovery, attribution, and cold-chain evidence), shall not retain personal/location data beyond its purpose, and shall support authorised, audited deletion aligned with applicable privacy law for the markets of use.
Rationale | The device monitors cargo, not people, but location traces and operator identities can create disproportionate privacy exposure; data-minimisation bounds that exposure and supports customer data- ownership/retention control (SYS-148, D4).
Priority | Should
Source | Regulatory
Verification Method | Inspection
Traceability | SYS-148 (privacy/minimisation), PRD-SEC-010, PRD-CLOUD-011, PRD-REG-009; → SRS-DATA-009/010, SRS-SEC-016. Standard: GDPR data-minimisation; EU GDP; privacy-by-design (revision to confirm).
Notes | Evidence-Pending — regional privacy-law mapping to confirm (SYS-148, D4).
SRS-DATA-009 — Data ownership and tenant isolation
Requirement Statement | The data software shall treat each customer's records and configuration as owned by that customer, shall isolate each tenant's data so that no tenant can access, infer, or affect another tenant's records, and shall apply retention, export, and deletion only within the owning tenant's authorised scope.
Rationale | Regulated customers require ownership and provable isolation of their cold-chain evidence; cross-tenant leakage would be both a security breach and a data-integrity failure (SYS-125).
Priority | Must
Source | Security
Verification Method | Test
Traceability | SYS-125 (ownership/isolation), PRD-CLOUD-011/009, PRD-SEC-005/007; → SRS-SEC-008/014, SRS-CLOUD (multi-tenancy). Standard: EU GDP data governance; ISO/IEC 27001 (revision to confirm).
Notes | Isolation depends on RBAC (SRS-SEC-014) and at-rest key separation (SRS-SEC-008).
7.4 Retention, backup, restore & recovery
SRS-DATA-010 — Configurable data retention
Requirement Statement | The data software shall retain trip records, events, and audit data for a configurable retention period, shall not delete or expire retained records except through an audited, authorised action, and shall record the applicable retention policy with the data.
Rationale | EU GDP and Part 11 impose record-retention obligations that vary by market and customer; retention must be configurable and deletion must never be silent or unauthorised (SYS-124).
Priority | Must
Source | Regulatory
Verification Method | Test, Inspection
Traceability | SYS-124 (retention), PRD-CLOUD-011, PRD-CLOUD-007, PRD-SEC-004; → SRS-DATA-008/012, SRS-SEC-019 (audit). Standard: EU GDP; 21 CFR Part 11 record retention (revision to confirm).
Notes | Evidence-Pending — retention periods per market/customer, set in the Regulatory stage (SYS-124).
SRS-DATA-011 — Cloud backup and durability
Requirement Statement | The data software shall maintain durable, recoverable backups of all ingested records, audit trails, and configuration, shall protect backups at rest with encryption, and shall preserve the tamper-evidence and attribution of backed-up data so that a backup is as defensible as the primary store.
Rationale | The cloud is where the record becomes authoritative evidence; a component or site failure must not lose it, and a backup that loses integrity/attribution is not an evidentiary backup (SYS-128).
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-128 (backup/DR), PRD-CLOUD-012/002, PRD-REL-006; → SRS-DATA-012, SRS-SEC-008/018. Standard: ALCOA+ enduring; ISO/IEC 27001 (revision to confirm).
Notes | Backup encryption ties SRS-SEC-008; integrity artefacts preserved per SRS-SEC-018.
SRS-DATA-012 — Restore and disaster recovery to an integrity-verified state
Requirement Statement | The data software shall be able to restore records, audit trails, and configuration from backup to a consistent, integrity-verified, attribution-preserving state within defined recovery-point and recovery-time objectives following a component or site failure, without silent loss or corruption, and shall verify integrity and attribution on restore rather than merely restoring availability.
Rationale | Restore must re-establish tamper-evidence and attribution, not just data presence; a restore that cannot prove the record's integrity produces inadmissible evidence (SYS-128).
Priority | Must
Source | Reliability
Verification Method | Test, Analysis
Traceability | SYS-128 (restore/DR), PRD-CLOUD-012, PRD-REL-006; → SRS-DATA-011/013, SRS-CLOUD (DR). Standard: ALCOA+ enduring/consistent (revision to confirm).
Notes | Evidence-Pending — RPO/RTO targets to confirm (SYS-128, SLA).
7.5 Export, import & validation
SRS-DATA-013 — Ingest-side data validation, quality and consistency checks
Requirement Statement | On ingest the data software shall validate each incoming record against its declared schema, verify integrity and attribution, check chronological consistency and de-duplication, and flag — never silently drop or repair — any record that is malformed, out-of-schema, unattributable, integrity-failing, or inconsistent, recording the anomaly in the trip record and audit trail.
Rationale | Integrity preservation across an intermittent link (SYS-020) requires the ingest boundary to detect corruption and gaps rather than paper over them; silent repair would fabricate evidence, violating ALCOA+ original/accurate (SYS-129).
Priority | Must
Source | Regulatory
Verification Method | Test
Traceability | SYS-129 (validation/quality), SYS-020 (idempotent ingest), PRD-CLOUD-002, PRD-SEC-001, PRD-FUN-017; → SRS-SEC-018, SRS-CONN-007, SRS-DATA-001/003/005. Standard: ALCOA+; EU GDP (revision to confirm).
Notes | Flags, never fixes; consumes the integrity artefact from SRS-SEC-018 and the schema from SRS-DATA-001.
SRS-DATA-014 — On-demand record export in an admissible, integrity-preserving format
Requirement Statement | The data software shall export a complete trip record on demand in a documented, human- and machine-readable format that preserves samples, events, gap markers, quality flags, provenance metadata, time base, and tamper-evidence, such that the record's ALCOA+ properties survive export.
Rationale | Evidence must move to where release/acceptance decisions and disputes live (ERP/QMS/partner systems) without losing integrity; an integrity-preserving export is that mechanism (SYS-126, SYS-027).
Priority | Must
Source | Regulatory
Verification Method | Test, Demonstration
Traceability | SYS-126 (export), SYS-027, PRD-FUN-018, PRD-CLOUD-005, PRD-REG-003; → SRS-DATA-003/013, SRS-CLOUD (export). Standard: 21 CFR Part 11 records; ALCOA+ (revision to confirm).
Notes | Evidence-Pending — concrete formats (e.g. PDF + CSV/JSON) to confirm (SYS-126). Export is a governed, audited action (SRS-SEC-014/019).
SRS-DATA-015 — Documented import / integration format
Requirement Statement | The data software shall provide a documented import/integration format for enterprise/OEM systems to ingest trip records and configuration, preserving samples, events, gaps, quality flags, provenance, time base, and tamper-evidence, and shall validate imported data against schema and integrity on entry, flagging any record that fails.
Rationale | Records must move both out of and into the platform without losing integrity; a documented, validated interchange format lets partner systems exchange evidence without corrupting it (SYS-126).
Priority | Should
Source | Architecture
Verification Method | Test, Demonstration
Traceability | SYS-126 (import/integration), PRD-CLOUD-010/005, PRD-REG-003; → SRS-DATA-013/014, SRS-CLOUD (APIs). Standard: 21 CFR Part 11; ALCOA+ (revision to confirm).
Notes | Imported data is validated exactly as ingested device data (SRS-DATA-013); API auth per SRS-SEC.
7.6 Storage sizing & historical storage
SRS-DATA-016 — Non-volatile storage sizing for a full 90-day trip
Requirement Statement | The data software shall size the on-device data footprint — sample and event log at the minimum (worst-case densest) logging interval, plus all events, gaps, provenance metadata, and integrity artefacts, with a defined design margin — to fit within the non-volatile storage capacity for a full 90-day armed trip without overwriting any unsent or unacknowledged record.
Rationale | Store-and-forward (D2) can require an entirely offline 90-day trip to be held locally; the data footprint is a function of sample size (SRS-DATA-001), compression (SRS-DATA-006), and interval, and the buffer must never run out or overwrite unsent data (SYS-127/018).
Priority | Must
Source | Reliability
Verification Method | Analysis, Test
Traceability | SYS-127 (storage sizing, Evidence-Pending), SYS-018 (buffer), PRD-FUN-012, PRD-PERF-006, PRD-REL-005; → SRS-DEV (NVM/buffer), SRS-DATA-001/006, SRS-OTA-010. Standard: — (evidence).
Notes | Evidence-Pending — the required byte budget is not invented here; it is computed by analysis and confirmed by test, and is the sizing figure cross-referenced by SYS-018/127. Storage medium/ buffer management owned by SRS-DEV.
SRS-DATA-017 — Historical / long-term storage and queryability
Requirement Statement | The data software shall retain historical trip records for the full configured retention period in a form that remains readable, integrity-verifiable, and queryable (by device, trip, time range, and excursion/event) for release, audit, trend, and dispute purposes, notwithstanding schema version evolution.
Rationale | Retained records are only useful if they remain decodable and findable years later; schema versioning must not strand old data (SYS-130, ALCOA+ enduring/available).
Priority | Should
Source | Regulatory
Verification Method | Test, Demonstration
Traceability | SYS-130 (historical storage), SYS-124, PRD-CLOUD-005/011/006; → SRS-DATA-001/002/010, SRS-CLOUD (storage/search). Standard: ALCOA+ enduring/available; EU GDP (revision to confirm).
Notes | Schema-version awareness (SRS-DATA-001/002) keeps old records decodable across evolution.
---
8. Non-functional Requirements
SRS-DATA-018 — Deterministic round-trip integrity across tiers
Requirement Statement | The data software shall guarantee that any record written on the device and transferred through serialisation, compression, encryption, transport, and cloud storage is reconstructable byte-for-byte in its logical content (values, flags, timestamps, metadata, integrity artefacts), such that a cross-tier round-trip introduces no alteration, reordering, or loss.
Rationale | The whole evidentiary claim rests on the record surviving the device→transport→cloud path unchanged; any non-determinism in the data pipeline would undermine tamper-evidence and ALCOA+ consistency (SYS-122/129).
Priority | Must
Source | Architecture
Verification Method | Test, Analysis
Traceability | SYS-122, SYS-129, SYS-020, PRD-CLOUD-002, PRD-SEC-001; → SRS-DATA-005/006/007/013, SRS-SEC-018, SRS-CONN-007. Standard: ALCOA+ consistent/original (revision to confirm).
Notes | This is the module's governing invariant; verified end-to-end by SRS-DATA-019.
---
Error Handling
SRS-DATA-019 — No silent data loss, drop, or repair
Requirement Statement | On any data-handling error — serialisation/deserialisation fault, compression failure, decryption failure, schema mismatch, or integrity/attribution failure — the data software shall preserve the affected record, flag the anomaly in the trip record and audit trail, and surface it for triage, and shall never silently discard, truncate, or auto-repair an evidentiary record.
Rationale | A silently dropped or "repaired" record fabricates or destroys evidence; every data error must degrade to flagged-and-preserved, never to lost or altered data (SYS-129/021, ALCOA+).
Priority | Must
Source | Reliability
Verification Method | Test
Traceability | SYS-129, SYS-021 (buffer-full behaviour), PRD-CLOUD-002, PRD-FUN-017; → SRS-DATA-013, SRS-CONN-023, SRS-SEC-027. Standard: ALCOA+ original/complete (revision to confirm).
Notes | Aligns with the no-silent-failure invariant across SRS-CONN/SRS-DEV/SRS-SEC.
---
Diagnostics
SRS-DATA-020 — Data-quality and integrity reporting
Requirement Statement | The data software shall report per-trip data-quality diagnostics — gap count and extent, flagged/implausible samples, failed integrity/attribution checks, de-duplication events, and schema versions encountered — to authorised fleet and audit tooling, so that the completeness and defensibility of a record are inspectable without manual reconstruction.
Rationale | A reviewer must be able to see at a glance whether a record is complete and integral; surfaced data-quality diagnostics support honest disposition and audit (SYS-129, SYS-017).
Priority | Should
Source | User Need
Verification Method | Test, Demonstration
Traceability | SYS-129, SYS-017 (honest incomplete indication), PRD-CLOUD-006, PRD-FUN-017; → SRS-DATA-013, SRS-DIAG. Standard: ALCOA+ complete (revision to confirm).
Notes | Feeds the honest incomplete-monitoring indication (SYS-017) and fleet dashboards.
---
Security Considerations
SRS-DATA-021 — Data-side confidentiality, integrity and governance
Requirement Statement | The data software shall ensure that every controlled datum is, throughout its lifecycle, confidentiality-protected (encryption, SRS-SEC), integrity-protected and attributable (tamper-evidence, SRS-SEC-018), governed by tenant ownership and RBAC (SRS-DATA-009 / SRS-SEC-014), and handled under data-minimisation and retention controls — with no data-handling path that bypasses these protections.
Rationale | Data security is only as strong as its weakest handling path; the data pipeline must not create an unprotected shortcut around encryption, integrity, isolation, or retention (SYS-125/132/138/148).
Priority | Must
Source | Security
Verification Method | Analysis, Test
Traceability | SYS-125, SYS-132, SYS-138, SYS-148, PRD-SEC-001/005/007/010, PRD-CLOUD-011; → SRS-SEC-006/ 008/014/016/018, SRS-DATA-007/008/009. Standard: ISO/IEC 27001; 21 CFR Part 11; ALCOA+ (revision to confirm).
Notes | Cryptographic and access mechanisms owned by SRS-SEC; this states the data-software governance posture.
---
Verification Requirements
SRS-DATA-022 — Data verification for schema, round-trip, retention and recovery
Requirement Statement | The data software shall be verified by schema-conformance and version-migration tests, lossless serialisation/compression round-trip tests, cross-tier deterministic-reconstruction tests, ingest-validation (flag-not-repair) tests, tenant-isolation tests, retention/deletion (audited) tests, backup/restore integrity-and-attribution tests, export/import fidelity tests, storage-sizing analysis against SYS-127, and historical-queryability tests, with results traced to the requirements above.
Rationale | The ALCOA+/defensible-record claims are only credible if verified end-to-end across the data pipeline under the boundary crossings and failure modes the record actually meets (SYS-122/127/128/129).
Priority | Must
Source | Regulatory
Verification Method | Test, Analysis, Inspection
Traceability | SYS-119…131 (esp. SYS-122/127/128/129), PRD-CLOUD-002/005/011/012, PRD-SEC-001/007, PRD-FUN-017/018; → SRS-VER, SRS-SEC-030. Standard: GAMP 5 CSV; 21 CFR Part 11; ALCOA+ (revision to confirm).
Notes | Storage sizing (SYS-127), retention periods (SYS-124), export formats (SYS-126), and RPO/RTO (SYS-128) are Evidence-Pending; this fixes the verification obligation that closes them.
---
Traceability
Every requirement above traces upward to at least one System Requirement (SYS-119…131, plus SYS-006/017/018/ 020/021/027) and Product Requirement (PRD-FUN-001/012/013/014/017/018/019, PRD-CLOUD-002/005/010/011/012, PRD-SEC-001/002/004/005/006/007/010, PRD-PERF-006, PRD-REL-005/006, PRD-REG-003/009, PRD-ARCH-007). Principal upstream anchors: SYS-119/120/121 (schemas & provenance), SYS-131 (time in data), SYS-122/123 (serialisation/ compression), SYS-124/125 (retention/ownership/isolation), SYS-127 (storage sizing, Evidence-Pending — cross-referenced by SRS-DATA-016), SYS-128 (backup/restore/DR), SYS-129 (validation/quality), SYS-126 (export/import), SYS-130 (historical storage). Cross-module: cryptographic integrity, at-rest/in-transit encryption, and access control → SRS-SEC; on-device NVM, buffer, and RTC → SRS-DEV; resumable transport and network time → SRS-CONN; data preservation across firmware update → SRS-OTA; ingest, multi-tenancy, storage, search, DR, and APIs → SRS-CLOUD; numeric performance/storage budgets → SRS-PERF / SYS-127; record-integrity hazards → SRS-SAFE / HAL (Stage 10); CSV/regulatory posture → SRS-REG. Evidence-Pending items — storage sizing (SYS-127), retention periods (SYS-124), export/import formats (SYS-126), RPO/RTO (SYS-128), and regional privacy-law mapping (SYS-148) — are carried as gaps, not invented, and close on their associated analyses/tests and SRS-DATA-022.

## 8 System Architecture
System Architecture
The blueprint: how the parts, information, power and behaviour are organised.
↩ Contents
In this section:
• System Architecture — Smart Cold Chain Monitor (SCM)
• Architecture Decision Records (ADRs) — Smart Cold Chain Monitor (SCM)
• Communication Architecture — Smart Cold Chain Monitor (SCM)
• Component List — Smart Cold Chain Monitor (SCM)
• Data-Flow & Control-Flow Architecture — Smart Cold Chain Monitor (SCM)
• Deployment Architecture — Smart Cold Chain Monitor (SCM)
• Fault-Management & Diagnostic Architecture — Smart Cold Chain Monitor (SCM)
• Functional Decomposition — Smart Cold Chain Monitor (SCM)
• Interface Control — Smart Cold Chain Monitor (SCM)
• System Architecture — Logical & Physical Architecture — Smart Cold Chain Monitor (SCM)
• Power / Energy Architecture — Smart Cold Chain Monitor (SCM)
• State / Mode Model — Smart Cold Chain Monitor (SCM)
System Architecture — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 08 — System Architecture · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Status: Authored — instantiated, product-specific.
This stage defines the SCM's logical and physical architecture, its functional and component decomposition, the interface-control overview, data/control/power/communication/fault/deployment architectures, the state/mode model, and the Architecture Decision Records for the D1–D4-driven choices. It realises the System Requirements (Stage 06) and Subsystem/SRS requirements (Stage 07) as an architecture, and is the basis for the Design Outputs (Stage 09, ORCH-013). Based on 00-concept/concept-and-decisions.md (D1–D4) and the Stage-05 design-architecture (AD-1…AD-8).
Architectural thesis: the SCM is an offline-first edge node with opportunistic connectivity (D2), built around one overriding constraint — survive 90 days at −30 °C without dying or losing data (D3, the top program risk) — with a duty-cycled energy governor as the central mechanism and a split trust boundary (edge no-login read vs platform governance) as the security spine.
---
1. Architecture element ID scheme
Element | Prefix | Range | Where
Functional decomposition | FUN-### | FUN-001…040 | [functional-decomposition.md](functional-decomposition.md)
Component list | CMP-### | CMP-001…055 | [component-list.md](component-list.md)
Interface-control | INT-### | INT-001…035 | [interface-control.md](interface-control.md)
Architecture decisions | ADR-### | ADR-001…012 | [architecture-decision-records.md](architecture-decision-records.md)
Subsystems (physical): S1 Enclosure/Mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile.
---
2. Document set
# | File | Contents | Diagrams
1 | [logical-physical-architecture.md](logical-physical-architecture.md) | Architectural intent; logical layers; physical three-tier + device domains; S1–S9 decomposition; layering rationale; drivers→mechanisms; neighbour flags | Mermaid logical↔physical map
2 | [functional-decomposition.md](functional-decomposition.md) | 40 functions (FUN-001…040) decomposing F1–F10, each allocated to subsystems and traced to SYS/FWR/SRS | Mermaid decomposition tree
3 | [component-list.md](component-list.md) | 55 components (CMP-001…055) — HW, FW, cloud/mobile services — traced to functions & requirements | —
4 | [interface-control.md](interface-control.md) | 35 interfaces (INT-001…035), typed PWR/DATA/RF/THRM/MECH, mapped to I-01…I-12 and IFR/SYS | master register
5 | [data-flow-control-flow.md](data-flow-control-flow.md) | Data-flow (sensor→cloud, offline & BLE paths); control-flow (excursion-alert loop, arm/pre-arm); ALCOA+ lineage | 3 Mermaid (flowchart + sequenceDiagram)
6 | [power-energy-architecture.md](power-energy-architecture.md) | Energy-governor intent; power tree; 90-day −30 °C energy architecture (illustrative, Evidence-Pending, validate by test); charge architecture; energy-allocation policy | Mermaid power tree
7 | [state-mode-model.md](state-mode-model.md) | Mode catalogue (Shipping/Sleep/Sampling-Logging/Alerting/Charging/OTA/Fault-Safe/Provisioning-Arming); transitions & invariants | Mermaid stateDiagram-v2
8 | [communication-architecture.md](communication-architecture.md) | LTE-M primary; store-and-forward; Wi-Fi/BLE offload; alert-path priority; PSM/eDRX; roaming; RF-in-container (Evidence-Pending) | Mermaid multi-bearer path
9 | [fault-management-diagnostics.md](fault-management-diagnostics.md) | Fault taxonomy; detection (POST/BIST/watchdog/plausibility); responses; safe-state; honest-gap principle; diagnostics interfaces | Mermaid fault→detect→response
10 | [deployment-architecture.md](deployment-architecture.md) | Device + cloud + mobile deployment; topology; CI/CD & environments; trust-split; data-residency (D4) | Mermaid deployment diagram
11 | [architecture-decision-records.md](architecture-decision-records.md) | ADR-001…012 for the D1–D4 / AD-1…AD-8 choices with alternatives, consequences, affected requirements, validation hooks | —
Total: 11 architecture files + this index. 12 Mermaid diagrams across the set (state machine, data flow, control flow, power tree, functional tree, comms, fault, deployment, logical↔physical).
---
3. How the architecture resolves the key tensions (severity-first)
Driver / risk | Architectural mechanism | ADR | Status
D3 — 90-day endurance at −30 °C | Duty-cycled energy governor; radios gated hard; low-temp chemistry | ADR-003, ADR-004 | Evidence-Pending — validate by test (PRD-VER-004)
D3 — cold charge safety | Charge-inhibit-below-0 °C interlock (HW+FW); self-heat evaluation | ADR-004 | Evidence-Pending / Preliminary
RF-in-container link | LTE-M primary + Wi-Fi/BLE offload; antenna strategy; store-and-forward tolerance | ADR-010 | Evidence-Pending — RF-in-container test (PRD-VER-005)
Last mile as black box | Offline-first store-and-forward; on-device verdict | ADR-002, ADR-007 | Baselined
Missed freeze | Two-sided detection with internal + external probe | ADR-005 | Preliminary (HAL)
Record integrity / disposition error | Tamper-evident append-only log; trusted time; split trust boundary | ADR-006 | Baselined / Preliminary
OTA bricking / unmonitored fleet | Signed, resumable, fail-safe OTA; never-brick invariant | ADR-008 | Baselined
Reuse TCO / reverse logistics | Reusable-fleet lifecycle (return→clean→recharge→triage→re-provision) | ADR-009 | Baselined
---
4. Evidence posture (honesty)
Per 00-governance/evidence-standard.md and hallucination-policy.md, every architectural claim carries a evidence label (Standard / Cited / Judgment / Evidence-Pending). The 90-day −30 °C energy budget is presented as an illustrative model, explicitly Evidence-Pending and validated by test — no battery capacity or chemistry is asserted as fact. The RF-in-container link budget and antenna strategy are the principal open architectural risk (Evidence-Pending), closed by the RF-in-container test. Safety-relevant architecture (charge-inhibit, safe-state, honest-gap, two-sided detection) is Preliminary and reconciled with the Hazard Analysis Library (HAL, Stage 10).
---
5. Traceability (chain position)
Backward: architecture elements (FUN/CMP/INT/ADR) reference the System (SYS-###) and Subsystem/SRS (HWR/MER/FWR/IFR/SRS-<CODE>) requirements they realise, and through them the PRD/UN/D1–D4. Forward: the architecture is the input to Design Outputs (DO-<CODE>-NNN, Stage 09), Hazard Analysis (Stage 10), ERFKL (Stage 11), and Verification/Validation (Stages 13/14). The complete bidirectional matrix is assembled by ORCH-018 / EM-ENG-012 (Stage 18).
Flow: PRD (05) → SYS (06) → Subsystem/SRS (07) → Architecture (08, this stage) → Design Outputs (09) → Hazard Analysis (10) / ERFKL (11) → Risk-Control Requirements (12) → Verification (13) / Validation (14). Outputs feed Design Outputs (Stage 09) and ORCH-018 / EM-ENG-012.
Architecture Decision Records (ADRs) — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · Owner: Chief Systems Engineer (EM-ENG-001) · Traceability: ORCH-018 / EM-ENG-012.
This file records the architecturally significant decisions for the SCM — the choices that shape subsystem allocation, interfaces, energy, connectivity, safety, security, and lifecycle, and that are expensive to reverse later. Each ADR resolves a foundational decision (D1–D4, see [00-concept/concept-and-decisions.md](../00-concept/concept-and-decisions.md)) and/or a key architectural decision (AD-1…AD-8) and the concept tensions C1–C9 captured in [05-product-requirements/design-architecture.md](../05-product-requirements/design-architecture.md).
Evidence & honesty posture (CLAUDE.md rules 1, 3; evidence-standard). Every decision below is labelled Standard / Cited / Judgment / Evidence-Pending. Risk is scored severity-first (Action Priority, not raw RPN). Two decisions — ADR-004 (low-temperature battery / charge safety) and ADR-010 (RF-in-container connectivity) — are the program's top open risks; they are Accepted (Evidence-Pending validation) and no battery capacity, chemistry, cold-endurance, or RF-link figure is asserted as fact anywhere in this file. Those numbers are set as requirements in Stage 06 and closed by test. Safety requirements remain Preliminary until reconciled with the Hazard Analysis Library (HAL, Stage 10); each carries its validation hook below.
Verification/validation vocabulary. SYS-### = system requirement (Stage 06). PRD-VER-### = product verification (does the design meet the requirement — test/analysis/inspection). PVL/PVaL = the product-verification-level / product-validation-level evidence event (validation: does the device meet the user need in use). HA-###/RCR-### = hazard record / risk-control requirement to be assigned in the HAL.
ADR | Title | Drives | Status
ADR-001 | Reusable, rechargeable multi-trip device | D1 | Accepted
ADR-002 | Offline-first store-and-forward architecture | D2 / AD-1 | Accepted
ADR-003 | Duty-cycled energy governor as the endurance mechanism | D3 / AD-2 | Accepted
ADR-004 | Low-temperature chemistry + charge-inhibit-below-0 °C + self-heat | D3 / AD-3 | Accepted (Evidence-Pending validation)
ADR-005 | Two-sided detection with internal + external probe | AD-4 | Accepted
ADR-006 | Split trust boundary: edge no-login read vs platform governance | AD-5 | Accepted
ADR-007 | Layered information: on-device verdict → cloud detail | AD-6 | Accepted
ADR-008 | Signed, resumable, fail-safe OTA | AD-7 | Accepted
ADR-009 | Reusable-fleet lifecycle loop | AD-8 / D1 | Accepted
ADR-010 | LTE-M primary bearer + Wi-Fi/BLE offload + RF antenna strategy | D2 / D4 | Accepted (Evidence-Pending validation)
ADR-011 | Multi-market regulatory posture | D4 | Accepted
ADR-012 | On-device OLED + single glove-button HMI | AD-6 / AD-5 | Accepted
---
ADR-001 — Reusable, rechargeable multi-trip device (not single-trip disposable)
Status | Accepted
Context | The concept explicitly includes USB-C charging, a rechargeable battery, and OTA firmware updates (concept §1), plus a 90-day endurance target and a cloud-dashboard fleet model — a bundle of features that only makes economic and functional sense for a device used across many trips. The forcing question (D1) was whether the SCM is a single-trip disposable logger (like many blister/PCM excursion tags) or a reusable multi-trip instrument. C6 states the underlying tension: reuse lowers per-trip cost (better TCO) but creates reverse-logistics, cleaning, and re-provisioning cost that disposables avoid.
Decision | The SCM is a reusable, rechargeable, multi-trip device whose whole lifecycle (return → clean → recharge → readiness-triage → re-provision) is a first-class design input. Every downstream subsystem is sized for repeated use, not one-shot operation.
Alternatives considered
1. Single-trip disposable logger — rejected. Contradicts the concept's own USB-C / rechargeable / OTA feature set; wastes the cloud-fleet investment; and a disposable cannot be recalibrated between trips, undermining measurement trust over time. It would also multiply lithium-battery e-waste and transport (UN 38.3) burden per shipment.
2. Hybrid: reusable body with a disposable/replaceable battery module — rejected for the baseline, retained as a serviceability option (see ADR-009). It complicates sealing (S1) and UN 38.3 handling and does not remove the reverse-logistics need that D1 accepts anyway.
Consequences
• Positive: lowest cost-per-trip at fleet scale; supports recalibration and OTA over a long service life; justifies the cloud registry/dashboard; reduces per-trip battery waste.
• Negative: introduces a return/clean/recharge/re-provision logistics loop the operator must run (ADR-009); demands capacity-retention-over-cycles engineering (SYS-091) and cleaning-chemical compatibility of the enclosure (S1).
• Residual risk: capacity fade over reuse cycles can cause a mid-life dead-in-cold event (feeds ADR-004 and MR-RISK-001); a device redeployed out-of-calibration produces untrustworthy records (controlled by ADR-009).
Requirements affected | SYS-088 (rechargeable source), SYS-091 (capacity retention over reuse cycles), SYS-093 (USB-C recharge to readiness), SYS-097 (SoC/SoH gauge), SYS-030 (readiness); PRD-ELEC-001/003, PRD-REL-002, PRD-FUN-020; UN-SVC, UN-FUN-017.
Validation / how we'll know | cycle-life test demonstrating SYS-086 endurance is still met (or reliably self-flagged and redeployment blocked) after the specified number of charge cycles (SYS-091, Test); depot readiness-triage demonstration (ADR-009). Closes as PRD-VER against SYS-091.
Evidence | Judgment (D1; based on the concept's own feature set).
---
ADR-002 — Offline-first store-and-forward architecture (not live cellular streaming)
Status | Accepted
Context | D2 and AD-1. The concept asks for continuous monitoring, excursion alerts, and scheduled uploads over 90 days. Continuous live cellular telemetry (e.g., a radio session every few minutes for 90 days) is energy-prohibitive against the D3 budget (Judgment, confirmed as the dominant load in the power budget). C2 (live data vs battery) and C7 (real-time expectation vs no-connectivity reality — the device spends much of a trip inside a metallized shipper with poor or no link) are the forces. A record that is lost because the radio drained the battery, or an architecture that stops working when the link drops, both fail the user's core need: a complete, auditable cold-chain record.
Decision | The SCM is an offline-first edge node: it logs locally at a configurable interval (default 5–10 min), uploads on a configurable cadence (default hourly), pushes an immediate excursion/shock alert when a bearer is available, and buffers to non-volatile storage and resumes with idempotent (exactly-once), in-order ingest across any number of connectivity interruptions. Live telemetry is treated as upside, never a dependency.
Alternatives considered
1. Live cellular streaming — rejected. Violates the 90-day energy budget (radios are the dominant load) and fails whenever the link is absent inside the shipper; it would trade the evidentiary record for an availability promise the RF environment cannot keep.
2. Pure local logger, read-out only at destination (no cellular) — rejected. Removes in-transit excursion alerts, eliminating the rescue window that is a primary user benefit; the concept explicitly requires cellular alerting.
Consequences
• Positive: the record survives connectivity loss and radio-off periods; energy is spent on batched, opportunistic uploads instead of constant sessions; alerts are captured the instant they are detected even when delivery is deferred.
• Negative: alert latency is connectivity-bound — an alert can be captured but not delivered until a bearer appears; this is a recognised hazard (delayed/missing alert).
• Residual risk: an excursion detected deep inside a shipper with no egress is recorded but its alert is delayed; mitigated by alert-path prioritisation (SYS-112) and multi-bearer offload (ADR-010), and carried to the HAL as the delayed-alert hazard.
Requirements affected | SYS-111 (store-and-forward, buffer/resume/no-loss), SYS-108 (opportunistic multi-bearer), SYS-112 (alert-path prioritisation), SYS-089 (radio gating), SYS-013/SYS-017/SYS-018/SYS-020 (offline capture/record integrity); PRD-CONN-004/008/009, PRD-FUN-012/013, PRD-ARCH-005.
Validation / how we'll know | buffer/resume/idempotency test proving zero record loss or duplication across repeated link interruptions (SYS-111, Test); alert-priority test proving an alert is delivered ahead of backlog on first available bearer (SYS-112). Delayed-alert hazard reconciled in HAL (Stage 10).
Evidence | Judgment / Standard (D2; ALCOA+/EU GDP completeness for the buffer-and-resume integrity, revision to confirm).
---
ADR-003 — Duty-cycled power-state governor as the endurance mechanism
Status | Accepted
Context | D3 / AD-2. The headline promise — 90 days at −30 °C from a single rechargeable charge — is the top program constraint. Chemistry alone cannot deliver it (see ADR-004); the architecture must earn endurance by controlling demand. C1/C2/C3 (cost, battery life, GPS appetite vs endurance) all resolve here. Radios and GNSS are the dominant loads; if they are allowed to run freely, the budget cannot close.
Decision | Endurance is delivered by an autonomous, duty-cycled power-state governor in the MCU firmware (S3) that holds the device in its lowest viable deep-sleep state between scheduled activities and hard-gates every high-drain load (LTE-M, Wi-Fi, BLE, GNSS) so that no radio or GNSS activity occurs outside its scheduled, energy-budgeted window — the single exception being a queued excursion alert. The 90-day cold budget is decomposed into a per-load-class energy allocation (sleep/quiescent, sensing, compute, storage, HMI, cellular, Wi-Fi, BLE, GNSS) that the governor enforces. The headline number is validated by test, not asserted.
Alternatives considered
1. Always-on / lightly-managed power — rejected. Cannot meet a 90-day budget with radios and GNSS as the dominant loads; quiescent-plus-active drain would exhaust any realistically sized cell.
2. Fixed low sample/upload rate with no dynamic gating — rejected. Simpler, but cannot adapt radio/GNSS energy to coverage or backlog and cannot shed discretionary functions gracefully as energy depletes (SYS-098); it wastes energy on doomed radio attempts inside the shipper.
Consequences
• Positive: makes the 90-day/−30 °C target designable and accountable (per-class budget with reserve); enables graceful degradation (shed uploads/GNSS before core sensing/logging); bounds the worst-case current profile.
• Negative: increases firmware complexity and adds a timing/latency cost — a device asleep between windows discovers and reacts to some events later; requires careful wake-logic and watchdog design so a governor bug cannot silently stop monitoring.
• Residual risk: per-class allocations (especially radio and GNSS energy) are unresolved pending the RF-in-container test (ties ADR-010); quiescent current and cold self-discharge are unresolved (SYS-092). Budget could fail to close if radio energy is worse than modelled.
Requirements affected | SYS-089 (governor + hard radio gating), SYS-090 (per-load energy-budget allocation), SYS-092 (quiescent/sleep current at −30 °C), SYS-086/087 (endurance & usable-capacity budget), SYS-098 (graceful degradation / function shedding), SYS-117 (GNSS duty-cycled within budget); PRD-ELEC-008, PRD-ARCH-004, PRD-PERF-005.
Validation / how we'll know | cold current-profile test measuring per-load energy and quiescent drain at −30 °C (SYS-089/090/092, Test); the 90-day cold-endurance test that closes SYS-086 (MR-VAL-006 / PVaL); depletion test proving discretionary-function shedding order (SYS-098).
Evidence | Judgment (D2/D3; quantified by current-profile and endurance test — Evidence-Pending on the numbers, but the mechanism is baselined).
---
ADR-004 — Low-temperature chemistry + charge-inhibit-below-0 °C + self-heat evaluation
Status | Accepted (Evidence-Pending validation) — top program risk (D3), validated by test
Context | D3 / AD-3 — the program's #1 technical risk (MR-RISK-001). The concept demands a rechargeable source (D1) that operates for 90 days across −30 °C to +60 °C. Standard Li-ion loses usable capacity in deep cold and cannot be charged below 0 °C — charging a cold cell drives lithium plating, an internal-short and thermal-runaway precursor. This is a genuine, unresolved physics tension: a safety hazard created by the operating envelope itself, not merely a performance shortfall. Severity-first, the outcome (fire / thermal event, or a silent dead-in- cold false-negative record) is Critical.
Decision | Adopt a three-part battery-safety architecture: (1) select a low-temperature-tolerant rechargeable chemistry (candidate classes: low-temp Li-ion or LTO-class — final chemistry to be selected against its cold datasheet, not asserted here); (2) the BMS shall hard-inhibit charging whenever cell temperature is below 0 °C (or the higher cell-defined safe limit) and clearly indicate the charge-inhibited state; (3) evaluate a self-heat-before-charge capability that warms a cold-returned cell to a permitted charge temperature — provided it can never override the charge-inhibit and never exceed safe thermal limits. No capacity, cold-endurance, or chemistry figure is asserted as fact; all are set as requirements and closed by test.
Alternatives considered
1. Standard Li-ion + duty-cycling only — rejected. Loses too much usable capacity at −30 °C and forbids depot recharge of cold returns without a warming step; high risk of missing the endurance target and of unsafe cold charging.
2. Primary (non-rechargeable) lithium cell (e.g., Li-SOCl₂, strong cold performance) — rejected for the baseline. It contradicts the reusable/rechargeable decision (D1, ADR-001), removes USB-C recharge and depot turnaround, and worsens per-trip battery waste. Retained only as a fallback if no rechargeable chemistry can close the cold budget.
3. Allow low-temperature charging with a reduced ("trickle") current — rejected on safety. Even reduced-rate cold charging risks lithium plating; severity-first, the charge-inhibit is non-negotiable and must not be defeated (including by self-heat).
Consequences
• Positive: directly attacks both legs of the D3 tension (cold endurance and charge safety); the charge-inhibit removes the dominant Li-ion cold-charging hazard; self-heat (if feasible) preserves fast depot turnaround without unsafe charging.
• Negative: low-temp chemistries may cost more, offer lower energy density, or have limited supply; self-heat costs energy and thermal-design complexity; charge-inhibit means a cold returned device cannot be charged until it warms, adding depot dwell time.
• Residual risk (KEY): whether any rechargeable chemistry meets 90 days at −30 °C with margin is unproven (SYS-086/087/088); cold self-discharge and cycle fade (SYS-091/092); self-heat feasibility (SYS-155/SYS-095) — all Evidence-Pending, carried as the top hazard to the HAL.
Requirements affected | SYS-088 (low-temp rechargeable source), SYS-086/087 (cold endurance & usable-capacity budget), SYS-094 & SYS-155 (charge-inhibit below safe cell temp), SYS-095 (self-heat-before-charge, under evaluation), SYS-096 & SYS-156 (OV/UV/OC/OT/short protection, no fire/thermal-runaway), SYS-091 (capacity retention over cycles), SYS-100 (charge-inhibit state indication), SYS-160 (dead-in-cold prevention); PRD-ELEC-001/002/005/006/007, PRD-REL-001/003, PRD-SAFE-007/008/012.
Validation / how we'll know | cold-capacity test on the selected cell (closes SYS-088); 90-day cold-endurance test at −30 °C (closes SYS-086, MR-VAL-006 / PVaL); charge-inhibit test proving charging cannot proceed below the safe limit (SYS-094/SYS-155); fault-injection / abuse test to IEC 62368-1 / UN 38.3 concepts proving no fire/thermal-runaway (SYS-096/SYS-156); self-heat feasibility test that must also prove it cannot defeat the inhibit (SYS-095). All reconciled with HAL (HA-###/RCR-### for the cold-charge and dead-in-cold hazards).
Evidence | Evidence-Pending (chemistry, capacity, and cold endurance — D3, MR-RISK-001); charge-inhibit and protection are Standard in principle (IEC 62368-1 / IEC 62133 concepts, final thresholds per cell datasheet; revision to confirm).
---
ADR-005 — Two-sided detection with internal + external probe (freeze detection as a safety requirement)
Status | Accepted (safety controls Preliminary pending HAL)
Context | AD-4. Many freeze-sensitive vaccines are damaged by freezing as much as by heat, yet freeze damage is silent and historically under-detected (Cited: WHO/UNICEF freezing reporting; STN-002, UN-SAFE-002). A single internal sensor measures the device's own cavity, not the product; and detecting only the high-side (overheat) excursion leaves the low-side (freeze) hazard uncontrolled. Severity-first: a missed freeze leads to a spoiled dose administered — a top loss cause and one of the two directions of disposition harm (SYS-149).
Decision | Implement two-sided detection — flagging freezing (low-side) excursions with detection reliability at least equal to overheating (high-side) — using both an internal temperature sensor and an external probe, where the external probe represents the product zone rather than the device cavity. Freeze detection is engineered as a safety risk control, not an optional feature.
Alternatives considered
1. Single internal sensor, high-side-biased detection — rejected. Under-detects the silent freeze hazard and measures the wrong thermal mass (device, not product); this is exactly the historical failure mode the product exists to fix.
2. External probe only — rejected. Loses the internal reference needed for sensor-health cross-checks, cold-junction/self-heat compensation, and detecting a disconnected/faulted probe; two-sided sensing also underpins no-false-reassurance (SYS-151).
Consequences
• Positive: controls the missed-freeze hazard; product-zone representation improves disposition accuracy in both directions; dual sensors enable sensor-fault detection and honest gap-marking.
• Negative: the external probe adds a sealing/retention challenge (S1: sealed cable entry, probe form and placement) and a use-error surface (probe misplacement) that must be controlled by design (SYS-153); adds calibration and BOM cost.
• Residual risk: probe placement/attachment variability in the field; freeze-onset detection timing target unvalidated until the two-sided detection test; alarm balance (SYS-152) must avoid both false freeze alarms and missed freezes.
Requirements affected | SYS-150 (two-sided/freeze detection as risk control), SYS-149 (do-not-mislead disposition anchor), SYS-151 (no false reassurance on incomplete record), SYS-152 (alarm balance), SYS-153 (use-error protection incl. probe misplacement); PRD-SAFE-002, PRD-FUN-006/007, and S4 sensing / S1 enclosure allocations.
Validation / how we'll know | two-sided detection validation test proving low-side detection reliability ≥ high-side for onset/direction/duration (SYS-150, Test); probe-seal and probe-retention verification (S1); summative usability covering probe placement (SYS-153). Reconciled with HAL as the missed-freeze hazard (HA-###/RCR-###).
Evidence | Cited (WHO/UNICEF freezing reporting; STN-002; revision to confirm) + Standard (ISO 14971 severity-dominant framing).
---
ADR-006 — Split trust boundary: edge no-login read vs platform governance
Status | Accepted
Context | AD-5, resolving C9 (security lockdown vs open, frictionless edge). The last-mile user is minimally trained, gloved, offline, and time-pressured and must be able to read the accept/reject verdict instantly — a login at the edge would be a use-error and safety hazard. Yet configuration, the evidentiary data, and access control must be locked down for data-integrity and regulatory reasons (21 CFR Part 11-style audit, RBAC/SSO, EU GDP). These two needs pull in opposite directions and must be separated architecturally.
Decision | Adopt a split trust boundary. The on-device verdict is readable with no login (the frictionless edge). Everything that changes state or exposes controlled data — configuration, provisioning, arming, controlled-data read, firmware — requires mutual authentication over whatever bearer is used (BLE/Wi-Fi/LTE-M), and the platform holds governance: device registry, RBAC/SSO, and an append-only Part 11-style audit trail. Read of the public verdict is not a controlled action; write and controlled-read are.
Alternatives considered
1. Fully locked device (login required even to read the verdict) — rejected. Creates a use-error/safety hazard for gloved offline users and defeats the product's core "glance and decide" value; a forgotten/failed login at the point of care could block a correct disposition.
2. Fully open device (no auth even for configuration) — rejected. Unauthorised configuration could disarm monitoring or corrupt a trip profile, and offers no data-integrity/audit basis for GDP/Part 11 — a compliance and safety failure.
Consequences
• Positive: frictionless, safe edge reading and governed configuration/data; the attack surface (provisioning/pairing) is authenticated while the low-value public verdict stays open; supports auditable access control.
• Negative: two authorization models to design, test, and document; requires a clear definition of exactly what is "public verdict" vs "controlled data" so the boundary cannot leak controlled information through the open read.
• Residual risk: the full threat model and key-management design are Evidence-Pending (cybersecurity engineer); a poorly drawn public/controlled boundary could expose more than the verdict.
Requirements affected | SYS-110 (secure provisioning / authenticated pairing over any bearer), SYS-154 (prevent unmonitored trip start / mis-arm), SYS-115 (region config integrity), SYS-118 (trusted time); plus platform RBAC/SSO/audit (S8); PRD-SEC-005/006, PRD-CONN-006, PRD-SAFE-006; UN-SEC-003/004.
Validation / how we'll know | security test proving configuration/controlled-read requires mutual auth on every bearer and that public-verdict read exposes no controlled data (SYS-110, Test); mis-arm/void-state test (SYS-154); Part 11 audit-trail inspection on the platform. Threat model to be completed and reconciled (cybersecurity, Evidence-Pending).
Evidence | Standard (21 CFR Part 11 access control/audit; TLS 1.2+/1.3; Bluetooth Core pairing — revisions to confirm) + Judgment on the public/controlled split; threat model Evidence-Pending.
---
ADR-007 — Layered information: on-device verdict → cloud detail
Status | Accepted
Context | AD-6, resolving C5 (rich data vs a simple, unambiguous verdict). The edge user needs one unambiguous decision — accept / reject / quarantine — under cold, glare, glove, and time pressure; but auditors, QA, and regulators need the full record (per-sample temperature/ humidity/shock, excursion timing, location, connectivity diagnostics, ALCOA+ metadata). Putting the full record on the tiny cold-legible display would overwhelm the edge user; hiding the record would fail auditability.
Decision | Layer the information. The device presents an unambiguous verdict (and honest status, including any monitoring gap — never an unqualified clean-pass on an incomplete record). The full evidentiary record lives in the cloud and in exports as the authoritative source. Edge = decision; cloud = detail and proof.
Alternatives considered
1. Full detail on the device — rejected. Not legible or safe on a small display under field conditions; invites misreading and slow decisions; the device cannot be the authoritative, access-controlled, auditable store the platform is.
2. Verdict only, no rich cloud record — rejected. Destroys auditability and the ability to investigate/justify a disposition; fails EU GDP / Part 11 data-integrity expectations.
Consequences
• Positive: fast, low-error edge decisions; complete auditable record where governance lives; the small display's energy and legibility budget is spent only on the verdict + honest status.
• Negative: the verdict logic must be trustworthy on its own (a wrong verdict is a direct disposition hazard, SYS-149); requires the record to reach the cloud reliably (ties store-and- forward, ADR-002) so the authoritative detail is actually available.
• Residual risk: an edge verdict shown while the corresponding detailed record is still buffered offline — mitigated by honest on-device status and by ADR-002's guaranteed eventual delivery.
Requirements affected | SYS-149 (do-not-mislead), SYS-151 (no false reassurance on incomplete record), SYS-150 (verdict reflects two-sided detection), and the S6 verdict display / S8 authoritative record allocations; PRD-FUN-006/010/011/017, PRD-UX, PRD-CLOUD export; UN-VIS-007, STN-004/027.
Validation / how we'll know | verdict-correctness validation (SYS-149, Test/Analysis); test that an incomplete/gapped trip never shows an unqualified accept (SYS-151); export/record completeness inspection against ALCOA+ (S8). Summative usability that the verdict is understood correctly at the edge (SYS-153).
Evidence | Judgment (AD-6; UN-VIS-007, STN-004/027) + Standard (ALCOA+/Part 11 for the authoritative record; revision to confirm).
---
ADR-008 — Signed, resumable, fail-safe OTA
Status | Accepted
Context | AD-7. A reusable fleet (ADR-001) must be updatable in the field and at depot over lossy, low-power links (ADR-010) without wasting energy on failed restarts, without bricking the fleet, and without losing or corrupting the evidentiary record across an update or reset. Firmware is also a security attack surface: an unsigned or tampered image could disarm monitoring or exfiltrate data. Three forces converge: security (authenticity), energy/reliability (resumability over bad links), and safety (fail-safe, no bricking, no data loss).
Decision | OTA is signed, resumable, and fail-safe. Images are verifiable-genuine (cryptographically signed; only verified images are applied — tie to secure boot / secure element), transfers are resumable over LTE-M/Wi-Fi/USB-C without restarting from the beginning, and the update process cannot brick the device (e.g., A/B or fallback image) and preserves the trip data intact across update/reset, returning to a defined safe state.
Alternatives considered
1. Unsigned or single-image in-place OTA — rejected. Unsigned firmware is an unacceptable security risk; single-image in-place update risks a bricked, unrecoverable device if power or link drops mid-write — catastrophic for a fielded reusable fleet.
2. USB-C-only update at depot (no OTA) — rejected. Cannot patch a fleet already in the field (including security fixes) and forces physical handling for every update; retained only as a fallback transport alongside OTA.
Consequences
• Positive: the fleet is securely and safely updatable over the real (lossy) network; no bricking and no record loss across updates; energy not wasted on failed restarts.
• Negative: requires dual-image/fallback storage overhead, signing-key management and a secure element, and a staged/verify-before-apply flow — more firmware and cloud-distribution complexity.
• Residual risk: signing-key management is part of the Evidence-Pending threat model (cybersecurity); an OTA that changes measurement/verdict behaviour must be re-validated before fleet rollout.
Requirements affected | SYS-114 (resumable firmware-update transport), SYS-032 (signing / verification / fail-safe application), SYS-158 (fail-safe on fault incl. OTA, no data loss), SYS-159 (defined safe state), SYS-099 (brown-out-safe committed writes); PRD-SEC-008, PRD-CONN-011, PRD-FUN-022, PRD-ARCH-008, PRD-REL-004/005; UN-SEC-005, UN-SAFE-010.
Validation / how we'll know | OTA test injecting mid-transfer link loss and mid-apply power loss proving resume, no-brick, and record preservation (SYS-114/SYS-158, Test/fault-injection); signature-verification test rejecting an unsigned/tampered image (SYS-032); brown-out integrity test (SYS-099).
Evidence | Standard (secure-update / signed-firmware principles, secure boot; revision to confirm) + Judgment on the fail-safe (A/B) mechanism; key management Evidence-Pending.
---
ADR-009 — Reusable-fleet lifecycle loop (return → clean → recharge → readiness-triage → re-provision)
Status | Accepted
Context | AD-8 / D1, resolving C6 (reuse TCO vs reverse-logistics cost). ADR-001 makes the SCM reusable; that only pays off if the between-trips loop is engineered, not left to chance. A device redeployed out-of-calibration, under-charged, on stale firmware, or unclean produces untrustworthy records or dies mid-trip — silently converting the reuse benefit into a safety and data-integrity liability.
Decision | Define and design for an explicit reusable-fleet lifecycle loop: return → clean → recharge → readiness triage (charge state, calibration validity, firmware version) → re-provision (new trip profile), with deployment blocked when a device is out-of-calibration, battery-health-failed, or otherwise not trip-ready. Readiness is a first-class, platform-tracked device state.
Alternatives considered
1. Ad-hoc reuse (recharge and redeploy, no triage gate) — rejected. Lets out-of-calibration, degraded-battery, or stale-firmware devices back into service — untrustworthy records and mid-life dead-in-cold events.
2. Full depot recalibration every trip — rejected as the baseline. Maximally safe but expensive and slow; replaced by calibration-validity triage (recalibrate only when validity lapses), which the readiness gate enforces.
Consequences
• Positive: guarantees only trip-ready, in-calibration, healthy-battery, current-firmware devices deploy; ties battery-health self-flagging (SYS-091) and readiness (SYS-030) into an enforceable gate; supports fleet TCO.
• Negative: requires depot process, tooling, and platform state tracking; cleaning imposes enclosure chemical-compatibility and re-sealing constraints (S1); adds dwell time (esp. cold returns that must warm before charge — ties ADR-004).
• Residual risk: human/process error at the depot bypassing triage — mitigated by the platform deploy-block; cleaning-chemical compatibility of seals is a mechanical-engineering open item.
Requirements affected | SYS-091 (capacity retention / self-flag & block redeploy), SYS-030 (readiness), SYS-097 (SoC/SoH gauge feeding triage), SYS-093 (USB-C recharge), SYS-114 (firmware currency), SYS-154 (re-provision/arm correctly); PRD-FUN-020, PRD-REL-002, PRD-ELEC-003/004; UN-SVC, UN-FUN-017; D1.
Validation / how we'll know | readiness-triage demonstration showing deploy is blocked for out-of-cal / battery-failed / stale-firmware units (SYS-030/091, Demonstration/Test); cleaning-chemical compatibility test on the enclosure and re-seal verification (S1); cycle-life test underpinning the health-flag threshold (SYS-091).
Evidence | Judgment (D1; UN-SVC, UN-FUN-017); cleaning/seal compatibility flagged to mechanical engineer (Evidence-Pending).
---
ADR-010 — LTE-M primary bearer + Wi-Fi/BLE offload + RF-in-container antenna strategy
Status | Accepted (Evidence-Pending validation) — principal open architectural risk
Context | D2 / D4, and the principal open architectural unknown: the device operates inside a metallized, insulated shipper (RF-hostile). The link budget, antenna strategy, and in-container GNSS performance are not yet quantified — RF attenuation inside such a shipper can exceed a low-power WAN radio's fade margin, so egress may fail. Simultaneously the device must be low-power (D3), multi-market (D4, different band plans), and must cross carrier/national boundaries mid-trip. No link-margin, attenuation, or delivery-success figure is asserted as fact.
Decision | Use LTE-M (3GPP Cat-M1) as the primary low-power wide-area bearer, with Wi-Fi and BLE as opportunistic offload, under an energy-aware, fail-over multi-bearer selection policy (prefer the lowest-energy path that meets the cadence, never require any single bearer). Pursue an antenna strategy of integral antennas for LTE-M/Wi-Fi/BLE/GNSS with provision for a qualified external-antenna option (via an IP-sealed connector) as a live architectural branch to be exercised if and only if the integral antenna cannot close the in-container link budget. Offline-first (ADR-002) means loss of all bearers degrades to buffered capture, never data loss; GNSS is duty-cycled, fails soft, and is never a safety function.
Alternatives considered
1. Single-bearer cellular (LTE-M only), integral antenna fixed — rejected. Brittle inside the shipper and at coverage gaps; no depot Wi-Fi backlog drain; commits to an antenna solution before the RF risk is measured.
2. NB-IoT primary — rejected as primary. Lower throughput and higher latency for buffered backlog and firmware; LTE-M's mobility and PSM/eDRX behaviour suit duty-cycled roaming lanes better. (Retained as a possible regional fallback.)
3. Committing now to an external antenna (or to integral-only) — rejected/deferred. The integral-vs-external decision is Evidence-Pending and must be driven by the RF-in-container test, not pre-judged; premature commitment risks either a sealing/UX penalty or a link failure.
Consequences
• Positive: resilient, energy-aware delivery across intermittent coverage and borders; depot Wi-Fi and local BLE drain backlog cheaply; keeps the external-antenna option open without committing to its sealing/cost penalty prematurely.
• Negative: multi-bearer + roaming + per-market band plans add modem/SIM and firmware complexity; an external-antenna branch imposes an IP-sealed connector on the enclosure (S1) and radiated-power re-conformance (regulatory).
• Residual risk (KEY): whether any antenna strategy closes the in-container link budget is unproven (SYS-104/SYS-109); in-container GNSS fix availability is unproven (SYS-117); per-market carrier coverage and the SIM/eSIM roaming strategy are unresolved (SYS-116). A delayed/undelivered alert is a safety hazard (ties ADR-002, SYS-112).
Requirements affected | SYS-104 (in-container link budget & delivery success), SYS-109 (antenna strategy incl. external option), SYS-117 (GNSS duty-cycled, in-container pending), SYS-105 (LTE-M primary), SYS-106 (Wi-Fi offload), SYS-107 (BLE provisioning/offload), SYS-108 (multi-bearer selection), SYS-112 (alert-path prioritisation), SYS-113 (connectivity diagnostics), SYS-116 (roaming), SYS-115 (regional radio config), SYS-090 (radio/GNSS energy allocation — ties ADR-003); PRD-CONN-001/002/003/004/005/007/010/012, PRD-ENV-009.
Validation / how we'll know | the RF-in-container qualification test (PRD-VER-005) — measures attenuation, achievable link margin, delivery-success rate, antenna efficiency / total radiated power in-container, and GNSS fix-rate/time-to-fix; this test fixes the delivery-success target and decides the integral-vs-external antenna branch (closes SYS-104/SYS-109/SYS-117). Multi-bearer fail-over test (SYS-108) and roaming test across carriers/regions (SYS-116). Delayed- alert hazard reconciled in HAL.
Evidence | Evidence-Pending (link budget, antenna performance, in-container GNSS, per-market coverage — concept §7, D2/D4); bearer choices are Standard (3GPP LTE-M/Cat-M1, IEEE 802.11, Bluetooth Core; revisions to confirm).
---
ADR-011 — Multi-market regulatory posture (WHO PQS + EU GDP + US + air/ground transport)
Status | Accepted (specific categories/classifications Evidence-Pending)
Context | D4. The SCM serves global vaccine distribution, so it must satisfy several regimes at once, and these obligations shape the architecture (data integrity, radio configurability, battery transport, EMC/environment) — they are not a late-stage paperwork exercise. The exact WHO PQS device category and the US FDA classification are not yet fixed and must be ingested during the regulatory stage; they are recorded as gaps, not invented.
Decision | Adopt a multi-market regulatory posture and build its obligations into the architecture: WHO PQS cold-chain monitor/logger performance specs; EU GDP Good Distribution Practice with ALCOA+ data-integrity and a 21 CFR Part 11-style audit trail; US FDA cold-chain/data-integrity expectations (device positioned as an information accessory, not a therapeutic — classification to confirm); and air/ground transport & radio/EMC: UN 38.3 / IATA DGR (lithium battery), FCC Part 15 / CE RED (BLE/Wi-Fi/LTE-M/GPS), IEC 62368-1 (safety), CISPR/IEC 61000 (EMC), IEC 60529 (IP) and IEC 60068 (environmental). Radio is region-configurable so the device transmits only within each market's permitted band plan.
Alternatives considered
1. Single-market first (e.g., US or EU only), globalise later — rejected. Retrofitting Part 11 audit, ALCOA+ integrity, per-region radio conformance, and UN 38.3 after the architecture is frozen is far costlier and risks a re-architecture; global lanes are the intended use from day one.
2. Treat regulatory as a downstream compliance test only — rejected. Data-integrity, radio-configurability, and battery-transport requirements are architectural (they touch S3, S5, S2, S8); deferring them would invalidate design choices late.
Consequences
• Positive: one architecture serving all target lanes; data-integrity and radio-conformance designed in, not bolted on; battery-transport compliance treated as a release gate from the start.
• Negative: the strictest requirement across regimes tends to govern (higher cost/complexity); some regions may require per-market hardware variants (band plans); multiple certification test-house engagements.
• Residual risk: exact WHO PQS category and US classification are Evidence-Pending and could shift specific acceptance criteria; standard revisions to confirm across the board.
Requirements affected | SYS-115 (regional radio configurability), SYS-116 (roaming), SYS-101/SYS-157 (UN 38.3 / IATA transport), SYS-103 (ESD/EMC on charge port), SYS-111/SYS-118 (ALCOA+ integrity, trusted time), SYS-162 (hazard/RC traceability), plus platform Part 11 audit (S8); PRD-REG-004/006, PRD-ELEC-013/015, PRD-CONN-012; D4, UN-REG-005/006, UN-SAFE-009.
Validation / how we'll know | WHO PQS performance verification (once category confirmed); UN 38.3 battery test report and IATA documentation (SYS-101/157, Test); FCC Part 15 / CE RED radio certification and band-plan inspection (SYS-115, Inspection/Test); IEC 62368-1 safety, CISPR/IEC 61000 EMC (incl. IEC 61000-4-2 ESD, SYS-103), IEC 60529 IP and IEC 60068 environmental qualification; Part 11 audit-trail inspection. Regulatory engineer (EM-ENG-006) confirms category/classification gaps.
Evidence | Standard (WHO PQS, EU GDP/ALCOA+, 21 CFR Part 11, UN 38.3/IATA, FCC Part 15/CE RED, IEC 62368-1, CISPR/IEC 61000, IEC 60529, IEC 60068 — all revision to confirm); specific WHO PQS category and US classification Evidence-Pending (concept §7).
---
ADR-012 — On-device OLED + single glove-button HMI (not phone-dependent)
Status | Accepted
Context | AD-6 / AD-5. The safety-critical decision happens at the last mile, where the user is minimally trained, wears gloves (cold-chain / clinical), works in cold and glare, is often offline, and cannot be assumed to have a paired, charged, logged-in phone. The verdict and honest monitoring status must be readable and the device operable on the device itself, with no login (ADR-006) and no dependency on external infrastructure. C-forces: rich capability vs a simple, robust, always-available edge interface.
Decision | Provide a self-contained on-device HMI — a cold-legible, glare-legible OLED verdict display plus a single glove-operable button and a low-energy status indicator — that shows the accept/reject/quarantine verdict and honest status (including charge-inhibited and any monitoring gap) and supports the essential edge action (e.g., arm/acknowledge), offline and login-free. The mobile/BLE app is a companion for provisioning, configuration, arming aid, and richer local read — an enhancement, never a prerequisite for the safety-critical edge read/decision.
Alternatives considered
1. Phone-dependent HMI (no on-device display/controls; verdict only via app) — rejected. Fails when the phone is absent, dead, unpaired, or the user is not logged in — precisely the last-mile conditions; makes the safety-critical read depend on external infrastructure and a login, a use-error/safety hazard.
2. Multi-button / touchscreen device HMI — rejected. Touchscreens and small multi-button layouts are unreliable with gloves, in cold, and under glare, and cost energy and legibility; they raise use-error risk for a minimally trained operator. A single large button + clear verdict minimises use error.
Consequences
• Positive: the safety decision is always available at the edge without phone, network, or login; a single button + unambiguous verdict minimises use-error for gloved, cold, low-training users; energy spent only on a cheap verdict/status indication.
• Negative: an on-device display and button add cost, an enclosure cutout to seal (S1), and an energy line-item (bounded by ADR-003); the small display constrains information to the verdict layer (consistent with ADR-007).
• Residual risk: display cold-legibility/glare-legibility and glove-operability are usability targets not yet validated; the charge-inhibited state must be visibly distinct from "dead" so a depot operator doesn't mistake it for a fault (SYS-100).
Requirements affected | SYS-153 (use-error protection for low-training gloved offline users), SYS-154 (unmistakable un-armed/mis-arm state), SYS-151 (honest status / no false reassurance), SYS-100 (charge / charge-inhibit state indication), SYS-098 (low-battery warning), SYS-107 (BLE companion app as enhancement); PRD-UX-001/004/007/008, PRD-SAFE-005/006, PRD-ELEC-012; UN-UX, UN-VIS-007, STN-004/027.
Validation / how we'll know | summative usability test with representative gloved, cold, offline, minimally trained users on the safety-critical tasks (SYS-153, Test/summative usability — feeds the Usability Engineering File); cold/glare legibility and glove-operability demonstration; state-indication demonstration distinguishing armed/fault/charging/charge-complete/charge-inhibited (SYS-100). Reconciled with HAL for use-error hazards.
Evidence | Standard (IEC 62366-1 use-related risk; revision to confirm) + Judgment (AD-6/AD-5; UN-UX, last-mile use context).
---
Traceability note: each ADR above cites the D1–D4 / AD-1…AD-8 / C-tension it resolves and the SYS-### requirements it constrains; validation hooks point to the Stage-06 verification methods, PRD-VER-005 (RF-in-container), the 90-day cold-endurance / MR-VAL-006 validation, and the HAL (Stage 10) for safety reconciliation. ADR-004 and ADR-010 remain the two Evidence-Pending top risks and are closed by test, not assertion. Assembled end-to-end by ORCH-018 / EM-ENG-012. Forward links: Design Outputs (Stage 09), Hazard Analysis Library (Stage 10), Verification (Stage 13) / Validation (Stage 14).
Communication Architecture — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 — System Architecture · Owner: Principal Systems Engineer (EM-ENG-001) with the Principal Systems/RF Engineer and Product Security Engineer (EM-ENG-008). This document specifies the end-to-end communication architecture that realises the device→cloud value chain capture → detect → warn → decide → prove on the wire. It is based on foundational decisions D2 (offline-first: local logging + scheduled-cadence upload + immediate excursion alerts; alert latency is connectivity-bound) and D4 (multi-market: WHO PQS / EU GDP / US / air-transport, FCC Part 15 / CE RED radio regimes), architectural decisions AD-1 (offline-first store-and-forward) and AD-7 (signed, resumable, fail-safe OTA), and the connectivity system requirements SYS-104…118. It decomposes to the subsystem/detailed connectivity requirements (SRS-CONN) and interfaces with the cloud ingest surface (SRS-CLOUD-003/004/019) and security mechanisms (SRS-SEC, EM-ENG-008).
Evidence posture (honesty). Radio standards are cited by designation only (3GPP LTE-M/Cat-M1, IEEE 802.11, Bluetooth Core, TLS 1.2+/1.3), revision to confirm; no clause or numeric figure is asserted as fact. The RF-in-container link budget and antenna strategy remain the principal open architectural risk and are held Evidence-Pending (SYS-104, SYS-109, SYS-117), closed by the RF-in-container qualification test (PRD-VER-005). Cadence intervals, PSM/eDRX timers, retry/backoff constants, and per-market carrier coverage are configurable/Evidence-Pending and are not invented here — the architecture fixes the behaviour, the power budget and PRD-VER-005 fix the numbers.
Allocation legend: S3 Compute/MCU+FW · S5 Connectivity · S7 Storage · S8 Cloud · S9 Mobile.
---
1. Architectural intent — a link is an opportunity, never a dependency
The SCM is an offline-first edge node with opportunistic connectivity (AD-1/D2), not a live-streaming telemetry endpoint. The overriding rule that shapes every decision below: the authoritative record is captured and committed on the device the instant it is measured (SYS-013/SYS-018); the network only moves a copy of it out. No sample, event, excursion, or alert is ever lost because a bearer was unavailable — it is buffered and delivered opportunistically (SYS-111). Two consequences follow:
1. Two logical planes. A bulk plane (routine cadence telemetry, backlog drain, diagnostics, OTA download) is energy-scheduled and can wait for a cheap bearer. An alert plane (excursion/shock, safety-classed) is delivered as soon as any bearer is available, ahead of bulk (SYS-112). The two planes share the radios but never share priority.
2. Energy is the scarce resource, not bandwidth. The radios are the dominant load against the 90-day / −30 °C budget (D3, AD-2); bearer choice is therefore an energy-aware decision (SYS-108), and the modem spends most of its life asleep in a power-saving state (§6).
2. Bearer stack and priority (LTE-M → Wi-Fi → BLE)
Three bearers, selected opportunistically (SYS-108); no single bearer is required for the device to function or to eventually deliver its record (offline-first, SYS-108/SYS-111).
Rank | Bearer | Standard (designation) | Role | Reach | Relative energy | SYS / SRS
1 — primary | LTE-M (Cat-M1) | 3GPP LTE-M | Wide-area store-and-forward upload and the primary alert egress path | Global lanes, in-transit | Highest per-session; mitigated by PSM/eDRX + batching | SYS-105 (Safety, primary alert path)
2 — offload | Wi-Fi (802.11) | IEEE 802.11 | Opportunistic bulk backlog drain at depots/facilities on a provisioned network | Local (depot, warehouse, clinic) | Low per-bit when present; preferred for backlog | SYS-106
3 — short-range | BLE | Bluetooth Core / LE | Provisioning, config, local record read, short-range offload to a paired app/gateway | ~metres, to phone/gateway | Lowest; but requires a person/gateway present | SYS-107, SYS-110
Priority rationale. LTE-M is ranked primary because it is the only bearer available in transit, which is exactly when a safety alert must egress; it is a low-power WAN suited to duty-cycled operation across global lanes (D2/D4, SYS-105). Wi-Fi and BLE are higher-value when present (lower energy per bit, higher throughput) but are only present at endpoints/handling points, so they are used opportunistically to drain backlog cheaply and to provision — never relied upon for in-transit alerting. The selection policy prefers the lowest-energy path that satisfies the applicable cadence/latency target (SYS-108): for a bulk backlog at a depot, Wi-Fi wins; for an in-transit excursion alert, LTE-M is used immediately because no other bearer exists there.
3. Store-and-forward buffering and opportunistic upload
Buffer (SYS-111, SYS-168). All samples, events, excursions, gap markers, and audit entries are written to the non-volatile log (S7) atomically at capture, under the tamper-evidence applied on-device (SYS-132); the log is sized for a full 90-day trip at the minimum sample interval with margin (S7 allocation). The upload queue is a view over the committed log, not a separate volatile buffer — so a reset or brown-out mid-upload never loses un-uploaded data (SYS-168, power-fail-safe atomic writes).
Forward (SYS-111, SYS-174, SRS-CLOUD-003). On connectivity loss the device continues buffering; on reconnection it resumes upload in intact chronological order with idempotent (exactly-once) ingest, so that no record is duplicated or lost across any number of interruptions. Delivery uses bounded retry with exponential backoff (SYS-174) to tolerate lossy links without wasting energy on tight retry loops, and a high-water-mark / acknowledgement cursor so an interrupted session resumes from the last cloud-confirmed record rather than restarting. The cloud enforces exactly-once semantics with an idempotency key on (device identity, record sequence) (SRS-CLOUD-003); the device and cloud reconcile stores end-to-end so any undelivered or divergent record is detected and re-delivered or flagged (SYS-176, SRS-CLOUD-002).
Opportunistic drain. When a low-energy high-throughput bearer appears (Wi-Fi at a depot, or a BLE gateway), the device drains its backlog on the cheapest available path (SYS-106/SYS-108), preserving the energy budget for in-transit LTE-M alerting. Buffer backlog depth is recorded as a connectivity diagnostic (SYS-113) so a fleet operator can see a device that is capturing but not draining.
4. Cadence path vs immediate-alert path
The two planes of §1 map to two distinct transport behaviours (D2):
• Cadence (bulk) path. Routine telemetry is logged at a configurable sample interval (default 5–10 min) and uploaded on a cadence (default hourly, configurable) in batches (SYS-019/D2). The modem wakes on the cadence timer (or on an opportunistic bearer), transmits the batch, and returns to a power-saving state. This path is energy-scheduled and deferrable — if no bearer is available, the batch simply waits in the buffer.
• Immediate-alert path. On detection of an excursion or shock event, the alert is captured to the log instantly (SYS-013) and the device attempts immediate delivery over whatever bearer is available, ahead of any queued bulk telemetry (SYS-112). If no bearer is available, the alert is buffered like any record and pushed the instant one appears — alert latency is therefore connectivity-bound, not logic-bound (D2; a recognised delayed-alert hazard, SYS-112, reconciled with the HAL). The alert is not discarded or downgraded; it simply becomes the first thing sent when connectivity returns.
This separation is the architectural expression of D2: live delivery is upside, never a dependency; the record and the alert both survive total loss of connectivity as buffered evidence.
5. Alert-path prioritisation
A rescue-enabling alert is never queued behind bulk data (SYS-112, Safety, Critical). Under constrained connectivity or energy, the transport layer enforces a strict priority:
1. Excursion / shock alerts (safety-classed) — delivered first, on the first available bearer.
2. Fault / readiness / liveness events (SYS-170, SRS-DIAG-008) — device-health events that a fleet operator needs, but subordinate to safety alerts.
3. Routine cadence telemetry batches (bulk) — drained opportunistically.
4. Diagnostics, metrics, OTA download (SRS-DIAG-009, SRS-DIAG-013) — strictly subordinate; diagnostics always yield the radio to the alert path and never delay a sample, detection, or alert (SRS-DIAG-013).
Prioritisation is a queue discipline on the upload cursor, not a separate radio: the alert is promoted to the head of the queue and, if a session is already open for bulk, it preempts the batch. The cloud correspondingly must not drop a safety-relevant event under load (SRS-CLOUD-008) and delivers/escalates it through the notification engine (SRS-CLOUD-009).
6. Connection lifecycle and LTE-M power saving (PSM / eDRX)
The modem spends the overwhelming majority of the 90-day trip not connected, to protect the energy budget (D3/AD-2). The LTE-M connection lifecycle is:
1. Deep sleep (idle). MCU and modem in their lowest-power states between events. The modem uses PSM (Power Saving Mode) — it remains registered on the network but is unreachable and draws near-zero power for a negotiated period, avoiding the energy cost of re-attach on each wake.
2. Wake trigger. Either the cadence timer fires (routine batch), an excursion/shock event fires (immediate alert, §4), or the modem's eDRX (extended Discontinuous Reception) paging window opens to receive any pending downlink (a command or OTA availability signal from the shadow/twin, SRS-CLOUD-002/005).
3. Attach / resume. The modem exits PSM, re-uses its retained context where possible, and establishes the mutually-authenticated secure session (§8).
4. Transact. Alerts first, then batched backlog (§5), bounded by the energy allocated to the session.
5. Teardown → PSM. The session is closed and the modem returns to PSM/deep sleep.
PSM vs eDRX trade (Judgment; timers Evidence-Pending, set by the power budget). PSM gives the deepest saving but makes the device unreachable for downlink between wakes — acceptable because the device is the initiator for uploads/alerts and downlink (commands, OTA) is not time-critical and is picked up on the next wake (SRS-CLOUD-002 shadow reconciliation guarantees no lost/duplicated intent). eDRX gives bounded downlink reachability at higher standby cost and is used where a market/deployment needs more responsive command delivery. The specific PSM active/periodic-TAU and eDRX cycle timers are Evidence-Pending, closed by the 90-day cold-endurance power budget (SYS-163, MR-VAL-006). Under graceful degradation the upload cadence is stretched and location/non-essential uploads are shed first, while core logging and alerting are preserved to the last available energy (SYS-171).
7. Roaming and multi-region (D4)
Global vaccine lanes cross carrier and national boundaries mid-trip (D4); the wide-area path must not break at each border.
• Roaming continuity (SYS-116). A roaming-capable SIM/eSIM provisioning strategy maintains WAN connectivity across carriers and regions, automatically selecting an available permitted network without manual intervention, and logs the network/carrier used with connectivity diagnostics (SYS-113). Per-market carrier coverage and the SIM/eSIM strategy are Evidence-Pending (SYS-116) — carried as a gap, not invented; flagged to Regulatory (EM-ENG-006) and the connectivity carrier program.
• Regional radio conformance (SYS-115). Radio operation (bands, power, regulatory profile) is configurable per market and the device shall not transmit outside the bands/limits permitted for its configured region (FCC Part 15 / CE RED; 3GPP band plans; revision to confirm). Some regions may require per-market hardware variants (PRD-REG-004) — a compliance-blocking, Critical constraint; flagged to Regulatory (EM-ENG-006). The antenna must cover LTE-M, Wi-Fi, BLE, and GNSS bands per market (SYS-109).
• Data-residency coupling. Which cloud region a device delivers to is a deployment/residency concern (GDPR / EU GDP, D4) handled in the deployment architecture; the communication layer must carry a region/tenant tag so ingest routes to the residency-correct region (SYS-125 isolation; SYS-148 privacy; see deployment-architecture.md).
8. Security of the link (mutual TLS)
All device↔cloud and app↔cloud communication is protected by authenticated, encrypted transport (TLS 1.2 or later / 1.3), with mutual authentication (SYS-137, SRS-CLOUD-019): the device authenticates itself using its hardware-rooted, non-cloneable identity (SYS-134, secure element) and validates the server's certificate; the platform admits data only from a registered, authenticated device (SRS-CLOUD-001). Downgraded, unauthenticated, or spoofed connections are rejected (SYS-137). Provisioning, configuration, and controlled-data read over any bearer (BLE/Wi-Fi/LTE-M) require mutual authentication first (SYS-110), so BLE provisioning is not a soft attack surface — while the login-free on-device edge read of the verdict is preserved and is never carried over the link at all (AD-5, SYS-147). Records carry their capture-time cryptographic tamper-evidence unbroken through transport (SYS-132), so the link secures confidentiality and authenticity in transit while integrity is anchored at capture. Cipher-suite selection, key lifecycle (SYS-140), and the formal threat model (SYS-143) are owned by SRS-SEC / EM-ENG-008 and are Evidence-Pending.
9. Time synchronisation
A defensible record and correct excursion timing require a trusted, UTC-referenced clock (SYS-118, SYS-006). The device disciplines its RTC to a trusted network time reference whenever a link is available, bounds and records the residual clock error/drift, and continues timestamping from its free-running clock (with drift disclosed) when offline (D2). On reconnection it re-disciplines and the disclosed drift is bound into the record so the audit trail remains attributable/accurate (21 CFR Part 11 / ALCOA+; revision to confirm). The cloud is the time authority reference (SRS-CLOUD-010 time base and sync/drift disclosure); the device↔cloud time relationship is itself part of the tamper-evident record (SYS-133/SYS-135).
10. RF-in-container attenuation — the principal open architectural risk (Evidence-Pending)
The device operates inside a metallized, insulated shipper, which is RF-hostile (system block diagram I-08; concept §7). This is the single largest open architectural unknown in the communication design and is held Evidence-Pending, closed by the RF-in-container qualification test PRD-VER-005:
• The problem. RF attenuation inside a metallized/insulated shipper can exceed the fade margin of a low-power WAN radio (SYS-104). If egress fails, offline capture (D2, SYS-111) still preserves the record — but alerts are deferred, which is a safety hazard contributor (delayed/undelivered alert). GNSS is the most attenuated band of all and each fix is energy-costly, so location must fail soft (record best-available / last-known with a quality flag, SYS-117). The numeric attenuation, achievable link margin, and per-trip delivery-success target are not yet quantified — they are the headline Evidence-Pending items (SYS-104, SYS-109, SYS-117).
• Antenna strategy (SYS-109, Evidence-Pending). The architecture carries a live branch: an integral multi-band antenna (LTE-M, Wi-Fi, BLE, GNSS) is the baseline, with provision for a qualified external-antenna option via an IP67-sealed connector (PRD-ENV-007) if the integral antenna cannot close the SYS-104 budget inside the shipper. This branch couples to Mechanical (enclosure seal, connector, antenna keep-out) and Regulatory (radiated-power conformance per region, SYS-115).
• Mitigations already fixed by architecture (independent of the pending number). Store-and-forward buffering (§3) guarantees no data loss regardless of attenuation; alert-path prioritisation (§5) delivers the alert the instant any margin exists; multi-bearer opportunism (§2) means a depot Wi-Fi or BLE gateway can drain a device that could not reach LTE-M in the shipper; connectivity diagnostics (SYS-113) make a silently-attenuated device visible to the fleet. These bound the consequence of attenuation; the link margin itself remains Evidence-Pending until PRD-VER-005.
• Neighbour flags (CLAUDE.md rule 2). RF/antenna engineer (link budget, antenna, in-container GNSS); Mechanical/materials engineer (sealed external-antenna connector, keep-out); Regulatory (EM-ENG-006) (per-market radiated-power conformance, carrier coverage). RF-in-container performance may drive the external-antenna decision and even per-market hardware variants.
11. Multi-bearer decision & upload path (diagram)
mermaid flowchart TD START([Event: cadence timer, excursion/shock, or PSM/eDRX wake]) --> CAP["Capture to NV log<br/>atomic + tamper-evident<br/>SYS-013 / SYS-132 / SYS-168"] CAP --> CLASS{"Alert-class event?<br/>excursion / shock"}
CLASS -->¦Yes¦ APRIO["Promote to head of queue<br/>alert plane · SYS-112"] CLASS -->¦No¦ BULKQ["Queue as bulk / cadence batch<br/>bulk plane · SYS-019"]
APRIO --> BEARER BULKQ --> DUE{"Upload due?<br/>cadence or opportunistic<br/>bearer present"} DUE -->¦No¦ SLEEP DUE -->¦Yes¦ BEARER{"Select lowest-energy<br/>bearer that meets target<br/>SYS-108"}
BEARER -->¦Depot / provisioned¦ WIFI["Wi-Fi offload<br/>drain backlog · SYS-106"] BEARER -->¦In transit¦ LTE["LTE-M primary<br/>SYS-105"] BEARER -->¦App / gateway near¦ BLE["BLE short-range<br/>SYS-107 / SYS-110"] BEARER -->¦None available¦ BUFFER["Keep buffered<br/>no data loss · SYS-111"]
WIFI --> TLS LTE --> TLS BLE --> TLS TLS["Mutual TLS session<br/>HW-rooted identity<br/>SYS-137 / SYS-134"] --> TX["Transmit: alerts first,<br/>then bulk backlog · SYS-112"]
TX --> ACK{"Cloud idempotent ACK?<br/>exactly-once ingest<br/>SRS-CLOUD-003"} ACK -->¦ACK¦ ADV["Advance delivery cursor<br/>reconcile · SYS-176"] ACK -->¦No / timeout¦ BACKOFF["Bounded retry + backoff<br/>SYS-174"] BACKOFF --> BUFFER
ADV --> TIME["Discipline RTC to network time<br/>disclose drift · SYS-118"] TIME --> SLEEP(["Teardown → PSM / eDRX deep sleep<br/>SYS-105 power saving"]) BUFFER --> SLEEP
SLEEP -.-> START
RFNOTE["RF-in-container attenuation<br/>Evidence-Pending · SYS-104/109/117<br/>closed by PRD-VER-005"] RFNOTE -.->¦degrades link margin¦ BEARER
Figure 1 — Multi-bearer decision and upload path. An event is captured to the tamper-evident NV log first (offline-first, AD-1/D2), then classified into the alert plane (promoted ahead of bulk, SYS-112) or the bulk cadence plane. Bearer selection prefers the lowest-energy path that meets the target (SYS-108); if no bearer is available the record stays buffered with zero loss (SYS-111). Delivery is over mutual TLS (SYS-137/SYS-134), alerts first, with idempotent cloud ACK (SRS-CLOUD-003), bounded retry/backoff on failure (SYS-174), RTC discipline on success (SYS-118), and return to PSM/eDRX deep sleep. RF-in-container attenuation (Evidence-Pending, PRD-VER-005) degrades link margin at the bearer-selection step but never threatens the buffered record.
---
12. Requirements and interface traceability
Bearers | SYS-105 (LTE-M primary/alert egress), SYS-106 (Wi-Fi offload), SYS-107 (BLE provisioning/short-range), SYS-108 (opportunistic multi-bearer selection).
Store-and-forward & robustness | SYS-111 (buffer/resume/no loss), SYS-174 (retry/backoff/no duplication), SYS-176 (end-to-end reconciliation), SYS-168 (power-fail-safe writes), SRS-CLOUD-003/004 (ordered idempotent integrity-preserving ingest).
Alert path | SYS-112 (alert-before-bulk prioritisation), SRS-CLOUD-008/009 (event processing, notification/escalation).
Power-saving lifecycle | SYS-105 (duty-cycled PSM/eDRX behaviour), SYS-171 (graceful degradation / load-shed order), SYS-163/164 (endurance / reserve, Evidence-Pending).
Roaming / multi-region | SYS-115 (per-market radio conformance), SYS-116 (roaming, Evidence-Pending), SYS-125/SYS-148 (tenant isolation / privacy for residency routing).
RF-in-container (Evidence-Pending) | SYS-104 (link budget), SYS-109 (antenna strategy), SYS-117 (in-container GNSS) — all closed by PRD-VER-005.
Link security & time | SYS-110 (authenticated provisioning any bearer), SYS-137 (mutual TLS), SYS-134 (device identity / secure boot), SYS-132/133/135 (tamper-evidence/attribution/audit), SYS-118 (network time sync), SYS-140/143 (key lifecycle / threat model, Evidence-Pending, SRS-SEC/EM-ENG-008).
Diagnostics on the link | SYS-113 (connectivity diagnostics / signal-quality logging), SYS-114 (resumable firmware transport), SRS-DIAG-009/013 (remote diagnostics, energy subordination).
Interfaces | I-08 (device↔WAN/GNSS RF, attenuated by shipper), I-09 (device↔BLE/mobile), I-10 (WAN↔cloud API), I-12 (OTA), I-06 (MCU↔storage buffer), I-01 (enclosure/antenna boundary).
Forward links: SRS-CONN (connectivity detailed requirements), SRS-CLOUD (ingest/alerting), SRS-SEC (crypto/keys/threat model), SRS-OTA (update campaign), SRS-PERF (latency/energy numbers); safety reconciliation → HAL (Stage 10). Evidence-Pending items (RF-in-container link budget & antenna strategy, PSM/eDRX timers, per-market carrier coverage/SIM strategy, cipher suites, key lifecycle, threat model) are carried as gaps, not invented, and close on PRD-VER-005, the power budget (MR-VAL-006), and SRS-SEC/CSV.
Component List — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · Owner: Principal Systems Engineer, integrating EM-ENG-001 and the Principal Electrical/Hardware Engineer. Enumerates the ~55 architecture components CMP-001 … CMP-055, grouped by subsystem S1–S7 (device), the device firmware components (S3-FW), and the logical cloud/mobile components (S8/S9). Each component names the function it realises (FUN-# from functional-decomposition.md and/or SYS-#) and the realising detailed requirement (HWR-/SYS-/SRS-), with an honest status/evidence label.
Type legend. HW = electronic hardware · mech = mechanical/enclosure · SW = firmware/software · service = cloud/backend or mobile service. Status/evidence. Baselined · Preliminary · Evidence-Pending (unresolved number/branch, closed by a named test) · Preliminary — HAL (reconcile with Hazard Analysis, Stage 10).
Evidence-Pending flag (chemistry & antenna). The final low-temperature cell chemistry (CMP-006) and the in-container antenna strategy — integral vs external (CMP-028/CMP-029) — are the two principal open architectural unknowns. No chemistry, capacity, datasheet number, or RF attenuation figure is invented; each is carried Evidence-Pending and closed by a named test (cold-capacity test; RF-in-container qualification, PRD-VER-005). This mirrors D3 (energy tension) and the RF-in-container risk (concept §7).
---
S1 — Enclosure / mechanical
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-001 | Sealed IP-rated enclosure | mech | Houses & thermally couples device inside shipper (FUN-001/019) | IP-rated sealed housing sized to standard shippers; −30…+60 °C; antenna keep-out; cleaning-chemical compatible | PRD-MECH, HWR-032, SYS-104 | Preliminary
CMP-002 | Seals / gaskets | mech | Preserve ingress rating at joints/port/probe (FUN-039) | Elastomer seals rated across temp range and cold-to-warm cycling; compression set within limit | PRD-MECH, HWR-038/040 | Preliminary
CMP-003 | Mount / retention feature | mech | Fit & retain device in shipper (FUN-001) | Retains position for probe placement; glove-serviceable; MECH interface I-02 | PRD-MECH, SYS-003 | Preliminary
CMP-004 | Sealed USB-C port + receptacle | mech/HW | Depot recharge & wired OTA/data path (FUN-039/037) | Sealed port + compliant USB-C sink receptacle; safe under reverse-insertion/OVP/short; preserves seal | HWR-039, SYS-093, SYS-114 | Baselined
CMP-005 | External-probe connector | mech/HW | Retain & seal probe channel (FUN-002) | Positive retention + ingress seal; detects partial-seat/loose (feeds probe-presence); insertion-cycle rated | HWR-040, HWR-028, SYS-003 | Preliminary
S2 — Power
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-006 | Rechargeable low-temp cell | HW | Primary energy source (FUN-038/039) | Chemistry (low-temp Li-ion or LTO-class) whose −30 °C usable capacity, after derating, meets SYS-087 budget with margin | HWR-001, SYS-088 | Evidence-Pending (chemistry)
CMP-007 | Battery pack assembly | HW/mech | Deliver rail voltage & capacity (FUN-038) | xSyP config fixed by SYS-090 budget; documented interconnect + balancing where needed; traceable to cell datasheet | HWR-002, SYS-090 | Evidence-Pending (config)
CMP-008 | Battery-protection IC / AFE (BMS) | HW | OV/UV/OC/OT/short protection, MCU-independent (FUN-040) | Autonomous protection persisting in any MCU/FW fault; thresholds per cell datasheet | HWR-003, SYS-096 | Preliminary — HAL
CMP-009 | Secondary non-resettable protector (fuse) | HW | Last-line over-current/over-voltage backstop (FUN-040) | Series fuse and/or secondary OVP; terminates hazard if primary IC fails (single-fault safety) | HWR-012, SYS-096 | Preliminary — HAL
CMP-010 | Charge-path / PMIC (charger + rails) | HW | Manage charge; regulate & gate rails (FUN-039/038) | Programmable charge profile per cell datasheet; USB-C sink; per-rail regulation + independently gateable load switches | HWR-004, HWR-043, SYS-093, SYS-099 | Evidence-Pending (profile)
CMP-011 | Fuel-gauge IC | HW | Temp-compensated SoC/SoH (FUN-039/028/035) | Coulomb-count/model gauge with temperature compensation; accuracy supports readiness + non-misleading low-batt warning | HWR-008, SYS-097 | Evidence-Pending (accuracy)
CMP-012 | Cell-temperature sensor (NTC) | HW | Charge-inhibit qualification (FUN-039) | Cell-coupled (not board-air) NTC routed to charger temp-qualified enable; resolves 0 °C inhibit boundary with margin | HWR-006, SYS-094 | Preliminary
CMP-013 | Self-heat element + drive | HW | Warm cell before charge for depot turnaround (FUN-039) | Heater + drive that cannot bypass charge-inhibit or exceed safe upper thermal limit; feasibility under evaluation | HWR-007, SYS-095 | Evidence-Pending (feasibility)
CMP-014 | Charge-input protection front-end | HW | Tolerate reverse-polarity/OVP/surge (FUN-040) | Ideal-diode/OVP front-end + TVS clamp ahead of PMIC; rejects/limits out-of-spec source | HWR-009, HWR-045, SYS-102 | Preliminary — HAL
S3 — Compute / MCU (hardware)
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-015 | Low-power MCU (+ brown-out supervisor) | HW | Host governor, detection, verdict (FUN-017/038) | −30…+60 °C rated; retention deep-sleep; wake-on-RTC/interrupt; brown-out detector holds reset below threshold | HWR-013, HWR-019, SYS-089 | Preliminary
CMP-016 | Secure element / root of trust | HW | Secure boot, keys, identity (FUN-027/037/030) | Verifies FW authenticity before execution; tamper-resistant key/identity storage; anchors signed OTA + auth | HWR-016, SYS-032/110 | Baselined
CMP-017 | RTC + backup source | HW | Trusted monotonic UTC time (FUN-005) | Monotonic UTC across sleep/brown-out/service; backup cell/supercap preserves time across interruption | HWR-015, SYS-006/118 | Baselined
CMP-018 | Watchdog | HW | Recover firmware hang to safe state (FUN-040/007) | Independent of main clock where feasible; resets to defined safe state without corrupting committed log/armed state | HWR-017, SYS-047/099 | Preliminary
S4 — Sensing
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-019 | Internal temp sensor + signal chain | HW | Internal temperature channel (FUN-001) | ±0.5 °C or better across range, 0.1 °C resolution; stable ADC/reference, low self-heating, aging-flagged | HWR-021, HWR-027, SYS-001/003 | Evidence-Pending (accuracy class)
CMP-020 | External temperature probe + lead | HW/mech | Product-zone temperature channel (FUN-002) | Independently identifiable; meets HWR-021 at probe tip; probe-presence/continuity detection; retention-featured | HWR-022, HWR-028, SYS-003 | Evidence-Pending (offset/time-constant)
CMP-021 | Humidity sensor | HW | Relative-humidity channel (FUN-003) | ±5 %RH across 10–95 %RH non-condensing; condensation-tolerant placement | HWR-023, SYS-004 | Preliminary
CMP-022 | 3-axis accelerometer | HW | Shock/impact capture (FUN-004/011) | Full-scale + bandwidth for per-profile shock; low-power interrupt/threshold mode | HWR-024, SYS-005 | Preliminary
CMP-023 | GNSS / GPS receiver | HW | Duty-cycled location fix (FUN-033) | Configurable/event-triggered acquisition; fully powered-off between fixes; fail-soft best/last-known + quality flag | HWR-025, SYS-117/029 | Evidence-Pending (in-container fix/energy)
S5 — Connectivity
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-024 | LTE-M modem (primary bearer) | HW | Wide-area upload + alert path (FUN-013/022/023) | 3GPP Cat-M1; PSM/eDRX low-power; market band plans; MCU can power-cycle/gate; RF front-end holds spec to −30 °C | HWR-029, HWR-035, SYS-105 | Baselined (cold RF margin EP)
CMP-025 | Wi-Fi radio | HW | Opportunistic bulk offload (FUN-022) | IEEE 802.11; fully power-gated; coexists with LTE-M/BLE per antenna plan | HWR-030, SYS-106 | Baselined
CMP-026 | BLE radio | HW | Provisioning, local read, short-range offload (FUN-027/022) | Bluetooth Core/BLE; authenticated pairing/bonding; low idle exposure when unpaired | HWR-031, SYS-107/110 | Baselined
CMP-027 | SIM / eSIM | HW | Roaming wide-area connectivity (FUN-022) | Roaming-capable provisioning; auto-selects permitted network across borders; tamper-protected in sealed enclosure | HWR-034, SYS-116/115 | Evidence-Pending (carrier/SIM strategy)
CMP-028 | Integral antenna(s) + matching | HW | RF egress for all bands (FUN-013/022/033) | Integral antennas + matching for LTE-M/Wi-Fi/BLE/GNSS; enclosure keep-out; tuned to close SYS-104 in-container budget | HWR-032, SYS-109/104 | Evidence-Pending (antenna)
CMP-029 | External-antenna option (IP-sealed connector) | HW/mech | Fallback RF egress if integral cannot close budget (FUN-013/022) | Qualified external antenna via IP67-sealed connector; preserves ingress; no degradation to other bands when unused | HWR-033, SYS-109/104 | Evidence-Pending (antenna branch)
S6 — HMI
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-030 | OLED display | HW | Render offline verdict + state (FUN-019) | Legible accept/reject/quarantine at −30 °C within responsiveness budget; display technology confirmed by cold test | HWR-046, SYS-015/041 | Evidence-Pending (OLED cold response)
CMP-031 | Glove-friendly button | HW/mech | Arm & query single-action (FUN-029/019) | Single primary button; gloved actuation force/travel/size; sealed; wakes MCU; debounced | HWR-047, HWR-050, SYS-023 | Preliminary
CMP-032 | Status LED | HW | Indicate armed/fault/charge/inhibit states (FUN-039/019) | Low-energy indicator distinguishing armed/fault/charging/complete/charge-inhibited; visible to depot operator | HWR-048, SYS-100/094 | Baselined
S7 — Storage
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-033 | Non-volatile log memory (flash/FRAM) | HW | Full-trip buffer + audit + config (FUN-007/012/021/031) | Stores full 90-day/min-interval log with ≥50 % spare; program/erase endurance + retention (at temp) exceed reuse life; no overwrite of unsent data | HWR-020, SYS-018/043 | Evidence-Pending (part/format/endurance)
Interconnect / PCB
ID | Name | Type | Function (FUN/SYS) | Key spec / selection criteria | Realising req | Status
CMP-034 | PCB (stack-up) + conformal coat | mech/HW | Carry signal/power/RF; environmental protection (all FUN) | Defined stack-up, controlled-impedance RF/high-speed nets, continuous ground; conformal coat for condensation with sensor/antenna/contact keep-outs; ESD/evidence discipline | HWR-037, HWR-038, HWR-041, HWR-042 | Preliminary
---
S3-FW — Device firmware components (SW)
CMP-035 — Bootloader / secure-boot loader
Type | SW · Subsystem: S3 · Function: FUN-037 (verify & stage firmware), boot chain of trust.
Key spec/selection | Verifies image authenticity/integrity against the secure element before execution; supports resumable staging and fail-safe A/B or rollback so an update cannot brick the device.
Realising req | SYS-032, SYS-114, HWR-016 · Status: Baselined.
CMP-036 — Scheduler / RTOS
Type | SW · Subsystem: S3 · Function: cadence for FUN-001…004, FUN-022, FUN-033.
Key spec/selection | Deterministic scheduling of sampling, upload, and location windows; wake-on-RTC/ interrupt; cooperates with the power-state manager to return to deep sleep between activities.
Realising req | SYS-001, SYS-089 · Status: Baselined.
CMP-037 — Power-state manager (energy governor)
Type | SW · Subsystem: S3 · Function: FUN-038 (duty-cycle & hard radio gating).
Key spec/selection | Holds lowest viable sleep state between activities; gates LTE-M/Wi-Fi/BLE/GNSS/ display/sensor rails; permits only a queued excursion alert outside the scheduled window; enforces per-class energy allocations. The single most load-bearing FW component for the D3 90-day/−30 °C budget.
Realising req | SYS-089, SYS-090, SYS-086 · Status: Baselined; budget numbers Evidence-Pending.
CMP-038 — Sensor drivers / acquisition
Type | SW · Subsystem: S3 · Function: FUN-001…004, FUN-006.
Key spec/selection | Drives temp/humidity/accel/GNSS peripherals; applies calibration/linearisation; runs plausibility checks and channel-fault flagging; gates each sensor rail to its acquisition window.
Realising req | SYS-001/007, HWR-026/028 · Status: Baselined.
CMP-039 — Detection engine
Type | SW · Subsystem: S3 · Function: FUN-008…011.
Key spec/selection | Two-sided threshold evaluation, product-stability excursion logic (time/cumulative), hysteresis/debounce/consolidation, shock classification; freeze detection reliability equals heat.
Realising req | SYS-008/009/010/014, SYS-005 · Status: Preliminary — HAL.
CMP-040 — Verdict engine
Type | SW · Subsystem: S3 · Function: FUN-017/018/020.
Key spec/selection | Computes accept/reject/quarantine from the trip record; suppresses false "accept" on gaps/faults; deterministic and consistent with the cloud recomputation.
Realising req | SYS-015/016/017 · Status: Preliminary — HAL.
CMP-041 — Storage manager
Type | SW · Subsystem: S3/S7 · Function: FUN-007/012/021/024/031.
Key spec/selection | Gap-free append with power-fail-safe writes; wear management; buffer-full prioritisation (excursion/audit first); tamper-evidence support; exactly-once send-marking.
Realising req | SYS-002/018/021/028, SYS-099 · Status: Baselined; sizing Evidence-Pending.
CMP-042 — Comms agent (bearer manager)
Type | SW · Subsystem: S3/S5 · Function: FUN-013/014/022/023.
Key spec/selection | Energy-aware multi-bearer selection (LTE-M→Wi-Fi→BLE); ordered/resumable/idempotent upload; alert-before-bulk prioritisation; connectivity diagnostics; roaming/network selection.
Realising req | SYS-019/020/108/111/112/113/116 · Status: Baselined; in-container link Evidence-Pending.
CMP-043 — OTA agent
Type | SW · Subsystem: S3/S5 · Function: FUN-037.
Key spec/selection | Resumable download over LTE-M/Wi-Fi/USB-C; hands only verified images to the bootloader; preserves trip data across update; fail-safe application with notice of unmonitored windows.
Realising req | SYS-032, SYS-114 · Status: Baselined.
CMP-044 — Security / crypto module
Type | SW · Subsystem: S3 · Function: FUN-027/030/031/037.
Key spec/selection | Uses the secure element for identity, key handling, authenticated pairing/sessions, tamper-evidence over records/audit; enforces login-free edge read vs controlled configuration (AD-5).
Realising req | SYS-110/026/028/032, HWR-016/018 · Status: Baselined.
CMP-045 — Diagnostics / health
Type | SW · Subsystem: S3 · Function: FUN-028/035, connectivity diagnostics.
Key spec/selection | Captures readiness (SoC/SoH, calibration, firmware, last-seen), pre-arm checks, connectivity/RSSI history, buffer backlog; feeds fleet health and RF-in-container evidence.
Realising req | SYS-024/030/113 · Status: Baselined.
---
S8 / S9 — Cloud platform & mobile (logical components)
CMP-046 — Ingest / store-and-forward endpoint
Type | service · Subsystem: S8 · Function: FUN-023 (idempotent ingest), FUN-013.
Key spec/selection | Exactly-once (idempotent) ingest of ordered, resumable uploads; accepts alerts and bulk telemetry; preserves attribution/time/gaps on receipt.
Realising req | SYS-020/111, SYS-011 · Status: Baselined.
CMP-047 — Device registry
Type | service · Subsystem: S8 · Function: FUN-035, FUN-015 (routing lookup).
Key spec/selection | Authoritative device identity, trip/owner binding, readiness state; drives deploy-block and alert routing.
Realising req | SYS-030/012 · Status: Baselined.
CMP-048 — Rules / alert engine
Type | service · Subsystem: S8 · Function: FUN-015/016.
Key spec/selection | Routes alerts to configured responders; escalates unacknowledged alerts per policy; cloud-side alarm consolidation.
Realising req | SYS-012/014 · Status: Baselined.
CMP-049 — Record store
Type | service · Subsystem: S8 · Function: FUN-030/020.
Key spec/selection | Stores the complete tamper-evident trip record; recomputes the cloud verdict for device↔cloud consistency; enduring/consistent per ALCOA+.
Realising req | SYS-026/016 · Status: Baselined.
CMP-050 — Fleet dashboard + reporting
Type | service · Subsystem: S8 · Function: FUN-036, FUN-032 (report view).
Key spec/selection | Fleet-wide status (location, readiness, active trips, alerts, last-seen); flags overdue/stray; reporting surface for release/acceptance.
Realising req | SYS-031/030 · Status: Baselined.
CMP-051 — OTA service
Type | service · Subsystem: S8 · Function: FUN-037.
Key spec/selection | Signs and distributes firmware campaigns; staged/resumable delivery; never pushes an unverifiable or fleet-bricking image.
Realising req | SYS-032, SYS-114 · Status: Baselined.
CMP-052 — API gateway
Type | service · Subsystem: S8 · Function: FUN-032 (export/API), enterprise integration.
Key spec/selection | Authenticated device and enterprise APIs; export in documented human/machine-readable format preserving attribution/time/gaps/tamper-evidence; integrates ERP/QMS/partner systems.
Realising req | SYS-027, PRD-CLOUD · Status: Baselined.
CMP-053 — RBAC / SSO / identity
Type | service · Subsystem: S8 · Function: FUN-027 (platform governance).
Key spec/selection | Role-based access + SSO for configuration/data/access control; enforces the trust-boundary split (login-free edge read; controlled platform).
Realising req | SYS-110, PRD-SEC · Status: Baselined.
CMP-054 — Audit store
Type | service · Subsystem: S8 · Function: FUN-031.
Key spec/selection | Append-only, attributable, time-stamped, un-alterable-without-detection audit of config/arming/firmware/calibration/record actions (Part 11-style).
Realising req | SYS-028 · Status: Baselined.
CMP-055 — Mobile / BLE app
Type | service/SW · Subsystem: S9 · Function: FUN-025/026/027/028, local read/notify.
Key spec/selection | Provisioning, trip-profile configuration + validation, arming aid/pre-arm guidance, local record read, notifications; edge read stays login-free while platform enforces governance.
Realising req | SYS-022/024/025/107/110, PRD-MOB · Status: Baselined.
---
Coverage & evidence note
CMP-001…CMP-055 cover every device subsystem (S1–S7), the eleven required firmware components (bootloader, scheduler, power-state manager, sensor drivers, detection engine, verdict engine, storage manager, comms agent, OTA agent, security/crypto module, diagnostics), and the ten cloud/mobile services (ingest, registry, rules/alert engine, record store, dashboard, OTA service, API gateway, RBAC/SSO, audit store, mobile app), with sequential IDs and no gaps. Every component names the function(s) it realises and a realising requirement.
Chemistry and antenna are Evidence-Pending. The cell chemistry (CMP-006) and the integral-vs-external antenna strategy (CMP-028/CMP-029), plus the dependent items — pack config (CMP-007), charge profile/gauge accuracy (CMP-010/011), self-heat feasibility (CMP-013), GNSS-in-container performance (CMP-023), cold RF margin (CMP-024), storage sizing (CMP-033), and OLED cold response (CMP-030) — carry Evidence-Pending status and are closed by named tests (cold-capacity, cycle-life, RF-in-container / PRD-VER-005, gauge-accuracy, cold-display). Safety-critical protection items (CMP-008/009/012/014) are severity-first and Preliminary — reconcile with HAL (Stage 10). No chemistry, capacity, datasheet, or RF figure is invented.
Data-Flow & Control-Flow Architecture — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · §8.x. Owner: Principal Systems/Interface Engineer (EM-ENG-001), coordinating with Compute/FW (S3), Storage (S7), Connectivity (S5), and Cloud/Platform (S8). This document traces how a measurement becomes a defensible record and how an excursion becomes an acknowledged intervention, across the interfaces defined in [interface-control.md](interface-control.md) (INT-001…INT-035) and based on the system requirements (SYS-001…SYS-068) and detailed interface requirements (IFR-001…IFR-035). It realises the SCM value chain capture → detect → warn → decide → prove and reflects the offline-first store-and-forward architecture (D2, AD-1) and the duty-cycled energy governor (D3/AD-2).
Two views follow: (a) the data-flow architecture (where data moves and rests) and (b) the control-flow architecture (how the excursion-alert loop and the arm/pre-arm-check loop sequence). Diagrams use valid Mermaid.
---
(a) Data-flow architecture
The primary path traces one sample/event from acquisition to the enterprise; the two branches show the offline buffering path (no link) and the BLE local-read path (no WAN, no login). Every arrow is an interface from the ICD.
mermaid flowchart TB %% ===== Acquisition on the device ===== subgraph DEV["SCM device (offline-first edge node)"] SENS["S4 Sensors<br/>int/ext temp, RH, accel, GNSS"] ACQ["S3 MCU acquisition<br/>channel-tag + RTC timestamp (UTC)<br/>plausibility flag"] DET["S3 Detection<br/>two-sided thresholds +<br/>product-stability logic"] LOG[("S7 NVM log<br/>append-only, tamper-evident<br/>power-fail-safe")] QUE["S3 Store-and-forward queue<br/>ordered, seq-numbered,<br/>ack-gated"] BEAR["S5 Bearer select<br/>LTE-M &gt; Wi-Fi &gt; BLE-offload"] BLEsvc["S5/S3 BLE GATT server<br/>verdict + record read"] end
%% ===== Cloud ===== subgraph CLD["S8 Cloud platform"] ING["Ingest (idempotent)<br/>de-dup on deviceId+seq"] REC[("Record store<br/>authoritative trip record")] RUL["Rules / excursion eval<br/>(consistent with edge)"] ALR["Alert routing + escalation"] DASH["Fleet dashboard +<br/>reporting / export"] end
%% ===== External ===== APP["S9 Mobile / BLE app"] ENT["Enterprise ERP / QMS"] RESP["Responder / quality owner"]
%% ---- capture path ---- SENS -- "INT-017/018/019 DATA" --> ACQ ACQ --> DET ACQ -- "INT-021 DATA (every sample)" --> LOG DET -- "INT-021 DATA (excursion event)" --> LOG LOG -- "enqueue unsent" --> QUE
%% ---- forward path (connected) ---- QUE -- "INT-025 host" --> BEAR BEAR -- "INT-032 MQTT/HTTPS over mTLS<br/>(thru shipper INT-031)" --> ING ING --> REC ING --> RUL RUL --> ALR REC --> DASH ALR --> RESP DASH -- "INT-034 REST/OpenAPI" --> ENT DASH -- "INT-034 export PDF/CSV/JSON" --> ENT
%% ---- offline branch ---- BEAR -. "no link: retain,<br/>never overwrite unsent" .-> LOG LOG -. "buffer full 90-day trip<br/>(INT-021, SYS-018)" .-> QUE
%% ---- BLE local-read branch ---- LOG -- "INT-021 read" --> BLEsvc DET -- "verdict" --> BLEsvc BLEsvc -. "INT-033 GATT (INT-029 RF)<br/>login-free summary /<br/>auth full record" .-> APP
Figure a — SCM data-flow architecture. A sample is acquired from the sensors (INT-017/018/019), channel-tagged and timestamped against the RTC (UTC) and plausibility-flagged at the MCU (SYS-001/002/ 003/006/007), evaluated for excursion (SYS-009/010), and written to the append-only, tamper-evident, power-fail-safe NVM log first (INT-021, SYS-013) — the record exists before any link is attempted. Unsent records enter the store-and-forward queue (ordered, sequence-numbered) and leave over the best available bearer (INT-025 → INT-032, LTE-M primary) through the attenuated in-container RF path (INT-031). The cloud ingests idempotently (de-dup on deviceId+seq), builds the authoritative record, runs the same excursion evaluation as the edge, routes/escalates alerts, and serves the dashboard/export and enterprise API (INT-034). When no link can be closed the forward path degrades to buffering (never overwriting unsent data) until a full 90-day trip is held (SYS-018). Independently, the BLE local-read branch (INT-033 over INT-029) serves the login-free verdict summary and the authenticated full record to the mobile app with no WAN and no cloud (AD-5/AD-6).
---
(b) Control-flow architecture
(b.1) Excursion-alert control loop
mermaid sequenceDiagram autonumber participant S4 as S4 Sensing participant S3 as S3 MCU/FW participant S7 as S7 NVM participant S5 as S5 Radios participant S8 as S8 Cloud participant R as Responder
S4->>S3: sample (INT-017/018) + timestamp S3->>S3: evaluate two-sided + stability logic (SYS-010) alt excursion detected S3->>S7: commit excursion event (INT-021, SYS-013) [always] S3->>S3: update on-device verdict (SYS-015/016) alt bearer available? S3->>S5: emit HIGH-PRIORITY alert ahead of backlog (INT-025) S5->>S8: alert msg over mTLS (INT-032/IFR-025) S8-->>S5: application ACK S5-->>S3: delivered (mark alert sent) S8->>R: route + escalate (SYS-012) R-->>S8: acknowledge S8->>S8: close / stop escalation timer else no link S3->>S3: queue alert (jumps bulk backlog) Note over S3,S5: retry on next window (PSM/eDRX);<br/>event already durable in S7 (SYS-013) end else in range S3->>S7: append routine sample (INT-021) S3->>S3: deep-sleep until next interval (AD-2) end
Figure b.1 — excursion-alert control loop. On detection the event is committed to NVM first and always (SYS-013) — connectivity is never a precondition for the record. The verdict is updated (SYS-015/016). If a bearer is available the device emits a high-priority alert that jumps ahead of any bulk telemetry backlog (INT-032/IFR-025, SYS-011) and requires a cloud application-ACK before treating it as delivered; the cloud routes and escalates until a responder acknowledges (SYS-012), running an escalation timer. If no link exists, the alert is queued to jump the backlog on the next PSM/eDRX window while the event remains durable locally. This loop is why an excursion is never lost even if egress is delayed (D2).
(b.2) Arm / pre-arm-check control flow
mermaid flowchart TB START([User initiates arm<br/>button INT-023 or app INT-033]) --> READY{"Pre-arm readiness<br/>(SYS-024, IFR-018)"} READY -->¦charge OK?¦ C1{"SoC adequate<br/>(INT-020)"} C1 -->¦no¦ FAILC["Guide: charge device<br/>(INT-016 inhibit if cold)"] C1 -->¦yes¦ C2{"Calibration valid?"} C2 -->¦no¦ FAILCAL["Block deploy<br/>(SYS-030)"] C2 -->¦yes¦ C3{"Firmware valid /<br/>known-good?"} C3 -->¦no¦ FAILFW["Require OTA<br/>(INT-035)"] C3 -->¦yes¦ C4{"Profile valid?<br/>low&lt;high, interval 1-60m<br/>(SYS-025, IFR-017)"} C4 -->¦no¦ FAILP["Reject profile,<br/>keep prior active"] C4 -->¦yes¦ C5{"Probe present /<br/>not open-short<br/>(INT-018)"} C5 -->¦no¦ FAILPR["Flag probe fault<br/>(SYS-007)"] C5 -->¦yes¦ ARM["ARM: nonce-checked,<br/>atomic (IFR-018)"] ARM --> CONF["Positive on-device<br/>confirmation (SYS-023)<br/>OLED INT-022 + LED INT-023"] CONF --> AUDIT[("Append-only audit:<br/>armed/config/actor/time<br/>(INT-021, SYS-028)")] AUDIT --> TRIP([Armed — monitoring]) FAILC --> READY FAILCAL --> STOP([Not armed]) FAILFW --> READY FAILP --> READY FAILPR --> READY
Figure b.2 — arm / pre-arm-check control flow. Arming is gated behind readiness pre-conditions (adequate charge via INT-020, valid calibration, known-good firmware, valid profile, probe present via INT-018) so a trip cannot silently start unmonitored (SYS-023/024, the mis-arm hazard). A failing condition guides the user rather than proceeding; profile writes are validated and atomic (SYS-025/IFR-017) and the arm command is nonce-checked against replay (IFR-018). Only on success does the device give positive on-device confirmation (INT-022/INT-023) and append an attributable audit entry (INT-021, SYS-028) before entering the armed/monitoring state.
---
3. Narrative — lineage, integrity, ordering, and alert prioritisation
3.1 Data lineage and integrity (ALCOA+)
Every record carries its provenance from the first boundary. At acquisition the MCU binds each reading to its source channel (internal vs external probe — INT-017/018, SYS-003/056), timestamps it against the trusted monotonic RTC referenced to UTC and records the time source/sync status (SYS-006), and attaches a data-quality/gap flag without discarding the raw value (SYS-002/007). This makes the record Attributable, Legible, Contemporaneous, Original, and Accurate at the source, and — because the write to NVM is atomic/power-fail-safe and tamper-evident (INT-021, SYS-057) — Enduring and Complete across duty-cycle rail-downs (INT-013/015) and brown-outs. The device→cloud message additionally carries explicit SI units, channel identity, and a per-record integrity value (IFR-023) so the cloud reconstructs the record without unit/channel/time ambiguity and can verify integrity end-to-end, and the enterprise export/API (INT-034, IFR-027) preserves attribution, time, gaps, and tamper-evidence through to the release decision. The on-device verdict is derived from the same excursion evaluation as the cloud record (SYS-016), and any divergence is detectable and flagged — two sources of truth are kept consistent rather than allowed to disagree.
3.2 Ordering, idempotency, and resume (store-and-forward, D2)
The store-and-forward path (INT-032, IFR-024) is the integrity spine of the offline-first architecture. Records leave the queue in intact chronological order tagged with monotonic sequence numbers; the cloud ingests exactly once by de-duplicating on (deviceId, sequence) so a retransmitted record is never stored twice; an interrupted upload resumes from the last cloud-acknowledged sequence without restarting; and the device marks buffered data safely delivered only on positive application-level ACK (MQTT v5 PUBACK/reason-code or REST 2xx). When no link can be closed (including the unresolved in-container RF case, INT-031, Evidence-Pending) the device buffers a full 90-day trip without overwriting unsent data (SYS-018) and, if capacity is approached, preserves excursion and audit records first and warns rather than silently discarding safety-relevant data (SYS-021). Ordering + idempotency + resume + ack-gated deletion preserve ALCOA+ completeness and consistency across any number of link interruptions.
3.3 Alert-path prioritisation
The alert path is deliberately decoupled from and prioritised above the bulk-telemetry path. Three mechanisms enforce this: (1) local-first capture — every excursion is committed to NVM at detection independent of connectivity (SYS-013, INT-021), so the alert path is an accelerator of intervention, not the record of the event; (2) queue precedence — the alert is a distinct high-priority message sent ahead of any bulk backlog on the selected bearer and delivered with acknowledged QoS (SYS-011/IFR-025, INT-032), so a rescue-enabling alert never waits behind hours of routine samples; and (3) reachability sizing — LTE-M PSM/eDRX timers are sized to the alert-latency target (IFR-010, INT-027) so the modem sleeps for the energy budget yet wakes inside the alert window (the D2/D3 reconciliation). The cloud then routes and escalates until acknowledged (SYS-012), converting a detected excursion into an accountable human intervention — the point at which monitoring becomes prevented loss.
3.4 Honest evidence
The critical open items in these flows are carried explicitly, never papered over: the in-container RF link (INT-031, I-08) is Evidence-Pending and closed by the RF-in-container qualification test (PRD-VER-005) — until then, upload/alert egress is never a dependency for the record (D2); the thermal probe offset/lag (INT-001…004), rail budgets (INT-012…015), and GNSS-in-container cadence (INT-019/030) are Evidence-Pending against named analyses/tests; and the error/timeout/ back-pressure and authenticated-identity governance (SYS-067/SYS-068) is Preliminary pending the Hazard Analysis Library (Stage 10) and the cybersecurity threat model. The data-flow and control-flow above are correct as an architecture; their quantitative closure is the work of Stages 09–14.
Deployment Architecture — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 — System Architecture · Owner: Principal Systems Engineer (EM-ENG-001) with the Principal Cloud/Platform Engineer, Product Security Engineer (EM-ENG-008), and Quality/Regulatory Engineer (EM-ENG-006). This document specifies how the SCM is deployed across its three tiers — device + cloud + mobile — including fleet provisioning, cloud service topology, mobile offline sync, CI/CD and environments, the split trust boundary (AD-5), and data-residency/regulatory deployment (D4). It is based on the cloud software requirements SRS-CLOUD (SRS-400), the security system requirements SYS-132…148, the reliability platform requirements SYS-175/176, and the architectural decisions AD-1 (offline-first store-and-forward), AD-5 (split trust boundary), and AD-7 (signed, resumable OTA).
Evidence posture (honesty). The platform specifies behaviour and obligations, not a specific cloud provider, product, or infrastructure choice (SRS-CLOUD §1). Concrete availability SLA, RPO/RTO targets, retention periods, export formats, per-market carrier coverage, and regional residency mapping are Evidence-Pending (SYS-124/126/128/175, SYS-148, SYS-116) and are not invented here — the architecture fixes the topology and the guarantees; the numbers close on their tests and CSV/validation. Standards are cited by designation (21 CFR Part 11 / ALCOA+, EU GDP, GAMP 5, SOC 2, ISO/IEC 27001, IEC 62443, OWASP ASVS, OAuth 2.0 / OIDC, SAML 2.0, TLS 1.2+/1.3) — revision to confirm; no clause is asserted as fact.
Tier legend: Device (S1–S7) · Cloud (S8) · Mobile (S9) · Enterprise (customer ERP/QMS/IdP).
---
1. Deployment intent — three tiers, one authoritative record, split trust
The SCM deploys as three cooperating tiers whose boundaries are set by AD-1 and AD-5:
1. Device tier (edge). An offline-first node inside the shipper that captures and commits the authoritative record locally (AD-1/D2) and exposes a login-free verdict read (AD-5). It depends on no continuous connection.
2. Cloud tier (platform). The governed side of the trust boundary (AD-5): where the record becomes authoritative evidence, where alerts route/escalate, and where all configuration, access, and administration are authenticated, role-authorised, and audited (SRS-CLOUD-025, SYS-147).
3. Mobile tier (companion). A BLE-paired app for provisioning, trip-profile configuration, arming aid, local read, and notifications (S9), which itself federates to the platform's identity/governance.
The enterprise tier (customer ERP/QMS/partner systems and the customer identity provider) is not part of the SCM product but is a first-class integration target (SRS-CLOUD-011/017/019). The unifying invariant: the record is captured on the device, becomes authoritative in the cloud, and is read frictionlessly at the edge — while configuration and data governance live only on the platform side of the trust boundary.
2. Device deployment
• In-shipper placement (S1). Each device is deployed inside the insulated shipper, thermally coupled to the payload with an external probe representing the product zone and mechanically retained (system block diagram I-01/I-02). This placement is what creates the RF-in-container attenuation problem (Evidence-Pending, SYS-104/109/117; see communication-architecture.md) and drives the offline-first/store-and-forward deployment model — a device that cannot reach a bearer in the shipper still captures the full record and drains it opportunistically.
• Fleet model (D1 — reusable/multi-trip). Devices are managed as a reusable fleet: return → clean → recharge → readiness triage (charge/cal/firmware) → re-provision → redeploy (AD-8). Each device has an authoritative registry entry (SRS-CLOUD-001) recording its cryptographic identity, durable serial/asset identity, active profile/firmware version, calibration status/validity, ownership/tenant, and lifecycle state. Deployment is blocked when the pre-arm readiness gate fails (out-of-calibration, low cold-SoH, invalid firmware/profile; SRS-DIAG-003, SYS-165/030).
• Provisioning (SYS-141, SRS-CLOUD-001). Each device's unique identity and initial keys are provisioned into the hardware root of trust through a controlled, authenticated, audit-logged process at manufacture (end-of-line, SRS-DIAG-010), establishing the device's birth record and enrolling it in the device registry. The reusable/refurbish loop (D1) requires authenticated field re-provisioning / key rotation on refurbishment or re-deployment, without exposing private key material in cleartext at any point (SYS-141) — provisioning process Evidence-Pending, owned by SRS-SEC. Production debug ports are locked before field deployment (SYS-144).
• Device digital twin (SRS-CLOUD-002). Because devices are frequently offline (D2), each has a cloud-side shadow/twin holding last-reported state (firmware, profile, readiness, connectivity health, last-seen) and desired state (target profile/firmware, pending commands). Desired changes are applied and confirmed when the duty-cycled device next connects, without loss or duplication of intent — this is the deployment mechanism for fleet configuration and OTA targeting of an intermittently-connected fleet.
3. Cloud deployment topology
The platform is deployed as independently deployable, horizontally scalable microservices on a container/orchestration substrate behind an API gateway (SRS-CLOUD-020, SRS-CLOUD-019). Services and their deployment roles:
Service (S8) | Role | Scaling / availability | SRS-CLOUD
API gateway | Terminates mutual TLS; authenticates/authorises every request; per-tenant rate-limit + input validation; sole external surface | Stateless, scale-out; front-line availability | -019
Ingest / store-and-forward endpoint | Accepts buffered payloads; ordered, idempotent (exactly-once) de-dup; verifies + preserves tamper-evidence; degrades to accept-and-queue under partial outage | Elastic to bursty depot backlog drain; no loss/reorder on scale-out | -003, -004, -021
Device registry / identity | Authoritative per-device identity, config, calibration, firmware, tenant, lifecycle; admits only registered authenticated devices | HA, durable | -001
Device shadow / twin | Desired vs reported state reconciliation for offline fleet | HA, durable | -002
Command service + message routing | Authenticated, role-authorised, at-most-once commands; tenant-isolated routing to record store / rules / analytics | Scale-out; no cross-tenant leakage | -005, -006
Rules / event engine | Configurable, versioned, audited alert/fault rules; detects excursion / device-fault / liveness events; consistent with on-device verdict | Scale-out; no dropped safety event under load | -007, -008
Notification / escalation engine | Multi-channel alert delivery (push/email/SMS/webhook); acknowledgement tracking; time-bounded escalation | HA (safety path) | -009
Record-assembly + durable record store | Assembles complete, attributable, tamper-evident Part 11-style trip record; durable, multi-tenant isolated, retention-governed storage | Durable, geo-redundant; no single-component loss | -010, -012
Search / analytics / reporting | Scoped, read-only query, analytics, reporting over authoritative records | Scale-out (read); never mutates source | -013, -014
Export service | On-demand integrity-preserving export (human + machine readable); every export an audited action | Scale-out | -011
Identity / access (RBAC + SSO) | User lifecycle; least-privilege RBAC; OIDC/SAML federation to enterprise IdP | HA | -015, -016, -017
Append-only audit service | Immutable, attributable audit of all governed actions; reconciles cloud ↔ uploaded device audit trail | Durable, tamper-evident | -018
OTA service | Orchestrates signed, staged, resumable OTA campaigns against the shadow/twin (AD-7) | Scale-out | -005, → SRS-OTA
Observability / security monitoring | Metrics/logs/traces; health + security-event detection; operator alerting | Cross-cutting | -024, SYS-142
• Availability & durability (SYS-175, SRS-CLOUD-021). The platform meets its defined availability target for ingest and alerting, and shall not lose or silently corrupt any ingested record under any single-component failure, degrading gracefully (accept-and-queue) rather than dropping data. Availability SLA is Evidence-Pending; the no-loss/no-silent-corruption obligation is fixed.
• Backup / DR (SYS-128/175, SRS-CLOUD-022). Durable, geographically redundant backups of records, audit trails, and configuration; restore to a consistent, integrity-verified, attribution-preserving state within defined RPO/RTO (Evidence-Pending) — restore must re-establish tamper-evidence and attribution, not merely availability.
• Multi-tenancy (SYS-125, SRS-CLOUD-012). Each customer owns and is isolated to its own records and configuration; no cross-tenant leakage in routing, query, analytics, or export.
• Multi-region for data-residency (D4, SRS-CLOUD-022). The platform supports multi-region operation where data-residency obligations require it (§7); regional residency mapping is Evidence-Pending.
4. Mobile deployment
• BLE-paired companion app (S9, SYS-107/110). The mobile app deploys on handlers'/operators' devices and pairs to the SCM over authenticated BLE for provisioning, trip-profile configuration, arming aid, local record read, and notifications. Provisioning/config require mutual authentication first (SYS-110); the app is the governed-side actor at the edge (SRS-MOB, SRS-CLOUD-017).
• Offline sync. The app operates where WAN is absent (D2): it can read the device's record and status over BLE with no network, and syncs to the platform when connectivity returns — carrying any short-range-offloaded device data and pushing config/audit actions to the cloud, where they are governed, attributed, and audited (SRS-CLOUD-005/018). The device↔app↔cloud path preserves the same ordered, idempotent, tamper-evident semantics as the direct device↔cloud path (SYS-111/132; SRS-CLOUD-003).
• Identity & governance (SYS-136/146, SRS-CLOUD-016/017). The app authenticates users, federates to the enterprise IdP via SSO (OIDC/SAML), enforces least-privilege RBAC on governed actions, and protects sessions/tokens — without imposing any authentication on the login-free on-device edge read (SYS-147).
• Notifications (SRS-CLOUD-009 → SRS-MOB-008). The app is a delivery channel for the cloud notification/escalation engine (push), tracked for acknowledgement and escalation.
5. CI/CD and environments
• Environments. Separate dev / test / prod environments for the cloud platform, with the same microservice topology, so integration, CSV, and performance testing run against production-like deployments before promotion (GAMP 5 CSV for regulated customers, SRS-CLOUD-026).
• CI/CD pipeline. Build → automated test → security/SBOM checks → controlled promotion. The pipeline verifies the authenticity and integrity of third-party and build artefacts and maintains a software bill of materials (SBOM) across device firmware, cloud, and mobile, enabling identification of affected releases when a component CVE is disclosed (SYS-145). Device firmware releases are built, signed against the hardware root of trust, and distributed as signed, resumable, anti-rollback, fail-safe OTA through the OTA service with staged rollout to protect the fleet (SYS-139/AD-7, SRS-CLOUD-005).
• Change governance. Every configuration and deployment change is authenticated, authorised, and audited (SRS-CLOUD-018); rule and profile changes are versioned (SRS-CLOUD-007). Blue/green or rolling deployment supports scaling and upgrade without downtime and without loss or reordering of in-flight records (SRS-CLOUD-020).
• Vulnerability management (SYS-143). An ongoing threat-model-driven vulnerability process (monitoring → triage → remediation via signed OTA) spans device, cloud, and mobile for the fielded life; threat model Evidence-Pending (owned by SRS-SEC / EM-ENG-008).
6. The split trust boundary (AD-5)
Deployment enforces two distinct trust domains (SYS-147, SRS-CLOUD-025):
• (a) Edge domain — login-free read. The accept/reject/quarantine verdict and trip status can be read on the device display with no network, phone, login, or credential (AD-5, concept C9). The edge read is strictly read-only with respect to governed state — it cannot alter, delete, or reconfigure any governed data or the record. This is deployed entirely device-side (S3/S6) and is never gated by the platform and never carried over the network.
• (b) Platform-governed domain. All configuration, data access, export, administration, and record modification require authenticated, role-authorised, and audited access (SYS-147, SRS-CLOUD-016/018). This spans cloud (S8) and mobile (S9), federated to the enterprise IdP (SSO), and is where RBAC, session protection, tenant isolation, at-rest/in-transit encryption, and the append-only audit live.
The boundary is a security requirement, not a UX convenience (SYS-147): the last-mile user has no phone, network, or training, so the verdict must be readable with zero friction, yet the authoritative record must remain governed. Tamper-evident records (SYS-132) make the edge read safely read-only by construction.
7. Data-residency and regulatory deployment (D4 / GDPR)
Global vaccine distribution (D4) makes deployment regionally constrained:
• Data residency (SRS-CLOUD-022, SYS-125/148). Where a market's data-residency obligation requires it, records are ingested, stored, and recovered in the region that satisfies residency; the communication layer carries a region/tenant tag so ingest routes to the residency-correct region (see communication-architecture.md §7). Regional residency mapping is Evidence-Pending (D4).
• Privacy / data-minimisation (SYS-148, GDPR / EU GDP). Collection, transmission, and retention of location and any personal/operator data are minimised to the operational purpose (asset recovery, attribution, cold-chain evidence); retention/deletion occurs only through authorised, audited action; handling aligns with applicable privacy law per market (regional mapping Evidence-Pending).
• Regulatory record governance (D4). The platform's record assembly, audit trail, export, and access controls realise 21 CFR Part 11 / ALCOA+ (records, signatures, audit), EU GDP (data governance), and GAMP 5 CSV for regulated customers (SRS-CLOUD-010/011/018/026); WHO PQS device-performance and US FDA classification expectations are carried per D4 (exact WHO PQS category / US classification Evidence-Pending, concept §7; flagged to Regulatory EM-ENG-006).
• Radio / transport regional deployment. Per-market radio conformance (SYS-115, possibly per-market hardware variants) and lithium-battery transport compliance (UN 38.3 / IATA DGR, SYS-157) constrain how and where physical devices deploy — flagged to Regulatory (EM-ENG-006).
8. Deployment diagram (device ↔ cloud ↔ enterprise ↔ mobile)
mermaid flowchart TB subgraph EDGE["Device tier — edge (S1-S7, in shipper)"] DEV["SCM device<br/>capture + commit record · offline-first<br/>AD-1 / D2 / SYS-111"] READ["Login-free verdict read<br/>OLED · no network/login<br/>AD-5 / SYS-147"] DEV --- READ end
subgraph MOBILE["Mobile tier (S9)"] APP["BLE companion app<br/>provision · config · arm · read · notify<br/>offline sync · SYS-107/110"] end
subgraph CLOUD["Cloud tier — platform (S8, microservices / containers)"] GW["API gateway<br/>mutual TLS · authz · rate-limit<br/>SRS-CLOUD-019"] ING["Ingest store-and-forward<br/>ordered · idempotent · integrity<br/>SRS-CLOUD-003/004"] REG["Device registry / shadow-twin<br/>SRS-CLOUD-001/002"] RULE["Rules + event + notification/escalation<br/>SRS-CLOUD-007/008/009"] REC["Record assembly + durable<br/>multi-tenant store<br/>SRS-CLOUD-010/012"] DASH["Dashboard · search · analytics · export<br/>SRS-CLOUD-011/013/014"] OTA["OTA service<br/>signed · staged · resumable · AD-7<br/>SYS-139"] IAM["RBAC + SSO (OIDC/SAML)<br/>SRS-CLOUD-016/017"] AUD["Append-only audit + observability<br/>SRS-CLOUD-018/024 · SYS-142"] end
subgraph ENT["Enterprise tier (customer)"] IDP["Enterprise IdP<br/>SSO federation"] ERP["ERP / QMS / partner APIs<br/>SRS-CLOUD-011/019"] end
subgraph GOV["Governance (D4)"] RES["Multi-region data residency<br/>Evidence-Pending · SYS-148/125<br/>SRS-CLOUD-022 · GDPR/EU GDP"] DR["Backup / DR · RPO/RTO<br/>Evidence-Pending · SYS-175<br/>SRS-CLOUD-022"] end
DEV -->¦LTE-M / Wi-Fi mutual TLS<br/>SYS-137 · I-08/I-10¦ GW DEV -.->¦BLE authenticated · I-09¦ APP APP -->¦governed sync · mutual TLS¦ GW
GW --> ING ING --> REG ING --> RULE ING --> REC REC --> DASH RULE -->¦alerts push/email/SMS¦ APP RULE -->¦alerts / webhook¦ ERP DASH -->¦export / API¦ ERP OTA -->¦signed FW via shadow · I-12¦ GW IAM --- APP IDP --> IAM AUD --- REC AUD --- IAM
RES -.-> REC DR -.-> REC
SPLIT["Split trust boundary AD-5 / SYS-147:<br/>edge read = login-free, read-only ·<br/>platform = authenticated, RBAC, audited"] SPLIT -.-> READ SPLIT -.-> GW
Figure 1 — SCM deployment across the three tiers. The device captures and commits the authoritative record offline in the shipper and exposes a login-free, read-only verdict at the edge (AD-5). Data and alerts reach the cloud over mutual TLS (LTE-M/Wi-Fi) directly, or over authenticated BLE via the mobile app, which itself syncs through the governed platform. The cloud microservice topology ingests store-and-forward data (ordered, idempotent, integrity-preserving), assembles the authoritative Part 11-style record in durable multi-tenant storage, routes/escalates alerts, serves dashboard/search/analytics/export, and pushes signed staged OTA. Enterprise systems federate identity (SSO) and integrate via export/API. Data residency and backup/DR (both Evidence-Pending) govern where and how durably the record lives; the split trust boundary separates the frictionless edge read from the governed platform domain.
---
9. Requirements traceability
Cloud topology & platform (SRS-CLOUD) | -001 (registry/identity), -002 (shadow/twin), -003 (ordered idempotent ingest), -004 (integrity-preserving ingest), -005 (commands), -006 (routing), -007/008/009 (rules/events/notification+escalation), -010 (record assembly), -011 (export), -012 (durable multi-tenant storage/retention), -013/014 (search/analytics), -015/016/017 (user mgmt / RBAC / SSO), -018 (audit), -019 (API gateway), -020 (microservice/container/horizontal scale), -021 (availability, no-loss), -022 (backup/DR/multi-region/RPO-RTO), -023 (graceful failure handling), -024 (observability), -025 (platform security posture / split-trust), -026 (cloud verification / CSV).
Security (SYS-132…148) | SYS-132/133/134/135 (tamper-evidence/attribution/identity/audit), SYS-136 (RBAC/SSO), SYS-137 (mutual TLS), SYS-138 (at-rest encryption), SYS-139 (signed/staged/fail-safe OTA), SYS-140/141 (key lifecycle / provisioning, Evidence-Pending), SYS-142 (security monitoring), SYS-143 (threat model, Evidence-Pending), SYS-144 (debug-port lockdown), SYS-145 (SBOM/supply chain), SYS-146 (session mgmt), SYS-147 (split trust boundary), SYS-148 (privacy/residency, Evidence-Pending).
Reliability / platform | SYS-175 (availability/durability/RPO-RTO, Evidence-Pending), SYS-176 (end-to-end delivery reconciliation), SYS-165/030 (readiness gating for redeploy), SRS-DIAG-003/010/011 (readiness gate / EOL / depot triage).
Architectural decisions | AD-1 (offline-first), AD-5 (split trust), AD-7 (OTA), AD-8 (reusable-fleet lifecycle); foundational D1 (reusable), D2 (offline-first), D4 (multi-market/residency).
Interfaces | I-08 (device↔WAN), I-09 (device↔BLE/mobile), I-10 (WAN↔cloud API), I-11 (cloud↔enterprise), I-12 (OTA).
Forward links: SRS-CLOUD (platform behaviour), SRS-SEC (crypto/keys/threat model, EM-ENG-008), SRS-OTA (campaign/packaging), SRS-MOB (mobile app), SRS-FLEET (fleet workflow), SRS-DATA (schema/retention), SRS-PERF (availability/latency/RPO-RTO numbers); regulatory reconciliation → Regulatory (EM-ENG-006). Evidence-Pending items (availability SLA, RPO/RTO, retention periods, export formats, regional residency mapping, per-market carrier coverage, key lifecycle, formal threat model, WHO PQS category / US classification) are carried as gaps, not invented, and close on their associated tests and CSV/validation.
Fault-Management & Diagnostic Architecture — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 — System Architecture · Owner: Principal Systems Engineer (EM-ENG-001) with the Safety/Risk Engineer, Reliability Engineer, and Embedded Software Engineer (EM-ENG-007). This document specifies how the SCM detects, contains, responds to, and annunciates faults across all three tiers, and how it proves — continuously and on demand — that it is actually monitoring. It is based on the safety system requirements SYS-149…162, the reliability/availability system requirements SYS-163…176, the diagnostics software requirements SRS-DIAG (SRS-1400), and the FWR/fault group of the device firmware requirements. It realises the org's overriding fault principle: honest gap, no false reassurance — a device is at all times either verifiably monitoring or visibly flagged as not, and a fault must never corrupt the record or leave the device silently non-monitoring (SYS-151, SYS-167, SYS-170).
Evidence posture (honesty). Safety requirements are Preliminary — reconcile with the HAL (Stage 10); Action Priority is severity-first (CLAUDE.md rule 3; ISO 14971), severity fixed by outcome, probability/detectability Evidence-Pending until the HAL. Reliability targets (MTBF, cold state-of-health, watchdog/self-check cadence, degradation thresholds) are Evidence-Pending, closed by the reliability program (SYS-173) and named tests. Standards are cited by designation (ISO 14971, IEC 62304, IEC 62368-1, IEC 62133) — revision to confirm; no clause is asserted as fact.
Allocation legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · SYS = system/multiple.
---
1. Architectural intent — faults are made visible, never hidden
The SCM's dominant safety contribution is informational (SYS-149): both disposition-error directions harm the beneficiary (a spoiled dose administered or a potent dose wasted). The most dangerous failure is therefore not a crash — it is a silent failure that reads as a clean pass: a device that gave out in the cold whose missing interval looks uneventful (SYS-160), a hung device believed to be monitoring (SYS-169), a mis-armed trip that was never monitored (SYS-154), or an incomplete record that shows an unqualified "accept" (SYS-151). Fault management is architected around one invariant:
No-false-reassurance invariant. Every fault either (a) is transparently recovered and logged as a
recovery event so the trip is never assumed continuous, or (b) is annunciated locally and to the fleet
and marked in the record as a gap. An "armed" device is never silently non-monitoring, and an incomplete
record never presents as a clean pass. (SYS-151, SYS-167, SYS-169, SYS-170; SRS-DIAG-014.)
Fault management spans a four-stage pipeline — detect → contain/respond → annunciate → record — applied to a defined fault taxonomy, and is instrumented by a diagnostics architecture (POST/BIST, health monitor, metrics/logging, crash capture) that makes each fault observable and each device provably ready.
2. Fault taxonomy
Each class carries its dominant hazard, its detection owner, and its typical response. Severity is outcome-fixed; probability/detectability are Evidence-Pending until the HAL.
# | Fault class | Description | Dominant risk | Detection (S) | Typical response
F-SNS | Sensor fault | Temp (int/ext probe), humidity, or accelerometer channel out of range, stuck, disconnected, or drifted | Wrong/missing excursion detection → wrong disposition (SYS-149/150) | S4/S3 | Plausibility flag → degrade/annunciate; block clean-pass (SYS-151)
F-STO | Storage fault | NV log write/read failure, corruption, wear-out, capacity exhaustion | Loss/corruption of the evidentiary record (SYS-168) | S7/S3 | Atomic-write reject torn record; safe-state; annunciate
F-COM | Comms loss | No bearer / attenuated link / roaming gap | Deferred/undelivered alert (safety, SYS-104/112) | S5/S3 | Buffer + store-and-forward (SYS-111); alert on reconnection
F-NRG | Low energy | State-of-charge depleting toward end-of-trip; cold voltage sag (D3) | Dead-in-cold silent record loss (SYS-160/163) | S2/S3 | Graceful degradation / load-shed (SYS-171); reserve for safe landing (SYS-164)
F-CHG | Charge fault | Charge attempt below safe cell temperature; over/under-V, over-C, over-T, charger fault | Lithium plating → internal short → thermal event (SYS-155/156) | S2 (BMS)/S3 | Charge-inhibit (SYS-155); protection to safe state (SYS-156); annunciate
F-CLK | Clock fault | RTC failure, unbounded drift, loss of trusted time | Untrustworthy timestamps → indefensible record (SYS-118/SYS-006) | S3 | Free-run with disclosed drift (SYS-118); flag; re-discipline on link
F-BRN | Brownout | Voltage collapse below operating threshold (heightened at −30 °C) | Torn write / reset mid-operation (SYS-168) | S2/S3 | Brown-out detect → atomic-write barrier; safe-state; logged reset
F-OTA | OTA failure | Interrupted / unverified / rolled-back firmware update | Bricked or silently-unmonitored fleet (SYS-139/AD-7) | S3/S8 | Resume; verify against RoT; fail-safe to prior image (SYS-139)
F-TMP | Tamper | Physical/logical tamper: enclosure breach, debug-port probe, integrity/signature failure, unauthorised config | Corrupted/repudiable record; compromised fleet (SYS-132/144) | S3/S1/S8 | Tamper-evident flag in record; debug-port lockdown (SYS-144); security event (SYS-142)
Two cross-cutting classes sit above the table: F-ARM (mis-arm / unmonitored-trip start, SYS-154) caught by the pre-arm readiness gate (§4/§7), and F-SW (firmware hang/crash) caught by the watchdog (§3).
3. Detection mechanisms
Detection is layered so that faults are caught at boot, continuously while armed, and on demand — nothing relies on a single check.
• POST — power-on self-test (SRS-DIAG-001). On every power-on and reset, exercises MCU core, NV storage, clock/RTC, sensing channels, radios, HMI, and power/charge path, recording a structured pass/fail per subsystem before declaring the device ready. Establishes a known, attested starting state for every trip and every recovery; must fit the cold-boot timing budget (SRS-PERF-001).
• BIST — built-in self-test on demand (SRS-DIAG-002). Authenticated, on-demand exercise of each sensing channel (against internal references/plausibility), storage read/write, clock discipline, each radio bearer, HMI, and the power/charge path, reporting a per-subsystem result with a diagnostic code per failure for actionable depot/manufacturing triage.
• Watchdog supervision (SYS-169, FWR fault group). A hardware watchdog returns a hung or unresponsive device to an armed monitoring state without user intervention, and records every watchdog recovery as an attributable, time-stamped event — a recovered device is never assumed to have monitored continuously (the false-reassurance defence).
• Plausibility & consistency checks (SYS-007, SRS-DIAG-004). Sensor readings are checked for range, rate-of-change, stuck-value, and inter-channel plausibility (e.g. internal vs external probe divergence); clock discipline and drift are bounded; storage-write success is confirmed. A cross-check between the on-device excursion verdict and the cloud re-evaluation detects divergence (SYS-016, SRS-CLOUD-008).
• Continuous health monitoring (SRS-DIAG-004) + availability self-check (SYS-170). While armed, the device continuously watches task liveness, memory high-water mark, storage-write success, sensor liveness/plausibility, clock discipline, and battery SoC/SoH, raising an attributable health event on any indicator crossing its bound. The availability self-check periodically verifies the whole sensing→logging chain and, on detecting an outage, flags the device as not monitoring — closing the silent-outage gap (SYS-170).
• BMS protection & brown-out detection (SYS-155/156/168). The BMS enforces charge-inhibit below the safe cell temperature and OV/UV/OC/OT/short protection at the hardware level; a brown-out detector arms the atomic-write barrier so a torn write fails the record rather than silently corrupting it (SYS-168).
• Integrity/signature & tamper checks (SYS-132/134/139/144). Every record carries capture-time tamper-evidence; secure boot verifies firmware against the hardware root of trust at every boot; OTA images are verified before application; production debug ports are locked; failed integrity/signature checks raise security events (SYS-142).
4. Fault responses — retry, degrade, safe-state, annunciate
Responses are severity-ordered: the more a fault threatens the record or safety, the more decisively the system fails safe.
1. Retry / recover (transient faults). Comms loss → buffer and re-attempt with bounded backoff (SYS-111/174); firmware hang → watchdog recovery with logged event (SYS-169); OTA interruption → resume, re-verify, and if unverifiable fail-safe to the prior image (SYS-139). Recovery is always logged so continuity is never assumed.
2. Degrade (resource-constrained faults). Under low energy the device sheds discretionary functions in a defined priority order — location and non-essential uploads first, then upload cadence, then non-essential HMI — while preserving core logging, excursion detection/alerting, and the last-committed record to the last available energy (SYS-171; shed order/thresholds Evidence-Pending). A degraded sensor channel degrades detection gracefully and flags reduced confidence rather than presenting a false verdict.
3. Safe-state (unsafe or record-threatening faults). On any fault from which continued normal operation would be unsafe or misleading, the system enters the defined safe state (§5): charge-fault → charge inhibited and BMS holds the cell non-hazardous (SYS-155/156); storage/brown-out fault → atomic-write barrier protects committed data (SYS-168); any fault → record preserved intact and honest status held (SYS-158/159/167).
4. Annunciate (all persistent faults). The fault is surfaced locally (OLED/LED honest status, SYS-170/SYS-151), to the fleet as a coded, attributable failure event (SRS-DIAG-008), and in the record as a gap/quality flag so the trip can never present an unqualified clean pass (SYS-151).
Arming-gate response (SYS-154, SRS-DIAG-003). Before a trip, the pre-arm readiness self-check composes POST/BIST and health state into a single go/no-go and blocks arming while any pre-condition fails (adequate SoC-for-trip, valid calibration, valid firmware/profile, probe present), guiding the user to correct it. This prevents a silently-doomed trip (mis-arm is a top hazard) and is the last defence before deployment (readiness gating also blocks a battery whose cold SoH can no longer meet 90 days, SYS-165).
5. Safe-state definition and entry (safety-critical)
The safe state is a first-class, verifiable design target (SYS-159), not an emergent behaviour. It is defined (SYS-158/159/167) as a known state that simultaneously:
• Preserves the record. All committed samples, events, excursions, gap markers, and audit entries remain intact; no torn/partial record survives as valid; no previously committed data is corrupted (SYS-168). The record is never lost as a side-effect of failing safe.
• Holds the energy/charge subsystem non-hazardous. The BMS holds the cell within OV/UV/OC/OT/short limits and charging is inhibited if the cell is below the safe low-temperature limit (SYS-155/156) — no fire, burn, or thermal-runaway hazard in normal use, foreseeable misuse, or single-fault conditions.
• Presents an honest status — never misleading. The device does not present a monitoring indication it cannot back (SYS-151/159/170): it shows either honest "monitoring" or explicit "not monitoring / fault", and on power restoration it exposes the record and any power-loss gap rather than resuming as if monitoring were continuous (SYS-160).
Entry conditions. The safe state is entered on detection of any fault from which continued normal operation would be unsafe or misleading (SYS-159) — charge/battery hazard (SYS-155/156), storage/brown-out threat to the record (SYS-168), firmware/reset/OTA fault (SYS-158/139), sensor/clock fault that would make the verdict or timestamp untrustworthy (SYS-007/118), or an availability-self-check outage (SYS-170). From the safe state the system attempts automatic recovery to normal monitoring and logs the transition, never remaining in a state that corrupts the record or falsely reads as monitoring (SYS-167). Safe-state behaviour and its entry transitions are verified by fault-injection test (SYS-158/159; SRS-DIAG-016). Preliminary — reconcile with HAL; each entry condition will trace to a HA-###/RCR-### record.
6. Honest-gap / no-false-reassurance principle
This is the governing safety principle for the whole module and the reason the taxonomy above always ends in annunciate + record, never silently recover:
• An honest gap is safer than false confidence (SYS-151). When the record contains a monitoring gap, sensor fault, power-loss period, or interrupted armed interval, the lapse is made visible in the record and in on-device status, and the trip never shows an unqualified clean-pass / "accept".
• Recovered ≠ continuous (SYS-169). Every watchdog recovery, reset, and brown-out is a logged, attributable event feeding the honest-monitoring indication (SYS-017); a recovered device is never assumed to have monitored across the gap.
• Armed ⇒ verifiably monitoring or visibly flagged (SYS-170). The availability self-check guarantees a device is never silently outaged while appearing alive.
• Dead-in-cold is a false-negative hazard, not just an endurance miss (SYS-160). A device that dies silently in deep cold produces a missing interval that reads as uneventful — reframed here as a disposition-safety anchor (ties SYS-149) and mitigated by the 90-day endurance requirement (Evidence- Pending, closed by test) plus mandatory gap-exposure on power restoration.
• Diagnostics must not become a failure mode (SRS-DIAG-014). A failure inside the diagnostics subsystem itself (log-write failure, dump-capture failure) is contained and flagged without corrupting/evicting trip records and without falsely reporting the device as healthy.
7. Diagnostics interfaces (on-device, remote, manufacturing, service)
The diagnostics-access layer (SRS-DIAG §5) exposes the same underlying self-test/health/metrics evidence through four access modes, each with mode-appropriate authentication (SRS-DIAG-015):
• On-device (SYS-170, SRS-DIAG-001/003). POST result and honest monitoring/fault status on the OLED/LED with no login (edge read, AD-5); the pre-arm readiness go/no-go guides the user. Read-only with respect to governed state (SYS-147).
• Remote / fleet (SRS-DIAG-008/009). Authenticated remote retrieval of diagnostic logs, metrics, crash dumps, and self-test results over any bearer, and authenticated invocation of a non-intrusive self- test — without loss of buffered trip data and within the energy budget; a remote self-test never disrupts active monitoring or the alert path (SRS-DIAG-013). Faults are reported to the fleet as coded, attributable events so an operator can locate a faulted/not-ready device without physical access.
• Manufacturing / end-of-line (SRS-DIAG-010). Full-coverage diagnostics mode exercising every sensor channel, storage, clock, each radio, HMI, and the power/charge path with pass/fail and measured values; records the EOL result and calibration/provisioning state as the device's birth record for reuse-life traceability. Access-controlled and locked before field deployment (SYS-144, SRS-SEC).
• Service / depot (SRS-DIAG-011). Authenticated triage running BIST, reporting battery SoH and readiness, surfacing stored crash dumps and failure history, and giving a clear serviceable / return / retire recommendation for the reuse loop (D1). SoH redeploy threshold depends on Evidence-Pending cold cycle-life data (SYS-165).
• Developer debug (SYS-144, SRS-DIAG-012). Trace/console/JTAG-class interface — authenticated, access-controlled, disabled or locked in production/field, and auditable, so it is present-but-locked, never present-and-exposed (an open debug port defeats secure boot, at-rest encryption, and provisioning).
8. Health / heartbeat telemetry and crash capture
• Metrics (SRS-DIAG-005). The device collects and persists performance/reliability metrics — boot/wake time, per-cycle compute time, RAM high-water mark, write throughput/latency, missed-deadline count, reset/watchdog/brown-out/crash counts, per-function energy estimates, and battery SoH — and makes them available to remote diagnostics. These feed the reliability program (SYS-173) and close the D3 energy budget (MR-VAL-006), both Evidence-Pending.
• Heartbeat / liveness (SYS-031, SYS-170, SRS-DIAG-004/008). A duty-cycled liveness/health signal lets the fleet distinguish offline-but-healthy from silently faulted (a stray/overdue device, SYS-031), ridden on the energy-budgeted diagnostics path that never competes with the alert path (SRS-DIAG-013).
• Structured logging (SRS-DIAG-006). Diagnostic events are recorded in a structured, timestamped, attributable log with severity and diagnostic code, under the same integrity guarantees as the record and subordinate to excursion/audit retention — diagnostics never evict safety-critical trip data (SYS-021).
• Crash / fault dump capture (SRS-DIAG-007). On a crash, exception, watchdog reset, or brown-out, the device captures a bounded diagnostic dump (fault cause/exception context, relevant state, recent event history) to NV storage for later retrieval, without corrupting committed trip data. The dump is the primary root-cause evidence and feeds the software DFMEA / ERFKL (Stage 11). Dump size is bounded to protect the energy/storage budget.
• Cloud-side observability (SRS-CLOUD-024, SYS-142). The platform emits metrics/logs/traces for ingest, processing, alerting, storage, and API paths, monitors service health and security-relevant events (auth failures, integrity/signature failures, backlog growth), and alerts operators on degradation — closing the loop so a fault anywhere across the three tiers is observable.
9. Fault classes → detection → response → annunciation (diagram)
mermaid flowchart LR subgraph FAULTS["Fault taxonomy (§2)"] FSNS["F-SNS Sensor"] FSTO["F-STO Storage"] FCOM["F-COM Comms loss"] FNRG["F-NRG Low energy"] FCHG["F-CHG Charge fault"] FCLK["F-CLK Clock"] FBRN["F-BRN Brownout"] FOTA["F-OTA OTA failure"] FTMP["F-TMP Tamper"] FARM["F-ARM Mis-arm"] FSW["F-SW FW hang/crash"] end
subgraph DETECT["Detection (§3)"] POST["POST / BIST<br/>SRS-DIAG-001/002"] PLAUS["Plausibility &<br/>consistency · SYS-007"] HEALTH["Health monitor +<br/>availability self-check<br/>SRS-DIAG-004 / SYS-170"] WDOG["Watchdog · SYS-169"] BMS["BMS + brown-out<br/>SYS-155/156/168"] INTEG["Integrity / secure boot /<br/>tamper · SYS-132/134/144"] GATE["Pre-arm readiness gate<br/>SRS-DIAG-003 / SYS-154"] end
subgraph RESP["Response (§4)"] RETRY["Retry / recover<br/>SYS-111/169/139"] DEGRADE["Degrade / load-shed<br/>SYS-171"] SAFE["SAFE STATE<br/>preserve record · hold cell safe ·<br/>honest status · SYS-158/159/167"] BLOCK["Block arming<br/>SYS-154/165"] end
subgraph ANN["Annunciate + record (§4/§6)"] LOCAL["On-device honest status<br/>SYS-170 / SYS-151"] FLEET["Fleet failure event<br/>SRS-DIAG-008 / SYS-142"] RECGAP["Gap / quality flag in record<br/>no clean-pass · SYS-151/160"] end
FSNS --> PLAUS FSTO --> HEALTH FCOM --> HEALTH FNRG --> HEALTH FCHG --> BMS FCLK --> PLAUS FBRN --> BMS FOTA --> INTEG FTMP --> INTEG FARM --> GATE FSW --> WDOG FSNS --> POST FSTO --> POST
POST --> RETRY PLAUS --> DEGRADE HEALTH --> DEGRADE HEALTH --> SAFE WDOG --> RETRY BMS --> SAFE INTEG --> SAFE INTEG --> RETRY GATE --> BLOCK
RETRY --> LOCAL DEGRADE --> LOCAL SAFE --> LOCAL BLOCK --> LOCAL RETRY --> FLEET DEGRADE --> FLEET SAFE --> FLEET LOCAL --> RECGAP FLEET --> RECGAP
RECGAP -.->¦no false reassurance invariant §6¦ SAFE
Figure 1 — Fault classes → detection → response → annunciation. Each fault class routes to one or more detection mechanisms; detection drives a severity-ordered response (retry → degrade → safe-state → block arming); every persistent fault is annunciated locally, to the fleet, and in the record as a gap/quality flag so the trip can never present an unqualified clean pass. The no-false-reassurance invariant (§6) closes the loop: honesty in the record is guaranteed by failing into a defined safe state.
---
10. Requirements traceability
Safety (Preliminary — reconcile with HAL) | SYS-149 (severity anchor), SYS-150 (two-sided detection), SYS-151 (no false reassurance on incomplete record), SYS-152 (alarm balance), SYS-153 (use-error), SYS-154 (mis-arm prevention), SYS-155 (charge-inhibit), SYS-156 (no fire/thermal-runaway), SYS-157 (transport safety), SYS-158 (fail-safe + recover w/o data loss), SYS-159 (defined safe state + entry), SYS-160 (dead-in-cold false-negative), SYS-161 (residual risk), SYS-162 (hazard traceability).
Reliability / availability | SYS-163/164/165 (cold endurance / reserve / cycle-life SoH, Evidence-Pending), SYS-166 (charge-safety over life), SYS-167 (safe state on any fault), SYS-168 (power-fail-safe atomic writes), SYS-169 (watchdog + logged recovery), SYS-170 (availability self-check / silent-outage prevention), SYS-171 (graceful degradation / load-shed order, Evidence-Pending), SYS-172 (robustness over reuse), SYS-173 (reliability program, Evidence-Pending), SYS-174 (comms robustness), SYS-175 (platform reliability/RPO/RTO, Evidence-Pending), SYS-176 (end-to-end delivery reconciliation).
Diagnostics (SRS-DIAG) | SRS-DIAG-001 (POST), -002 (BIST), -003 (pre-arm readiness gate), -004 (health monitoring), -005 (metrics), -006 (structured logging), -007 (crash dump), -008 (fleet failure reporting), -009 (remote diagnostics), -010 (manufacturing/EOL), -011 (service/depot triage), -012 (secured debug), -013 (energy/priority subordination), -014 (diagnostics never masks/evicts record), -015 (authenticated access), -016 (coverage/fault-injection verification).
FWR fault group (device firmware) | watchdog supervision, brown-out handling, atomic-write barrier, secure-boot verification, safe-state arbiter (realise SYS-158/159/167/168/169 on-device).
Cloud fault handling | SRS-CLOUD-008 (event processing, no dropped safety event), SRS-CLOUD-009 (notification/escalation), SRS-CLOUD-023 (graceful logged handling / dead-letter, never silent discard), SRS-CLOUD-024 (observability), SYS-142 (security monitoring/incident logging).
Interfaces | I-05 (sensors→MCU), I-06 (MCU→storage), I-07 (HMI status), I-02/I-03 (power/charge), I-08/I-10 (fleet failure reporting, comms-loss context), I-12 (OTA).
Forward links: HAL (Stage 10) — each SYS-149…162 safe-state/annunciation requirement traces to a HA-###/RCR-### record; ERFKL (Stage 11) — crash dumps and coverage feed the software/hardware DFMEA; Verification (Stage 13) — fault-injection, watchdog, self-check, charge-inhibit, and coverage tests. Evidence-Pending items (cold endurance/reserve/SoH thresholds, watchdog/self-check cadence, degradation order, MTBF, platform SLA/RPO/RTO) are carried as gaps, not invented, and close on the reliability program (SYS-173) and named tests.
Functional Decomposition — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · Owner: Principal Systems Engineer, integrating EM-ENG-001. Decomposes the top-level product functions F1–F10 (design-architecture.md §4) into ~40 child functions FUN-001 … FUN-040, each allocated to the subsystems of the system block diagram (S1–S9) and traced to the realising system / hardware / software requirements. Based on D1–D4 and the value chain capture → detect → warn → decide → prove under the D3 constraint: survive 90 days at −30 °C without dying or losing data.
Subsystem legend. S1 Enclosure/mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud+dashboard · S9 Mobile/BLE app.
Evidence labels. Standard (cited by designation, revision to confirm) · Judgment (engineering inference stated as such) · Evidence-Pending (number/branch unresolved, closed by a named test). The D3 energy governor (FUN-038) and the RF-in-container branch (FUN-013/022/033) carry Evidence-Pending flags per the source libraries.
---
Decomposition tree
mermaid flowchart TB ROOT["SCM value chain<br/>capture → detect → warn → decide → prove"]
ROOT --> F1["F1 Measure & timestamp<br/>(S4+S3+S7)"] ROOT --> F2["F2 Detect excursions<br/>(S3)"] ROOT --> F3["F3 Alert & escalate<br/>(S3+S5+S8)"] ROOT --> F4["F4 Decide — offline verdict<br/>(S3+S6)"] ROOT --> F5["F5 Store & forward<br/>(S7+S5)"] ROOT --> F6["F6 Configure, provision & arm<br/>(S9+S6+S3)"] ROOT --> F7["F7 Record, prove & export<br/>(S7+S8)"] ROOT --> F8["F8 Locate & attribute<br/>(S4+S3+S8)"] ROOT --> F9["F9 Fleet, manage & OTA<br/>(S8+S5)"] ROOT --> F10["F10 Power, charge & self-protect<br/>(S2+S3)"]
F1 --> FUN001["FUN-001 Acquire internal temp"] F1 --> FUN002["FUN-002 Acquire external-probe temp"] F1 --> FUN003["FUN-003 Acquire humidity"] F1 --> FUN004["FUN-004 Acquire 3-axis acceleration"] F1 --> FUN005["FUN-005 Trusted timestamp & clock discipline"] F1 --> FUN006["FUN-006 Sensor plausibility & fault flag"] F1 --> FUN007["FUN-007 Gap-free sample logging"]
F2 --> FUN008["FUN-008 Two-sided threshold eval"] F2 --> FUN009["FUN-009 Product-stability excursion logic"] F2 --> FUN010["FUN-010 Alarm-balance (hyst/debounce)"] F2 --> FUN011["FUN-011 Shock/impact event detection"] F2 --> FUN012["FUN-012 Local excursion capture (offline)"]
F3 --> FUN013["FUN-013 Excursion alert emit/queue"] F3 --> FUN014["FUN-014 Alert-path prioritization"] F3 --> FUN015["FUN-015 Cloud alert routing"] F3 --> FUN016["FUN-016 Alert escalation"]
F4 --> FUN017["FUN-017 Offline verdict computation"] F4 --> FUN018["FUN-018 Incomplete-monitoring qualification"] F4 --> FUN019["FUN-019 Verdict rendering on HMI"] F4 --> FUN020["FUN-020 Device↔cloud verdict consistency"]
F5 --> FUN021["FUN-021 Full-trip NVM buffering"] F5 --> FUN022["FUN-022 Multi-bearer selection"] F5 --> FUN023["FUN-023 Ordered resumable de-dup upload"] F5 --> FUN024["FUN-024 Buffer-full graceful degradation"]
F6 --> FUN025["FUN-025 Trip-profile authoring"] F6 --> FUN026["FUN-026 Configuration validation"] F6 --> FUN027["FUN-027 Secure provisioning & pairing"] F6 --> FUN028["FUN-028 Pre-arm readiness checks"] F6 --> FUN029["FUN-029 Arm with positive confirmation"]
F7 --> FUN030["FUN-030 Tamper-evident record assembly"] F7 --> FUN031["FUN-031 Append-only audit trail"] F7 --> FUN032["FUN-032 Record export"]
F8 --> FUN033["FUN-033 GNSS fix acquisition (duty-cycled)"] F8 --> FUN034["FUN-034 Event time/location attribution"]
F9 --> FUN035["FUN-035 Device readiness reporting"] F9 --> FUN036["FUN-036 Fleet visibility & stray recovery"] F9 --> FUN037["FUN-037 Signed fail-safe OTA"]
F10 --> FUN038["FUN-038 Energy governor / radio gating"] F10 --> FUN039["FUN-039 Charge mgmt + cold charge-inhibit"] F10 --> FUN040["FUN-040 Battery protection & power-fail-safe"]
---
F1 — Measure & timestamp (S4 + S3 + S7)
FUN-001 — Acquire internal temperature
Description | Sample the internal (on-board) temperature channel each scheduled interval and hand a linearised, calibrated reading to the sample assembler as one of two independent temperature channels.
Parent function | F1
Allocated subsystem(s) | S4 (sensor + signal chain), S3 (schedule, linearisation)
Realising requirements | SYS-001, SYS-003, HWR-021, HWR-027
Inputs | Enclosure thermal exposure (I-01); acquisition trigger from scheduler (FUN-038); calibration coefficients.
Outputs | Internal-temperature value + resolution/units to FUN-007; quality context to FUN-006.
Trigger/cadence | Configurable base interval 1–60 min (default 5 min) for the armed trip.
Evidence | Judgment on default interval per WHO PQS E006 logger practice; accuracy class ±0.5 °C Evidence-Pending (HWR-021).
FUN-002 — Acquire external-probe temperature
Description | Sample the external-probe temperature channel — the product-zone representation — as a distinct, independently identifiable channel in every record, with probe-continuity context.
Parent function | F1
Allocated subsystem(s) | S4 (probe front-end), S1 (probe body/lead/connector), S3
Realising requirements | SYS-003, HWR-022, HWR-028, HWR-040
Inputs | Product-zone thermal coupling via probe (I-01, I-05); probe-presence/continuity state.
Outputs | External-probe temperature value + channel-ID to FUN-007; probe-fault flag to FUN-006.
Trigger/cadence | Same base logging interval as FUN-001.
Evidence | Judgment (probe represents product zone; two-sided detection depends on it, AD-4); probe offset/time-constant Evidence-Pending (HWR-022).
FUN-003 — Acquire relative humidity
Description | Measure relative humidity across the operating range at the logging interval as a secondary condition indicator, tolerant of cold-to-warm condensation.
Parent function | F1
Allocated subsystem(s) | S4 (humidity sensor), S3
Realising requirements | SYS-004, HWR-023
Inputs | Ambient humidity at the sensor aperture (I-05); acquisition trigger.
Outputs | RH value (10–95 %RH non-condensing) to FUN-007.
Trigger/cadence | Base logging interval.
Evidence | Judgment (secondary indicator; ±5 %RH target, HWR-023).
FUN-004 — Acquire 3-axis acceleration
Description | Read 3-axis acceleration in a low-power/interrupt-driven mode to support shock-event detection without an always-on high-rate sampler.
Parent function | F1
Allocated subsystem(s) | S4 (accelerometer), S3 (threshold/interrupt config)
Realising requirements | SYS-005, HWR-024, HWR-026
Inputs | Mechanical acceleration (I-05); per-profile shock threshold.
Outputs | Acceleration samples / hardware event interrupt to FUN-011.
Trigger/cadence | Interrupt-driven on threshold crossing; else quiescent (energy-gated by FUN-038).
Evidence | Judgment (interrupt/threshold mode avoids alarm-fatigue noise, C8; HWR-024).
FUN-005 — Trusted timestamp & clock discipline
Description | Stamp every sample and event against a monotonic, UTC-referenced RTC; discipline the RTC to network time when a link is available and free-run with bounded, disclosed drift otherwise; record time source and sync status with each record set.
Parent function | F1
Allocated subsystem(s) | S3 (RTC + backup), S5 (transport for time), S8 (time authority)
Realising requirements | SYS-006, SYS-118, HWR-015
Inputs | RTC tick; network time reference on connection (I-10); RTC valid/running status.
Outputs | UTC timestamp + drift/sync metadata attached to every sample/event.
Trigger/cadence | Per sample/event; re-discipline on each successful connection.
Evidence | Standard (21 CFR Part 11 attributable/accurate; revision to confirm).
FUN-006 — Sensor plausibility & fault flagging
Description | Evaluate each reading against physically plausible bounds and channel-fault indications (e.g. probe open/short), attach a data-quality flag, and never discard the raw value — a flagged reading beats a silently-wrong one.
Parent function | F1
Allocated subsystem(s) | S3, S4
Realising requirements | SYS-007, HWR-028
Inputs | Raw readings (FUN-001…004); probe-continuity state; plausibility bounds.
Outputs | Per-channel data-quality flag to FUN-007 and FUN-018.
Trigger/cadence | Every sample.
Evidence | Judgment (no false reassurance; SYS-007).
FUN-007 — Gap-free sample logging
Description | Assemble the multi-channel sample (temps, RH, accel context, quality flags, timestamp) and commit it to non-volatile storage such that no armed-trip interval is silently absent; any unrecorded interval is represented by an explicit gap marker.
Parent function | F1
Allocated subsystem(s) | S3, S7
Realising requirements | SYS-002, SYS-001, HWR-020, SYS-099
Inputs | Assembled sample fields (FUN-001…006); NVM write interface (I-06).
Outputs | Committed sample record; gap marker on miss → to FUN-018, FUN-021, FUN-030.
Trigger/cadence | Every logging interval, for the whole armed trip.
Evidence | Standard (ALCOA+ completeness, EU GDP; revision to confirm); power-fail-safe writes tie to SYS-099/HWR-019.
F2 — Detect excursions (S3)
FUN-008 — Two-sided threshold evaluation
Description | Compare each temperature channel against independent low (freeze) and high (heat) thresholds, and RH/shock values against their thresholds, from the active trip profile.
Parent function | F2
Allocated subsystem(s) | S3 (evaluation); S8/S9 author thresholds
Realising requirements | SYS-008, SYS-009
Inputs | Calibrated samples (FUN-007); active-profile thresholds (FUN-025).
Outputs | Per-channel threshold-breach candidates to FUN-009.
Trigger/cadence | Every sample.
Evidence | Judgment (freeze- and heat-sensitive products need distinct limits; SYS-008).
FUN-009 — Product-stability excursion logic
Description | Convert threshold breaches into confirmed excursions using configurable time-over/under-threshold and/or cumulative-exposure logic rather than instantaneous transients; record direction, onset time, and duration. Freeze detection reliability shall equal heat detection.
Parent function | F2
Allocated subsystem(s) | S3
Realising requirements | SYS-010, SYS-009
Inputs | Breach candidates (FUN-008); stability parameters from profile.
Outputs | Confirmed excursion object (direction/onset/duration) to FUN-010/012/013/017.
Trigger/cadence | On breach; evaluated against accumulated exposure window.
Evidence | Judgment; Preliminary — reconcile with HAL (SYS-010). Freeze parity Cited (WHO/UNICEF freezing reporting).
FUN-010 — Alarm-balance (hysteresis / debounce / consolidation)
Description | Apply configurable hysteresis, debounce, and alert-consolidation so the alerting rate is tunable to avoid alarm fatigue without suppressing genuine excursions.
Parent function | F2
Allocated subsystem(s) | S3, S8
Realising requirements | SYS-014
Inputs | Confirmed excursions (FUN-009); alarm-balance parameters.
Outputs | De-duplicated/consolidated excursion notifications to FUN-013.
Trigger/cadence | On confirmed excursion / state change.
Evidence | Standard (alarm-management principles; revision to confirm); Preliminary — HAL.
FUN-011 — Shock/impact event detection
Description | Detect and classify acceleration events exceeding the per-profile threshold, capturing peak magnitude and timestamp; ignore routine-handling accelerations below threshold.
Parent function | F2
Allocated subsystem(s) | S4 (interrupt), S3 (classification)
Realising requirements | SYS-005, HWR-024
Inputs | Accelerometer event/interrupt (FUN-004); shock threshold.
Outputs | Shock-event record (peak, time) to FUN-012/013.
Trigger/cadence | Event-driven on threshold crossing.
Evidence | Judgment (product-relevant mishandling without noise; SYS-005).
FUN-012 — Local excursion capture (offline)
Description | Write every detected excursion/shock event to non-volatile storage at the instant of detection, independent of connectivity, so no excursion depends on a live link to exist in the record.
Parent function | F2
Allocated subsystem(s) | S3, S7
Realising requirements | SYS-013, SYS-021
Inputs | Confirmed excursion/shock objects (FUN-009/011); NVM write (I-06).
Outputs | Persisted event record → FUN-013, FUN-021, FUN-030.
Trigger/cadence | Immediately on detection.
Evidence | Judgment (offline-first; last mile is never a black box, D2; SYS-013).
F3 — Alert & escalate (S3 + S5 + S8)
FUN-013 — Excursion alert emit / queue
Description | On detection, emit an alert to the cloud within the latency target when connectivity is available; otherwise queue it for delivery on next connection. Local capture (FUN-012) already guarantees the record exists.
Parent function | F3
Allocated subsystem(s) | S3, S5, S8
Realising requirements | SYS-011, SYS-112, SYS-105
Inputs | Consolidated excursion (FUN-010/012); bearer availability (FUN-022).
Outputs | Alert message on bearer (I-08/I-10) or queued alert.
Trigger/cadence | Event-driven; alert is the one radio activity permitted outside the scheduled window (FUN-038).
Evidence | Judgment (converts monitoring into prevented loss, D2). RF egress inside shipper Evidence-Pending (SYS-104).
FUN-014 — Alert-path prioritization
Description | When connectivity or energy is constrained, deliver excursion/shock alerts ahead of routine bulk telemetry over the selected bearer, so a rescue-enabling alert is never queued behind backlog.
Parent function | F3
Allocated subsystem(s) | S3 (queue policy), S5 (transport), S8 (alert ingest)
Realising requirements | SYS-112, SYS-108
Inputs | Alert queue + telemetry backlog; energy state (FUN-038).
Outputs | Ordered transmit queue (alerts first).
Trigger/cadence | On each connection opportunity under constraint.
Evidence | Judgment; Preliminary — reconcile with HAL (delayed-alert hazard).
FUN-015 — Cloud alert routing
Description | Route each received alert to the configured responder(s) for the device's trip/owner via the cloud rules engine.
Parent function | F3
Allocated subsystem(s) | S8
Realising requirements | SYS-012
Inputs | Ingested alert (I-10); trip/owner routing configuration; registry (FUN-035).
Outputs | Notifications to responders (email/SMS/push/app) and dashboard.
Trigger/cadence | On alert ingest.
Evidence | Judgment (an alert reaching no accountable human is not a risk control; SYS-012).
FUN-016 — Alert escalation
Description | Escalate an alert per configurable policy if it is not acknowledged within a defined time, widening the responder set until acknowledged.
Parent function | F3
Allocated subsystem(s) | S8
Realising requirements | SYS-012
Inputs | Unacknowledged-alert timer; escalation policy; acknowledgement events.
Outputs | Escalated notifications; escalation entry in audit trail (FUN-031).
Trigger/cadence | On acknowledgement-timeout.
Evidence | Judgment (SYS-012).
F4 — Decide, offline verdict (S3 + S6)
FUN-017 — Offline verdict computation
Description | Compute an unambiguous accept / reject / quarantine verdict on the device from the trip record against the active profile, requiring no network, phone, login, or manual action.
Parent function | F4
Allocated subsystem(s) | S3
Realising requirements | SYS-015, HWR-013
Inputs | Trip record (FUN-007/012); active profile (FUN-025).
Outputs | Verdict state + basis to FUN-018/019/020.
Trigger/cadence | On button query and on trip-end; within the MCU latency budget.
Evidence | Judgment (edge verdict is the last-mile decision point, AD-6; SYS-015).
FUN-018 — Incomplete-monitoring qualification
Description | When the record contains gaps, sensor faults, or an interrupted monitoring period, suppress any unqualified "accept" and indicate that monitoring was incomplete — no false pass.
Parent function | F4
Allocated subsystem(s) | S3, S6
Realising requirements | SYS-017, SYS-007
Inputs | Gap markers (FUN-007); quality flags (FUN-006); verdict candidate (FUN-017).
Outputs | Qualified verdict / "incomplete monitoring" state to FUN-019.
Trigger/cadence | At each verdict computation.
Evidence | Judgment; Preliminary — reconcile with HAL (SYS-017).
FUN-019 — Verdict rendering on HMI
Description | Render the verdict and device state legibly on the OLED (and status LED) at the cold extreme, within the responsiveness budget, so a gloved last-mile user can read it offline.
Parent function | F4
Allocated subsystem(s) | S6 (OLED/LED), S3
Realising requirements | SYS-015, HWR-046, HWR-048, HWR-050
Inputs | Qualified verdict (FUN-018); button wake (FUN-029/HMI).
Outputs | On-screen accept/reject/quarantine + state indication (I-07).
Trigger/cadence | On button press / defined events; display gated otherwise (FUN-038).
Evidence | Judgment; OLED cold response Evidence-Pending (HWR-046).
FUN-020 — Device↔cloud verdict consistency
Description | Ensure the on-device verdict derives from the same excursion evaluation as the cloud trip record so both agree for a given trip/profile; detect and flag any divergence.
Parent function | F4
Allocated subsystem(s) | S3, S8
Realising requirements | SYS-016
Inputs | Device verdict (FUN-017); cloud-recomputed verdict (FUN-030).
Outputs | Consistency confirmation or divergence flag → dashboard/audit.
Trigger/cadence | On upload / cloud record assembly.
Evidence | Judgment (two disagreeing sources of truth destroy disposition trust; SYS-016).
F5 — Store & forward (S7 + S5)
FUN-021 — Full-trip NVM buffering
Description | Buffer the complete sample + event log for a full 90-day trip at the minimum logging interval with defined margin, without overwriting unsent data.
Parent function | F5
Allocated subsystem(s) | S7 (NVM), S3
Realising requirements | SYS-018, HWR-020
Inputs | Committed samples/events (FUN-007/012).
Outputs | Persisted, ordered buffer for upload (FUN-023).
Trigger/cadence | Continuous over the armed trip.
Evidence | Judgment on sizing (≈1.3×10⁵ samples/channel at 1-min/90-day, HWR-020); Evidence-Pending storage sizing (SYS-018, closed by SYS-127).
FUN-022 — Multi-bearer selection
Description | Select among LTE-M (primary), Wi-Fi, then BLE to deliver buffered data/alerts, preferring the lowest-energy path that meets the cadence/latency target; require no single bearer to be present.
Parent function | F5
Allocated subsystem(s) | S5 (radios), S3 (policy), S8 (ingest)
Realising requirements | SYS-019, SYS-108, SYS-105, SYS-106, SYS-107, HWR-029/030/031
Inputs | Bearer availability/signal quality; energy state (FUN-038); backlog depth.
Outputs | Chosen bearer + session to FUN-023; diagnostics (SYS-113).
Trigger/cadence | On scheduled upload cadence (default hourly) and on alert.
Evidence | Judgment (energy-aware fail-over, AD-1/D2). In-container link Evidence-Pending (SYS-104).
FUN-023 — Ordered, resumable, de-duplicated upload
Description | Upload buffered records in intact time order, resume an interrupted upload without loss, and ensure exactly-once (idempotent) cloud ingest across any number of interruptions.
Parent function | F5
Allocated subsystem(s) | S3, S5, S8
Realising requirements | SYS-020, SYS-111
Inputs | Ordered buffer (FUN-021); selected bearer (FUN-022); cloud ack/idempotency keys (I-10).
Outputs | Delivered records marked sent; resume checkpoint.
Trigger/cadence | During each upload session; resume on reconnection.
Evidence | Standard (ALCOA+ enduring/consistent; revision to confirm).
FUN-024 — Buffer-full graceful degradation
Description | As NVM capacity is approached, prioritise retention of excursion events and audit records, warn of the capacity limit, and never silently discard excursion or audit data.
Parent function | F5
Allocated subsystem(s) | S3, S7
Realising requirements | SYS-021
Inputs | Buffer-depth watermark (FUN-021); record-class priorities.
Outputs | Capacity warning; prioritised retention decision.
Trigger/cadence | On buffer high-watermark.
Evidence | Judgment (degrade preserving most safety-relevant data first; SYS-021).
F6 — Configure, provision & arm (S9 + S6 + S3)
FUN-025 — Trip-profile authoring / selection
Description | Let an authorised user select or define a trip profile (thresholds, excursion logic, logging interval, upload cadence, alert routing) over BLE/mobile or on-device and apply it before arming.
Parent function | F6
Allocated subsystem(s) | S9 (app), S3 (apply), S8 (profile authority)
Realising requirements | SYS-022, SYS-008
Inputs | User profile selections (I-09/I-10); profile library.
Outputs | Active profile to FUN-008/009/017; profile-change audit entry (FUN-031).
Trigger/cadence | Pre-arm, per deployment.
Evidence | Judgment (configurable across products without hardware change; SYS-022).
FUN-026 — Configuration validation
Description | Validate a profile against constraints (e.g. low < high threshold, interval within bounds) and reject/flag an invalid or inconsistent configuration before it can be armed.
Parent function | F6
Allocated subsystem(s) | S3, S9, S8
Realising requirements | SYS-025
Inputs | Candidate profile (FUN-025); constraint set.
Outputs | Validated profile or rejection with reason.
Trigger/cadence | On profile apply, before arm.
Evidence | Standard (IEC 62366-1 use-error prevention; revision to confirm).
FUN-027 — Secure provisioning & authenticated pairing
Description | Permit configuration, provisioning, and controlled-data read only after mutual authentication of the requesting party over any bearer (BLE/Wi-Fi/LTE-M); keep login-free edge verdict read (AD-5) while configuration stays access-controlled.
Parent function | F6
Allocated subsystem(s) | S3 (auth), S5 (secure transport), S9 (app), S8 (identity/keys)
Realising requirements | SYS-110, SYS-107, HWR-016
Inputs | Requesting party credentials; device identity/keys in secure element.
Outputs | Authenticated session or denial; provisioning audit entry.
Trigger/cadence | On provisioning/config/controlled-read request.
Evidence | Standard (TLS 1.2+/1.3; Bluetooth Core pairing; revision to confirm).
FUN-028 — Pre-arm readiness checks
Description | Before confirming armed, verify readiness pre-conditions — adequate charge, valid calibration, valid firmware, valid profile, probe present — and guide the user to correct any failure.
Parent function | F6
Allocated subsystem(s) | S3, S6, S9
Realising requirements | SYS-024, SYS-030, HWR-028
Inputs | SoC/SoH (FUN-039), calibration validity, firmware version, profile status, probe-presence.
Outputs | Pass/fail per pre-condition + guidance; block-arm on failure.
Trigger/cadence | On arm request.
Evidence | Judgment (a device armed while un-ready produces a worthless record; SYS-024).
FUN-029 — Arm with positive confirmation
Description | Require an explicit arm action and give positive on-device confirmation that the device is armed, configured, and placed, so a trip cannot silently start unmonitored.
Parent function | F6
Allocated subsystem(s) | S3, S6 (button/OLED)
Realising requirements | SYS-023, HWR-047
Inputs | Passed readiness (FUN-028); button actuation (I-07).
Outputs | Armed state + on-screen/LED confirmation; arm audit entry.
Trigger/cadence | On user arm action.
Evidence | Judgment; Preliminary — reconcile with HAL (mis-arm is a top hazard; SYS-023).
F7 — Record, prove & export (S7 + S8)
FUN-030 — Tamper-evident trip record assembly
Description | Assemble, per trip, a complete record — all samples, events, excursions, gaps, configuration, and lifecycle actions — each attributable to an actor/source and time, protected by tamper-evidence, on device and in cloud.
Parent function | F7
Allocated subsystem(s) | S7, S8, S3
Realising requirements | SYS-026, SYS-016, HWR-016
Inputs | Logged samples/events (FUN-007/012); config/lifecycle events; trusted time (FUN-005).
Outputs | Complete tamper-evident trip record → FUN-020/032; audit linkage.
Trigger/cadence | Continuous accretion; finalised at trip-end.
Evidence | Standard (EU GDP, 21 CFR Part 11, ALCOA+; revision to confirm).
FUN-031 — Append-only audit trail
Description | Maintain an append-only audit trail (device and cloud) of configuration, arming, firmware, calibration, and record actions, each attributable, time-stamped, and un-alterable without detection.
Parent function | F7
Allocated subsystem(s) | S3, S7, S8
Realising requirements | SYS-028
Inputs | Lifecycle/config/alert actions (FUN-016/025/027/029/037).
Outputs | Immutable audit entries → cloud audit store.
Trigger/cadence | On every auditable action.
Evidence | Standard (Part 11 audit trail; revision to confirm).
FUN-032 — Record export
Description | Export a complete trip record on demand in a documented, human- and machine-readable format that preserves attribution, time, gaps, and tamper-evidence for release/acceptance and audit.
Parent function | F7
Allocated subsystem(s) | S8
Realising requirements | SYS-027
Inputs | Finalised trip record (FUN-030); export request (I-10/I-11).
Outputs | Exported record file/API payload to enterprise/quality systems.
Trigger/cadence | On demand.
Evidence | Standard (Part 11 records; revision to confirm).
F8 — Locate & attribute (S4 + S3 + S8)
FUN-033 — GNSS fix acquisition (duty-cycled)
Description | Acquire a GNSS/GPS fix on a configurable and event-triggered cadence within the location energy allocation, record fix quality per attempt, power the receiver fully off between fixes, and degrade gracefully (best-available / last-known + quality flag) when no fix is obtainable inside the shipper.
Parent function | F8
Allocated subsystem(s) | S4 (GNSS), S3 (scheduling/gating), S5 (assist data)
Realising requirements | SYS-117, SYS-029, HWR-025
Inputs | GNSS RF (I-08); location cadence/events; energy budget (FUN-038).
Outputs | Location fix + quality flag to FUN-034.
Trigger/cadence | Configurable cadence + on events (excursion, arrival); hard-gated.
Evidence | Judgment (fail-soft, never safety). In-container fix rate/energy Evidence-Pending (SYS-117; may drive external-antenna branch, HWR-033).
FUN-034 — Event time/location attribution
Description | Attribute each significant event (excursion, shock, arm, arrival) with time and best-available location for stray recovery and last-mile attribution.
Parent function | F8
Allocated subsystem(s) | S3, S8
Realising requirements | SYS-029, SYS-031
Inputs | Event objects (FUN-012); latest fix (FUN-033); trusted time (FUN-005).
Outputs | Time/location-stamped events → record (FUN-030) and fleet view (FUN-036).
Trigger/cadence | On each significant event.
Evidence | Judgment (SYS-029).
F9 — Fleet, manage & OTA (S8 + S5)
FUN-035 — Device readiness reporting
Description | Report each device's readiness — state of charge, calibration validity/date, firmware version, last-seen — to the fleet platform, and block deployment of a device that fails readiness.
Parent function | F9
Allocated subsystem(s) | S3 (report), S8 (registry/policy)
Realising requirements | SYS-030, SYS-091, HWR-008
Inputs | SoC/SoH (FUN-039), calibration state, firmware version, connectivity.
Outputs | Readiness record in registry; deploy-block flag for out-of-spec devices.
Trigger/cadence | On connection / depot check-in.
Evidence | Judgment (D1 reuse loop; out-of-cal/under-charged must not deploy; SYS-030).
FUN-036 — Fleet visibility & stray recovery
Description | Present fleet-wide status (location, readiness, active trips, alerts, last-seen) and flag devices that are overdue, stray, or unreported for recovery.
Parent function | F9
Allocated subsystem(s) | S8
Realising requirements | SYS-031, SYS-113
Inputs | Readiness (FUN-035); location/attribution (FUN-034); connectivity diagnostics.
Outputs | Fleet dashboard views; stray/overdue recovery flags.
Trigger/cadence | Continuous / dashboard refresh.
Evidence | Judgment (a reusable fleet must be trackable to control TCO/loss, D1; SYS-031).
FUN-037 — Signed, fail-safe OTA
Description | Accept and apply only firmware whose authenticity and integrity are verified; download resumably over LTE-M/Wi-Fi/USB-C; apply without loss of trip data; never leave a fielded device bricked or unmonitored-without-notice.
Parent function | F9
Allocated subsystem(s) | S3 (verify/stage/apply), S5 (transport), S8 (distribution)
Realising requirements | SYS-032, SYS-114, HWR-016, HWR-017
Inputs | Signed firmware image (I-12); root of trust (secure element); resume checkpoint.
Outputs | Verified applied firmware or safe rollback; OTA audit entry (FUN-031).
Trigger/cadence | On OTA campaign / depot update.
Evidence | Standard (secure-update principles, IEC 62443 concepts; revision to confirm; AD-7).
F10 — Power, charge & self-protect (S2 + S3)
FUN-038 — Energy governor / duty-cycle & radio gating
Description | Enforce the autonomous power-state governor: hold the device in its lowest viable sleep state between scheduled activities and hard-gate every high-drain load (LTE-M, Wi-Fi, BLE, GNSS, display, sensor rails) so no radio/GNSS activity occurs outside its scheduled, energy-budgeted window — the one exception being a queued excursion alert. This is the enabler that makes 90-day/−30 °C endurance achievable (D3, AD-2).
Parent function | F10
Allocated subsystem(s) | S3 (governor), S2 (rails/load switches), S5, S4, S6
Realising requirements | SYS-089, SYS-090, SYS-086, SYS-092, HWR-014, HWR-026, HWR-049
Inputs | Schedule; energy budget/allocations; wake events (RTC, button, accel interrupt).
Outputs | Rail enable/gate decisions; sleep/wake transitions to all subsystems.
Trigger/cadence | Continuous; wake-on-RTC/interrupt then return to deep sleep.
Evidence | Judgment (radios dominate load, must be bounded hard). 90-day/−30 °C budget and per-class allocations Evidence-Pending (SYS-086/090; closed by cold-endurance + RF-in-container tests).
FUN-039 — Charge management with cold charge-inhibit
Description | Manage USB-C recharge per the cell's charge profile and inhibit charging in hardware whenever cell temperature is below 0 °C (or the cell-defined safe limit), signalling the charge-inhibited state; optionally self-heat the cell to a permitted charge temperature without ever overriding the inhibit; report SoC/SoH via the fuel gauge for readiness and low-battery warning.
Parent function | F10
Allocated subsystem(s) | S2 (PMIC/BMS/gauge/heater), S4 (cell-temp sense), S3, S6 (indication)
Realising requirements | SYS-093, SYS-094, SYS-095, SYS-097, HWR-004, HWR-005, HWR-006, HWR-007, HWR-008
Inputs | USB-C source (I-03/I-04); cell temperature; SoC/SoH; charge state.
Outputs | Regulated charge current or inhibit; charge/inhibit indication (I-07); SoC/SoH to FUN-028/035.
Trigger/cadence | On depot charge connection; temperature-gated continuously during charge.
Evidence | Standard (IEC 62368-1 / IEC 62133 / JEITA charge-temperature concepts; revision to confirm). Charge-inhibit is severity-first Critical/Safety (AD-3); threshold Preliminary — HAL; self-heat feasibility Evidence-Pending (HWR-007).
FUN-040 — Battery protection & power-fail-safe operation
Description | Protect the cell against overvoltage, undervoltage, overcurrent, over-temperature, and short in normal use, foreseeable misuse, and single-fault conditions (independent of MCU/firmware), with a secondary non-resettable backstop; provide brown-out protection for clean safe-state entry with no corrupted write; shed discretionary functions (uploads, GNSS) before core sensing/logging on low battery, and never silently stop monitoring an armed device.
Parent function | F10
Allocated subsystem(s) | S2 (protection IC/fuse/OVP), S3 (brown-out/degradation), S7
Realising requirements | SYS-096, SYS-098, SYS-099, SYS-102, SYS-103, HWR-003, HWR-009, HWR-012, HWR-019
Inputs | Cell/rail voltages, currents, temperatures; charge-input transients; SoC (FUN-039).
Outputs | Protective safe-state actions; brown-out reset gating; graceful degradation order.
Trigger/cadence | Continuous, autonomous; on fault/brown-out/depletion.
Evidence | Standard (IEC 62368-1 / IEC 62133 concepts; revision to confirm). Severity-first Critical/Safety; several thresholds Preliminary — reconcile with HAL (SYS-096/102/103).
---
Coverage note
FUN-001…FUN-040 fully cover the ten top-level functions F1–F10 with no ID gaps. Every function traces to at least one SYS-/HWR- requirement. Evidence-Pending flags are carried honestly from the source libraries: the D3 endurance/energy budget (FUN-038), RF-in-container link and antenna branch (FUN-013/022/033), storage sizing (FUN-021), OLED cold response (FUN-019), and cold charge-safety numbers (FUN-039) are unresolved and closed by named tests, never invented. Preliminary/HAL-reconcile items (FUN-009/010/014/018/029/039/040) forward to the Hazard Analysis Library (Stage 10).
Interface Control — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · §8.x Interface Control Document (ICD) overview. Owner: Principal Systems/Interface Engineer (EM-ENG-001), coordinating with Power (S2), RF/Connectivity (S5), Cloud/Platform (S8), Mobile (S9), and Cybersecurity (EM-ENG-008). This document is the architecture-stage interface register: it names every realised interface (INT-001…INT-035), fixes its type, endpoints, direction, medium, key parameters, and traces each one down to the detailed interface requirements (IFR-001…IFR-035) and system interface requirements (SYS-051…SYS-068) that realise it, and up to the typed interface inventory I-01…I-12 of [../05-product-requirements/system-block-diagram.md](../05-product-requirements/system-block-diagram.md) §3.
How to read this document. Each INT-### is one physical or logical boundary in the built system. The INT register is a superset of the block-diagram inventory: the twelve typed interfaces I-01…I-12 are each decomposed into the concrete connectors, buses, rails, radio links, and protocol endpoints that actually exist in hardware and software. Nothing here invents a value: every quantified parameter is either a requirement already baselined upstream, a cited standard designation (revision to confirm), or an explicit Evidence-Pending gap closed by a named test.
---
1. Conventions
Interface types (per block-diagram §1): PWR power/energy · DATA digital data/signal (bus, log, config, protocol) · RF radio-frequency link · THRM thermal coupling · MECH mechanical mount/fit/retention.
Subsystem legend: S1 Enclosure · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile · EXT external world (payload, shipper, charger, networks, GNSS, enterprise).
Criticality follows CLAUDE.md rule 3 (severity-first): anything that authenticates a party, protects record integrity, gates cell charging, carries the excursion-alert path, or feeds the disposition verdict is Critical. Status is one of Baselined (contract fixed), Preliminary (to reconcile with the Hazard Analysis Library, Stage 10), or Evidence-Pending (value unknown, closed by a named test — never invented). Standards are cited by designation only, revision to confirm.
---
2. Master interface register (INT-001 … INT-035)
INT | Name | Type | From → To | Dir | Medium / protocol | Maps to I-0N | Realising req | Crit | Status
INT-001 | Payload → external probe tip | THRM | Payload zone (EXT) → probe (S1/S4) | → | Conduction/convection | I-01 | SYS-051, IFR-003 | Critical | Evidence-Pending
INT-002 | Payload/shipper → enclosure wall | THRM | Payload+shipper (EXT) → enclosure (S1) | → | Conduction/radiation | I-01 | SYS-051 | High | Evidence-Pending
INT-003 | Enclosure micro-env → internal temp sensor | THRM | Enclosure (S1) → internal T sensor (S4) | → | Conduction/convection | I-01 | SYS-051, SYS-003 | Critical | Evidence-Pending
INT-004 | Probe body → external-temp sense element | THRM | Probe body (S1) → probe T element (S4) | → | Conduction | I-01 | SYS-051, IFR-003 | Critical | Evidence-Pending
INT-005 | Shipper ↔ enclosure fit/retention | MECH | Shipper (EXT) ↔ enclosure (S1) | ↔ | Form-fit / retention feature | I-02 | SYS-052, IFR-035 | High | Preliminary
INT-006 | Enclosure ↔ probe lead entry & strain relief | MECH | Enclosure (S1) ↔ probe lead (S1) | ↔ | Sealed gland / strain relief | I-02 | SYS-052, IFR-003 | High | Preliminary
INT-007 | Enclosure ↔ sealed USB-C port aperture | MECH | Enclosure (S1) ↔ port (S1) | ↔ | Sealed aperture / gasket | I-02 | SYS-052, SYS-054, IFR-001 | High | Baselined
INT-008 | USB-C source → port VBUS | PWR | Charger (EXT) → port (S1) | → | USB-C VBUS, 5 V (≤9 V PD cap) | I-03 | SYS-053, SYS-054, IFR-001, IFR-006 | High | Baselined
INT-009 | Port → PMIC input | PWR | Port (S1) → PMIC (S2) | → | VBUS conductor + input protection | I-03 | SYS-053, IFR-006 | High | Baselined
INT-010 | PMIC → BMS → cell charge path | PWR | PMIC (S2) → BMS → cell (S2) | → | Charge current, inhibit-gated | I-03 | SYS-053, IFR-007 | Critical | Baselined
INT-011 | Cell ↔ BMS protection/gauge | PWR | Cell (S2) ↔ BMS (S2) | ↔ | Cell terminals + sense | I-03 | SYS-053, IFR-007, IFR-008 | Critical | Baselined
INT-012 | BMS → MCU / always-on rail | PWR | BMS (S2) → MCU (S3) | → | Regulated rail (always-on) | I-04 | SYS-055, IFR-009 | High | Evidence-Pending
INT-013 | BMS → sensing rail (gated) | PWR | BMS (S2) → sensing (S4) | → | Regulated rail, FW-gated | I-04 | SYS-055, IFR-009 | High | Evidence-Pending
INT-014 | BMS → connectivity rail (gated) | PWR | BMS (S2) → radios (S5) | → | Regulated rail, FW-gated | I-04 | SYS-055, IFR-009 | High | Evidence-Pending
INT-015 | BMS → HMI + storage rails (gated) | PWR | BMS (S2) → HMI (S6)/NVM (S7) | → | Regulated rail(s), FW-gated | I-04 | SYS-055, SYS-057, IFR-009 | High | Evidence-Pending
INT-016 | Charge-inhibit control signal | DATA/CTRL | BMS (S2) ↔ PMIC (S2) ↔ MCU (S3) | ↔ | Discrete + status register | I-03/I-04 | SYS-053, IFR-007, IFR-008 | Critical | Baselined
INT-017 | Internal sensor bus → MCU | DATA | Sensors (S4) → MCU (S3) | → | I2C 100/400 kHz / SPI | I-05 | SYS-056, SYS-003, IFR-004 | Critical | Baselined
INT-018 | External probe channel → MCU | DATA | Probe front-end (S4) → MCU (S3) | → | Analog/1-Wire/I2C + presence | I-05 | SYS-056, SYS-003, IFR-003, IFR-004 | Critical | Preliminary
INT-019 | GNSS receiver → MCU | DATA | GNSS rx (S4) → MCU (S3) | → | UART/I2C (NMEA/binary) | I-05 | SYS-056, SYS-029, IFR-013 | Medium | Evidence-Pending
INT-020 | Fuel-gauge → MCU | DATA | Fuel-gauge (S2) → MCU (S3) | → | I2C/SMBus | I-04 | SYS-055, SYS-024, IFR-008 | High | Baselined
INT-021 | MCU → NVM (log/config/audit) | DATA | MCU (S3) → NVM (S7) | → | SPI, power-fail-safe writes | I-06 | SYS-057, SYS-018, IFR-035 | Critical | Baselined
INT-022 | MCU → OLED display | DATA | MCU (S3) → OLED (S6) | → | SPI/I2C, framebuffer | I-07 | SYS-058, SYS-015, IFR-035 | High | Baselined
INT-023 | MCU → status LED / button → MCU | DATA | MCU (S3) ↔ LED+button (S6) | ↔ | GPIO / PWM + debounced input | I-07 | SYS-058, SYS-023, IFR-035 | High | Baselined
INT-024 | MCU ↔ secure element | DATA | MCU (S3) ↔ SE (S3) | ↔ | SPI/I2C, keys/secure boot | I-06 | SYS-057, SYS-068, IFR-032 | Critical | Baselined
INT-025 | MCU ↔ radio host interface | DATA | MCU (S3) ↔ radios (S5) | ↔ | UART/SPI/SDIO + AT/host API | I-08/I-09 | SYS-062, SYS-019, IFR-010, IFR-034 | High | Baselined
INT-026 | Radio ↔ antenna feed network | RF | Radios (S5) ↔ antennas (S1/S5) | ↔ | RF traces/coax, matching net | I-08 | SYS-059, IFR-014 | High | Evidence-Pending
INT-027 | LTE-M ↔ cellular network | RF | LTE-M (S5) ↔ NET (EXT) | ↔ | 3GPP Cat-M1, PSM/eDRX | I-08 | SYS-059, SYS-019, IFR-010 | High | Evidence-Pending
INT-028 | Wi-Fi ↔ access point | RF | Wi-Fi (S5) ↔ AP (EXT) | ↔ | IEEE 802.11, WPA2/WPA3 | I-08 | SYS-059, SYS-019, IFR-011 | Medium | Evidence-Pending
INT-029 | BLE ↔ mobile app radio link | RF | BLE (S5) ↔ app (S9) | ↔ | Bluetooth LE, LESC pairing | I-09 | SYS-060, IFR-012, IFR-016 | High | Baselined
INT-030 | GNSS constellation → receiver | RF | GNSS (EXT) → rx antenna (S4) | → | L1 multi-constellation + A-GNSS | I-08 | SYS-029, IFR-013 | Medium | Evidence-Pending
INT-031 | In-container RF link (thru shipper) | RF | Antennas (S1/S5) ↔ NET/GNSS (EXT) | ↔ | Attenuated path thru shipper | I-08 | SYS-059, IFR-014 | Critical | Evidence-Pending
INT-032 | Device ↔ cloud (store-and-forward + alert) | DATA | Device (S3/S5) ↔ ingest (S8) | ↔ | MQTT v5 / HTTPS over mTLS | I-10 | SYS-062, SYS-020, SYS-011, IFR-021…IFR-026 | Critical | Baselined
INT-033 | Device ↔ BLE app (GATT) | DATA | Device (S3/S5) ↔ app (S9) | ↔ | BLE GATT, versioned service | I-09 | SYS-061, SYS-022, IFR-015…IFR-020 | High | Baselined
INT-034 | Cloud ↔ enterprise (REST/API + export) | DATA | Cloud (S8) ↔ ERP/QMS (EXT) | ↔ | REST/OpenAPI, OAuth 2.0, JSON/PDF/CSV | I-11 | SYS-064, SYS-027, IFR-027…IFR-030 | High | Baselined
INT-035 | OTA service → device (signed FW) | DATA | OTA (S8) → device updater (S3/S5) | → | Signed package, resumable, A/B | I-12 | SYS-065, SYS-032, IFR-031…IFR-033 | Critical | Baselined
Note — internal test boundary. The device also carries an internal debug/EOL-test port (SWD/JTAG + UART, IFR-005). It is deliberately not assigned an INT-### as a fielded interface because it is non-user-accessible and locked in production (readout-protection/lifecycle fuse set, authenticated debug). It is governed under §4 (ICD governance) as a manufacturing/security interface, not a system run-time interface.
---
3. Interface detail blocks
Each block gives the endpoints, direction, medium/protocol, key parameters, the realising requirement, the block-diagram interface it maps to, its criticality, and its status/evidence.
INT-001 — Payload → external probe tip
Type | THRM · Endpoints: cold-chain payload zone (EXT) → external temperature probe (S1 lead / S4 element) · Direction: payload → probe (one-way thermal exposure).
Medium/parameters | conduction/convection coupling between the product zone and the probe tip. The probe must track the product-zone temperature within a bounded, characterised offset and time-constant (both to confirm by thermal-characterisation test); placement, thermal mass, and self-heating set the error.
Realises | SYS-051, IFR-003 · Maps to: I-01 · Criticality: Critical (freeze/heat verdict rests on this proxy, AD-4).
Status/evidence | Evidence-Pending — probe offset/lag budget closed by the thermal-characterisation test (forward to PVL). Judgment; no numeric offset asserted.
INT-002 — Payload/shipper → enclosure wall
Type | THRM · Endpoints: payload + insulated shipper (EXT) → sealed enclosure wall (S1) · Direction: → (exposure).
Medium/parameters | conduction/radiation through the enclosure into the internal micro-environment; the shipper is simultaneously the thermal barrier being ridden inside. Sets the lag seen by the internal sensor (INT-003).
Realises | SYS-051 · Maps to: I-01 · Criticality: High.
Status/evidence | Evidence-Pending — enclosure thermal path characterised with INT-001/INT-003. Judgment.
INT-003 — Enclosure micro-environment → internal temperature sensor
Type | THRM · Endpoints: enclosure interior air/mass (S1) → internal temperature sensor (S4) · Direction: →.
Medium/parameters | conduction/convection to the on-board sensor. The internal channel represents the enclosure micro-environment (one of the two sides of the two-sided detection, AD-4); its relationship to the external-probe channel is what the characterisation test bounds.
Realises | SYS-051, SYS-003 · Maps to: I-01 · Criticality: Critical.
Status/evidence | Evidence-Pending (offset/lag). Judgment.
INT-004 — Probe body → external-temperature sense element
Type | THRM · Endpoints: probe body/tip (S1) → sensing element inside the probe (S4) · Direction: →.
Medium/parameters | conduction from the probe body to the thermistor/digital sense element; the element's self-heating and the tip mass are part of the INT-001 error budget.
Realises | SYS-051, IFR-003 · Maps to: I-01 · Criticality: Critical.
Status/evidence | Evidence-Pending; probe technology (thermistor vs digital) to confirm with S4 (IFR-003 Preliminary). Judgment.
INT-005 — Shipper ↔ enclosure fit and retention
Type | MECH · Endpoints: standard insulated vaccine shipper (EXT) ↔ enclosure (S1) · Direction: ↔ (fit/retention).
Medium/parameters | form-fit within the defined shipper set and a tool-free retention feature; the device must not displace payload, defeat the thermal barrier, or shift/float in transit (retention is the precondition for representative thermal coupling INT-001 and a clean shock signature INT-017).
Realises | SYS-052, IFR-035 · Maps to: I-02 · Criticality: High.
Status/evidence | Preliminary (reconcile with HAL); shipper set defined by the target fleet (D1/D4). Judgment.
INT-006 — Enclosure ↔ probe lead entry and strain relief
Type | MECH · Endpoints: enclosure sealed body (S1) ↔ external-probe lead (S1) · Direction: ↔.
Medium/parameters | IP-sealed gland/grommet and strain relief where the probe lead crosses the sealed boundary; keyed, sealed probe connector (pinout {sense+, sense−, ID/presence, shield/GND}, IFR-003). The seal integrity here is part of the enclosure IP rating.
Realises | SYS-052, IFR-003 · Maps to: I-02 · Criticality: High.
Status/evidence | Preliminary; sealing method and connector keying flagged to Mechanical/Materials engineer. Judgment.
INT-007 — Enclosure ↔ sealed USB-C port aperture
Type | MECH · Endpoints: enclosure (S1) ↔ USB-C receptacle aperture (S1) · Direction: ↔.
Medium/parameters | sealed aperture/gasket around the USB-C receptacle maintaining the enclosure IP rating while allowing charging; the electrical contract of the port is INT-008.
Realises | SYS-052, SYS-054, IFR-001 · Maps to: I-02 · Criticality: High.
Status/evidence | Baselined as a sealed-port requirement; seal solution flagged to Mechanical/Materials (cleaning-chemical compatibility). Judgment.
INT-008 — USB-C source → port VBUS
Type | PWR · Endpoints: depot USB-C charger (EXT) → USB-C receptacle (S1) · Direction: → (sink).
Medium/parameters | full 24-pin Type-C receptacle wired sink-only (UFP): VBUS+GND on all four power pairs, single CC with Rd (5.1 kΩ ±10 %) advertising sink role and orientation. Input envelope: default 5 V at the current the source advertises (up to 3.0 A); where USB-PD is present, negotiate no PDO above a bounded cap (≤9 V / W to-confirm) and reject higher-voltage PDOs; remain safe on a legacy 5 V charger and under reverse-insertion/OV/short. Never sources VBUS, never presents Rp.
Realises | SYS-053, SYS-054, IFR-001, IFR-006 · Maps to: I-03 · Criticality: High.
Status/evidence | Baselined (sink-only role, Rd) with the PD power/voltage cap Evidence-Pending (IFR-006, closed with S2 charge-rate/thermal budget). Standard (USB Type-C / USB-PD; IEC 62368-1; revision to confirm).
INT-009 — Port → PMIC input
Type | PWR · Endpoints: USB-C port (S1) → PMIC / charge-path input (S2) · Direction: →.
Medium/parameters | VBUS conductor into the input-protection stage (OVP/reverse/ESD) feeding the charger front-end; input clamp within the device's input-protection rating (IFR-006).
Realises | SYS-053, IFR-006 · Maps to: I-03 · Criticality: High.
Status/evidence | Baselined; input-protection rating tied to the PD cap (INT-008, Evidence-Pending). Standard.
INT-010 — PMIC → BMS → cell charge path
Type | PWR · Endpoints: PMIC/charger (S2) → BMS → low-temp Li-ion cell (S2) · Direction: →.
Medium/parameters | cell charge current, gated by charge-inhibit (INT-016) whenever cell temperature is below the low-temperature threshold (0 °C nominal) or above the high limit; optional self-heat-before-charge (AD-3). Charge never proceeds into a cold cell.
Realises | SYS-053, IFR-007 · Maps to: I-03 · Criticality: Critical (charge-safety leg of D3).
Status/evidence | Baselined (charge-inhibit is the safety control). Standard (Li-ion charge-temperature limits per cell datasheet; IEC 62368-1; revision to confirm).
INT-011 — Cell ↔ BMS protection/gauge
Type | PWR · Endpoints: Li-ion cell (S2) ↔ BMS + fuel-gauge (S2) · Direction: ↔ (charge + discharge + sense).
Medium/parameters | cell terminals plus voltage/current/temperature sense for protection (OV/UV/OC/OT) and fuel-gauging; discharge to the rails (INT-012…INT-015) and charge from INT-010 both cross this boundary. Reported to the MCU over INT-020.
Realises | SYS-053, IFR-007, IFR-008 · Maps to: I-03 · Criticality: Critical.
Status/evidence | Baselined; cell chemistry and protection thresholds flagged to Battery/Power engineer (D3 top risk). Standard.
INT-012 — BMS → MCU / always-on rail
Type | PWR · Endpoints: BMS/regulator (S2) → MCU always-on domain (S3) · Direction: →.
Medium/parameters | regulated always-on rail sustaining the MCU deep-sleep + RTC domain between samples (the energy governor lives here, AD-2); nominal voltage/tolerance to confirm by the S2 energy-budget analysis.
Realises | SYS-055, IFR-009 · Maps to: I-04 · Criticality: High.
Status/evidence | Evidence-Pending (rail voltages/tolerances closed by endurance analysis). Judgment.
INT-013 — BMS → sensing rail (gated)
Type | PWR · Endpoints: BMS/regulator (S2) → sensing subsystem (S4) · Direction: →.
Medium/parameters | firmware-gated rail so sensors draw energy only during a scheduled sample; gate transitions must never corrupt an in-flight NVM write (coordinated with INT-021, SYS-057).
Realises | SYS-055, IFR-009 · Maps to: I-04 · Criticality: High.
Status/evidence | Evidence-Pending (rail budget). Judgment.
INT-014 — BMS → connectivity rail (gated)
Type | PWR · Endpoints: BMS/regulator (S2) → radios (S5) · Direction: →.
Medium/parameters | hard-gated rail to the dominant energy load (radios); powered only during scheduled upload/alert windows (AD-2). Inrush/hold-up sized for modem transmit bursts (to confirm).
Realises | SYS-055, IFR-009 · Maps to: I-04 · Criticality: High.
Status/evidence | Evidence-Pending. Judgment.
INT-015 — BMS → HMI + storage rails (gated)
Type | PWR · Endpoints: BMS/regulator (S2) → OLED/LED (S6) and NVM (S7) · Direction: →.
Medium/parameters | gated HMI rail (OLED on only during a read/verdict interaction) and a storage rail whose down-transition is interlocked against active writes (SYS-057). Cold legibility of the OLED is a PRD-UX constraint flagged to HMI.
Realises | SYS-055, SYS-057, IFR-009 · Maps to: I-04 · Criticality: High.
Status/evidence | Evidence-Pending (rail budget). Judgment.
INT-016 — Charge-inhibit control signal
Type | DATA/CTRL · Endpoints: BMS (S2) ↔ PMIC (S2) ↔ MCU (S3) · Direction: ↔.
Medium/parameters | discrete inhibit line (BMS→PMIC gating cell charge current to zero) plus a status/cause reported to the MCU over the fuel-gauge link (INT-020). Trips below 0 °C nominal / above the high-temp limit; the inhibited state and its cause are made legible (SYS-024) rather than presenting as a silent failure to charge.
Realises | SYS-053, IFR-007, IFR-008 · Maps to: I-03/I-04 · Criticality: Critical.
Status/evidence | Baselined (AD-3, D3 charge-safety control). Standard.
INT-017 — Internal sensor bus → MCU
Type | DATA · Endpoints: internal temperature/humidity/accelerometer (S4) → MCU (S3) · Direction: →.
Medium/parameters | I2C at 100/400 kHz (Standard/Fast) with a fixed documented 7-bit address per device, and/or SPI with a dedicated chip-select per device. Each reading is bound to its source address/CS; a NAK, CRC/PEC failure, or bus timeout returns a defined error to acquisition rather than a stale/substituted value (feeds SYS-007 plausibility flagging). No two devices share an address (address map fixed in the ICD).
Realises | SYS-056, SYS-003, IFR-004 · Maps to: I-05 · Criticality: Critical.
Status/evidence | Baselined. Standard (I2C spec addressing/PEC; SPI de-facto; revision to confirm).
INT-018 — External probe channel → MCU
Type | DATA · Endpoints: external-probe sensing front-end (S4) → MCU acquisition (S3) · Direction: →.
Medium/parameters | analog thermistor excitation/return or a 1-Wire/I2C digital sensor bus on the keyed probe connector, plus a presence/ID line so an absent, wrong-type, or open/short-circuit probe is detected and flagged (SYS-007/SYS-024) rather than read as a valid product-zone temperature.
Realises | SYS-056, SYS-003, IFR-003, IFR-004 · Maps to: I-05 · Criticality: Critical (an undetected open probe reading room temperature would fabricate a false "accept").
Status/evidence | Preliminary — probe technology (thermistor vs digital) to confirm with S4 (IFR-003). Judgment.
INT-019 — GNSS receiver → MCU
Type | DATA · Endpoints: GNSS receiver (S4) → MCU (S3) · Direction: →.
Medium/parameters | UART or I2C (NMEA/binary) carrying fixes and per-attempt fix quality (fix type, satellites, HDOP); the MCU schedules duty-cycled/event-triggered fixes within the energy budget and records last-known/best-available with a quality flag when no fix is obtained inside the shipper.
Realises | SYS-056, SYS-029, IFR-013 · Maps to: I-05 · Criticality: Medium (attribution support, never a safety function).
Status/evidence | Evidence-Pending — in-container fix availability/TTFF/per-fix energy closed by PRD-VER-005. Judgment.
INT-020 — Fuel-gauge → MCU
Type | DATA · Endpoints: BMS/fuel-gauge (S2) → MCU (S3) · Direction: →.
Medium/parameters | I2C/SMBus (gauge address per the INT-017 map) reporting at least state-of-charge, voltage, cell temperature, charge/discharge state, and charge-inhibit/fault flags (INT-016) on a defined polling cadence. Enables pre-arm charge readiness (SYS-024), the energy governor (SYS-055/SYS-108), and honest low-battery indication; an implausible/NAK read is flagged (SYS-007), not trusted.
Realises | SYS-055, SYS-024, IFR-008 · Maps to: I-04 · Criticality: High.
Status/evidence | Baselined; gauge part/protocol fixed in S2 detailed design. Judgment.
INT-021 — MCU → NVM (log / config / audit)
Type | DATA · Endpoints: MCU (S3) → non-volatile storage (S7) · Direction: →.
Medium/parameters | SPI to non-volatile memory carrying the append-only sample/event log, configuration, and audit trail. Writes are atomic / power-fail-safe — a write is either fully committed or detectably incomplete on power loss (no torn/partial silent records) — and records carry tamper-evidence enabling later detection of alteration/deletion. A rail-down (INT-013/INT-015) or brown-out during a write must not fabricate or lose a record (ALCOA+ enduring).
Realises | SYS-057, SYS-018, IFR-035 · Maps to: I-06 · Criticality: Critical (this is where the evidentiary record is made durable).
Status/evidence | Baselined. Standard (ALCOA+, 21 CFR Part 11; revision to confirm).
INT-022 — MCU → OLED display
Type | DATA · Endpoints: MCU (S3) → OLED (S6) · Direction: →.
Medium/parameters | SPI/I2C framebuffer driving the accept / reject / quarantine verdict and the incomplete-monitoring qualifier (SYS-017). The displayed verdict must be consistent with the MCU's authoritative state, and a stuck/failed element must be detectable so a failed segment never turns "reject" into "accept" (SYS-058). Cold- and glare-legibility are PRD-UX constraints (HMI).
Realises | SYS-058, SYS-015, IFR-035 · Maps to: I-07 · Criticality: High (last-mile decision surface, AD-6).
Status/evidence | Baselined. Judgment.
INT-023 — MCU → status LED / button → MCU
Type | DATA · Endpoints: MCU (S3) ↔ status LED and single glove-friendly button (S6) · Direction: ↔.
Medium/parameters | GPIO/PWM to the indicator and a debounced GPIO input from the single button; a single actuation maps to a single unambiguous action (SYS-058), including the explicit arm action and wake-to-read. No multi-press ambiguity.
Realises | SYS-058, SYS-023, IFR-035 · Maps to: I-07 · Criticality: High.
Status/evidence | Baselined. Judgment.
INT-024 — MCU ↔ secure element
Type | DATA · Endpoints: MCU (S3) ↔ secure element (S3) · Direction: ↔.
Medium/parameters | SPI/I2C to the secure element holding the per-device keys, the mTLS client cert/key (INT-032), the OTA trust anchor (INT-035), and secure-boot measurements. Keys never leave the SE; this boundary underpins device identity and non-repudiation (SYS-068).
Realises | SYS-057, SYS-068, IFR-032 · Maps to: I-06 · Criticality: Critical.
Status/evidence | Baselined; key-management/threat-model detail flagged to Cybersecurity (Evidence-Pending threat model). Standard (IEC 62443 concepts; revision to confirm).
INT-025 — MCU ↔ radio host interface
Type | DATA · Endpoints: MCU (S3) ↔ LTE-M / Wi-Fi / BLE radios (S5) · Direction: ↔.
Medium/parameters | the on-board host control link (UART/SPI/SDIO + AT-command or vendor host API) by which firmware drives the radios, selects the bearer (LTE-M primary, Wi-Fi/BLE offload), and manages PSM/eDRX sleep. Distinct from the over-the-air links (INT-026…INT-031). Bounded timeout/retry per IFR-034.
Realises | SYS-062, SYS-019, IFR-010, IFR-034 · Maps to: I-08/I-09 · Criticality: High.
Status/evidence | Baselined; host protocol fixed in S5 detailed design. Judgment.
INT-026 — Radio ↔ antenna feed network
Type | RF · Endpoints: radios (S5) ↔ antenna(s) (S1/S5) · Direction: ↔.
Medium/parameters | RF traces/coax and matching network from each radio to the antenna(s), including provision for a qualified IP-sealed external-antenna connector option (AD-1). Feed loss and match are inputs to the in-container link budget (INT-031).
Realises | SYS-059, IFR-014 · Maps to: I-08 · Criticality: High.
Status/evidence | Evidence-Pending — integral-vs-external antenna unresolved; closed by PRD-VER-005. Judgment.
INT-027 — LTE-M ↔ cellular network
Type | RF · Endpoints: LTE-M modem (S5) ↔ cellular network (EXT) · Direction: ↔.
Medium/parameters | 3GPP Cat-M1 with a market-selectable band set, operator APN(s), and PSM (TAU/Active-Timer) and/or eDRX sized to the upload cadence and alert-latency target — the modem sleeps between sessions to preserve the 90-day budget yet stays reachable within the alert window (D2/D3). Must not transmit outside region-permitted bands/limits.
Realises | SYS-059, SYS-019, IFR-010 · Maps to: I-08 · Criticality: High (primary alert/upload bearer).
Status/evidence | Evidence-Pending — per-market band plan, PSM/eDRX values, carrier/APN strategy unresolved. Standard (3GPP LTE-M/Cat-M1; revision to confirm).
INT-028 — Wi-Fi ↔ access point
Type | RF · Endpoints: Wi-Fi radio (S5) ↔ depot access point (EXT) · Direction: ↔.
Medium/parameters | IEEE 802.11 with at minimum WPA2-PSK, and where available WPA3-SAE and WPA2/WPA3-Enterprise (EAP). Credentials accepted only over the authenticated provisioning path (BLE GATT INT-033 or authenticated cloud config); on association/auth failure the device fails closed (retains buffered data for another bearer) rather than joining an open network.
Realises | SYS-059, SYS-019, IFR-011 · Maps to: I-08 · Criticality: Medium (offload bearer).
Status/evidence | Evidence-Pending — enterprise-Wi-Fi/EAP method support to confirm. Standard (IEEE 802.11, WPA2/WPA3; revision to confirm).
INT-029 — BLE ↔ mobile app radio link
Type | RF · Endpoints: BLE radio (S5) ↔ mobile app (S9) · Direction: ↔.
Medium/parameters | Bluetooth LE peripheral advertising a defined service UUID with advertising interval/TX power tuned to depot/last-mile range; connection parameters (interval, slave latency, supervision timeout) sized for provisioning throughput while energy-bounded; advertising/connectable window constrained when idle/armed so the radio is not a standing exposure surface. Carries the logical GATT interface INT-033.
Realises | SYS-060, IFR-012, IFR-016 · Maps to: I-09 · Criticality: High.
Status/evidence | Baselined (advertising/connection params); pairing method reconcile with threat model. Standard (Bluetooth Core LE; revision to confirm).
INT-030 — GNSS constellation → receiver
Type | RF · Endpoints: GNSS constellation (EXT) → receiver antenna (S4) · Direction: →.
Medium/parameters | L1 multi-constellation (GPS + at least one of GLONASS/Galileo/BeiDou) with A-GNSS assist data delivered over the cloud link (INT-032) to shorten TTFF and per-fix energy. GNSS is the most attenuated band inside a metallized shipper.
Realises | SYS-029, IFR-013 · Maps to: I-08 · Criticality: Medium.
Status/evidence | Evidence-Pending — in-container fix availability/TTFF/energy closed by PRD-VER-005. Judgment.
INT-031 — In-container RF link (through the shipper)
Type | RF · Endpoints: device antennas (S1/S5) ↔ cellular/Wi-Fi/GNSS (EXT) through the metallized insulated shipper · Direction: ↔.
Medium/parameters | the aggregate attenuated RF path the device must close from inside a representative shipper to deliver buffered telemetry, excursion alerts, and location fixes at the delivery-success target set by the RF-in-container test. On failure to close a link the device degrades to store-and-forward buffering (INT-021), never record loss (D2, SYS-013).
Realises | SYS-059, IFR-014 · Maps to: I-08 · Criticality: Critical (principal open architectural risk).
Status/evidence | Evidence-Pending — attenuation, achievable margin, and integral-vs-external antenna unresolved; closed by the RF-in-container qualification test (PRD-VER-005). Until then, upload and alert egress are never a dependency for the record. This is the honest headline gap of the whole interface set. Evidence-Pending.
INT-032 — Device ↔ cloud (store-and-forward + alert)
Type | DATA · Endpoints: device client (S3/S5) ↔ cloud ingest (S8) · Direction: ↔.
Medium/parameters | MQTT v5 (or HTTPS for record/OTA fetch) over TLS 1.2/1.3 with mutual authentication — device presents a per-device X.509 client cert/key held in the secure element (INT-024) and validates the server against a pinned CA/trust anchor. Topic namespace scm/{tenant}/{deviceId}/{telemetry¦alert¦status¦cmd¦ota}; every message carries schema-version, deviceId, monotonic sequence number. Store-and-forward semantics (D2): ordered, exactly-once ingest by de-duplicating on (deviceId, sequence), resume from last acknowledged sequence, and ack-gated deletion (buffered data marked delivered only on PUBACK/2xx). Excursion alerts are a distinct high-priority message sent ahead of bulk backlog and ack-required (SYS-011/IFR-025), while the same event is always already committed locally (SYS-013). Malformed/unknown-version messages are rejected/quarantined, and bidirectional back-pressure (MQTT v5 receive-maximum / backoff+jitter) is honoured (IFR-026/IFR-034). Failed handshake → retain data, retry, never cleartext fallback.
Realises | SYS-062, SYS-020, SYS-011, SYS-063, SYS-068, IFR-021…IFR-026 · Maps to: I-10 · Criticality: Critical.
Status/evidence | Baselined (transport, mutual-auth, store-and-forward, alert precedence). Encoding (JSON vs CBOR) Evidence-Pending against the RF payload budget (IFR-023). Standard (MQTT v5, TLS 1.2+/1.3, ALCOA+; revision to confirm).
INT-033 — Device ↔ BLE app (GATT)
Type | DATA · Endpoints: device GATT server (S3/S5) ↔ mobile app (S9), carried on INT-029 · Direction: ↔.
Medium/parameters | a versioned custom GATT service (128-bit vendor UUID) with fixed characteristics: Protocol/Schema-Version (read), Device-Info/Identity (read), Trip-Profile-Write (write, authenticated — atomic, validated, CRC/hash, length-prefixed), Arm/Disarm-Control (write, authenticated — nonce/sequence anti-replay, readiness-gated), Local-Verdict/Record-Read (read/notify — login-free summary per AD-5, authenticated full record read as an ordered, resumable notification stream with an end-of-record marker), and Status/Notification (notify). LE Secure Connections pairing with an MITM-protected method (numeric-comparison/passkey, not Just Works) is required before any authenticated characteristic; writes to unknown handles or out-of-contract values are rejected with a defined GATT error. ATT MTU target ≥185 B, DLE/2M PHY where supported, with acknowledged-indication flow control and bounded timeouts.
Realises | SYS-060, SYS-061, SYS-022, SYS-023, SYS-068, IFR-015…IFR-020 · Maps to: I-09 · Criticality: High (Critical for the authenticated arm/profile paths).
Status/evidence | Baselined (GATT structure, atomic profile write, verdict read); pairing/ association method Preliminary (reconcile with cybersecurity threat model). Standard (Bluetooth GATT; revision to confirm).
INT-034 — Cloud ↔ enterprise (REST/API + export)
Type | DATA · Endpoints: cloud platform (S8) ↔ enterprise ERP/QMS (EXT) · Direction: ↔.
Medium/parameters | a documented REST API specified by a published OpenAPI definition exposing trip records, alerts, and fleet/device status in a versioned, machine-readable format (JSON, with PDF/CSV export) preserving attribution, time, gaps, and tamper-evidence. OAuth 2.0 bearer tokens scoped to least privilege, integrated with the customer's SSO/IdP, with RBAC on every resource and audit-logging of every access; cursor-based pagination for deterministic full export, documented rate limits (429/Retry-After), and signed retryable webhooks for pushed alerts. Explicit API versioning with a deprecation window. This control does not gate the login-free on-device edge verdict (AD-5).
Realises | SYS-064, SYS-027, SYS-068, IFR-027…IFR-030 · Maps to: I-11 · Criticality: High (Critical for auth/RBAC, IFR-028).
Status/evidence | Baselined. Standard (OpenAPI, OAuth 2.0, EU GDP, 21 CFR Part 11; revision to confirm).
INT-035 — OTA service → device (signed firmware)
Type | DATA · Endpoints: OTA service (S8) → device updater (S3/S5) · Direction: → (fetch/apply), status ← reported over INT-032 · Medium/parameters: a defined update package + manifest (target hardware/model, from/to version, per-image digest) + signature block; the device determines applicability from the manifest and rejects wrong-model/downgrade packages before writing flash. Each image's signature is verified against a trust anchor in the secure element (INT-024) and its digest checked before activation — failed verification aborts and leaves the running known-good image untouched. Resumable (byte-range) transfer over LTE-M/Wi-Fi/USB-C, fail-safe A/B apply with automatic rollback on failed activation/self-test, trip data preserved, and per-device status (downloading/verified/applied/ rolled-back/failed-with-reason) reported to the fleet campaign. Never bricks a device or leaves it unmonitored-without-notice (AD-7).
Realises | SYS-065, SYS-032, SYS-068, IFR-031…IFR-033 · Maps to: I-12 · Criticality: Critical (fleet-wide integrity and fail-safety surface).
Status/evidence | Baselined. Standard (secure-boot / signed-update, A/B rollback, IEC 62443 concepts; revision to confirm).
---
4. Interface conformance & ICD governance
Single authoritative ICD (SYS-066 / IFR-035). This register, together with IFR-001…IFR-035, is the authoritative Interface Control Document. Every realised subsystem interface shall conform to the typed inventory I-01…I-12 — matching the declared type (PWR/DATA/RF/THRM/MECH), direction, signals, format, and contract — and no undocumented interface shall exist between subsystems or across the device/cloud/ mobile trust boundaries. An undocumented interface is treated as both an integration defect and an ungoverned attack surface.
Traceability closure. Each INT-### traces down to at least one IFR-### and up to exactly one (or a small named set of) I-0N. The mapping in §2 is bidirectional with the block diagram: the twelve I-0N interfaces each expand into one or more INT-### rows, and every INT-### names its parent I-0N. Coverage of I-01…I-12 by this register: I-01 → INT-001…004; I-02 → INT-005…007; I-03 → INT-008…011, INT-016; I-04 → INT-012…015, INT-016, INT-020; I-05 → INT-017…019; I-06 → INT-021, INT-024; I-07 → INT-022…023; I-08 → INT-025…028, INT-030, INT-031; I-09 → INT-025, INT-029, INT-033; I-10 → INT-032; I-11 → INT-034; I-12 → INT-035. (INT-016 and INT-025 intentionally span two typed interfaces; this is recorded, not hidden.)
Change control. Any interface change shall be governed through an ICD revision with a backward-compatibility assessment (IFR-035, IFR-030). Wireless and cross-tier protocol changes (INT-032/INT-033/INT-034/INT-035) additionally require schema/version-negotiation review so fielded devices and cloud stay interoperable across staged OTA version skew (SYS-063/SYS-065, IFR-022).
Error/timeout/back-pressure uniformity (SYS-067 / IFR-034). Every data, RF, and power interface (INT-012…INT-035 as applicable) carries a documented timeout / retry-with-backoff+jitter / back-pressure contract with a defined safe, energy-bounded fallback (buffer, degrade, or safe-state) on exhaustion — an interface that can hang or spin retries is a data-integrity and an energy hazard under the D3 budget. This contract is Preliminary pending HAL reconciliation (Stage 10).
Authenticated identity at wireless/cross-tier boundaries (SYS-068). INT-029/INT-032/INT-033/INT-034/ INT-035 require each party to cryptographically authenticate the other before privileged operations (configuration, arming, ingest, firmware acceptance), keyed from the secure element (INT-024), so every such interaction is attributable (non-repudiation, 21 CFR Part 11). This is Preliminary pending the cybersecurity threat model (Evidence-Pending).
Internal test interface. The debug/EOL-test port (IFR-005) is governed here as a manufacturing/security interface: it enables factory programming/calibration/test and is locked in fielded units (debug disabled or authentication-gated, readout-protection/lifecycle fuse set) so a fielded device exposes no unauthenticated code, key, or record readout.
Open-gap ledger (honest evidence). The interfaces carried Evidence-Pending are, and remain until their named test closes them: INT-001…004 (thermal probe offset/lag — thermal-characterisation test), INT-012…015 (rail voltages/tolerances — S2 energy-budget analysis), INT-019/INT-030 (GNSS-in-container cadence/TTFF — PRD-VER-005), INT-026/INT-027/INT-028/INT-031 (RF-in-container link budget and antenna strategy — PRD-VER-005, the principal open architectural risk, AD-1), part of INT-008 (USB-PD power cap — S2 thermal budget), and part of INT-032 (telemetry encoding — RF payload budget). Preliminary interfaces (INT-005/006, INT-018, the SYS-067/SYS-068 governance contracts, and the BLE pairing method) reconcile with the Hazard Analysis Library and the cybersecurity threat model. No Evidence-Pending value is asserted as fact anywhere in this document.
System Architecture — Logical & Physical Architecture — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 08 — System Architecture · Owner: Chief Systems Engineer (EM-ENG-001) · Status: Authored — instantiated, product-specific · Traceability: ORCH-018 / EM-ENG-012.
This file is the authoritative architecture description for the SCM. It converts the Stage-05 product architecture (../05-product-requirements/design-architecture.md, AD-1…AD-8; and ../05-product-requirements/system-block-diagram.md, S1–S9 / I-01…I-12) and the Stage-06 system requirements (../06-system-requirements/, SYS-001…SYS-190) into logical and physical views, a subsystem decomposition, and the traceability from architecture drivers to architectural mechanisms. It is based on ../00-concept/concept-and-decisions.md (foundational decisions D1–D4).
Companion Stage-08 files, authored in parallel, hold the elemental detail this file references:
• Functional decomposition — functional-decomposition.md (FUN-###): the leaf functions of the logical layers below.
• Component list — component-list.md (CMP-###): the physical parts realising the subsystems below.
• Interface control — interface-control.md (INT-###): the realised contracts for I-01…I-12.
Evidence posture (honesty). Every claim below is labelled Standard (a cited normative document, revision to confirm), Cited (a named external source), Judgment (engineering reasoning stated as such), or Evidence-Pending (a quantity or decision not yet resolved and closed only by test/analysis). No standard clause number is asserted as fact and no unresolved quantity is invented (per 00-governance/evidence-standard.md). The two architecturally load-bearing open risks — RF-in-container link performance (I-08) and the −30 °C / 90-day energy budget (D3) — are carried explicitly as Evidence-Pending and flagged throughout.
---
1. Architectural intent
The SCM is architected as an offline-first edge node with opportunistic connectivity — deliberately not a live-streaming telemetry endpoint (D2, AD-1). It is organised around one value chain, capture → detect → warn → decide → prove, under a single overriding constraint:
Survive 90 days at −30 °C without dying or losing data (D3 — the top program risk).
Every architectural decision is a resolution of that constraint against the other drivers (measurement accuracy, offline verdict, auditability, security, reusability). Three properties follow directly and shape the whole architecture:
1. The core decision path must be edge-autonomous. Capturing, timestamping, detecting excursions, storing the record, and rendering the accept/reject/quarantine verdict must all complete on the device, with no network, phone, login, or manual (AD-6; SYS-015). The last-mile user has none of those things. Connectivity is upside — it delivers alerts faster and moves the record to the cloud — but it is never a dependency (D2, AD-1; SYS-013/018). (Judgment, based on D2/AD-1.)
2. Energy is the master resource, and time is how we spend it. At −30 °C a rechargeable cell delivers a fraction of its rated capacity and cannot be charged at all below 0 °C (D3). The architecture therefore treats the duty-cycle — how long each subsystem is powered per unit time — as the primary design variable, and introduces an energy governor (the MCU power-state machine, S3) as a cross-cutting concern that gates every other subsystem (AD-2; SYS-055). The 90-day budget is met by duty-cycling, not chemistry alone (D3, AD-2). (Judgment; the numeric budget is Evidence-Pending, closed by the endurance test — SYS-086/090/092.)
3. Trust is split at the tier boundary. The on-device read is frictionless (no login) so the edge works for an untrained last-mile user; governance (RBAC/SSO, configuration control, 21 CFR Part 11-style audit) lives on the platform (AD-5; SYS-068). This split is an architectural fault line, not just a policy — it dictates where identity is proven and where records become authoritative.
The device is a monitoring/information device: it does not control temperature and is not the primary means of preserving the vaccine. Its safety contribution is informational — it supports the accept/reject disposition decision, in which both error directions harm the beneficiary (a spoiled dose administered, or a potent dose discarded). This makes no-false-reassurance an architectural invariant: an incomplete or faulted record must never render as a clean "accept" (SYS-002/007/017). (Judgment, based on the product description §1.)
---
2. Architectural views
Two complementary views describe the same system: the logical architecture (what the system does, as functional layers independent of hardware) and the physical architecture (what the system is made of, as three tiers and, within the device, six physical domains). The Mermaid figure at the end of this section maps one onto the other.
2.1 Logical architecture (functional layers)
The SCM's logical architecture is a layered pipeline with two constructs that do not belong to any one layer: the energy governor (a cross-cutting concern that meters every layer) and the split trust boundary (a fault line that runs vertically through the stack, separating edge-autonomous layers from platform-governed ones). Each layer is defined by responsibility, not by hardware.
# | Logical layer | Responsibility (what it does) | Realised primarily by | Key SYS refs
L1 | Sensing | Transduce the physical world — internal & external-probe temperature, humidity, 3-axis shock, and (duty-cycled) GNSS — into raw electrical/digital readings, each tied to a source channel. | S4 (+ S1 probe/thermal path) | SYS-001/003/004/005; SYS-051/056
L2 | Acquisition & timestamp | Sample on the configured cadence, bind each reading to a trusted UTC-referenced monotonic time base, tag channel identity and data-quality/plausibility status, and hand structured samples upward. | S3 (MCU/RTC), S4 | SYS-001/006/007; SYS-056
L3 | Detection & verdict logic | Apply two-sided (freeze/heat), product-relevant excursion logic (threshold × time/cumulative), flag excursions with direction/onset/duration, and compute the accept/reject/quarantine verdict — the edge-autonomous decision core. | S3 | SYS-008/009/010/014/015/016/017
L4 | Storage | Persist samples, events, excursions, gaps, config, and lifecycle actions to tamper-evident, power-fail-safe, append-only non-volatile memory, sized for a full 90-day trip. | S7 (+ S3 write logic) | SYS-002/013/018/021/026/028; SYS-057
L5 | Comms / store-and-forward | Buffer when offline; upload ordered, resumable, de-duplicated batches on cadence over the best available bearer; jump alerts ahead of bulk backlog when connected. | S5 (+ S3 scheduler) | SYS-011/019/020; SYS-059/062/067
L6 | Cloud services | Ingest (idempotent store-and-forward), device registry, rules/alert routing & escalation, trip-record assembly, audit trail, signed OTA, RBAC/SSO, export/API. | S8 | SYS-012/026/027/028/031/032; SYS-062/063/064/065/068
L7 | Presentation | Render decisions to humans: the edge presentation is the on-device OLED verdict + single button (offline, no-login); the platform presentation is the fleet dashboard, notifications, and the mobile provisioning/read app. | S6 (edge), S8/S9 (platform) | SYS-015/017/023/024; SYS-058
Cross-cutting concern — the energy governor. The MCU power-state machine (S3) is not a layer; it is a cross-cutting controller that meters L1–L5 on the device. It holds the device in deep sleep between samples, wakes the sensing/acquisition path only long enough to take and store a reading, and gates the high-draw rails (connectivity, sensing, HMI) so a subsystem draws energy only when the active power state requires it (AD-2; SYS-055). Radios (S5) are the dominant load and are gated hardest — this is why comms is scheduled/batched (L5) rather than continuous. The governor also owns the safety interlock that a rail-down must never corrupt an in-flight NVM write (L4; SYS-057/067). (Judgment; budget Evidence-Pending — SYS-086/090/092.)
Cross-cutting fault line — the split trust boundary (AD-5). A vertical boundary runs through the stack: L1–L4 and the edge half of L7 are edge-autonomous and login-free (frictionless read), while L6 and the platform halves of L5/L7 are platform-governed (identity proven, access controlled, records made authoritative, audit enforced). Every crossing of this line (BLE provisioning I-09, device↔cloud I-10, OTA I-12) must be mutually authenticated before any privileged operation and must be attributable in the audit trail (SYS-068; AD-5/AD-7). (Standard — IEC 62443 concepts, 21 CFR Part 11 attributable; revision to confirm. Threat model Evidence-Pending.)
Verdict consistency across the boundary. Because both the edge (L3/S3) and the cloud (L6/S8) compute a verdict, the architecture requires them to derive from the same excursion evaluation, and requires any divergence to be detectable and flagged — two disagreeing sources of truth would destroy trust in disposition (SYS-016; AD-6). (Judgment.)
2.2 Physical architecture
The physical architecture has three tiers, and the device tier is internally partitioned into six physical domains so that power, noise, RF, thermal, and trust concerns are physically separated.
Tier 1 — Device (S1–S7), inside the shipper. The edge node. Sits inside the insulated shipping container, thermally coupled to the payload (I-01) and mechanically retained (I-02), RF-attenuated by the shipper (I-08). Owns L1–L5 and the edge half of L7.
Tier 2 — Cloud (S8). The platform. Store-and-forward ingest, registry, rules/escalation, record assembly + audit, OTA, RBAC/SSO, dashboard, export/API. Owns L6 and the platform half of L5/L7.
Tier 3 — Mobile (S9). The provisioning/read companion. BLE provisioning, trip-profile configuration, arming aid, local read, notifications. A platform-governed actor at the edge that participates in L7 and in configuration (crosses the trust boundary over I-09).
Device-internal physical partitioning (six domains). Within Tier 1 the physical layout is partitioned into domains that align with S1–S7 but are organised by physical concern:
Physical domain | Subsystems | What it physically is / why it is a separate domain
Power domain | S2 | Cell + BMS + PMIC/charge path + fuel gauge + protection. Separated because it is the energy source, the charge-safety authority (charge-inhibit <0 °C, AD-3), and the highest-current path — kept away from sensitive sensing/RF.
Compute domain | S3 | Low-power MCU + RTC + secure element. The energy governor and the trust root physically live here; it is the master that clocks and gates every other domain.
Sensing domain | S4 (+ S1 probe) | Internal temp, external-probe temp channel, humidity, accelerometer, GNSS front-end. Physically separated for thermal quiet and low self-heating so the probe/enclosure represent the product zone (SYS-051).
RF domain | S5 | LTE-M, Wi-Fi, BLE, GNSS radios + antenna(s). Isolated for EMC and to manage the shipper-attenuated aperture — the antenna placement/aperture is shared with S1 and is the principal RF risk (I-08).
HMI domain | S6 | OLED + single glove-friendly button + status LED. Physically at the enclosure surface for cold/glare-legible, glove-operable, offline read.
Storage domain | S7 | Non-volatile memory for log + config + audit. Physically distinct so tamper-evident, power-fail-safe writes survive rail-downs and brown-outs (SYS-057).
(Enclosure/mechanical) | S1 | Not an electronic domain but the physical envelope that seals, retains, thermally couples (probe), and provides the RF aperture — it hosts elements of the sensing, power (port), RF (antenna), and HMI domains.
2.3 Logical-to-physical mapping (Mermaid)
The figure maps the logical layers (L1–L7, left) onto the physical tiers and device domains (right). The energy governor and the trust boundary are shown as the two cross-cutting constructs. Interfaces are the typed I-01…I-12 contracts (see interface-control.md, INT-###).
mermaid flowchart LR %% ================= LOGICAL VIEW ================= subgraph LOGICAL["Logical architecture (functional layers)"] direction TB L1["L1 Sensing"] L2["L2 Acquisition &amp; timestamp"] L3["L3 Detection &amp; verdict logic<br/>(edge-autonomous core)"] L4["L4 Storage<br/>(tamper-evident, fail-safe)"] L5["L5 Comms / store-and-forward"] L6["L6 Cloud services"] L7E["L7 Presentation - EDGE<br/>(OLED verdict, no login)"] L7P["L7 Presentation - PLATFORM<br/>(dashboard, app, alerts)"]
L1 --> L2 --> L3 --> L4 --> L5 --> L6 L3 --> L7E L6 --> L7P end
%% Cross-cutting concerns GOV{{"Energy governor (AD-2)<br/>cross-cutting: meters L1-L5"}} GOV -.gates.-> L1 GOV -.gates.-> L2 GOV -.gates.-> L3 GOV -.gates.-> L4 GOV -.gates.-> L5
%% ================= PHYSICAL VIEW ================= subgraph PHYS["Physical architecture (tiers &amp; device domains)"] direction TB subgraph DEV["Tier 1 - Device (inside shipper)"] PWRD["Power domain (S2)<br/>cell + BMS + PMIC"] CMPD["Compute domain (S3)<br/>MCU + RTC + secure element"] SNSD["Sensing domain (S4 + S1 probe)"] RFD["RF domain (S5)<br/>LTE-M / Wi-Fi / BLE / GNSS + antenna"] HMID["HMI domain (S6)<br/>OLED + button + LED"] STOD["Storage domain (S7)<br/>non-volatile log"] end subgraph CLD["Tier 2 - Cloud (S8)"] CLOUD["Ingest / registry / rules /<br/>record + audit / OTA / RBAC-SSO"] end subgraph MOB["Tier 3 - Mobile (S9)"] APP["BLE provisioning /<br/>config / local read"] end end
%% ---- logical layer -> physical realiser ---- L1 --- SNSD L2 --- CMPD L3 --- CMPD L4 --- STOD L5 --- RFD L6 --- CLOUD L7E --- HMID L7P --- CLOUD L7P --- APP GOV --- CMPD PWRD -.powers all domains.- CMPD
%% ---- typed cross-boundary interfaces ---- RFD -. "I-08 RF thru shipper (Evidence-Pending)" .-> CLOUD APP -. "I-09 BLE (trust boundary)" .-> CMPD RFD -. "I-10 device to cloud (trust boundary)" .-> CLOUD CLOUD -. "I-12 signed OTA (trust boundary)" .-> CMPD
%% ---- trust boundary annotation ---- TB["Split trust boundary (AD-5):<br/>edge = login-free read; platform = RBAC/SSO + Part 11 audit"] TB -.-> APP TB -.-> CLOUD
Figure 1 — SCM logical layers mapped to physical tiers and device domains. The edge-autonomous core (L1–L4 + edge L7) lives entirely in the device tier; the energy governor (S3) gates L1–L5; the split trust boundary (AD-5) is crossed only by the authenticated interfaces I-09/I-10/I-12; RF-in-container (I-08) is the principal open risk. See system-block-diagram.md Figure 1 for the physical block wiring and interface-control.md for the realised interface contracts.
---
3. Subsystem decomposition (S1–S9)
Each subsystem below is decomposed into its constituent blocks (the elements that will become CMP-### in component-list.md), with the requirement families it owns and its principal design decisions. Interface contracts (INT-###) realise the I-01…I-12 typed boundaries; leaf functions (FUN-###) live in functional-decomposition.md.
Subsystem | Role | Constituent blocks (→ CMP-###) | Owns (SYS / requirement families) | Key design decisions & evidence
S1 Enclosure / mechanical | Seals, retains, thermally couples, and provides the RF aperture | Sealed IP-rated housing; external temperature probe + lead; mount/retention; sealed USB-C port; antenna aperture/window; durable labeling | SYS-051/052 (THRM/MECH boundaries); SYS-054 (port); part of SYS-059 (antenna placement); PRD-MECH, PRD-PHY, part of PRD-ENV | Probe represents the product zone (AD-4); sealed port and cleanable housing serve the reuse loop (D1); antenna aperture is the mechanical half of the I-08 RF risk. (Judgment; probe offset/lag budget Evidence-Pending — SYS-051.)
S2 Power | Stores/manages energy for 90-day cold endurance; safe charge | Low-temp Li-ion cell; BMS + protection (OV/UV/OC/OT); PMIC/charge path; fuel gauge; charge-inhibit-<0 °C control | SYS-086…SYS-103 (power/energy); SYS-053/054 (PWR interfaces); SYS-055 (rails); PRD-ELEC, PRD-REL, part of PRD-SAFE | Low-temp-tolerant chemistry + charge-inhibit <0 °C + self-heat-before-charge under evaluation (D3/AD-3); 90-day budget met by duty-cycling not chemistry (AD-2). (Standard — IEC 62368-1, UN 38.3; cell low-temp limits per datasheet; revision to confirm. Chemistry & budget Evidence-Pending.)
S3 Compute / MCU + FW | Runs capture→detect→decide; is the energy governor and trust root | Low-power MCU; RTC (trusted clock); secure element/secure boot; watchdog; firmware application (power-state machine, detection logic, storage/comms schedulers) | SYS-006 (time), SYS-007…SYS-017 (detect/decide), SYS-055 (rail control), SYS-057 (writes), SYS-032/065 (OTA), SYS-068 (auth); PRD-FUN, PRD-SEC (boot/keys), PRD-PERF, PRD-REL | MCU power-state machine is the cross-cutting energy governor (AD-2); RTC gives trusted timestamps (SYS-006); secure element roots the trust split (AD-5) and signed OTA (AD-7). (Judgment + Standard — Part 11, IEC 62443 concepts; revision to confirm.)
S4 Sensing | Measures the monitored environment | Internal temp sensor; external-probe temp channel; humidity sensor; 3-axis accelerometer; GNSS/GPS receiver | SYS-001/003/004/005/029 (acquire); SYS-056 (bus); part of SYS-051; PRD-FUN, PRD-PERF, part of PRD-MECH | Dual-channel temperature for two-sided detection and product-zone representation (AD-4); GNSS duty-cycled/event-triggered inside the energy budget (D3). (Judgment; ±0.5 °C accuracy target and GNSS-in-container cadence Evidence-Pending — SYS-029.)
S5 Connectivity | Moves data/alerts out; provisioning bearer | LTE-M modem (primary); Wi-Fi; BLE; antenna(s) (shared aperture with S1) | SYS-011/019/020 (upload/alert), SYS-104…SYS-118 (connectivity/RF), SYS-059/060/061/062 (RF/data interfaces); PRD-CONN, part of PRD-ENV (RF) | Multi-bearer opportunistic upload, no single bearer required (AD-1/D2); radios gated hardest by the governor (AD-2); antenna strategy fights shipper attenuation. (Judgment; RF-in-container link budget & antenna strategy Evidence-Pending — SYS-059, principal open risk.)
S6 HMI | On-device status/verdict + single-action control | OLED display; single glove-friendly button; status LED | SYS-015/017/023/024 (verdict/arm/readiness), SYS-058 (HMI interface); PRD-UX, part of PRD-MECH | Layered information: unambiguous edge verdict, detail in cloud (AD-6); cold/glare-legible, glove-operable, offline, no-login (AD-5 edge half). A failed segment must not turn "reject" into "accept" (SYS-058). (Judgment.)
S7 Storage | Retains the full trip log tamper-evidently | Non-volatile memory (log + config + append-only audit) | SYS-002/013/018/021/026/028 (record/buffer/audit), SYS-057 (fail-safe writes); PRD-FUN, PRD-SEC, PRD-REL | Sized for full 90-day trip at min interval with margin; append-only, tamper-evident, power-fail-safe writes (AD-6/AD-7 record integrity). (Standard — ALCOA+, Part 11; revision to confirm. Storage sizing Evidence-Pending — SYS-018/127.)
S8 Cloud + dashboard | Fleet registry, ingest, routing, record assembly, audit, OTA, APIs | Device registry; store-and-forward ingest; rules/alerting/escalation; fleet dashboard + reporting/export; OTA service; RBAC/SSO + audit log | SYS-012/026/027/028/030/031/032 (cloud functions), SYS-062/063/064/065/068 (cloud/OTA/auth interfaces); PRD-CLOUD, PRD-SEC, part of PRD-REG | Platform-governed half of the trust split (AD-5); idempotent ingest + ordered/resumable transfer (SYS-020/062); signed fail-safe OTA (AD-7); Part 11-style audit trail. (Standard — EU GDP, 21 CFR Part 11, IEC 62443 concepts; revision to confirm.)
S9 Mobile / BLE app | Provisioning, config, arming aid, local read, notifications | BLE pairing/bonding; trip-profile config UI; arming aid; local read; notification client | SYS-022/023/024/025 (config/arm), SYS-060/061/068 (BLE/auth); PRD-MOB, part of PRD-UX/PRD-CONN | Edge read stays login-free while the platform enforces governance (AD-5); provisioning writes authenticated/versioned so a stale/rogue app cannot misconfigure a device (SYS-061/068). (Standard — Bluetooth GATT; IEC 62443 concepts; revision to confirm.)
Evidence note. Subsystem allocation mirrors design-architecture.md §2 and the SYS allocation targets; the block decomposition is the Stage-08 refinement that seeds component-list.md (CMP-###). Numeric budgets (energy, RF link, storage sizing, probe offset) are set as requirements and closed by verification/validation — they are Evidence-Pending, never invented.
---
4. Layering & separation-of-concerns rationale
Why an edge-autonomous core (L1–L4 + edge L7). The decisive user — the last-mile receiver — has no phone, no network, no login, and no manual, yet must make an accept/reject/quarantine call (AD-6; SYS-015). If any part of the capture→detect→decide→store chain depended on connectivity, the last mile would become a black box exactly where the decision is made. So the architecture places all four core layers, plus the edge presentation, entirely inside the device tier and makes connectivity strictly additive (D2/AD-1; SYS-013/018). This is also a reliability argument: an offline-autonomous core has no external run-time dependency to fail, and a store-and-forward buffer converts an intermittent link from a failure mode into a latency variation. (Judgment, based on D2/AD-1 and SYS-013/015/018.)
Why a duty-cycled energy governor as a cross-cutting concern (not a layer). The 90-day/−30 °C budget (D3) cannot be met by any single layer's local optimisation; it is a system property of how long each subsystem is powered. Making the governor a cross-cutting controller (the S3 power-state machine that gates L1–L5) rather than a peer layer means every layer is designed to be interruptible and gateable — sensing wakes only to sample, comms batches, radios (the dominant load) are gated hardest (AD-2; SYS-055). The alternative — continuous live telemetry — was rejected in concept as energy-prohibitive (D2). The governor also owns the interlock that gating a rail must never tear an NVM write (SYS-057/067), which is why the storage layer's fail-safety is co-designed with the governor rather than bolted on. (Judgment; budget Evidence-Pending — SYS-086/090/092.)
Why split trust at the tier boundary (AD-5). Two requirements pull in opposite directions: the edge must be frictionless (no login, or the untrained last-mile user is locked out — C9), while the record and configuration must be governed (RBAC/SSO, Part 11 audit) or the compliance value evaporates. A single uniform trust posture cannot satisfy both. The architecture therefore draws the trust boundary at the tier line: edge read is open; everything that changes configuration, ingests data, or delivers firmware is authenticated and audited (SYS-068). Placing the fault line at the tier boundary (rather than inside a tier) keeps the security surface small and explicit — only three interfaces cross it (I-09, I-10, I-12), and each is a defined, mutually authenticated, attributable contract. (Judgment + Standard — IEC 62443 concepts, 21 CFR Part 11 attributable; revision to confirm; threat model Evidence-Pending.)
Why layer information (verdict → detail) — AD-6. The edge shows one unambiguous verdict; the full record lives in the cloud/export. This keeps the edge decision surface simple and glove/cold/glare-legible (SYS-058) while preserving the rich evidentiary record where it can be audited (L6). It also localises the no-false-reassurance invariant: the edge verdict logic (L3) is the single place that must refuse an unqualified "accept" on an incomplete record (SYS-017). (Judgment.)
Why a typed, single-ICD interface discipline. Layers and domains are only cleanly separable if their boundaries are explicit. All inter-subsystem and cross-tier contracts conform to the typed inventory I-01…I-12 and a single authoritative ICD, with no undocumented interface (SYS-066). This is both an integration-defect control (interface drift between disciplines is a leading defect source) and a security control (an undocumented interface is an ungoverned attack surface). (Standard — ISO/IEC/IEEE 29148; revision to confirm.)
---
5. Architecture qualities / drivers → mechanisms (traceability)
Each architectural driver traces to the concrete mechanism(s) that deliver it and the requirements/decisions that govern it. This is the "does the architecture actually satisfy its drivers" table.
Driver (quality) | Architectural mechanism(s) | Governing decisions & SYS | Evidence / status
Endurance (90 days @ −30 °C) | Duty-cycled energy governor (S3 power-state machine) gating L1–L5; radios gated hardest; batched store-and-forward; individually gateable rails; low-temp chemistry + charge-inhibit <0 °C | D3, AD-2, AD-3; SYS-055, SYS-086/090/092, SYS-019 | Evidence-Pending — budget validated by endurance test (Judgment on mechanism; the headline risk)
Measurement accuracy | Dual-channel sensing (internal + external probe); characterised thermal path (probe represents product zone); plausibility/data-quality flagging; trusted timestamp | AD-4; SYS-001/003/006/007, SYS-051/056 | ±0.5 °C target Evidence-Pending (Judgment; WHO PQS E006 class to confirm); probe offset/lag Evidence-Pending
Offline verdict (edge autonomy) | Edge-autonomous core L1–L4 + edge L7; on-device accept/reject/quarantine with no network/login; local capture of every excursion | D2, AD-1, AD-6; SYS-013/015/017, SYS-058 | Baselined mechanism (Judgment); verdict-consistency guard SYS-016
Auditability | Append-only tamper-evident storage (S7); power-fail-safe writes; complete gap-flagged trip record; cloud audit trail; admissible export/API | AD-6/AD-7; SYS-002/026/027/028, SYS-057, SYS-064 | Standard — ALCOA+, EU GDP, 21 CFR Part 11 (revision to confirm)
Security | Split trust boundary (edge open / platform governed); secure element + secure boot; mutual authentication at I-09/I-10/I-12; signed, resumable, fail-safe OTA; RBAC/SSO | AD-5, AD-7; SYS-032/065/068, SYS-061/063 | Standard concepts (IEC 62443, Part 11; revision to confirm); threat model Evidence-Pending
Reusability (fleet lifecycle) | USB-C recharge + charge-inhibit; readiness triage (charge/cal/firmware) gating deployment; re-provisioning; fleet visibility + stray recovery; OTA | D1, AD-8; SYS-030/031/032, SYS-024, SYS-053 | Baselined mechanism (Judgment); reverse-logistics carried from D1
Data delivery under intermittent coverage | Store-and-forward buffer (full-trip capacity); multi-bearer opportunistic upload; ordered/resumable/idempotent transfer; alerts jump the queue | D2, AD-1; SYS-018/019/020/062, SYS-011 | Mechanism Baselined; storage sizing and RF-in-container Evidence-Pending (SYS-018/059)
No false reassurance (safety invariant) | Gap-marking; sensor-fault flagging; edge verdict refuses unqualified "accept" on incomplete/faulted record; verdict/record consistency | AD-6; SYS-002/007/017, SYS-016 | Safety-classified; Preliminary — reconcile with HAL (Stage 10)
---
6. Neighbouring-discipline flags (CLAUDE.md rule 2)
This architecture stays in the systems lane and names the disciplines that must weigh in where a risk touches an adjacent domain. Each flag names the decision/interface at stake.
• Battery / power-electronics engineer — owns cell chemistry selection, BMS design, charge-inhibit-<0 °C and self-heat-before-charge, protection, and the 90-day/−30 °C energy budget (S2; D3, AD-2, AD-3; SYS-086/090/092, SYS-053). This is the top program risk and is Evidence-Pending until closed by test.
• RF / antenna engineer — owns the RF-in-container link budget and antenna strategy (S5 + S1 aperture; AD-1, I-08; SYS-059), including whether an external-antenna option is required. This is the principal open architectural risk and is Evidence-Pending, closed by the RF-in-container qualification test (PRD-VER-005).
• Mechanical / materials engineer — owns enclosure sealing (incl. USB-C port), IP rating, cleaning- chemical compatibility for the reuse loop, probe form/retention and its thermal path, and shipper fit/ retention (S1; AD-4; SYS-051/052/054). Probe offset/time-constant budget is Evidence-Pending.
• Cybersecurity engineer — owns the threat model, key management, secure boot, mutual authentication at the trust-boundary crossings, and OTA signing (S3/S8; AD-5, AD-7; SYS-032/061/063/065/068). Threat model is Evidence-Pending; the trust-split architecture is only real once it is completed.
• Quality / Regulatory engineer (EM-ENG-006) — owns WHO PQS category, US FDA classification, EU GDP / 21 CFR Part 11 data-integrity conformance, and transport/radio/EMC/environmental certification (D4; PRD-REG; SYS-177…SYS-190). Named regulatory gaps are Evidence-Pending (concept §7).
• Hazard analysis (ORCH-017 / EM-ENG-…, Stage 10) — all safety-classified architectural properties here (no-false-reassurance, arm-confirmation, two-sided detection reliability, fail-safe OTA) are Preliminary and must be reconciled with the Hazard Analysis Library; each will trace to an HA-###/RCR-### record.
---
7. Cross-references
• Product/design architecture (AD-1…AD-8): [../05-product-requirements/design-architecture.md](../05-product-requirements/design-architecture.md)
• System block diagram + typed interfaces I-01…I-12: [../05-product-requirements/system-block-diagram.md](../05-product-requirements/system-block-diagram.md)
• Product description (subsystems S1–S9, drivers): [../05-product-requirements/product-description.md](../05-product-requirements/product-description.md)
• System requirements (SYS-001…SYS-190): [../06-system-requirements/index.md](../06-system-requirements/index.md)
• Concept & foundational decisions (D1–D4): [../00-concept/concept-and-decisions.md](../00-concept/concept-and-decisions.md)
Companion Stage-08 files (parallel) | [functional-decomposition.md](functional-decomposition.md) (FUN-###) · [component-list.md](component-list.md) (CMP-###) · [interface-control.md](interface-control.md) (INT-###)
Evidence posture (honesty): architectural decisions cite D1–D4 and AD-1…AD-8; every claim is labelled Standard / Cited / Judgment / Evidence-Pending. The two load-bearing open risks — RF-in-container (I-08, SYS-059) and the −30 °C/90-day energy budget (D3, SYS-086/090/092) — are carried as Evidence-Pending and closed only by test. No standard clause number is asserted as fact; no unresolved quantity is invented.
Power / Energy Architecture — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · Author: Principal Power / Systems Engineer, integrating with EM-ENG-001 (Chief Systems Engineer), the Electrical/Hardware and Firmware engineers, and the Reliability and Safety engineers. This file is the architectural realisation of the power/energy system requirements SYS-086 … SYS-103, the hardware power group HWR-001 … HWR-012 (+ rails/charge-input HWR-036/039/043/045), and the firmware power-state group FWR-007 … FWR-014. It describes HOW the SCM's energy is sourced, gated, budgeted, and protected so the device can meet the headline promise — 90 days of continuous monitoring across −30 °C to +60 °C from a rechargeable source — which is the program's top risk (D3).
Evidence posture (CLAUDE.md rule 1; concept §evidence). No cell chemistry, capacity, current, or energy number is asserted as fact in this document. Every quantitative energy figure below is an ILLUSTRATIVE MODEL marked Evidence-Pending / illustrative model, validated by test (PRD-VER-004) and exists only to show the shape and sensitivities of the budget, not to claim a result. The final low-temperature chemistry and its cold datasheet are Evidence-Pending (concept §7). The charge-inhibit-below-0 °C interlock and battery protection are safety-critical; they are stated as design invariants, not tunable performance. Standards are cited by designation only, revision to confirm.
---
1. Energy architecture intent
The SCM is architected as a duty-cycled energy governor wrapped around an offline-first edge node. The governing physics is simple and unforgiving: over a 90-day trip the device is asleep for the overwhelming majority of the time, so average power — not peak power — decides whether the device survives the trip, and average power is dominated by (a) how little the device draws while asleep (quiescent drain) and (b) how rarely and how briefly the high-drain loads — the radios and GNSS — are allowed to turn on. This is the resolution of D3 (the −30 °C / 90-day tension) and the reason the architecture exists.
Three intents flow from this:
1. Radios are the dominant load and are gated hard (AD-2). LTE-M, Wi-Fi, BLE, and GNSS are treated as scheduled, energy-budgeted events, never as always-on services. Firmware (FWR-009) keeps every radio powered down or in its lowest state outside its scheduled window, and hardware (HWR-026, HWR-043) makes each high-drain rail independently gateable so a radio can draw energy only during its acquisition window. The single permitted exception is a queued excursion alert (FWR-009 / FWR-027 / FWR-043) — the one activity whose safety value justifies breaking the schedule. This directly realises SYS-089 (duty-cycled governor with hard radio gating).
2. The lowest viable sleep state is the default (AD-2). Between scheduled activities the MCU and peripherals return to their deepest retention-sleep state (FWR-008), retaining only the RTC, wake logic, battery protection, and log integrity. Quiescent drain at −30 °C (SYS-092 / HWR-010 / HWR-014) is a first-class budget line, not an afterthought, because it is what a 90-day idle multiplies.
3. Charge is a safety-gated, opportunistic, depot-side activity — never a field dependency. The device is offline-first and self-powered for the trip; charging happens at the depot (D1 reuse loop) and is hard-interlocked against cold charging (AD-3 / SYS-094 / HWR-005) because charging a cold Li-ion cell drives lithium plating and a thermal-runaway pathway. The charge architecture is therefore designed around refusing to charge more than around charging fast.
The energy governor is the single, deterministic power-state machine (FWR-007); its states and transitions are specified in the companion file [state-mode-model.md](state-mode-model.md).
---
2. Power tree
The power tree traces energy from the USB-C source, through the charge/protection chain, to the cell, and out through the regulated, individually-gateable rails to each load domain. The charge-inhibit-below-0 °C gate (AD-3) sits in the charge path before the cell and cannot be bypassed by firmware.
mermaid flowchart TD %% ---------- Charge source & input protection ---------- USBC["USB-C receptacle (sink)<br/>HWR-039 · SYS-093<br/>I-04 charge input"] INPROT["Input protection front-end<br/>reverse-polarity / OVP / ideal-diode<br/>HWR-009 · SYS-102"] OVPCLAMP["OV clamp / TVS transient protector<br/>HWR-045 · ESD HWR-041/SYS-103"] USBC --> INPROT --> OVPCLAMP
%% ---------- Charge management + cold-charge interlock ---------- PMIC["Charge-management IC / PMIC<br/>programmable charge profile<br/>HWR-004 · SYS-093"] OVPCLAMP --> PMIC
CELLTEMP["Cell-temperature sense (cell-coupled NTC)<br/>HWR-006 · SYS-094"] INHIBIT{"CHARGE-INHIBIT GATE<br/>T_cell &lt; 0 °C (or cell safe limit)?<br/>AD-3 · HWR-005 · SYS-094<br/>HW interlock + FW enforce FWR-013"} HEATER["Self-heat-before-charge element<br/>(feasibility under eval)<br/>HWR-007 · SYS-095"] CELLTEMP --> INHIBIT PMIC --> INHIBIT HEATER -.warm cell first.-> INHIBIT
%% ---------- BMS / protection / cell ---------- BMS["Battery-protection IC / AFE<br/>OV/UV/OC/OT/short — MCU-independent<br/>HWR-003 · SYS-096"] FUSE["Secondary non-resettable backstop<br/>fuse / 2nd OVP — HWR-012"] INHIBIT -->¦charge allowed only if T_cell in window¦ BMS BMS --> FUSE CELL["Battery cell / pack (xSyP)<br/>low-temp-capable chemistry — TBD<br/>HWR-001/002 · SYS-088<br/>Evidence-Pending"] FUSE --- CELL CELL --- BMS
%% ---------- Fuel gauge ---------- GAUGE["Fuel gauge — temp-compensated SoC/SoH<br/>HWR-008 · SYS-097"] CELL --- GAUGE
%% ---------- Regulation ---------- PROT["Protected cell node (system power)"] FUSE --> PROT RAILS["Regulated rails + sequencing + decoupling<br/>each rail independently gateable<br/>HWR-043 · SYS-055/099"] PROT --> RAILS
%% ---------- Always-on housekeeping domain ---------- HK["Always-on housekeeping domain<br/>RTC + backup · protection · gauge · wake logic<br/>HWR-010/015 · SYS-092 (quiescent budget)"] PROT --> HK
%% ---------- Load domains (all gated by FW governor) ---------- RAILS --> MCU["MCU / compute<br/>deep-sleep retention · watchdog · brown-out<br/>HWR-013/014/017/019 · SYS-089"] RAILS --> SENSE["Sensing rail(s) — per-sensor gated<br/>int/ext temp · humidity · accel<br/>HWR-021-028 · HWR-026 · SYS-090"] RAILS --> NVM["Non-volatile log / storage<br/>power-fail-safe writes<br/>HWR-020 · SYS-099"] RAILS --> LTE["LTE-M modem (primary WAN)<br/>peak-current rail + decoupling<br/>HWR-029/036 · SYS-089/090"] RAILS --> WIFI["Wi-Fi radio (opportunistic offload)<br/>HWR-030 · gated"] RAILS --> BLE["BLE radio (provision / local read)<br/>HWR-031 · gated"] RAILS --> GNSS["GNSS receiver (duty-cycled fix)<br/>HWR-025 · SYS-048 · gated"] RAILS --> OLED["OLED display + status LED<br/>event-driven, near-zero idle<br/>HWR-046/048/049 · SYS-090"]
%% ---------- Governor control plane ---------- GOV["FW energy governor / power-state machine<br/>FWR-007-013 · SYS-089<br/>gates every high-drain rail"] GOV -. gate/enable .-> LTE GOV -. gate/enable .-> WIFI GOV -. gate/enable .-> BLE GOV -. gate/enable .-> GNSS GOV -. gate/enable .-> SENSE GOV -. gate/enable .-> OLED GOV -. sleep/wake .-> MCU
Reading the tree. Charge energy enters only through the USB-C sink (HWR-039); it must survive reverse-polarity/OVP (HWR-009) and transient (HWR-045) protection before reaching the PMIC (HWR-004). The PMIC's output to the cell passes through the cold-charge interlock (HWR-005), which is qualified by cell-coupled temperature sensing (HWR-006) and is a hardware gate that firmware cannot defeat (FWR-013 is a redundant enforcement layer, not the primary interlock). The autonomous protection AFE (HWR-003) and a non-resettable backstop (HWR-012) sit around the cell so protection persists in any MCU or firmware fault. Discharge energy leaves the protected cell node into the regulated rails (HWR-043); every high-drain rail is individually gateable so the governor (FWR-007…013) can enforce SYS-089's hard gating. The always-on housekeeping domain (RTC, protection, gauge, wake logic) is the load that never sleeps and therefore owns the quiescent budget (SYS-092 / HWR-010).
---
3. 90-day / −30 °C energy architecture (the headline)
This is the D3 headline and the top program risk. It is presented as an illustrative energy-budget model that demonstrates the structure of the budget and where the sensitivities lie. It asserts no result.
### EVIDENCE BANNER — read before every number below
Every current, duty-cycle, energy, and total figure in this section is an ILLUSTRATIVE MODEL, marked
Evidence-Pending / illustrative model, validated by test (PRD-VER-004). The numbers are placeholder
stand-ins chosen only to make the arithmetic of the budget legible; the real values are unresolved and
come from the cold current-profile test, the RF-in-container test, and the 90-day cold-endurance test
(MR-VAL-006). No battery capacity, chemistry, or per-mode current is claimed as fact. Per SYS-087,
SYS-090, SYS-092, FWR-012 the per-class allocations are explicitly Evidence-Pending.
3.1 Per-mode current model (illustrative)
Per-mode current estimates for the device power states (aligned to the FWR-007 state set). All values illustrative; symbol ≈ emphasises the placeholder nature.
Mode / activity | Dominant loads active | Illustrative avg current while in mode | Illustrative duration per occurrence | Evidence
Deep-sleep / quiescent | RTC, protection AFE, fuel gauge, wake logic, MCU retention | ≈ tens of µA (order-of-magnitude only) | continuous between events | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-092, HWR-010/014
Sample + log | MCU active, sensor rails (int/ext temp, RH, accel), NVM write | ≈ few mA | ≈ seconds | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-037, FWR-015
On-device detect / verdict compute | MCU active only | ≈ few mA | ≈ sub-second–seconds | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-038/042
LTE-M upload (cadence) | LTE-M modem TX/RX peaks, MCU | ≈ tens–hundreds mA avg; peak far higher | ≈ tens of seconds per session | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-090, HWR-029/036; RF-in-container (SYS-049) dominates
Excursion alert (immediate) | LTE-M (or best bearer), MCU | ≈ as upload, single short session | ≈ tens of seconds, event-driven | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-011/039, FWR-043
GNSS fix | GNSS receiver acquisition, MCU | ≈ tens of mA | ≈ tens of seconds–minutes (cold-start, attenuated) | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-048, HWR-025; in-container fix availability unresolved
HMI / OLED read | OLED + MCU, event-driven | ≈ few–tens mA | ≈ seconds per user action | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-041, HWR-046/049
Charging (depot only) | PMIC + charge current (not a trip load) | n/a to trip budget | depot | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-093, HWR-004
3.2 Duty-cycle assumptions (illustrative)
The governor's schedule (FWR-010) sets how often each active mode occurs. Illustrative defaults, taken from the baselined functional cadences (SYS-001 default 5-min sample; SYS-019 default hourly upload):
Scheduled activity | Illustrative cadence (configurable) | Occurrences per 90 days (illustrative) | Evidence
Sample + log | every 5 min (range 1–60, SYS-001) | ≈ 25,900 | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
Cadence upload | hourly (SYS-019) | ≈ 2,160 sessions | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
GNSS fix | e.g. every 6–24 h + on event (SYS-048) | ≈ 90–360 | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
Excursion alert | event-driven (rare, but unbudgeted-in-time) | trip-dependent | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
HMI read | occasional (depot / last-mile) | trip-dependent | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
Deep-sleep | balance of all time (≈ >99 % of wall-clock) | continuous | Evidence-Pending / illustrative model, validated by test (PRD-VER-004)
3.3 How it rolls up toward a 90-day target (illustrative structure)
The budget is an energy balance: supply-side usable energy at cold, less derating and reserve, must exceed demand-side consumption over the worst-case 90-day cold-duty profile, with positive margin (SYS-087). The roll-up structure — not the values — is:
Demand (per 90 days) = Q_sleep = I_sleep × (total sleep time at −30 °C) ← usually the LARGEST term + Q_sample = E_sample × (number of samples) ← small each, very many + Q_upload = E_upload × (number of upload sessions) ← few, but expensive each + Q_gnss = E_gnss × (number of fixes) ← very expensive each, cadence-sensitive + Q_alert = E_alert × (number of excursion alerts) ← event-driven + Q_hmi = E_hmi × (number of reads) ← minor
Supply (usable at −30 °C) = C_rated × k_coldderate × k_agefade − Reserve
Endurance requirement (SYS-086/087/090): Supply ≥ Demand + Margin
Budget line (illustrative roll-up) | Illustrative share of 90-day demand | Why it matters | Evidence
Quiescent / sleep (Q_sleep) | typically the dominant share | 90 days × always-on µA; cold raises leakage & self-discharge | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-092, HWR-010
Cadence uploads (Q_upload) | major share | radios are the dominant active load (AD-2); cold + attenuation raise TX time/retries | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-090, SYS-049
Sampling + logging (Q_sample) | moderate share | many small events; sensor settling dominates each | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-037
GNSS (Q_gnss) | swing factor | each fix costly, cadence-sensitive; in-container fixes may fail/lengthen | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-048, HWR-025
Alerts (Q_alert) | small but unbounded-in-time | event-driven; the one override of gating | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-011/039
HMI (Q_hmi) | minor | event-driven, gated near-zero idle | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — HWR-049
Reserve + margin | held back, not spent | protects against cold under-estimation & fade | Evidence-Pending / illustrative model, validated by test (PRD-VER-004) — SYS-087 margin unresolved
3.4 The low-temperature capacity-derating tension
The core of D3: usable capacity is not a constant. A cell rated at room temperature delivers materially less usable energy at −30 °C (higher internal resistance, voltage sag under load, reduced accessible capacity), and it cannot be safely charged below 0 °C at all. So the supply side of the balance (C_rated × k_coldderate) shrinks exactly in the conditions where the device must survive longest, while the demand side grows (radios work harder against RF attenuation, oscillators/PA/LNA drift — HWR-035; sensor settling and MCU/OLED wake times lengthen at cold — SYS-041). This two-sided squeeze is why the 90-day/−30 °C combination is a genuine tension (D3) and the top program risk, and why the endurance requirements (SYS-086/087/045/046) are Evidence-Pending and validated by test, never extrapolated from room temperature.
Mitigations (architected, not assumed):
1. Chemistry selection (HWR-001 / SYS-088) — a low-temperature-tolerant chemistry (e.g. low-temp Li-ion or LTO-class) whose datasheet specifies −30 °C usable capacity; the number is Evidence-Pending (concept §7) and closed by cold-capacity test.
2. Aggressive duty-cycling (AD-2 / SYS-089 / FWR-008/009/010) — the primary lever: minimise active time so the shrunken cold budget is enough. This is what makes the promise achievable without betting the program on an ideal cell.
3. Charge-inhibit below 0 °C (AD-3 / HWR-005 / SYS-094) — removes cold charging as a hazard path entirely (safety mitigation, see §4).
4. Self-heat-before-charge evaluation (HWR-007 / SYS-095) — a depot mitigation for turnaround of cold returns, feasibility under evaluation; must never defeat the inhibit.
5. Honest degradation (SYS-098 / §5) — if the cold budget is nonetheless exhausted, shed discretionary loads first and preserve the record, so a cold under-estimate fails safe rather than losing data.
3.5 Sensitivities that dominate the budget
The model's purpose is to expose which variables the design and the verification must control. In order:
• Radio on-time (LTE-M) — highest active-energy sensitivity. Because radios are the dominant active load (AD-2), every extra second of TX — driven by RF attenuation inside the metallized/insulated shipper (SYS-049, the principal open architectural risk), cold RF-front-end degradation (HWR-035), retries (FWR-044), and upload cadence (SYS-019) — multiplies across ~2,160 sessions. Bounded by hard gating (SYS-089/FWR-009) and bounded back-off (FWR-044).
• Quiescent current at −30 °C — highest total-energy sensitivity. Because sleep is ~>99 % of wall-clock, a small µA error, multiplied by 90 days and worsened by cold leakage/self-discharge, can dominate the entire budget (SYS-092 / HWR-010/014).
• GNSS cadence — the swing factor. Each fix is expensive and in-container fixes may fail or lengthen (HWR-025); location is fleet/attribution support, never safety, so it is hard-gated and fails soft (SYS-048 / FWR-009). Cadence is the single largest discretionary knob.
• Cold capacity derating & cycle fade — the supply-side sensitivity. k_coldderate and k_agefade (SYS-091 capacity retention over reuse) set how much of the rated capacity is actually available; both are Evidence-Pending and can silently cause a mid-life dead-in-cold event.
These four are exactly the quantities the verification program closes: cold current-profile test (sleep + per-mode currents), RF-in-container test (radio/GNSS energy, SYS-049/PRD-VER-005), cold-capacity/cycle-life tests (supply side), and the integrating 90-day cold-endurance test (MR-VAL-006 / PRD-VER-004).
---
4. Charge architecture
Charging is a depot-side, safety-gated activity (D1 reuse loop), not a field function. The architecture prioritises never charging unsafely over charging fast.
• USB-C standard sink (SYS-093 / HWR-039 / I-04). Recharge is from any standard compliant USB-C source, no proprietary charger. The receptacle presents a compliant sink contract (defined pin roles, CC-as-sink, bounded voltage/current) and is sealed to the enclosure ingress rating (S1). Charge time to trip-readiness is a target (SYS-050) that is lengthened by the cold interlock and any self-heat step — Evidence-Pending.
• Charge-inhibit-below-0 °C interlock (AD-3 — SAFETY-CRITICAL). This is the safety spine of the charge path and a design invariant:
• Hardware primary (HWR-005 / SYS-094). A temperature-qualified charge-enable (JEITA-class) driven by cell-coupled temperature sensing (HWR-006) prevents any charge current reaching the cell when T_cell < 0 °C (or the cell-datasheet safe lower limit, whichever is higher) — even under a single MCU, firmware, or communication fault. Firmware cannot defeat it.
• Firmware enforcement layer (FWR-013). In charging state the firmware independently reads cell temperature and commands the charge path to remain inhibited while cold, and reflects the state to HMI and readiness. FWR-013 is defence-in-depth, not the sole means (per its own note) — the primary interlock is hardware.
• Rationale (physics). Charging a cold Li-ion cell drives lithium plating → dendrite growth → internal short → thermal event. This is a hazard, not a performance loss (severity-first Critical, CLAUDE.md rule 3). Cited: IEC 62368-1 / IEC 62133 / JEITA charge-temperature concepts, revision to confirm; reconcile with HAL (Stage 10).
• Self-heat-before-charge evaluation (HWR-007 / SYS-095). Where feasible, a controllable heater warms the cell to a permitted charge temperature before charging a cold-returned device, to improve depot turnaround. It is arranged so it cannot bypass or override HWR-005 and cannot exceed the cell's safe upper thermal limit. Feasibility is a live D3 branch — Evidence-Pending, validated by test.
• Protection (SYS-096 / HWR-003 / HWR-012 / HWR-009 / HWR-045). Multi-layer, MCU-independent:
• Autonomous protection AFE (OV / UV / OC / OT / short) that operates with firmware dead (HWR-003).
• A secondary non-resettable backstop (fuse and/or 2nd OVP) so loss of the primary protection still terminates a hazardous condition into a safe state (HWR-012) — single-fault safety.
• Charge-input fault tolerance: reverse-polarity / OVP front-end (HWR-009 / SYS-102) and OV clamp/TVS (HWR-045) reject or safely limit a bad depot source before it reaches the charge path.
• ESD/isolation on the port (HWR-041 / SYS-103) so handling discharges do not corrupt the record or damage the device.
• UN 38.3 / IATA transport qualification of cell and pack (HWR-011 / SYS-101) — a compliance gate (D4), Evidence-Pending test summary.
• Fuel-gauge / SoC & SoH (SYS-097 / HWR-008). A temperature-compensated coulomb-counting/model gauge reports SoC and SoH across the operating range with accuracy sufficient for readiness decisions (SYS-030/091) and a non-misleading low-battery warning (SYS-098). Cold shifts the voltage/SoC relationship, so temperature compensation is required to avoid a confidently-wrong "ready" or a silent depletion — temperature-dependent accuracy target Evidence-Pending.
---
5. Energy allocation policy (graceful degradation)
When usable energy runs low, the architecture protects the safety-relevant record and honest status last, shedding discretionary functions first (SYS-098 / FWR — power-mgr, HMI, comms). The policy is a strict priority order; the exact thresholds are Evidence-Pending (SYS-098 degradation order unresolved) but the order is a design invariant:
Priority (keep longest → shed first) | Function class | Behaviour as energy depletes | Evidence
1 — never shed while armed | Core sensing + gap-free logging + on-device excursion detection + committing the record | Preserved to the last; an armed device never silently stops monitoring (only Fault-with-safe-log may pause, and never silently) | SYS-098, SYS-047, SYS-013, FWR-015/021/025
2 — keep | Honest status + charge-inhibit indication + safe-state guarantees + audit trail | Kept so the device is never falsely "OK"; low-battery warned (SYS-098) | SYS-098, SYS-100, SYS-017/FWR-030
3 — degrade | Excursion alert delivery (radio) | Still attempted (the one gating override) but with bounded retry/back-off to protect the budget | SYS-011, FWR-043/044
4 — shed early | Routine cadence uploads (LTE-M / Wi-Fi / BLE) | Suspended before core functions; data still buffered offline (store-and-forward) | SYS-098, SYS-019, FWR-041
5 — shed first | GNSS / location and discretionary HMI | First to go; location is never safety (fails soft, last-known + quality flag) | SYS-098, SYS-048, HWR-025
Invariants (safety-critical):
• Losing an upload is acceptable; losing the log is not — the record is preserved over convenience (SYS-098 / SYS-021 buffer-full prioritises excursion + audit data).
• No false reassurance — as the device degrades it must keep reporting honestly that it is monitoring or flag that it is not (SYS-047 / SYS-017 / FWR-054); a depleting device must never present a clean "accept" over an incomplete record.
• Clean brown-out safe-state — a collapsing rail triggers a deterministic safe-state entry with no corruption of committed data (SYS-099 / HWR-019 / FWR-035), so depletion ends in a preserved record, not a torn one.
The state/mode realisation of this degradation ladder, the safe-state definition, and the full set of transitions and guards are specified in [state-mode-model.md](state-mode-model.md).
---
Traceability
Concept / decisions | D1 (reusable/rechargeable), D2 (store-and-forward), D3 (−30 °C / 90-day tension — top risk), D4 (multi-market transport/EMC).
Architecture decisions | AD-2 (duty-cycled energy governor), AD-3 (charge-inhibit + self-heat eval).
System requirements | SYS-086 … SYS-103 (power/energy/charge/protection), with SYS-045/046 (cold endurance), SYS-055/099 (rails/brown-out), SYS-048 (GNSS energy), SYS-049 (RF-in-container).
Hardware (power group) | HWR-001 … HWR-012, HWR-036 (radio peak current), HWR-039 (USB-C), HWR-041/045 (ESD/OVP), HWR-043 (rails), HWR-046/048/049 (HMI power).
Firmware (power-state group) | FWR-007 … FWR-014 (governor, gating, charge-inhibit enforcement, shipping mode), FWR-035 (power-fail-safe writes), FWR-041/043/044 (comms energy discipline).
Verification | cold current-profile test, RF-in-container test (PRD-VER-005), cold-capacity / cycle-life tests, integrating 90-day cold-endurance test (PRD-VER-004 / MR-VAL-006). All energy figures in §3 are illustrative and closed by these tests.
Neighbours to convene (CLAUDE.md rule 2) | Battery/power-electronics engineer (chemistry, BMS, self-heat), RF/antenna engineer (SYS-049 in-container link → radio energy), Reliability (cycle fade), Safety/Hazard (Stage 10 reconciliation of AD-3, SYS-094/096, safe-state).
Evidence: every energy quantity is an illustrative model marked Evidence-Pending and validated by test; charge-inhibit and battery protection are safety-critical design invariants, not tunable performance; chemistry and capacity are unresolved and never asserted (concept §7).
State / Mode Model — Smart Cold Chain Monitor (SCM)
SCM · Stage 08 (System Architecture) · Author: Principal Power / Systems Engineer with the Firmware and Safety engineers, integrating with EM-ENG-001. This file specifies the device's state/mode model — the behavioural realisation of the duty-cycled energy governor (AD-2) and the safety invariants around arming, charge-inhibit, OTA, and the safe state. It is the companion to [power-energy-architecture.md](power-energy-architecture.md) and the direct architectural expression of the firmware power-state group FWR-007 … FWR-014 and the arming/fault group FWR-031/032, FWR-053/054, plus SYS-159 (safe-state), SYS-032 (fail-safe OTA), SYS-091/094 (charge-inhibit).
Evidence posture. The state set and transitions are baselined design (FWR-007 is Baselined); guards that depend on numeric thresholds (charge-inhibit temperature, low-battery thresholds, self-check cadence) are Evidence-Pending / reconcile with HAL (Stage 10) and are stated here as conditions, not values. Charge-inhibit and the safe state are safety-critical invariants, not tunable behaviour. Standards cited by designation, revision to confirm.
---
1. Mode catalogue
The device is, at any instant, in exactly one power state of the deterministic governor's state set {deep-sleep(Sleep), idle, logging, alert, charging, OTA, fault, shipping} (FWR-007), overlaid by an arming state {unarmed, arming, armed, voided} (FWR-031) that determines whether the monitoring modes run. The catalogue below groups these into the operator-facing modes.
Mode | Purpose | Active subsystems | Power profile | Entry conditions | Exit conditions | Allowed transitions to
Shipping / Storage | Ultra-low-power dormancy for depot/transit of an unarmed device; protect shelf-life charge; must never look like monitoring | Always-on housekeeping only (RTC+backup, protection AFE, fuel gauge, wake logic); radios & sensors OFF; no scheduled sampling | Lowest — minimises self-discharge & quiescent drain (below even Sleep, no sample schedule) | Explicit "enter shipping" command (FWR-014); device unarmed | Explicit wake/arm action (button or provisioning); charger insert | Provisioning/Arming, Charging
Sleep (deep-sleep / quiescent) | Default low-power state between scheduled activities of an armed (or idle-configured) device; the dominant-time state | Housekeeping + RTC/timer wake armed; MCU in retention deep-sleep; all high-drain rails gated OFF | Lowest operational — owns the SYS-092 quiescent budget; ~>99 % of wall-clock (illustrative) | Scheduled task complete → return to lowest viable sleep (FWR-008) | RTC/timer tick (sample/upload due); accel shock IRQ; charger insert; button; excursion-alert hand-off | Sampling/Logging, Alerting, Charging, OTA, Provisioning/Arming, Fault/Safe
Sampling / Logging | Acquire multi-channel sample, timestamp, calibrate, detect excursions, commit gap-free record | MCU active; sensor rails (int/ext temp, RH, accel) gated on for the window; NVM write; detect-engine | Brief active burst (illustrative seconds); returns to Sleep on completion | RTC/timer scheduled sample due (FWR-010); or accel shock IRQ (event capture) | Sample committed; excursion declared → Alerting; else → Sleep | Sleep, Alerting, Fault/Safe
Alerting | Emit an excursion alert with priority over routine upload; the one radio activity permitted outside its window | Detect-engine + comms + best available bearer (LTE-M primary, else Wi-Fi/BLE); MCU active | Short high-drain radio session, event-driven (the permitted gating override, FWR-009) | Excursion committed (FWR-025) with connectivity available (FWR-027/043) | Alert delivered or queued (offline → queue-and-forward); back to Sleep | Sleep, Sampling/Logging, Fault/Safe
Charging | Depot recharge to trip-readiness under the cold-charge safety interlock | PMIC/charge path, cell-temp sense, BMS, fuel gauge, HMI/LED indication; sensing/logging continue if armed | Externally powered; charge current gated by T_cell interlock | Charger insertion detected (FWR-011); valid USB-C source | Charger removed; charge complete; charge-inhibited (cold) → indicate & hold | Sleep, Shipping, Provisioning/Arming, Fault/Safe
OTA | Apply signed, resumable, fail-safe firmware update without disturbing monitoring or bricking the device | OTA-agent, comms, security (secure element), storage (inactive bank); sensing/logging continue | Radio-active during download windows, within energy budget; no monitoring suspension | Authenticated update available in a connected window (FWR-046/047) | Verified image activated (A/B) & confirmed, or auto-rollback to known-good | Sleep (resumed prior state), Fault/Safe (on failed activation → rollback)
Fault / Safe | Defined safe state on any unsafe/misleading fault; preserve record, hold charge non-hazardous, honest status, attempt recovery | Fault arbiter, storage (preserve committed data), power-mgr (charge held non-hazardous), HMI (honest fault indication) | Minimal safe operation; watchdog/brown-out driven | Watchdog reset, brown-out, POST/init failure, storage/sensor/power/OTA fault (FWR-005/053) | Automatic recovery to prior monitoring state, or hold with honest fault indication | Sleep/Sampling (on recovery to prior armed state), Shipping (service)
Provisioning / Arming | Configure trip profile, run pre-arm readiness gate, explicitly arm with positive confirmation | BLE (or on-device HMI), scheduler (arming SM + config), sensor-driver (probe/POST), power-mgr (SoC check), HMI | Interactive, moderate; depot/last-mile | From Shipping/Sleep on button or BLE provisioning session (FWR-031) | Armed (all pre-arm checks pass) → monitoring; or blocked (readiness fail) / voided | Sleep (→ armed monitoring), Shipping, Fault/Safe
Arming overlay (FWR-031). Independent of power state, the arming state {unarmed → arming → armed → voided} gates the monitoring modes: Sampling/Logging and Alerting run their gap-free monitoring only while armed. An unarmed / arming-incomplete / voided state is made unmistakable on the device so a trip cannot silently start unmonitored (SYS-023 / FWR-031). Shipping mode is always unarmed.
---
2. State machine (Mermaid stateDiagram-v2)
mermaid stateDiagram-v2 [] --> Shipping
state "Shipping / Storage" as Shipping state "Provisioning / Arming" as Provision state "Sleep (deep-sleep)" as Sleep state "Sampling / Logging" as Sampling state "Alerting" as Alerting state "Charging" as Charging state "OTA Update" as OTA state "Fault / Safe" as Fault
note right of Shipping Ultra-low-power, UNARMED (FWR-014). Never presented as monitoring. Explicit wake/arm required. end note
%% ----- Provisioning & arming ----- Shipping --> Provision: button / BLE provisioning session Provision --> Provision: configure profile (FWR-004/031) Provision --> Sleep: ARMED - pre-arm readiness gate PASSED (FWR-032) Provision --> Provision: arm BLOCKED - a pre-arm check failed (FWR-032) Provision --> Shipping: cancel / void (FWR-031)
%% ----- Core duty-cycle loop (ARMED) ----- Sleep --> Sampling: RTC/timer scheduled sample due (FWR-010) Sleep --> Sampling: accel shock interrupt (FWR-011/020) Sampling --> Sleep: sample committed, no excursion (FWR-015/021) Sampling --> Alerting: excursion declared (FWR-025) Sleep --> Alerting: queued excursion alert hand-off (FWR-027/009) Alerting --> Sleep: alert delivered OR queued offline (FWR-043)
%% ----- Cadence upload runs inside a gated radio window ----- Sleep --> Sleep: scheduled cadence upload window (FWR-041)
%% ----- Charging (depot) with cold-charge interlock ----- Sleep --> Charging: charger inserted (FWR-011) Provision --> Charging: charger inserted Charging --> Charging: charge INHIBITED when cell below safe limit (AD-3/HWR-005/FWR-013) Charging --> Charging: charging when cell in permitted window Charging --> Sleep: charger removed, armed Charging --> Shipping: charger removed, unarmed
%% ----- OTA (never bricks, never stops monitoring) ----- Sleep --> OTA: authenticated update available (FWR-046/047) OTA --> Sleep: verified image activated and confirmed, data preserved (FWR-048/049) OTA --> Fault: activation failed, auto-rollback to known-good (FWR-048)
%% ----- Fault / Safe state (from anywhere) ----- Sleep --> Fault: watchdog / brown-out / fault (FWR-052/053) Sampling --> Fault: sensor/storage/power fault (FWR-053) Alerting --> Fault: fault (FWR-053) Charging --> Fault: charge/thermal fault (FWR-053) Provision --> Fault: init/POST failure (FWR-005) Fault --> Sleep: auto-recovery to prior ARMED state, gap logged (FWR-053/006) Fault --> Shipping: service / unrecoverable, honest fault indication held
note right of Fault SAFE STATE (SYS-159): preserve committed data, hold charge non-hazardous, present honest non-misleading status, attempt recovery. Never a silent monitoring stop. end note
---
3. Mode-transition rules & invariants
These rules are design invariants the state machine must satisfy. Safety-critical invariants dominate (CLAUDE.md rule 3, severity-first) and are marked [SAFETY].
3.1 Monitoring continuity
• [SAFETY] Logging never stops silently while armed. An armed device is at all times either verifiably monitoring or visibly flagged as not (SYS-047 / FWR-054). The only path that may pause monitoring is Fault/Safe, and even there the pause is never silent: the reset cause and the interrupted interval are logged as an explicit gap (FWR-006 / FWR-038) and the verdict is qualified (SYS-017 / FWR-030). There is no transition from Armed→(any state) that both stops sampling and leaves the device presenting a normal "monitoring" status.
• Gap-free logging across transitions. Every scheduled interval not recorded (sleep overrun, reset, fault, power-loss) is represented by an explicit gap marker, never silently absent (SYS-002 / FWR-021 / FWR-038). Wake→sample→commit→sleep commits before returning to sleep so no acquired sample is lost to a later power event (FWR-015).
3.2 Alert pre-emption
• [SAFETY] Alerts pre-empt sleep and jump the queue. A committed excursion (FWR-025) hands off a high-priority alert that may fire outside the scheduled radio window — the single permitted override of hard radio gating (FWR-009 / FWR-027) — and is transmitted ahead of routine cadence data (FWR-043). When offline, the alert is queued to NVM and delivered ahead of buffered data on reconnection (SYS-040), so no excursion depends on a live link to exist (SYS-013).
3.3 Charge-inhibit (cold-charge safety)
• [SAFETY] Charging is inhibited whenever the cell is below the safe low-temperature limit. In the Charging mode the transition to actively charging is guarded by T_cell ≥ 0 °C (or the cell-datasheet safe lower limit, whichever is higher) — AD-3 / SYS-094. The primary interlock is hardware (HWR-005, temperature-qualified charge-enable on cell-coupled sensing HWR-006) and holds under a single MCU/firmware/comms fault; firmware enforcement (FWR-013) is a redundant layer that must never defeat it. The charge-inhibited state is unambiguously indicated (SYS-100 / HWR-048) so a cold, charge-inhibited device is not mistaken for dead. Self-heat-before-charge (HWR-007), if provided, may only enable the guard by warming the cell — it can never bypass the inhibit.
3.4 OTA never bricks, never disarms silently
• [SAFETY] OTA is signed, resumable, and fail-safe (SYS-032 / AD-7). An update image is activated only after signature + integrity + rollback-protection verification (FWR-046); it installs into an inactive bank and activates via A/B with confirmation, so a failed or non-booting update auto-rolls back to the last known-good image (FWR-048) — the device is never left bricked or non-monitoring without notice. All committed data, configuration, calibration, keys, and audit trail are preserved across install / activation / rollback, and the pre-update monitoring/arming state is re-established (FWR-049). OTA download does not suspend scheduled sensing/logging (FWR-047).
3.5 Arming integrity
• [SAFETY] A trip cannot silently start unmonitored. Transition to Armed requires an explicit user arm action with positive on-device confirmation (SYS-023 / FWR-031) and is gated by the pre-arm readiness check (FWR-032): adequate SoC, valid & unexpired calibration, valid firmware, valid profile, probe present, POST passed. Any failed pre-condition blocks arming and guides correction (SYS-024 / SYS-030). Shipping/Storage is always unarmed and must never be mistaken for a monitoring device (FWR-014).
3.6 Safe-state definition (SYS-159 / SYS-167 / FWR-053)
[SAFETY] The safe state, entered on any fault from which continued normal operation would be unsafe or misleading, shall simultaneously
1. Preserve committed data — no torn/corrupt record; power-fail-safe atomic writes (SYS-099 / HWR-019 brown-out detector / FWR-035).
2. Hold the charge subsystem non-hazardous — protection (HWR-003/012) and charge-inhibit (HWR-005) remain in force independent of firmware.
3. Present honest, non-misleading status — never a normal "ready/monitoring" state over a fault; an incomplete record never reads as a clean "accept" (SYS-017 / FWR-005 / FWR-030).
4. Attempt automatic recovery to the prior armed/monitoring state, logging the reset cause and any monitoring gap (FWR-006 / FWR-052 / FWR-053).
• A watchdog serviced only on a verified-healthy main loop (FWR-052) turns a silent hang into a logged recovery rather than a false-reassurance outage.
3.7 Determinism & energy budget
• Exactly one power state at any instant (FWR-007); only documented transitions are permitted; undefined/overlapping states are prohibited (they would defeat both auditability and the energy budget).
• Lowest viable sleep is the default between activities (FWR-008); no state is held longer than needed to complete its task; every high-drain rail is gated OFF outside its scheduled window (FWR-009 / SYS-089).
• Wake sources are exhaustively defined and debounced (FWR-011): RTC/timer, accel shock IRQ, charger-insertion, user-button — no unmasked or chattering wake may drain the budget.
3.8 Graceful degradation overlay
• As energy depletes, the device sheds discretionary functions first and protects the record last (SYS-098): shed GNSS/HMI → suspend cadence uploads → keep alert delivery (bounded retry) → keep core sensing/logging/detection and honest status to the end. This ladder rides on top of the state machine (it changes which scheduled activities run, not the state set) and is specified in [power-energy-architecture.md](power-energy-architecture.md) §5. Losing an upload is acceptable; losing the log is not.
---
Traceability
Concept / decisions | D1 (reusable — shipping/charging modes), D2 (offline-first — alert queueing, store-and-forward), D3 (energy governor — Sleep default, hard gating).
Architecture decisions | AD-2 (duty-cycled governor), AD-3 (charge-inhibit), AD-7 (fail-safe OTA).
System requirements | SYS-089 (governor), SYS-091/094 (charge-inhibit), SYS-032 (fail-safe OTA), SYS-023/024/030 (arming & readiness), SYS-047 (no silent outage), SYS-159 (safe state), SYS-002/017 (gap-free / honest incomplete), SYS-098 (graceful degradation), SYS-011/013/040 (alerts).
Firmware | FWR-007 (state machine), FWR-008/009/010/011 (sleep/gating/schedule/wake), FWR-013 (charge-inhibit enforcement), FWR-014 (shipping mode), FWR-031/032 (arming SM + pre-arm gate), FWR-046/047/048/049 (OTA), FWR-005/052/053/054 (safe-mode / watchdog / fault arbiter / self-check), FWR-006/021/038 (reset cause / interval accounting / gap markers), FWR-035 (power-fail-safe writes).
Hardware | HWR-005/006 (cold-charge interlock + cell-temp sense), HWR-003/012 (protection), HWR-017/019 (watchdog / brown-out), HWR-048 (charge-inhibit indication).
Verification | state-transition & fault-injection tests (charge-inhibit sweep, interrupted-update / rollback, brown-out integrity, hang injection, arming/void), integrating with the 90-day endurance test (PRD-VER-004 / MR-VAL-006). Preliminary guards reconcile with the HAL (Stage 10).
Evidence: the state set and transitions are baselined; numeric guard thresholds (charge-inhibit temperature, low-battery, self-check cadence) are Evidence-Pending and reconcile with the HAL; charge-inhibit, arming integrity, OTA fail-safety, and the safe state are safety-critical invariants, not tunable performance.

## 9 Design Outputs
Design Outputs
The actual set of drawings, models, software and specifications from which the product can be reproduced.
↩ Contents
In this section:
• Design Outputs (Stage 09, ORCH-013) — Smart Cold Chain Monitor (SCM)
• DOL-000 — Design Output Governance (DO-GOV) — SCM
• DOL-100 — Product Configuration Baseline (DO-CFG) — SCM
• DOL-1000 — Connectivity & Interface Design Outputs (DO-IF) — SCM
• DOL-1100 — Cybersecurity Design Outputs (DO-SEC) — SCM
• DOL-1200 — Safety & Risk-Control Design Outputs (DO-SAFE) — SCM
• DOL-1300 — Human Factors & UI Design Outputs (DO-HF) — SCM
• DOL-1400 — Manufacturing Design Outputs (DO-MFG) — SCM
• DOL-1500 — Inspection & Production-Test Design Outputs (DO-TST) — SCM
• DOL-1600 — Packaging, Labeling & Logistics Design Outputs (DO-LBL) — SCM
• DOL-1700 — Installation, Service & Maintenance Design Outputs (DO-SVC) — SCM
• DOL-1800 — Regulatory & Compliance Design Outputs (DO-REG) — SCM
• DOL-1900 — Verification-Enabling Design Outputs (DO-VER) — SCM
• DOL-200 — System Architecture Outputs (DO-SYS) — SCM
• DOL-2000 — Design Transfer & Release Baseline Design Outputs (DO-XFER) — SCM
• DOL-300 — Mechanical Design Outputs (DO-MECH) — SCM
• DOL-400 — Electrical & Electronic Design Outputs (DO-EE) — SCM
• DOL-500 — PCB & PCBA Design Outputs (DO-PCB) — SCM
• DOL-600 — Embedded Firmware Design Outputs (DO-FW) — SCM
• DOL-700 — Device Software Outputs (DO-SW) — SCM
• DOL-800 — Cloud Platform Design Outputs (DO-CLD) — SCM
• DOL-900 — Mobile & Web Application Design Outputs (DO-APP) — SCM
• Design Input–Output Traceability Matrix — Smart Cold Chain Monitor (SCM)
• Design Output Register (Consolidated) — Smart Cold Chain Monitor (SCM)
• Product Design Output Master Specification — Smart Cold Chain Monitor (SCM)
Design Outputs (Stage 09, ORCH-013) — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 09 — Design Outputs · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Status: Draft baseline (pre-DR-DO) · Traceability: ORCH-018 / EM-ENG-012.
Navigation and status dashboard for the SCM Design Output Library (DOL) — the results of the design effort (CAD, drawings, schematics, layouts, source, configs, specs, manufacturing/test/labeling/service/ regulatory data, and acceptance criteria) from which the approved SCM can be reproduced. The controlling document is [Product-Design-Output-Master-Specification.md](Product-Design-Output-Master-Specification.md); the consolidated one-line-per-output view is [Design-Output-Register.md](Design-Output-Register.md); the requirement/risk/V&V/mfg chains are in [Design-Input-Output-Traceability-Matrix.md](Design-Input-Output-Traceability-Matrix.md).
Baseline at a glance: 272 controlled design outputs across 21 modules (DOL-000…DOL-2000); 103 flagged Critical. All entries are Draft/In-review and all revisions TBD-rev (Evidence Pending) in this pre-DR-DO baseline. All of DOL-1200 (Safety) is Preliminary — reconcile with HAL Stage 10.
---
Module table (counts read from the module files)
Module | Prefix | File | Outputs | Critical | Owner
DOL-000 Design Output Governance | DO-GOV | [DOL-000-governance.md](DOL-000-governance.md) | 9 | 0 | EM-ENG-001
DOL-100 Product Configuration Baseline | DO-CFG | [DOL-100-config-baseline.md](DOL-100-config-baseline.md) | 12 | 3 | EM-ENG-001
DOL-200 System Architecture Outputs | DO-SYS | [DOL-200-system-architecture-outputs.md](DOL-200-system-architecture-outputs.md) | 12 | 6 | EM-ENG-001
DOL-300 Mechanical Design Outputs | DO-MECH | [DOL-300-mechanical.md](DOL-300-mechanical.md) | 20 | 3 | Principal Mechanical Eng
DOL-400 Electrical & Electronic Outputs | DO-EE | [DOL-400-electrical.md](DOL-400-electrical.md) | 22 | 11 | Principal Electrical Eng (+ Battery/Power)
DOL-500 PCB & PCBA Outputs | DO-PCB | [DOL-500-pcb.md](DOL-500-pcb.md) | 13 | 2 | Principal PCB Eng
DOL-600 Embedded Firmware Outputs | DO-FW | [DOL-600-firmware.md](DOL-600-firmware.md) | 20 | 13 | EM-ENG-007
DOL-700 Device Software Outputs | DO-SW | [DOL-700-device-software.md](DOL-700-device-software.md) | 8 | 2 | EM-ENG-007
DOL-800 Cloud Platform Outputs | DO-CLD | [DOL-800-cloud.md](DOL-800-cloud.md) | 16 | 6 | Principal Cloud Eng
DOL-900 Mobile & Web Application Outputs | DO-APP | [DOL-900-mobile.md](DOL-900-mobile.md) | 10 | 2 | Principal Mobile Eng
DOL-1000 Connectivity & Interface Outputs | DO-IF | [DOL-1000-interfaces.md](DOL-1000-interfaces.md) | 12 | 7 | EM-ENG-001
DOL-1100 Cybersecurity Design Outputs | DO-SEC | [DOL-1100-security.md](DOL-1100-security.md) | 12 | 8 | EM-ENG-008
DOL-1200 Safety & Risk-Control Outputs | DO-SAFE | [DOL-1200-safety-risk-controls.md](DOL-1200-safety-risk-controls.md) | 14 | 14 | Reliability/Safety Eng
DOL-1300 Human Factors & UI Outputs | DO-HF | [DOL-1300-human-factors.md](DOL-1300-human-factors.md) | 10 | 3 | Human Factors Eng
DOL-1400 Manufacturing Design Outputs | DO-MFG | [DOL-1400-manufacturing.md](DOL-1400-manufacturing.md) | 14 | 3 | Manufacturing Eng
DOL-1500 Inspection & Production-Test Outputs | DO-TST | [DOL-1500-inspection-test.md](DOL-1500-inspection-test.md) | 14 | 5 | Quality Eng
DOL-1600 Packaging, Labeling & Logistics | DO-LBL | [DOL-1600-packaging-labeling.md](DOL-1600-packaging-labeling.md) | 12 | 3 | Manufacturing Eng (+ EM-ENG-006)
DOL-1700 Installation, Service & Maintenance | DO-SVC | [DOL-1700-service.md](DOL-1700-service.md) | 10 | 2 | Service Eng
DOL-1800 Regulatory & Compliance Outputs | DO-REG | [DOL-1800-regulatory.md](DOL-1800-regulatory.md) | 12 | 4 | EM-ENG-006
DOL-1900 Verification-Enabling Outputs | DO-VER | [DOL-1900-verification-enabling.md](DOL-1900-verification-enabling.md) | 10 | 4 | Quality Eng
DOL-2000 Design Transfer & Release Baseline | DO-XFER | [DOL-2000-design-transfer.md](DOL-2000-design-transfer.md) | 10 | 2 | EM-ENG-001
Total |  | 21 files | 272 | 103
---
Applicability Register (scope never silently dropped)
Item | Status | Rationale
DOL-000…DOL-2000 (all 21 modules) | Applicable | Globally distributed, regulated, reusable, connected, safety-significant IoT device spanning device + firmware + cloud + mobile; every discipline produces controlled outputs.
DOL-700 Device Software | Applicable (lean) | No rich on-device OS/app; DO-SW covers host tools, provisioning/calibration configs, SBOM, and factory-image packaging (8 outputs), not an application OS image.
SCM-100-XA external-antenna branch | Applicable — Evidence Pending | Emitted as a configured variant (DO-EE-021, DO-MECH-016); hardware branch confirmed only if the integral antenna cannot close the in-container link budget (AD-1, PRD-VER-005).
AI/ML, audio/video, payments output classes | Not Applicable | No such scope in the SCM (concept; Stage 07 §3). Recorded, not silently dropped.
---
Status & top release blockers (Master Spec §9)
Pre-DR-DO: none Released. The honest, severity-first blockers:
• RA-05 — HAL reconciliation. All 14 DOL-1200 DO-SAFE outputs are Critical and Preliminary; each must map to a baselined HAL control (Stage 10). Gate: DO-XFER-008. Top blocker.
• RA-07 — configuration freeze. Battery chemistry (DO-EE-003) and antenna (DO-EE-018) are Evidence Pending, blocking AML (DO-CFG-004) and downstream mfg/test readiness.
• RA-10 — regulatory classification. WHO PQS device category and US FDA classification (DO-REG-001/005) are Evidence Pending, blocking markings/UDI (DO-LBL-002/004) and submission.
• RA-08 / RA-11 — RF-in-container ICDs/test config (DO-IF-003, DO-VER-009) Evidence Pending, closed by PRD-VER-005.
---
Traceability note (chain position)
Design outputs are the results of the design effort; they are built against, and verified to, the design inputs of Stages 05–08.
Flow: PRD (05) → SYS (06) → Subsystem/SRS (07) → Architecture (08) → Design Outputs (09, this stage) → Hazard Analysis (10) / ERFKL (11) → Risk-Control Requirements (12) → Verification (13) / Validation (14) → Manufacturing & Design Transfer (15) → Release (16) / Field (17) → Complete Lifecycle Traceability (18).
Forward consumers of this baseline:
• Hazard Analysis (Stage 10, ORCH-017) — DOL-1200 (DO-SAFE-001…014) supplies the "Implementation Reference" for each risk control; every inherent-design/protective HAL control must appear there.
• Verification (Stage 13, ORCH-014) — DOL-1900 (DO-VER) supplies the measurable characteristics, limits, observability, test modes, and objective pass/fail that seed each PV-<CODE>-NNN; DOL-1500 (DO-TST) supplies the production acceptance tests.
• Manufacturing & Design Transfer (Stage 15) — DOL-1400/1500/1600/1700/2000 supply the process, test, packaging/labeling, service, and release-baseline outputs; DO-XFER-010 authorizes production.
• End-to-end lifecycle matrix (Stage 18, ORCH-018 / EM-ENG-012) — every output's per-module Input/Risk/Verification traceability fields resolve into the complete bidirectional matrix.
Evidence & honesty per 00-governance/evidence-standard.md and hallucination-policy.md: every entry is labelled Cited / Standard / Judgment / Evidence Pending; revisions are TBD-rev; no part numbers, datasheet values, standard clause numbers, or classifications are invented. Safety/charge/excursion/record outputs are Preliminary — reconcile with HAL Stage 10. </content>
DOL-000 — Design Output Governance (DO-GOV) — SCM
Module: DOL-000 · Output prefix: DO-GOV · Owner: Chief Systems Engineer (EM-ENG-001) · Discipline: Configuration Management / Design Controls · Status: Draft baseline (ORCH-013). Part of the SCM Design Output Library; governed by Product-Design-Output-Master-Specification.md.
1. Purpose
Define how every SCM design output is identified, numbered, versioned, reviewed, approved, released, stored, and changed — the rules the other 20 DOL modules obey so a "file" only counts as a controlled design output when it is complete, released, and traceable.
2. Scope
Output taxonomy & numbering, revision control, approval workflow, authoritative formats, derived-vs-source rules, configuration management, variant applicability, change control, interface ownership, supplier-/third-party-controlled outputs, SOUP handling, obsolete-output retention, archival & retrieval, review criteria, and the output-vs-record distinction. Applies to all SKUs (SCM-100 family).
3. Applicable standards & references
ISO 13485:2016 §7.3.4 / §4.2 (document & record control); FDA QMSR (21 CFR 820, eff. 2026-02-02); ISO 10007 (configuration management); IEC 62304 (software lifecycle, SOUP); ISO 14971 (risk file linkage); 00-governance/design-output-standard.md. Standard — designations cited; revisions to confirm.
4. Governance & ownership
EM-ENG-001 owns this module and the CM system. Each discipline owner (master §7) owns their module's outputs; EM-ENG-006 concurs on regulatory-critical outputs; EM-ENG-008 on security-critical outputs. All changes route through the ECO workflow (DO-GOV-004) and Design Review Board.
5. Output categories
Governance controlled documents (DO-DOC): numbering/CM plan, approval workflow, format policy, change control, supplier-output control, SOUP register, archival/retention, review checklist, output-vs-record policy.
6. Design Output Register (this module)
DO-GOV-001 — Output Identification & Numbering Scheme
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-DOC-001 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown/DOCX · Released format: PDF (controlled) · Repository: QMS/PLM document control · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-002, DO-CFG-011 (config-identification rules)
Input trace | design-output-standard.md; ISO 13485 §4.2.4 · Risk-control trace: — · Verification trace: DR-DO review record
Critical | No · Change class: Major
Acceptance criteria | DO-<CODE>-NNN and SCM-<AREA>-<TYPE>-NNN schemes fully defined; no ID collisions; every AREA/TYPE mapped to an owner.
Mfg use | Indirect (governs all mfg output IDs) · Notes: Judgment; instantiates the DOL standard.
DO-GOV-002 — Revision & Approval Workflow Definition
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board / Quality
Source format | Markdown/DOCX · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-001, DO-GOV-004
Input trace | ISO 13485 §7.3.4 (output review/approval before release) · Risk-control trace: — · Verification trace: Approval audit trail
Critical | No · Change class: Major
Acceptance criteria | Draft→In-review→Released→Superseded/Obsolete states defined with named approver roles; no output releases without recorded approval.
Mfg use | Indirect · Notes: Standard (ISO 13485 §7.3.4). Enforced electronically in PLM (Evidence Pending — tool selection).
DO-GOV-003 — Authoritative File-Format & Derived-vs-Source Policy
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-DOC-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-MECH (STEP/native), DO-PCB (Gerber/ODB++), DO-FW (binary/hash)
Input trace | design-output-standard.md · Risk-control trace: — · Verification trace: Format audit at DR-DO
Critical | No · Change class: Minor
Acceptance criteria | For each output class, the authoritative source format and the released/derived format are named; derived outputs (e.g., Gerber from layout) are marked derived and regenerable.
Mfg use | Direct (defines what mfg receives) · Notes: Judgment.
DO-GOV-004 — Engineering Change Control (ECO/ECN) Procedure
Type | DO-PROC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-PRC-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board / Quality
Source format | Markdown/DOCX · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-002, DO-CFG-011, DO-XFER-004
Input trace | ISO 13485 §7.3.9 (design changes); QMSR · Risk-control trace: Change-impact on risk file (ISO 14971) · Verification trace: ECO records; re-verification triggers
Critical | No · Change class: Major
Acceptance criteria | Every change classified (Major/Minor/Editorial), impact-assessed against requirements/risk/verification, approved, and propagated to affected outputs before release.
Mfg use | Direct (change propagation to line) · Notes: Standard.
DO-GOV-005 — Change-Classification Definition (Major/Minor/Editorial)
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Quality
Source format | Markdown · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-004
Input trace | design-output-standard.md (Change Classification field) · Risk-control trace: — · Verification trace: Sampled ECO audit
Critical | No · Change class: Minor
Acceptance criteria | Rules unambiguously classify a change; form/fit/function and safety/regulatory triggers force Major and re-verification.
Mfg use | Indirect · Notes: Judgment.
DO-GOV-006 — Supplier- & Third-Party-Controlled Output Register
Type | DO-DOC · Discipline: CM/Purchasing (SYS) · SKU: All · Subsystem/Component: AML items (S2/S3/S4/S5)
File # | SCM-GOV-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Quality/Purchasing
Source format | Markdown/XLSX · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-CFG-004 (AML), DO-GOV-007 (SOUP)
Input trace | ISO 13485 §7.4 (purchasing) · Risk-control trace: Critical-component control (BMS, SE, cell) · Verification trace: Incoming inspection (DO-TST)
Critical | No · Change class: Major
Acceptance criteria | Every supplier-controlled output (cell datasheet, BMS IC datasheet, modem certs) identified with the control mechanism (spec-in, PCN notification, second source).
Mfg use | Direct · Notes: Judgment; part numbers Evidence Pending (do not invent).
DO-GOV-007 — SOUP / Third-Party Software Component Register
Type | DO-DOC · Discipline: Software CM (S3/S8/S9) · SKU: All · Subsystem/Component: FW/cloud/mobile stacks
File # | SCM-GOV-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-007 / EM-ENG-008
Source format | Markdown/XLSX · Released format: PDF + machine-readable SBOM refs · Repository: Software repo / QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-FW-018 (FW SBOM), DO-SW-006, DO-CLD-015, DO-SEC-011
Input trace | IEC 62304 (SOUP); SRS-SEC · Risk-control trace: Dependency-vulnerability control · Verification trace: SBOM/vuln-scan gate
Critical | No · Change class: Major
Acceptance criteria | Every RTOS/stack/library (SOUP) listed with version, function, known-anomaly handling, and update path.
Mfg use | Indirect · Notes: Standard (IEC 62304).
DO-GOV-008 — Archival, Retention & Retrieval Procedure
Type | DO-PROC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-PRC-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Quality
Source format | Markdown · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-009, DO-XFER-001
Input trace | ISO 13485 §4.2.5 (record/document retention) · Risk-control trace: — · Verification trace: Retrieval test at audit
Critical | No · Change class: Minor
Acceptance criteria | Retention periods, obsolete-output retention, and retrieval SLA defined; superseded outputs retained and marked; retrievable within audit window.
Mfg use | Indirect · Notes: Standard.
DO-GOV-009 — Output-vs-Record Distinction Policy
Type | DO-DOC · Discipline: CM/Quality (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-GOV-DOC-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-006 · Approval: Quality
Source format | Markdown · Released format: PDF · Repository: QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-VER (verification-enabling), DO-XFER-002
Input trace | design-output-standard.md (related-but-distinct artifacts) · Risk-control trace: — · Verification trace: Sampled audit
Critical | No · Change class: Minor
Acceptance criteria | Clear rule separating design outputs from verification evidence, validation evidence, manufacturing records, and quality records; e.g., threat model = analysis artifact, implemented controls = outputs (DOL-1100).
Mfg use | Indirect · Notes: Standard.
7. Traceability notes
This module has few upstream requirement parents (it is meta-governance); its parents are the DOL standard and ISO 13485/QMSR clauses. It governs the traceability integrity of all other modules. Forward, it feeds DOL-2000 (design transfer / configuration audit) and the Stage-18 matrix.
8. Release / adequacy notes
Adequate when DO-GOV-001…009 are Released and the ECO workflow is operational in the chosen PLM/QMS tool (tool selection Evidence Pending). No Critical outputs; blocking dependency for DR-DO is that the workflow (DO-GOV-002/004) must be live before any output can be formally Released. </content>
DOL-100 — Product Configuration Baseline (DO-CFG) — SCM
Module: DOL-100 · Output prefix: DO-CFG · Owner: Chief Systems Engineer (EM-ENG-001) · Discipline: Configuration Management · Status: Draft baseline (ORCH-013).
1. Purpose
Define the SCM product family tree, SKU/option matrix, regional variants, the cross-tier compatibility matrix (HW/FW/cloud/mobile), the approved-parts baseline (AML), replaceable/consumable items, packaging configurations, configuration rules, and serial-number/identification conventions — the "what configuration is the approved product" anchor for every other output.
2. Scope
SCM-100 family (SCM-100, -NA, -EU, -ROW, -XA). Device HW config, firmware/cloud/mobile compatibility, AML, FRUs, consumables (none device-side beyond battery as non-user-replaceable), packaging configs, config rules, serialization/UDI-DI linkage.
3. Applicable standards & references
ISO 10007 (CM); ISO 13485 §7.5.8 (identification/traceability); FDA UDI rule (21 CFR 830) — DI linkage; GS1 (UDI carrier); D1–D4; Stage 08 component-list (CMP-001…055). Standard/Cited — revisions to confirm.
4. Governance & ownership
EM-ENG-001 owns the baseline; EM-ENG-006 owns UDI-DI/regulatory-config linkage; Purchasing owns AML with Quality concurrence. Chemistry (CMP-006) and antenna (CMP-028/029) are Evidence Pending and gate the AML freeze (RA-07).
5. Output categories
Product family tree (DO-DOC), SKU/option matrices (DO-DOC/DO-DATA), AML (DO-BOM/DO-DOC), compatibility matrix (DO-DATA), config rules (DO-DOC/DO-CFG), serialization convention (DO-DOC).
6. Design Output Register (this module)
DO-CFG-001 — SCM Product Family Tree (top-level)
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product (S1–S9)
File # | SCM-CFG-DOC-001 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown/PLM tree · Released format: PDF + PLM BOM · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-CFG-002, DO-MECH-000, DO-EE-000, DO-BOM (all)
Input trace | PRD-ARCH, PRD-VIS; SYS decomposition; CMP-001…055 · Risk-control trace: — · Verification trace: Config audit (DO-XFER-003)
Critical | No · Change class: Major
Acceptance criteria | Complete indented product tree from finished-device down to S1–S9 assemblies; every CMP-### mapped to a tree node; no orphan components.
Mfg use | Direct (top-level BOM structure) · Notes: Judgment; based on Stage-08 component list.
DO-CFG-002 — Model & SKU Matrix
Type | DO-DATA · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-CFG-DAT-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Product/Regulatory
Source format | XLSX/CSV · Released format: Controlled CSV + PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-CFG-003, DO-REG-002, DO-LBL-002
Input trace | D4 (multi-market); PRD-REG; SYS-177…190 · Risk-control trace: — · Verification trace: Regulatory config review
Critical | No · Change class: Major
Acceptance criteria | SCM-100/-NA/-EU/-ROW/-XA each defined by the delta from base (band plan, markings, data-residency, antenna); each SKU maps to one regulatory config (DO-REG-002).
Mfg use | Direct · Notes: Judgment.
DO-CFG-003 — Option & Accessory Matrix
Type | DO-DATA · Discipline: CM (SYS) · SKU: All · Subsystem/Component: External probe, external antenna, USB-C cable/charger, mount
File # | SCM-CFG-DAT-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Product
Source format | XLSX · Released format: CSV + PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-MECH-014 (probe), DO-EE-021 (ext-antenna), DO-SVC-007 (spares)
Input trace | PRD-MECH, PRD-SVC; CMP-020/029; SYS-003 · Risk-control trace: Probe-presence detection (HAL) · Verification trace: Accessory compat test
Critical | No · Change class: Major
Acceptance criteria | Each option (external probe SKUs, external-antenna kit, charger, replacement mount) has ordering identity, compatibility rule, and detection behavior.
Mfg use | Direct · Notes: Judgment; external-antenna option gated by AD-1 (Evidence Pending).
DO-CFG-004 — Approved Manufacturer / Parts List (AML)
Type | DO-BOM · Discipline: CM/Purchasing (SYS) · SKU: All · Subsystem/Component: All electronic + mechanical parts
File # | SCM-CFG-BOM-004 · Rev: TBD-rev (Evidence Pending — chemistry/antenna) · Status: Draft · Owner: EM-ENG-001/Purchasing · Approval: Quality/Purchasing
Source format | PLM BOM/XLSX · Released format: Controlled BOM export · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-EE-020 (electrical BOM), DO-PCB-013 (PCBA BOM), DO-MECH-019 (mech BOM), DO-GOV-006
Input trace | All HWR/MER; CMP-006/008/016/024 (critical parts) · Risk-control trace: Cell/BMS/SE approval control (HAL) · Verification trace: Incoming inspection (DO-TST-001)
Critical | Yes (approved-parts integrity for safety-critical cell/BMS/SE) · Change class: Major
Acceptance criteria | Every part has approved manufacturer(s) and rule for alternates; critical parts (cell, BMS/AFE, secure element, PMIC, fuse) are single-source-controlled with PCN notification; no invented part numbers — placeholders until selection.
Mfg use | Direct (procurement) · Notes: Evidence Pending — low-temp cell (CMP-006) and antenna branch (CMP-028/029) unresolved; AML cannot freeze until closed (RA-07).
DO-CFG-005 — HW / FW / Cloud / Mobile Compatibility Matrix
Type | DO-DATA · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Cross-tier
File # | SCM-CFG-DAT-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | XLSX · Released format: CSV + PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-FW-019 (release notes), DO-CLD-016, DO-APP-009, DO-IF-002 (version negotiation)
Input trace | SYS-063/065 (version skew); IFR-022; SRS-OTA · Risk-control trace: OTA never-brick (HAL) · Verification trace: Interop/version-skew test
Critical | Yes (fielded interoperability across staged OTA) · Change class: Major
Acceptance criteria | Defines supported (HW rev × FW version × cloud API version × app version) combinations and the minimum-compatible matrix across OTA version skew; incompatible combos are blocked, not silently degraded.
Mfg use | Indirect · Notes: Standard/Judgment.
DO-CFG-006 — Regional Variant Configuration Definitions
Type | DO-CFG · Discipline: CM/Regulatory (SYS) · SKU: -NA/-EU/-ROW · Subsystem/Component: Radios (S5), cloud region (S8)
File # | SCM-CFG-CFG-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-006 · Approval: Regulatory
Source format | Config schema/JSON · Released format: Signed config bundle · Repository: PLM/config repo · Parent assembly: DO-CFG-002
Related | DO-EE-018 (band plan), DO-CLD-013 (data residency), DO-REG-002
Input trace | D4; SYS-115/116 (roaming/SIM); PRD-REG · Risk-control trace: — · Verification trace: Regional RF/regulatory test
Critical | No · Change class: Major
Acceptance criteria | Each region's band plan, permitted-band enforcement, carrier/APN strategy, data-residency region, and marking set are defined and machine-enforceable.
Mfg use | Direct (config-at-programming) · Notes: Judgment; band plans/carrier strategy Evidence Pending (INT-027).
DO-CFG-007 — Replaceable-Component / FRU Definition
Type | DO-DOC · Discipline: CM/Service (S1/S2) · SKU: All · Subsystem/Component: External probe, mount, seals; battery (service-level only)
File # | SCM-CFG-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Service · Approval: Service/Quality
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-SVC-004 (replacement procedures), DO-SVC-007 (spares)
Input trace | PRD-SVC; D1 (reusable); MER-### (serviceable seals) · Risk-control trace: Battery service = authorized-only (HAL) · Verification trace: Service procedure validation
Critical | No · Change class: Major
Acceptance criteria | FRU list with field-vs-depot-vs-factory replaceability; sealed battery is not a field FRU (safety); external probe/mount/seal kits defined.
Mfg use | Indirect · Notes: Judgment (D1 reusable lifecycle).
DO-CFG-008 — Consumables & Cleaning-Materials Baseline
Type | DO-DOC · Discipline: CM/Service (S1) · SKU: All · Subsystem/Component: Enclosure/seals (cleaning compatibility)
File # | SCM-CFG-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: Service · Approval: Materials/Quality
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-MECH-010 (material spec), DO-SVC-006 (cleaning procedure)
Input trace | D1 (clean-between-trips); MER-### (chemical compatibility) · Risk-control trace: — · Verification trace: Cleaning-chemical compatibility test
Critical | No · Change class: Minor
Acceptance criteria | Approved cleaning agents defined and compatible with enclosure/seal materials; no consumables that void IP rating.
Mfg use | Indirect · Notes: Judgment.
DO-CFG-009 — Packaging Configuration Matrix
Type | DO-DATA · Discipline: CM/Packaging (SYS) · SKU: All · Subsystem/Component: Retail/bulk/service packaging
File # | SCM-CFG-DAT-009 · Rev: TBD-rev · Status: Draft · Owner: Manufacturing · Approval: Manufacturing/Logistics
Source format | XLSX · Released format: CSV + PDF · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-LBL-008 (packaging drawing), DO-LBL-010 (transport)
Input trace | PRD-MFG; UN 38.3/IATA (Li battery transport, D4) · Risk-control trace: Li-battery transport safety · Verification trace: Packaging/transport test
Critical | No · Change class: Major
Acceptance criteria | Single-unit, multi-pack, and depot-bulk configs defined with the lithium-battery transport classification per config.
Mfg use | Direct · Notes: Standard (UN 38.3/IATA — revision to confirm).
DO-CFG-010 — Serial-Number & Device-Identity Convention
Type | DO-DOC · Discipline: CM (S3) · SKU: All · Subsystem/Component: Device identity (SE/CMP-016)
File # | SCM-CFG-DOC-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Quality/Security
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-MFG-009 (serialization procedure), DO-LBL-003 (serial label), DO-SEC-008 (identity provisioning), DO-CLD-004 (registry)
Input trace | SYS-030 (registry); ISO 13485 §7.5.8; UDI · Risk-control trace: Device-identity integrity · Verification trace: Serialization verification (DO-TST-011)
Critical | Yes (identity binds record attribution + UDI + registry) · Change class: Major
Acceptance criteria | Serial format, uniqueness rule, and binding of serial ↔ SE device identity ↔ UDI-PI ↔ cloud registry are defined; no duplicate serials; identity is one-time-programmed.
Mfg use | Direct (serialization/programming) · Notes: Standard/Judgment.
DO-CFG-011 — Configuration-Identification & Baseline Rules
Type | DO-DOC · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-CFG-DOC-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-GOV-001, DO-XFER-003 (config audit)
Input trace | ISO 10007; design-output-standard.md · Risk-control trace: — · Verification trace: Configuration audit
Critical | No · Change class: Major
Acceptance criteria | Rules define what constitutes a released baseline, how baselines are named/frozen, and how a functional/physical configuration audit is performed at transfer.
Mfg use | Indirect · Notes: Standard (ISO 10007).
DO-CFG-012 — Product Configuration Rules & Compatibility Constraints
Type | DO-DATA · Discipline: CM (SYS) · SKU: All · Subsystem/Component: Cross-subsystem
File # | SCM-CFG-DAT-012 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | XLSX/rules · Released format: CSV + PDF · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-CFG-005, DO-CFG-006
Input trace | PRD-ARCH; SYS interfaces · Risk-control trace: — · Verification trace: Config-rule validation
Critical | No · Change class: Major
Acceptance criteria | Valid/invalid option combinations (e.g., external-antenna SKU excludes certain integral-antenna tuning; region ⇔ band plan) are machine-checkable; no buildable invalid configuration.
Mfg use | Direct (config validation at order) · Notes: Judgment.
7. Traceability notes
Parents: PRD-ARCH/REG/VIS/MFG/SVC; SYS-030/063/065/115/116/177…190; the full CMP list (Stage 08). Forward: gates the AML for DOL-400/500/300; feeds DOL-1600 (labeling/UDI), DOL-1800 (regulatory config), and DOL-2000 (configuration audit). DO-CFG-004 (AML) is the single most blocking output for release (RA-07).
8. Release / adequacy notes
Adequate when the family tree, SKU matrix, AML, and compatibility matrix are Released. Blocked by Evidence-Pending chemistry (CMP-006) and antenna branch (CMP-028/029) — the AML (DO-CFG-004) and hence the configuration freeze cannot complete until the cold-capacity and RF-in-container tests (PRD-VER-004/005) close. Two Critical outputs (DO-CFG-004 AML, DO-CFG-010 identity; DO-CFG-005 compatibility also Critical). </content>
DOL-1000 — Connectivity & Interface Design Outputs (DO-IF) — SCM
Module: DOL-1000 · Output prefix: DO-IF · Owner: Chief Systems Engineer (EM-ENG-001, Interface integrator) · Discipline: Interfaces / Connectivity (cross-subsystem) · Status: Draft baseline.
1. Purpose
Identify and govern the controlled Interface-Control Documents (ICDs) and protocol specifications that fix every SCM interface contract: internal buses/rails/connectors, wireless RF, BLE-GATT provisioning, device↔ cloud store-and-forward, cloud↔enterprise API, OTA, and the ICD-governance rules. These are the authoritative contracts DOL-400/600/800/900 implement.
2. Scope
ICDs derived from INT-001…035 and IFR-001…035, grouped by boundary. All SKUs; RF ICDs carry the antenna Evidence-Pending flag (AD-1).
3. Applicable standards & references
USB Type-C/PD; I2C/SPI; 3GPP LTE-M (Cat-M1); IEEE 802.11; Bluetooth Core/GATT; MQTT v5; TLS 1.2/1.3; OpenAPI; OAuth 2.0; ALCOA+/21 CFR Part 11. Standard — revisions to confirm.
4. Governance & ownership
EM-ENG-001 owns the ICD set (SYS-066/IFR-035); EM-ENG-008 concurs on authenticated/wireless boundaries; HAL (Stage 10) reconciles charge-inhibit, alert, and record interfaces. No undocumented interface may exist (DO-SYS-004).
5. Output categories
ICDs (DO-ICD), protocol/message specs (DO-DATA/DO-ICD), ICD-governance doc (DO-DOC).
6. Design Output Register (this module)
DO-IF-001 — Internal Electrical/Bus/Connector ICD (INT-006…024)
Type | DO-ICD · Discipline: Interfaces (S1–S7) · SKU: All · Subsystem/Component: Internal buses/rails/connectors
File # | SCM-IF-ICD-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE/Elec
Source format | Markdown table · Released format: PDF + CSV · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-001/002/007/013/014, DO-FW-010, DO-MECH-007/008
Input trace | IFR-003/004/006/007/008/009/035; SYS-051…057; INT-006…024 · Risk-control trace: Charge-inhibit (INT-016), probe presence (INT-018), NVM write-safety (INT-021) — reconcile with HAL Stage 10 · Verification trace: Interface conformance test
Critical | Yes (carries charge-inhibit, probe-presence, record-write contracts) · Change class: Major
Acceptance criteria | Every internal interface's pinout, bus params (I2C address/CS map, SPI), rail voltages, and error behavior fixed; NAK/CRC/timeout returns defined error; charge-inhibit signal + status contract explicit.
Mfg use | Direct (harness/connector) · Notes: Rail voltages Evidence Pending (INT-012…015).
DO-IF-002 — USB-C Charge/Data Interface ICD (INT-008/009)
Type | DO-ICD · Discipline: Interfaces (S1/S2) · SKU: All · Subsystem/Component: USB-C / CMP-004
File # | SCM-IF-ICD-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE/Power
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-009 (sink), DO-MECH-008 (aperture)
Input trace | IFR-001/006; SYS-053/054/114; INT-008/009; USB Type-C/PD · Risk-control trace: Charge-input abuse tolerance (HAL) · Verification trace: USB-C/PD conformance
Critical | No · Change class: Major
Acceptance criteria | Sink-only (UFP, Rd) contract; PD PDO cap and rejection rule; wired-OTA/data mode defined; safe under reverse/OVP/short.
Mfg use | Direct · Notes: PD cap Evidence Pending (INT-008).
DO-IF-003 — Wireless RF Interface ICD (INT-025…031)
Type | DO-ICD · Discipline: Interfaces/RF (S5) · SKU: All (band-plan per region) · Subsystem/Component: Radios/antenna
File # | SCM-IF-ICD-003 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: EM-ENG-001/RF · Approval: CSE/Regulatory
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-017/018/021, DO-FW-011, DO-REG-004
Input trace | IFR-010/011/012/013/014; SYS-059/104…118; INT-025…031; 3GPP/802.11/BLE · Risk-control trace: In-container link → delayed/missing alert (HAL) · Verification trace: RF-in-container test (PRD-VER-005) + radio conformance
Critical | Yes (alert/upload egress + region-permitted-band) · Change class: Major
Acceptance criteria | LTE-M band plan/PSM/eDRX, Wi-Fi (WPA2/WPA3, fail-closed), BLE link params, GNSS constellation/A-GNSS, and the in-container attenuated-path budget fixed; region-permitted bands enforced.
Mfg use | Indirect · Notes: Evidence Pending (in-container budget, antenna, band plan — AD-1).
DO-IF-004 — BLE-GATT Provisioning ICD (INT-033)
Type | DO-ICD · Discipline: Interfaces/Security (S5/S9) · SKU: All · Subsystem/Component: GATT service
File # | SCM-IF-ICD-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Security · Approval: CSE/EM-ENG-008
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-APP-003 (client), DO-FW-016 (crypto), DO-SEC-003 (pairing)
Input trace | IFR-015…020; SYS-060/061/022/023/068; INT-029/033 · Risk-control trace: Authenticated arm/profile write; anti-replay; login-free record read (reconcile with HAL Stage 10) · Verification trace: GATT contract + security test
Critical | Yes (authenticated arm/config surface) · Change class: Major
Acceptance criteria | Versioned 128-bit GATT service with fixed characteristics (schema-version, identity, atomic authenticated trip-profile write, anti-replay arm/disarm, login-free verdict + authenticated resumable record-read, status/notify); LE Secure Connections (MITM-protected); unknown handles rejected.
Mfg use | Indirect · Notes: Pairing method reconcile with threat model.
DO-IF-005 — Device↔Cloud Store-and-Forward Protocol ICD (INT-032)
Type | DO-ICD · Discipline: Interfaces (S3/S5/S8) · SKU: All · Subsystem/Component: Device↔cloud
File # | SCM-IF-ICD-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE/EM-ENG-008
Source format | Markdown + schema · Released format: PDF + schema · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-FW-011 (comms agent), DO-CLD-002 (ingest), DO-CLD-003 (schema), DO-SEC-003 (mTLS)
Input trace | IFR-021…026; SYS-011/020/062/063/068; INT-032; MQTT v5/TLS/ALCOA+ · Risk-control trace: Exactly-once ingest; alert-before-bulk; no cleartext fallback (reconcile with HAL Stage 10) · Verification trace: Protocol + store-and-forward + security test
Critical | Yes (record/alert transport integrity) · Change class: Major
Acceptance criteria | MQTT v5/HTTPS over mutual-TLS; topic namespace; schema-version/deviceId/sequence; exactly-once + resume + ack-gated deletion; alert precedence; malformed/unknown-version rejected; no cleartext fallback.
Mfg use | Indirect · Notes: Encoding (JSON/CBOR) Evidence Pending (RF payload budget).
DO-IF-006 — Cloud↔Enterprise API ICD (INT-034)
Type | DO-ICD · Discipline: Interfaces (S8) · SKU: All · Subsystem/Component: Enterprise API / CMP-052
File # | SCM-IF-ICD-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Cloud · Approval: CSE/EM-ENG-008
Source format | OpenAPI + Markdown · Released format: Versioned OpenAPI + PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-CLD-009 (API impl), DO-CLD-013 (RBAC)
Input trace | IFR-027…030; SYS-027/064/068; INT-034; OpenAPI/OAuth 2.0 · Risk-control trace: RBAC/auth + export integrity (HAL) · Verification trace: API contract + auth test
Critical | No · Change class: Major
Acceptance criteria | OpenAPI-defined REST with OAuth 2.0/SSO, RBAC, pagination, rate limits, versioning/deprecation, signed webhooks; export preserves attribution/time/gaps/tamper-evidence; does not gate edge verdict.
Mfg use | Indirect · Notes: Standard.
DO-IF-007 — OTA Update Interface ICD (INT-035)
Type | DO-ICD · Discipline: Interfaces/Security (S3/S5/S8) · SKU: All · Subsystem/Component: OTA
File # | SCM-IF-ICD-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Security · Approval: EM-ENG-008/CSE
Source format | Markdown + manifest schema · Released format: PDF + schema · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-FW-015 (agent), DO-CLD-010 (service), DO-SEC-005 (signing)
Input trace | IFR-031…033; SYS-032/065/068/114; INT-035 · Risk-control trace: Reject wrong-model/downgrade; verified before activation; never brick (reconcile with HAL Stage 10) · Verification trace: OTA campaign + rollback test
Critical | Yes (fleet-wide integrity + fail-safety) · Change class: Major
Acceptance criteria | Package+manifest (target HW/model, from/to version, per-image digest)+signature; applicability/downgrade rules; verify-against-SE before activation; resumable A/B with rollback; status reporting.
Mfg use | Indirect · Notes: Standard.
DO-IF-008 — Sensor/Probe Signal-Chain ICD (INT-017/018)
Type | DO-ICD · Discipline: Interfaces/Sensing (S4) · SKU: All · Subsystem/Component: Sensor buses / CMP-019/020
File # | SCM-IF-ICD-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-013/014, DO-FW-007
Input trace | IFR-003/004; SYS-003/007/056; INT-017/018 · Risk-control trace: Probe-presence/ID + bus-error contract prevents false accept (reconcile with HAL Stage 10) · Verification trace: Sensor-fault-injection
Critical | Yes (probe/sensor fault detection contract) · Change class: Major
Acceptance criteria | I2C address/CS map, probe pinout {sense+, sense−, ID/presence, shield/GND}, and defined error returns on NAK/CRC/open/short/absent; no stale/substituted value.
Mfg use | Direct (probe connector) · Notes: Probe tech to confirm (IFR-003).
DO-IF-009 — HMI Interface ICD (INT-022/023)
Type | DO-ICD · Discipline: Interfaces/HMI (S6) · SKU: All · Subsystem/Component: OLED/LED/button
File # | SCM-IF-ICD-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE/HF
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-017 (OLED), DO-HF-002 (verdict), DO-HF-004 (button)
Input trace | IFR-035; SYS-015/023/058; INT-022/023 · Risk-control trace: Verdict consistent with authoritative state; single-action button (HAL/HF) · Verification trace: HMI contract test
Critical | No · Change class: Major
Acceptance criteria | OLED framebuffer contract; displayed verdict consistent with authoritative state; stuck-segment detectable; single debounced actuation → single unambiguous action.
Mfg use | Direct · Notes: Judgment.
DO-IF-010 — Radio Host-Control Interface ICD (INT-025)
Type | DO-ICD · Discipline: Interfaces (S3/S5) · SKU: All · Subsystem/Component: MCU↔radios
File # | SCM-IF-ICD-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-FW-011 (comms agent), DO-EE-017 (radios)
Input trace | IFR-010/034; SYS-019/062/067; INT-025 · Risk-control trace: Timeout/retry/back-pressure energy-bounded fallback (HAL) · Verification trace: Host-interface test
Critical | No · Change class: Major
Acceptance criteria | Host control link (UART/SPI/SDIO + host API), bearer selection, PSM/eDRX control, and bounded timeout/retry with energy-bounded fallback defined.
Mfg use | Indirect · Notes: Standard (SYS-067 back-pressure Preliminary — HAL).
DO-IF-011 — Debug/EOL-Test Interface ICD (governed, non-fielded)
Type | DO-ICD · Discipline: Interfaces/Security (S3) · SKU: All · Subsystem/Component: SWD/JTAG+UART / IFR-005
File # | SCM-IF-ICD-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Security · Approval: EM-ENG-008/Mfg
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-EE-012 (debug port), DO-SEC-009 (lock), DO-MFG-008 (programming)
Input trace | IFR-005; SRS-SEC; SYS-132-line · Risk-control trace: Locked in production — no unauthenticated key/record readout (reconcile with HAL Stage 10) · Verification trace: Debug-lock verification (DO-TST-012)
Critical | Yes (security — production lock) · Change class: Major
Acceptance criteria | Factory programming/calibration/test access defined; fielded units locked (readout-protection/lifecycle fuse, authenticated debug); not a fielded system interface.
Mfg use | Direct · Notes: Standard (IEC 62443 concepts).
DO-IF-012 — ICD Governance & Interface-Conformance Rules
Type | DO-DOC · Discipline: Interfaces (SYS) · SKU: All · Subsystem/Component: All interfaces
File # | SCM-IF-DOC-012 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-004
Related | DO-GOV-004 (change control), all DO-IF-###
Input trace | IFR-035; SYS-066/067/068; interface-control.md §4 · Risk-control trace: No undocumented interface; authenticated cross-tier identity (HAL/security) · Verification trace: ICD-conformance audit
Critical | No · Change class: Major
Acceptance criteria | Every interface maps to I-01…I-12 with backward-compat change control; wireless/cross-tier changes trigger version-negotiation review; authenticated-identity + uniform timeout/back-pressure rules stated; no undocumented interface.
Mfg use | Indirect · Notes: SYS-067/068 governance Preliminary — reconcile with HAL/threat model.
7. Traceability notes
Parents: IFR-001…035; SYS-051…068 + the interface-touching SYS reqs; INT-001…035; DO-SYS-004 (ICD overview). Forward: DO-IF ICDs are the contracts implemented by DOL-400 (electrical), DOL-600 (FW), DOL-800 (cloud), DOL-900 (mobile); charge/alert/record/OTA ICDs reconcile with HAL (Stage 10) and feed DOL-1500 (conformance test), DOL-1800 (radio certs).
8. Release / adequacy notes
Adequate when all ICDs are Released, internally consistent, and no undocumented interface exists (DO-IF-012 audit). Six Critical (DO-IF-001 internal, DO-IF-003 RF, DO-IF-004 GATT, DO-IF-005 device↔cloud, DO-IF-007 OTA, DO-IF-008 sensor/probe, DO-IF-011 debug-lock — seven flagged). Blocker: RF ICD (DO-IF-003) Evidence Pending (AD-1); SYS-067/068 governance Preliminary — reconcile with HAL. </content>
DOL-1100 — Cybersecurity Design Outputs (DO-SEC) — SCM
Module: DOL-1100 · Output prefix: DO-SEC · Owner: Product Security Engineer (EM-ENG-008) · Discipline: Product Security (cross-tier) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the implemented security design outputs (not the threat-model analysis artifact): security architecture/trust boundaries, identity & authN/authZ design, key management & certificate lifecycle, encryption config, secure-boot & signed-update config, credential storage, audit-log/security-event design, debug locking, production-key provisioning, dependency/SBOM controls, and recovery/revocation.
2. Scope
Cross-tier security controls (device S3, cloud S8, mobile S9). All SKUs. The threat model itself is a design-analysis artifact governed under DOL-000 §output-vs-record and is Evidence Pending; DO-SEC outputs are the implemented controls that trace to it once baselined.
3. Applicable standards & references
IEC 62443 concepts (industrial security), NIST/IoT security concepts, TLS 1.2/1.3, X.509 PKI, 21 CFR Part 11 (attributable/non-repudiable actions), OWASP concepts. Standard — revisions to confirm.
4. Governance & ownership
EM-ENG-008 owns; pairs with DOL-600 (FW crypto), DOL-800 (cloud identity/RBAC/audit), DOL-900 (app security); HAL (Stage 10) reconciles security controls that also mitigate safety-relevant record/alert hazards. Threat model (Evidence Pending) is the parent analysis.
5. Output categories
Security architecture doc (DO-DOC), config (DO-CFG), key/cert lifecycle procedure (DO-PROC), audit/event design (DO-DOC/DO-DATA), dependency controls/SBOM policy (DO-DATA/DO-DOC).
6. Design Output Register (this module)
DO-SEC-001 — Security Architecture & Trust-Boundary Definition
Type | DO-DOC · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: Whole product
File # | SCM-SEC-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Markdown + diagrams · Released format: PDF · Repository: PLM/repo · Parent assembly: DO-SYS-001
Related | DO-SEC-002…011, DO-CLD-013, DO-FW-016
Input trace | PRD-SEC; SRS-SEC; SYS-132…148; AD-5 (split trust) · Risk-control trace: Split trust boundary; record/alert integrity (reconcile with HAL Stage 10) · Verification trace: Security architecture review + pen-test
Critical | Yes (defines the security spine) · Change class: Major
Acceptance criteria | Trust boundaries (edge login-free read vs governed platform), control allocation across tiers, and attack surface defined; consistent with the ICD authenticated-boundary rules (DO-IF-012).
Mfg use | Indirect · Notes: Threat model Evidence Pending; reconcile with HAL.
DO-SEC-002 — Identity, Authentication & Authorization Design
Type | DO-DOC · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: Identity/authZ / CMP-016/053
File # | SCM-SEC-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Markdown · Released format: PDF · Repository: PLM/repo · Parent assembly: DO-SEC-001
Related | DO-CLD-013 (RBAC/SSO), DO-IF-006 (API auth), DO-CFG-010 (identity)
Input trace | SRS-SEC; SYS-110/068; PRD-SEC · Risk-control trace: Authenticated privileged actions (non-repudiation, Part 11) · Verification trace: AuthN/authZ test
Critical | Yes (device/user identity + least privilege) · Change class: Major
Acceptance criteria | Per-device X.509 identity in the SE; user auth via SSO/OAuth; RBAC least-privilege; every privileged action (config/arm/ingest/firmware-accept) authenticated + attributable.
Mfg use | Indirect · Notes: Standard.
DO-SEC-003 — Encryption & Transport-Security Configuration
Type | DO-CFG · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: TLS/BLE crypto
File # | SCM-SEC-CFG-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Config + design · Released format: Versioned config + PDF · Repository: repo/PLM · Parent assembly: DO-SEC-001
Related | DO-IF-005 (mTLS), DO-IF-004 (BLE LESC), DO-FW-016
Input trace | SRS-SEC; SYS-032/110; INT-032/033; TLS 1.2/1.3 · Risk-control trace: No cleartext fallback; MITM-protected pairing (reconcile with HAL Stage 10) · Verification trace: TLS/BLE security test
Critical | Yes (data-in-transit + pairing security) · Change class: Major
Acceptance criteria | Mutual-TLS (pinned trust anchor) for device↔cloud; LE Secure Connections (MITM-protected, not Just Works) for BLE; approved cipher suites; at-rest encryption where applicable; no downgrade/cleartext fallback.
Mfg use | Indirect · Notes: Standard.
DO-SEC-004 — Secure-Boot & Signed-Update Configuration
Type | DO-CFG · Discipline: Security (S3) · SKU: All · Subsystem/Component: Secure boot / CMP-016/035
File # | SCM-SEC-CFG-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Config + design · Released format: Versioned config + PDF · Repository: repo/PLM · Parent assembly: DO-SEC-001
Related | DO-FW-014 (bootloader), DO-FW-015 (OTA), DO-IF-007
Input trace | SRS-SEC/OTA; SYS-032/114; INT-024/035 · Risk-control trace: Only authentic FW executes; signed OTA (reconcile with HAL Stage 10) · Verification trace: Secure-boot/signed-update test
Critical | Yes (code-integrity root of trust) · Change class: Major
Acceptance criteria | Secure-boot chain anchored in the SE; signed-update verification config; downgrade protection; trust-anchor lifecycle defined.
Mfg use | Direct (fuse/anchor set at programming) · Notes: Standard.
DO-SEC-005 — Key-Management & Certificate-Lifecycle Procedure
Type | DO-PROC · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: PKI / CMP-016
File # | SCM-SEC-PRC-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE/Quality
Source format | Procedure + design · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-001
Related | DO-SEC-007 (provisioning), DO-SEC-006 (credential storage), DO-CLD-010 (OTA signing)
Input trace | SRS-SEC; SYS-068; X.509 PKI · Risk-control trace: Signing-key compromise/revocation (HAL/security) · Verification trace: Key-lifecycle audit
Critical | Yes (signing/identity key protection) · Change class: Major
Acceptance criteria | Generation, storage (HSM), distribution, rotation, and revocation of device identity, mTLS, and firmware-signing keys/certs defined; keys never leave the SE/HSM.
Mfg use | Direct (provisioning) · Notes: Standard.
DO-SEC-006 — Credential/Record Tamper-Evidence & Secure-Storage Design
Type | DO-DOC · Discipline: Security (S3/S8) · SKU: All · Subsystem/Component: Record/credential storage
File # | SCM-SEC-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE/EM-ENG-006
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-SEC-001
Related | DO-FW-010 (storage), DO-CLD-003 (schemas), DO-SEC-010 (audit)
Input trace | SRS-SEC/DATA; SYS-028/018; ALCOA+/Part 11 · Risk-control trace: Record alteration/deletion undetected (reconcile with HAL Stage 10) · Verification trace: Tamper-evidence test
Critical | Yes (evidentiary-record integrity) · Change class: Major
Acceptance criteria | Records/audit carry tamper-evidence enabling detection of alteration/deletion; credentials stored in the SE; consistent device + cloud integrity model.
Mfg use | Indirect · Notes: Standard; reconcile with HAL.
DO-SEC-007 — Production Key/Identity Provisioning Design
Type | DO-PROC · Discipline: Security (S3) · SKU: All · Subsystem/Component: Provisioning
File # | SCM-SEC-PRC-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE/Mfg
Source format | Procedure + design · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-005
Related | DO-SW-001 (programming tool), DO-CFG-010 (identity), DO-CLD-004 (registry)
Input trace | SRS-SEC; SYS-030/068; IEC 62443 concepts · Risk-control trace: Unique-identity integrity; no shared keys (HAL/security) · Verification trace: Provisioning verification (DO-TST-012)
Critical | Yes (per-device identity/keys) · Change class: Major
Acceptance criteria | Each device provisioned with a unique identity/key set bound to serial and registry; no shared/default keys; provisioning attested and auditable.
Mfg use | Direct · Notes: Standard.
DO-SEC-008 — Debug-Lock & Anti-Tamper Design (fielded)
Type | DO-DOC · Discipline: Security (S3) · SKU: All · Subsystem/Component: Debug lock / IFR-005
File # | SCM-SEC-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-001
Related | DO-EE-012 (debug port), DO-IF-011 (test ICD), DO-MFG-008
Input trace | SRS-SEC; SYS-132-line; INT (debug) · Risk-control trace: No unauthenticated key/record readout in field (reconcile with HAL Stage 10) · Verification trace: Debug-lock verification (DO-TST-012)
Critical | Yes (fielded readout protection) · Change class: Major
Acceptance criteria | Fielded units have debug disabled/authenticated and readout-protection/lifecycle fuse set; anti-tamper detection where feasible; no unauthenticated code/key/record readout.
Mfg use | Direct (lock at EOL) · Notes: Standard.
DO-SEC-009 — Security Audit-Log & Security-Event Design
Type | DO-DOC · Discipline: Security (S8) · SKU: All · Subsystem/Component: Audit / CMP-054
File # | SCM-SEC-DOC-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE/EM-ENG-006
Source format | Markdown + schema · Released format: PDF + schema · Repository: repo/PLM · Parent assembly: DO-SEC-001
Related | DO-CLD-003 (audit schema), DO-SEC-006, DO-CLD-014 (monitoring)
Input trace | SRS-SEC; SYS-028; 21 CFR Part 11 · Risk-control trace: Non-repudiation of config/arming/firmware/record actions (HAL/Part 11) · Verification trace: Audit-log test
Critical | No · Change class: Major
Acceptance criteria | Append-only, attributable, time-stamped, un-alterable-without-detection audit of config/arming/firmware/calibration/record actions; security events surfaced to monitoring.
Mfg use | Indirect · Notes: Standard.
DO-SEC-010 — Vulnerability-Management & Security-Update Design
Type | DO-DOC · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: Vuln mgmt
File # | SCM-SEC-DOC-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-001
Related | DO-FW-015 (OTA), DO-CLD-010, DO-SEC-011 (SBOM)
Input trace | SRS-SEC; SYS-032; IEC 62443 concepts · Risk-control trace: Timely security patching via signed OTA (HAL/security) · Verification trace: Security-update test
Critical | No · Change class: Major
Acceptance criteria | Coordinated-disclosure intake, vuln triage, and a security-update path (signed OTA) with defined SLAs; incident-response hooks.
Mfg use | Indirect · Notes: Standard.
DO-SEC-011 — Dependency-Security Controls & Consolidated SBOM Policy
Type | DO-DATA · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: All software deps
File # | SCM-SEC-DAT-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Policy + SBOM refs · Released format: PDF + machine-readable SBOM index · Repository: repo/PLM · Parent assembly: DO-GOV-007
Related | DO-FW-018, DO-SW-006, DO-CLD-015, DO-APP-010
Input trace | SRS-SEC; IEC 62304 (SOUP) · Risk-control trace: Dependency-vulnerability control across tiers · Verification trace: Vuln-scan gates (all tiers)
Critical | No · Change class: Major
Acceptance criteria | Consolidated SBOM index across FW/device-SW/cloud/mobile; dependency-control policy; vuln-scan gates enforced at each tier's release.
Mfg use | Indirect · Notes: Standard.
DO-SEC-012 — Key/Certificate Revocation & Recovery Design
Type | DO-DOC · Discipline: Security (all tiers) · SKU: All · Subsystem/Component: Revocation/recovery
File # | SCM-SEC-DOC-012 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-005
Related | DO-SEC-005 (key lifecycle), DO-CLD-004 (registry), DO-SW-005 (recovery)
Input trace | SRS-SEC; SYS-068; X.509 PKI · Risk-control trace: Compromised-device/key containment (HAL/security) · Verification trace: Revocation/recovery test
Critical | No · Change class: Major
Acceptance criteria | Revoke a compromised device/key and recover it to a trusted state without exposing keys; registry reflects revocation; OTA/ingest reject revoked identities.
Mfg use | Indirect · Notes: Standard.
7. Traceability notes
Parents: PRD-SEC; SRS-SEC/OTA/DATA; SYS-032/068/110/132…148; CMP-016/044/053/054; INT-024/032/033/035; the (Evidence-Pending) threat model. Forward: implemented controls pair with DOL-600 (FW crypto/boot/OTA), DOL-800 (identity/RBAC/audit), DOL-900 (app security), DOL-1000 (authenticated boundaries), DOL-1200 (where a security control also mitigates a safety hazard), DOL-1800 (cybersecurity documentation outputs).
8. Release / adequacy notes
Adequate when the security architecture, identity/authZ, crypto/transport, secure-boot/signed-update, key lifecycle/provisioning, tamper-evidence, and debug-lock outputs are Released and pen-tested. Seven Critical (DO-SEC-001, -002, -003, -004, -005, -006, -007, -008). Blocker: the threat model is Evidence Pending — DO-SEC controls are drafted but must reconcile to it and to HAL (Stage 10) before release. </content>
DOL-1200 — Safety & Risk-Control Design Outputs (DO-SAFE) — SCM
Module: DOL-1200 · Output prefix: DO-SAFE · Owner: Reliability/Safety Engineer · Discipline: Safety / Risk Controls (cross-subsystem) · Status: Draft baseline — PRELIMINARY, reconcile with HAL (Stage 10, ORCH-017).
1. Purpose
Identify the implemented risk-control design outputs. Each DO-SAFE output is the "Implementation Reference" for a Hazard-Analysis-Library risk control (HA-<CODE>-NNN / RCR-###, Stage 10). Every HAL control classified as inherent-design or protective must appear here. These are not new designs but pointers to the concrete safety controls implemented in DOL-300/400/600/800 (with a safety-focused acceptance criterion and HAL linkage).
2. Scope
Hardware & software safety controls, protective measures, alarms/interlocks/limits, monitoring, fault detection, safe states, degraded modes, warnings, and control independence. All SKUs. The HAL is not yet baselined; all HA-###/RCR-### references below are placeholders (Evidence Pending) to be reconciled at Stage 10.
3. Applicable standards & references
ISO 14971 (risk management, severity-first Action Priority per CLAUDE.md rule 3); IEC 62368-1 (electrical/ battery safety); UN 38.3 (Li-battery); IEC 62304 (software safety class). Standard — revisions to confirm.
4. Governance & ownership
Reliability/Safety Engineer owns; each control's implementing output is owned by its discipline module (Power/FW/Cloud). This module is the safety cross-reference layer. Every entry is flagged Critical and "reconcile with HAL Stage 10." Severity dominates (Action Priority).
5. Output categories
Risk-control implementation references (DO-DOC pointing to DO-EE/DO-FW/DO-CLD/DO-MECH outputs) with a safety-acceptance criterion and HAL/RCR linkage.
6. Design Output Register (this module)
DO-SAFE-001 — Charge-Inhibit-Below-0 °C Interlock (implementation reference)
Type | DO-DOC · Discipline: Safety/Power (S2/S3) · SKU: All · Subsystem/Component: Charge-inhibit / CMP-008/010/012
File # | SCM-SAF-DOC-001 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE (EM-ENG-006)
Source format | Markdown · Released format: PDF · Repository: PLM (linked to risk file) · Parent assembly: DO-SYS-009
Related | Implements DO-EE-007 (HW) + DO-FW-005 (FW) · DO-IF-001 (INT-016)
Input trace | D3; SYS-094/095; PRD-SAFE; ADR-004 · Risk-control trace: HA-BATT-### / RCR-### (cold-charge cell damage/hazard) — reconcile with HAL Stage 10 · Verification trace: Cold-charge-inhibit fault-injection test
Critical | Yes · Change class: Major
Acceptance criteria | Charging is prevented below 0 °C nominal by an MCU-independent HW interlock, with FW defense-in-depth; inhibit persists in any FW fault; inhibited state + cause reported. Primary D3 charge-safety control.
Mfg use | Indirect · Notes: Reconcile with HAL Stage 10 (severity-dominant).
DO-SAFE-002 — Battery Protection (OV/UV/OC/OT/short) + Secondary Fuse
Type | DO-DOC · Discipline: Safety/Power (S2) · SKU: All · Subsystem/Component: BMS+fuse / CMP-008/009
File # | SCM-SAF-DOC-002 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-EE-005 (BMS) + DO-EE-010 (fuse)
Input trace | SYS-096; PRD-SAFE; IEC 62368-1 · Risk-control trace: HA-BATT-### / RCR-### (cell OV/UV/OC/OT/short → thermal event) — reconcile with HAL Stage 10 · Verification trace: Protection fault-injection (DO-TST-009)
Critical | Yes · Change class: Major
Acceptance criteria | Autonomous MCU-independent protection with a non-resettable secondary backstop; single-fault-safe (hazard terminated if primary IC fails).
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-003 — Charge-Input Abuse Protection (reverse/OVP/surge/short)
Type | DO-DOC · Discipline: Safety/Power (S2) · SKU: All · Subsystem/Component: Input protection / CMP-014
File # | SCM-SAF-DOC-003 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-EE-009 (USB-C sink) + DO-EE-010
Input trace | SYS-102; PRD-SAFE; USB-C/PD; IEC 62368-1 · Risk-control trace: HA-CHG-### / RCR-### (out-of-spec charge source) — reconcile with HAL Stage 10 · Verification trace: Reverse/OVP/short test
Critical | Yes · Change class: Major
Acceptance criteria | Device rejects/limits out-of-spec source; safe under reverse-insertion/OVP/short; PD PDO capped.
Mfg use | Indirect · Notes: Reconcile with HAL; PD cap Evidence Pending.
DO-SAFE-004 — Two-Sided Excursion-Detection Control (missed-freeze/heat)
Type | DO-DOC · Discipline: Safety/FW (S3) · SKU: All · Subsystem/Component: Detection / CMP-039
File # | SCM-SAF-DOC-004 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-FW-006 (detection) + DO-EE-013 (probe front-end)
Input trace | SYS-008/009/010/014; PRD-SAFE; ADR-005 · Risk-control trace: HA-DET-### / RCR-### (missed excursion → wrong disposition → compromised vaccine) — reconcile with HAL Stage 10 · Verification trace: Detection-logic test (profiled excursions)
Critical | Yes · Change class: Major
Acceptance criteria | Freeze detection reliability equals heat; two-sided product-stability logic; no undetected excursion path. Excursion logic — top HAL headline control.
Mfg use | Indirect · Notes: Reconcile with HAL Stage 10.
DO-SAFE-005 — Probe-Fault / Sensor-Plausibility Control (false-accept prevention)
Type | DO-DOC · Discipline: Safety/Sensing (S3/S4) · SKU: All · Subsystem/Component: Probe/sensor faults / CMP-020/038
File # | SCM-SAF-DOC-005 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-FW-007 (drivers/plausibility) + DO-EE-013 (presence detect) + DO-IF-008
Input trace | SYS-007/024; PRD-SAFE · Risk-control trace: HA-SENS-### / RCR-### (open/absent probe reads room temp → false accept) — reconcile with HAL Stage 10 · Verification trace: Probe-fault-injection test
Critical | Yes · Change class: Major
Acceptance criteria | Absent/wrong/open/short probe and implausible sensor reads are detected and flagged, never read as a valid product-zone temperature.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-006 — Verdict-Integrity Control (suppress false "accept")
Type | DO-DOC · Discipline: Safety/FW (S3) · SKU: All · Subsystem/Component: Verdict / CMP-040
File # | SCM-SAF-DOC-006 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-FW-008 (verdict) + DO-CLD-006 (recompute) + DO-HF-002 (display)
Input trace | SYS-015/016/017; PRD-SAFE; ADR-006 · Risk-control trace: HA-VER-### / RCR-### (false accept on gaps/faults → released compromised product) — reconcile with HAL Stage 10 · Verification trace: Verdict-consistency + gap-handling test
Critical | Yes · Change class: Major
Acceptance criteria | Verdict suppresses false accept on data gaps/faults; incomplete-monitoring qualifier shown; device and cloud verdicts consistent.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-007 — Safe-State / Fault-Safe Mode Control
Type | DO-DOC · Discipline: Safety/FW (S3) · SKU: All · Subsystem/Component: Safe state / CMP-037/045
File # | SCM-SAF-DOC-007 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-006
Related | Implements DO-FW-003 (mode) + DO-FW-013 (fault handler) + DO-SYS-009
Input trace | SYS-047/099; PRD-SAFE · Risk-control trace: HA-SW-### / RCR-### (FW hang/fault corrupts record or armed state) — reconcile with HAL Stage 10 · Verification trace: Fault-injection + watchdog test
Critical | Yes · Change class: Major
Acceptance criteria | Any fault transitions to a defined safe state without corrupting the committed log or armed state; watchdog-backed recovery.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-008 — Record-Integrity / Power-Fail-Safe Storage Control
Type | DO-DOC · Discipline: Safety/FW (S3/S7) · SKU: All · Subsystem/Component: Storage / CMP-041
File # | SCM-SAF-DOC-008 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE (EM-ENG-006)
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-005
Related | Implements DO-FW-010 (storage) + DO-EE-002 (rail interlock) + DO-SEC-006 (tamper-evidence)
Input trace | SYS-018/021/028/099; PRD-SAFE; ALCOA+ · Risk-control trace: HA-DATA-### / RCR-### (torn/lost/fabricated record → wrong disposition) — reconcile with HAL Stage 10 · Verification trace: Power-fail-during-write test
Critical | Yes · Change class: Major
Acceptance criteria | No torn/lost/fabricated record on power loss; tamper-evident; no overwrite of unsent data.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-009 — Excursion-Alert Delivery Control (delayed/missing alert)
Type | DO-DOC · Discipline: Safety/FW/Cloud (S3/S5/S8) · SKU: All · Subsystem/Component: Alert path / CMP-042/048
File # | SCM-SAF-DOC-009 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-005
Related | Implements DO-FW-011 (comms) + DO-CLD-007 (rules) + DO-CLD-008 (notify)
Input trace | D2; SYS-011/012/013/014; PRD-SAFE · Risk-control trace: HA-COMM-### / RCR-### (delayed/missing alert → no time to act) — reconcile with HAL Stage 10 · Verification trace: Alert-latency + store-and-forward test
Critical | Yes · Change class: Major
Acceptance criteria | Excursion event always committed locally; alert sent ahead of bulk, ack-required, escalated if unacknowledged; connectivity loss degrades to buffering, never record loss.
Mfg use | Indirect · Notes: Reconcile with HAL; in-container link Evidence Pending.
DO-SAFE-010 — Battery Mechanical/Thermal Isolation Control
Type | DO-DOC · Discipline: Safety/Mech (S1/S2) · SKU: All · Subsystem/Component: Battery retention / CMP-007
File # | SCM-SAF-DOC-010 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-MECH-012 (retention) + DO-EE-004 (pack)
Input trace | SYS-090; PRD-SAFE; UN 38.3 · Risk-control trace: HA-BATT-### / RCR-### (cell crush/puncture under shock → thermal event) — reconcile with HAL Stage 10 · Verification trace: Drop/vibe + UN 38.3
Critical | Yes · Change class: Major
Acceptance criteria | Cell retained without crush/puncture through drop/vibration/transport; thermal/electrical clearance maintained.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-011 — Ingress/Condensation Protection Control
Type | DO-DOC · Discipline: Safety/Mech/PCB (S1/S3) · SKU: All · Subsystem/Component: Seal + coat / CMP-002/034
File # | SCM-SAF-DOC-011 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-MECH-005 (seal) + DO-PCB-010 (conformal coat)
Input trace | SYS-069; PRD-SAFE; IEC 60529 · Risk-control trace: HA-ENV-### / RCR-### (condensation/ingress → leakage/corrosion → record-integrity electronics fault) — reconcile with HAL Stage 10 · Verification trace: IP + cold-cycle + cleanliness test
Critical | Yes · Change class: Major
Acceptance criteria | Seal + conformal coat protect record-integrity electronics through −30↔+60 °C condensation cycling over reuse life.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-012 — Energy-Reserve / Loss-of-Monitoring Control
Type | DO-DOC · Discipline: Safety/FW (S2/S3) · SKU: All · Subsystem/Component: Energy governor / CMP-037
File # | SCM-SAF-DOC-012 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-007
Related | Implements DO-FW-004 (governor) + DO-EE-008 (gauge) + DO-SYS-007
Input trace | D3; SYS-086/097; PRD-SAFE/REL · Risk-control trace: HA-PWR-### / RCR-### (energy exhaustion mid-trip → loss of monitoring, silent) — reconcile with HAL Stage 10 · Verification trace: 90-day cold endurance (PRD-VER-004)
Critical | Yes · Change class: Major
Acceptance criteria | Energy governor + honest low-battery indication prevent silent loss of monitoring; reserve preserves record commit and alert on low battery.
Mfg use | Indirect · Notes: Reconcile with HAL Stage 10; budget Evidence Pending (D3).
DO-SAFE-013 — Charge/Fault-State Legibility Control (no silent failure)
Type | DO-DOC · Discipline: Safety/HMI (S3/S6) · SKU: All · Subsystem/Component: HMI/LED / CMP-032
File # | SCM-SAF-DOC-013 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE/HF
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-HF-005 (states) + DO-EE-008 (gauge) + DO-IF-001 (INT-016/020)
Input trace | SYS-024/094/100; PRD-SAFE/UX · Risk-control trace: HA-HMI-### / RCR-### (silent charge-inhibit/fault → device deployed unready) — reconcile with HAL Stage 10 · Verification trace: State-legibility test
Critical | Yes · Change class: Major
Acceptance criteria | Charge-inhibited/fault/unready states are made legible (LED/OLED) with cause; no silent failure to charge or arm.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-SAFE-014 — OTA Never-Brick / Unmonitored-Window Control
Type | DO-DOC · Discipline: Safety/FW/Cloud (S3/S8) · SKU: All · Subsystem/Component: OTA / CMP-043/051
File # | SCM-SAF-DOC-014 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Safety Eng · Approval: Safety/CSE/EM-ENG-008
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-009
Related | Implements DO-FW-014/015 + DO-CLD-010 + DO-IF-007
Input trace | SYS-032/114; PRD-SAFE; ADR-008 · Risk-control trace: HA-OTA-### / RCR-### (bricked/unmonitored fleet after update) — reconcile with HAL Stage 10 · Verification trace: OTA rollback + version-skew test
Critical | Yes · Change class: Major
Acceptance criteria | Fail-safe A/B apply with automatic rollback; trip data preserved; unmonitored windows disclosed; never bricks a device.
Mfg use | Indirect · Notes: Reconcile with HAL.
7. Traceability notes
Parents: PRD-SAFE (Preliminary); the SYS-safety set (SYS-149…162); ADR-003/004/005/006/008. Each output is the "Implementation Reference" for a HAL risk-control record (HA-###/RCR-###, Stage 10). Forward: feeds HAL (Stage 10) directly, Risk-Control Requirements (Stage 12), and risk-control verification (Stage 13, PVL). The implementing outputs live in DOL-300/400/600/800/1000/1300.
8. Release / adequacy notes
All 14 outputs are Critical and PRELIMINARY. This module cannot be released until the HAL (Stage 10) is baselined and every inherent-design/protective HAL control has a matching DO-SAFE Implementation Reference (release criterion RA-05). The D3 charge-safety trio (DO-SAFE-001/002/003), the excursion/verdict trio (DO-SAFE-004/005/006), and the record/alert/energy set (DO-SAFE-008/009/012) are the severity-dominant controls. Reconcile with HAL Stage 10. </content>
DOL-1300 — Human Factors & UI Design Outputs (DO-HF) — SCM
Module: DOL-1300 · Output prefix: DO-HF · Owner: Human Factors Engineer · Discipline: Human Factors / HMI (S6, S9, S8-dashboard) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the human-factors design outputs: on-device HMI/verdict presentation, warning/alarm presentation, iconography, display characteristics, physical-control design, use-error mitigations, and user-assistance content — the outputs that make the SCM's critical tasks (arm, read verdict, respond to alert, provision) safe and unambiguous.
2. Scope
On-device OLED/LED/button (S6), mobile-app critical flows (S9, HF aspects), fleet-dashboard reporting HF (S8). All SKUs. Excludes app implementation (DOL-900) and cloud services (DOL-800) beyond HF specs.
3. Applicable standards & references
IEC 62366-1 (usability engineering concepts), IEC 60601-1-8 style alarm concepts (non-medical adaptation), WCAG (accessibility), ISO 9241 (ergonomics). Standard — revisions to confirm.
4. Governance & ownership
HF Engineer owns; pairs with DOL-900 (app), DOL-1600 (label content), DOL-1200 (use-error safety controls); HAL (Stage 10) reconciles use-error mitigations.
5. Output categories
UI/HF specs (DO-DOC), iconography/asset specs (DO-DOC), physical-control spec (DO-SPEC), user-assistance content (DO-MAN/DO-DOC).
6. Design Output Register (this module)
DO-HF-001 — On-Device HMI Specification (information architecture)
Type | DO-DOC · Discipline: HF (S6) · SKU: All · Subsystem/Component: OLED/LED/button
File # | SCM-HF-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE
Source format | Markdown + wireframes · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-HF-002…005, DO-IF-009 (HMI ICD)
Input trace | PRD-UX; SRS-DEV; SYS-015/019/023 · Risk-control trace: Legible critical info (HF/HAL) · Verification trace: Usability test
Critical | No · Change class: Major
Acceptance criteria | Complete on-device information architecture (states, verdict, charge/fault, arm) with screen hierarchy and interaction model.
Mfg use | Indirect · Notes: Judgment.
DO-HF-002 — Verdict Presentation Design (accept/reject/quarantine)
Type | DO-DOC · Discipline: HF/Safety (S6) · SKU: All · Subsystem/Component: OLED verdict / CMP-030
File # | SCM-HF-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Safety
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-FW-008 (verdict), DO-SAFE-006, DO-IF-009
Input trace | SYS-015/017; PRD-UX/SAFE; ADR-06 · Risk-control trace: Verdict + incomplete-monitoring qualifier unambiguous (reconcile with HAL Stage 10) · Verification trace: Verdict-comprehension usability test
Critical | Yes (last-mile disposition surface) · Change class: Major
Acceptance criteria | Accept/reject/quarantine and incomplete-monitoring qualifier are unambiguous, legible at −30 °C and under glare, and cannot be misread as accept when reject/quarantine; distinct without color-only coding.
Mfg use | Indirect · Notes: OLED cold response Evidence Pending; reconcile with HAL.
DO-HF-003 — Iconography, Symbols & Status-Indicator Design
Type | DO-DOC · Discipline: HF (S6/S9) · SKU: All · Subsystem/Component: Icons/LED semantics
File # | SCM-HF-DOC-003 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE
Source format | Design + spec · Released format: PDF + assets · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-HF-005 (states), DO-APP-004, DO-LBL-004 (symbols)
Input trace | PRD-UX; ISO 7000/IEC 60417 symbol concepts · Risk-control trace: Consistent status semantics (HF/HAL) · Verification trace: Symbol-comprehension test
Critical | No · Change class: Minor
Acceptance criteria | Icon/LED semantics defined and consistent across device/app/label; comprehensible without training; standardized symbols where they exist.
Mfg use | Indirect · Notes: Standard.
DO-HF-004 — Physical Control (glove-friendly button) HF Specification
Type | DO-SPEC · Discipline: HF (S6) · SKU: All · Subsystem/Component: Button / CMP-031
File # | SCM-HF-SPC-004 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Mech
Source format | Spec · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-MECH-014 (button mech), DO-IF-009 (INT-023)
Input trace | SYS-023; PRD-UX; ISO 9241 · Risk-control trace: Single-action, no multi-press ambiguity (HF/HAL) · Verification trace: Gloved-actuation usability test
Critical | No · Change class: Major
Acceptance criteria | Force/travel/size support gloved actuation at −30 °C; single actuation maps to a single unambiguous action (arm / wake-to-read); no multi-press ambiguity.
Mfg use | Indirect · Notes: Judgment.
DO-HF-005 — Device-State & Warning Presentation Design
Type | DO-DOC · Discipline: HF/Safety (S6) · SKU: All · Subsystem/Component: States/warnings
File # | SCM-HF-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Safety
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-SAFE-013 (legibility control), DO-EE-008 (gauge), DO-IF-009
Input trace | SYS-024/094/100; PRD-UX/SAFE · Risk-control trace: Charge-inhibit/fault/unready legible — no silent failure (reconcile with HAL Stage 10) · Verification trace: State-legibility usability test
Critical | Yes (unready/fault visibility) · Change class: Major
Acceptance criteria | Armed/charging/complete/charge-inhibited/fault/low-battery/unready states presented with cause; a not-ready device is never mistaken for ready.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-HF-006 — Provisioning & Arming Task-Flow (use-error mitigation) Design
Type | DO-DOC · Discipline: HF/Safety (S9) · SKU: All · Subsystem/Component: App critical flows
File # | SCM-HF-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Safety
Source format | Markdown + task-flow · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-APP-002 (screens), DO-APP-003 (GATT), DO-SAFE-006
Input trace | SYS-022/024/025; PRD-UX/SAFE; IEC 62366-1 · Risk-control trace: Pre-arm checks prevent deploying unready/misconfigured device (reconcile with HAL Stage 10) · Verification trace: Task-flow usability + use-error test
Critical | Yes (arming/provisioning use-errors) · Change class: Major
Acceptance criteria | Provisioning/trip-config/arming flows include pre-arm checks (readiness, calibration, config validation) and confirmations that mitigate the identified critical use-errors.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-HF-007 — Fleet-Dashboard Reporting HF Specification
Type | DO-DOC · Discipline: HF (S8) · SKU: All · Subsystem/Component: Dashboard / CMP-050
File # | SCM-HF-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-CLD-012
Related | DO-CLD-012 (dashboard), DO-HF-003
Input trace | SYS-031; PRD-CLOUD/UX · Risk-control trace: Clear release/acceptance decision support (HF/HAL) · Verification trace: Reporting usability test
Critical | No · Change class: Major
Acceptance criteria | Alerts, readiness, and disposition are presented so a responder can act correctly under time pressure; overdue/stray flagged; no misleading "all clear."
Mfg use | Indirect · Notes: Judgment.
DO-HF-008 — Alarm/Notification Presentation & Prioritization Design
Type | DO-DOC · Discipline: HF/Safety (S8/S9) · SKU: All · Subsystem/Component: Alert presentation
File # | SCM-HF-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Safety
Source format | Markdown + design · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-CLD-007/008, DO-SAFE-009
Input trace | SYS-012/014; PRD-UX/SAFE · Risk-control trace: Excursion alert salient + not habituating (HF/HAL) · Verification trace: Alert-comprehension test
Critical | No · Change class: Major
Acceptance criteria | Excursion alerts are salient, prioritized over informational messages, and designed to avoid alarm fatigue; acknowledgment path clear.
Mfg use | Indirect · Notes: Judgment.
DO-HF-009 — Quick-Start / User-Assistance Content
Type | DO-MAN · Discipline: HF (S1/S9) · SKU: All (localized) · Subsystem/Component: User assistance
File # | SCM-HF-MAN-009 · Rev: TBD-rev · Status: Draft · Owner: HF Eng · Approval: CSE/Regulatory
Source format | Content + layout · Released format: PDF · Repository: PLM · Parent assembly: DO-LBL-006 (IFU)
Related | DO-LBL-006 (IFU), DO-SVC-001 (install)
Input trace | PRD-UX/SVC; SYS-025; IEC 62366-1 · Risk-control trace: Correct arm/read/respond guidance (HF/HAL) · Verification trace: Comprehension test
Critical | No · Change class: Major
Acceptance criteria | Quick-start covers place/arm/read-verdict/respond-to-alert/recharge tasks; comprehensible; localizable.
Mfg use | Indirect · Notes: Judgment.
DO-HF-010 — Display-Characteristics & Cold-Legibility Specification
Type | DO-SPEC · Discipline: HF (S6) · SKU: All · Subsystem/Component: OLED / CMP-030
File # | SCM-HF-SPC-010 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: HF Eng · Approval: CSE
Source format | Spec · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-001
Related | DO-EE-017 (OLED), DO-MECH-014 (window)
Input trace | SYS-015/041; PRD-UX · Risk-control trace: Verdict legibility at −30 °C (HF/HAL) · Verification trace: Cold-display test
Critical | No · Change class: Major
Acceptance criteria | Contrast, viewing angle, response time, and character size legible at −30 °C within the responsiveness budget.
Mfg use | Indirect · Notes: Evidence Pending (OLED cold response, CMP-030).
7. Traceability notes
Parents: PRD-UX/SAFE; SRS-DEV/MOB; SYS-015/017/019/022/023/024/025/031; IEC 62366-1. Forward: pairs with DOL-900 (app), DOL-1600 (label/IFU content), DOL-1200 (use-error safety controls); use-error mitigations reconcile with HAL (Stage 10).
8. Release / adequacy notes
Adequate when HMI/verdict/state/task-flow specs are Released and usability-validated. Three Critical (DO-HF-002 verdict, DO-HF-005 states, DO-HF-006 arming/provisioning). Blocker: OLED cold legibility Evidence Pending; use-error mitigations reconcile with HAL. </content>
DOL-1400 — Manufacturing Design Outputs (DO-MFG) — SCM
Module: DOL-1400 · Output prefix: DO-MFG · Owner: Manufacturing Engineer · Discipline: Manufacturing / Process (all subsystems) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the manufacturing design outputs from which the SCM can be built at a contract manufacturer: manufacturing BOM/AML, assembly drawings, work instructions, process/tooling/fixture specs, soldering/adhesive/torque parameters, calibration/programming/serialization procedures, cleanliness/handling controls, and critical-process-parameter definitions.
2. Scope
Full build (SMT, enclosure assembly, sealing, programming, calibration, serialization, final assembly). All SKUs; regional/antenna deltas as populate/config options. Feeds Design Transfer (Stage 15).
3. Applicable standards & references
IPC-A-610 / J-STD-001 (assembly/soldering), IPC-CC-830 (coat), ISO 13485 §7.5 (production controls), ISO 14644 concepts (cleanliness if needed), UN 38.3 handling. Standard — revisions to confirm.
4. Governance & ownership
Manufacturing Engineer owns; consumes DOL-300/400/500/600/700 outputs; pairs with DOL-1500 (test) and DOL-1600 (packaging); HAL (Stage 10) reconciles safety-critical process parameters (charge-inhibit verification at build, seal integrity, programming/lock).
5. Output categories
Manufacturing BOM/AML (DO-BOM), assembly drawings (DO-DRW), work instructions (DO-WI), process/tooling specs (DO-SPEC/DO-TOOL), procedures (DO-PROC), critical-process-parameter defs (DO-SPEC).
6. Design Output Register (this module)
DO-MFG-001 — Manufacturing BOM & Approved-Manufacturer List (mBOM/AML)
Type | DO-BOM · Discipline: Mfg (all) · SKU: All · Subsystem/Component: Full build
File # | SCM-MFG-BOM-001 · Rev: TBD-rev (Evidence Pending — chemistry/antenna) · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | PLM mBOM · Released format: BOM export + PDF · Repository: PLM · Parent assembly: DO-CFG-004
Related | DO-EE-020, DO-PCB-013, DO-MECH-019, DO-CFG-004
Input trace | All discipline BOMs; PRD-MFG · Risk-control trace: Critical-part control (cell/BMS/SE/fuse) · Verification trace: Incoming inspection (DO-TST-001)
Critical | No · Change class: Major
Acceptance criteria | Consolidated mBOM with AML, reference designators, and consumables; buildable at CM; consistent with all discipline BOMs.
Mfg use | Direct · Notes: Evidence Pending (AML freeze blocked by chemistry/antenna, RA-07).
DO-MFG-002 — SMT / PCBA Assembly Work Instructions
Type | DO-WI · Discipline: Mfg (S3) · SKU: All (populate options) · Subsystem/Component: PCBA / CMP-034
File # | SCM-MFG-WI-002 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-PCB-008
Related | DO-PCB-008/009/012, DO-MFG-003
Input trace | DO-PCB-008; IPC-A-610 · Risk-control trace: Correct placement of cell/BMS/SE (HAL) · Verification trace: AOI/SPI + first-article
Critical | No · Change class: Major
Acceptance criteria | Step-by-step SMT assembly with placement, orientation, and inspection points; populate/no-populate per SKU.
Mfg use | Direct · Notes: Standard.
DO-MFG-003 — Reflow / Soldering Process Parameter Specification
Type | DO-SPEC · Discipline: Mfg (S3) · SKU: All · Subsystem/Component: PCBA / CMP-034
File # | SCM-MFG-SPC-003 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | Spec + profile · Released format: PDF · Repository: PLM · Parent assembly: DO-MFG-002
Related | DO-PCB-009 (stencil), DO-PCB-010 (coat)
Input trace | J-STD-001; DO-PCB-003 · Risk-control trace: Solder-joint reliability under thermal cycling (HAL) · Verification trace: Cross-section + thermal-cycle test
Critical | No · Change class: Major
Acceptance criteria | Reflow profile, paste, and process window give reliable joints for fine-pitch/QFN parts surviving −30↔+60 °C cycling; a critical process parameter (CPP).
Mfg use | Direct · Notes: Standard.
DO-MFG-004 — Enclosure Assembly & Sealing Work Instruction
Type | DO-WI · Discipline: Mfg (S1) · SKU: All (+ -XA aperture) · Subsystem/Component: Enclosure / CMP-001/002
File # | SCM-MFG-WI-004 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-MECH-000
Related | DO-MECH-005 (seal), DO-MFG-006 (torque), DO-TST-006 (IP)
Input trace | DO-MECH-000/005; MER · Risk-control trace: Seal integrity at assembly (reconcile with HAL Stage 10) · Verification trace: IP test (DO-TST-006)
Critical | Yes (seal integrity is a safety control) · Change class: Major
Acceptance criteria | Seal seating, groove-fill, and closure sequence specified so every unit meets IP after assembly; probe gland and USB-C aperture sealed per spec.
Mfg use | Direct · Notes: Reconcile with HAL.
DO-MFG-005 — Conformal-Coat Application Work Instruction
Type | DO-WI · Discipline: Mfg (S3) · SKU: All · Subsystem/Component: PCBA / CMP-034
File # | SCM-MFG-WI-005 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-PCB-010
Related | DO-PCB-010 (coat spec), DO-TST-005 (cleanliness)
Input trace | DO-PCB-010; IPC-CC-830 · Risk-control trace: Condensation protection (HAL) · Verification trace: Coat coverage + cleanliness test
Critical | No · Change class: Major
Acceptance criteria | Coat coverage and keep-outs (sensors/antenna/contacts) applied per spec; cleanliness before coat within limits.
Mfg use | Direct · Notes: Standard.
DO-MFG-006 — Torque / Fastening & Adhesive Parameter Specification
Type | DO-SPEC · Discipline: Mfg (S1) · SKU: All · Subsystem/Component: Enclosure fasteners
File # | SCM-MFG-SPC-006 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Mech
Source format | Spec · Released format: PDF · Repository: PLM · Parent assembly: DO-MFG-004
Related | DO-MECH-004 (housing), DO-MFG-004
Input trace | MER; DO-MECH-000 · Risk-control trace: Seal-compression consistency (HAL) · Verification trace: Torque audit
Critical | No · Change class: Major
Acceptance criteria | Torque values/sequence and any adhesive cure parameters ensure consistent seal compression without over-stress; a CPP.
Mfg use | Direct · Notes: Judgment.
DO-MFG-007 — Battery Handling & Installation Work Instruction
Type | DO-WI · Discipline: Mfg/Safety (S2) · SKU: All · Subsystem/Component: Battery / CMP-007
File # | SCM-MFG-WI-007 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Safety
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-EE-004
Related | DO-MECH-012 (retention), DO-SAFE-010, DO-LBL-010 (transport)
Input trace | DO-EE-004; UN 38.3; IEC 62368-1 · Risk-control trace: Cell abuse prevention in handling/install (reconcile with HAL Stage 10) · Verification trace: Handling audit + charge-safety test
Critical | Yes (cell handling safety) · Change class: Major
Acceptance criteria | ESD-safe, no-crush/puncture handling and installation; state-of-charge for shipping controlled; retention verified; line-safety controls defined.
Mfg use | Direct · Notes: Reconcile with HAL.
DO-MFG-008 — Firmware Programming & Provisioning Work Instruction
Type | DO-WI · Discipline: Mfg/Security (S3) · SKU: All (per-SKU config) · Subsystem/Component: Programming
File # | SCM-MFG-WI-008 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/EM-ENG-008
Source format | WI doc + scripts · Released format: PDF + controlled scripts · Repository: PLM/MES · Parent assembly: DO-SW-004
Related | DO-SW-001 (tool), DO-FW-020 (pkg), DO-SEC-007/008, DO-EE-012 (lock)
Input trace | DO-SW-004; SRS-SEC; SYS-032 · Risk-control trace: Identity/key provisioning + production debug-lock (reconcile with HAL Stage 10) · Verification trace: Programming/lock verification (DO-TST-012)
Critical | Yes (identity/keys + lock) · Change class: Major
Acceptance criteria | Programs signed factory image + config + unique identity/keys, sets debug lock, and refuses to pass an incompletely provisioned unit; auditable per unit.
Mfg use | Direct · Notes: Reconcile with HAL/security.
DO-MFG-009 — Serialization & UDI-Marking Procedure
Type | DO-PROC · Discipline: Mfg/Regulatory (S1) · SKU: All · Subsystem/Component: Serial/UDI
File # | SCM-MFG-PRC-009 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/EM-ENG-006
Source format | Procedure · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-CFG-010
Related | DO-CFG-010 (identity), DO-LBL-003 (serial label), DO-CLD-004 (registry)
Input trace | DO-CFG-010; UDI; ISO 13485 §7.5.8 · Risk-control trace: Identity/attribution integrity (HAL) · Verification trace: Serialization verification (DO-TST-011)
Critical | No · Change class: Major
Acceptance criteria | Each unit serialized/UDI-marked, bound to SE identity and registry; no duplicates; readable and durable marking.
Mfg use | Direct · Notes: Standard.
DO-MFG-010 — Calibration Process Work Instruction
Type | DO-WI · Discipline: Mfg/Metrology (S4) · SKU: All · Subsystem/Component: Sensors / CMP-019/020
File # | SCM-MFG-WI-010 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-SW-002
Related | DO-SW-002 (cal station), DO-FW-009 (cal schema), DO-TST-008
Input trace | DO-FW-009; SYS-003; ISO/IEC 17025 concepts · Risk-control trace: Measurement-accuracy integrity (HAL) · Verification trace: Calibration verification (DO-TST-008)
Critical | No · Change class: Major
Acceptance criteria | Reference-traceable calibration of internal + probe channels; coefficients written with metadata; cal limits enforced; results recorded.
Mfg use | Direct · Notes: Judgment; reference-standard traceability required.
DO-MFG-011 — Final-Assembly & Line-Clearance Work Instruction
Type | DO-WI · Discipline: Mfg (all) · SKU: All · Subsystem/Component: Final assembly
File # | SCM-MFG-WI-011 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | WI doc · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-MECH-000
Related | DO-MFG-004, DO-TST-013 (final), DO-LBL-007 (pack)
Input trace | DO-MECH-000; PRD-MFG · Risk-control trace: Correct-config build; mixed-SKU prevention · Verification trace: Line-clearance record
Critical | No · Change class: Major
Acceptance criteria | Final-assembly sequence, line-clearance between SKUs, and handling defined; prevents mixed-config builds.
Mfg use | Direct · Notes: Judgment.
DO-MFG-012 — Assembly Fixtures & Tooling Design
Type | DO-TOOL · Discipline: Mfg (S1/S3) · SKU: All (+ -XA) · Subsystem/Component: Fixtures
File # | SCM-MFG-TL-012 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MFG-004, DO-TST-014 (test fixtures)
Input trace | DO-MECH-000; PRD-MFG · Risk-control trace: Consistent seal/assembly (HAL) · Verification trace: Fixture qualification
Critical | No · Change class: Major
Acceptance criteria | Assembly/sealing/press fixtures give repeatable results; poka-yoke for orientation/SKU; qualified before use.
Mfg use | Direct · Notes: Judgment.
DO-MFG-013 — Injection-Mould Tooling Specification (enclosure)
Type | DO-TOOL · Discipline: Mfg (S1) · SKU: All (+ -XA cavity) · Subsystem/Component: Housing tools / CMP-001
File # | SCM-MFG-TL-013 · Rev: TBD-rev · Status: Draft · Owner: Mfg/Mech Eng · Approval: Mfg/Mech
Source format | Spec + tool CAD · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-001/003
Related | DO-MECH-001/003/010, DO-MECH-016 (-XA)
Input trace | DO-MECH-001/003/010; MER · Risk-control trace: Seal-surface CTQ moulded correctly (HAL) · Verification trace: Tool qualification + first-article
Critical | No · Change class: Major
Acceptance criteria | Mould tooling delivers housings meeting CTQ seal surfaces and material spec; -XA cavity/insert defined; process window documented.
Mfg use | Direct · Notes: Judgment; final resin grade Evidence Pending.
DO-MFG-014 — Rework / Repair & Handling Procedure
Type | DO-PROC · Discipline: Mfg (all) · SKU: All · Subsystem/Component: Rework
File # | SCM-MFG-PRC-014 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Mfg/Quality
Source format | Procedure · Released format: PDF · Repository: PLM/MES · Parent assembly: DO-MFG-001
Related | DO-MFG-002/004, DO-SVC-004
Input trace | PRD-MFG; IPC-7711/7721 concepts · Risk-control trace: No rework that compromises seal/cell/coat (HAL) · Verification trace: Post-rework re-test
Critical | No · Change class: Major
Acceptance criteria | Allowed rework limits and re-test requirements; prohibited rework on cell/seal defined; ESD/handling controls; re-test after rework mandatory.
Mfg use | Direct · Notes: Judgment.
7. Traceability notes
Parents: PRD-MFG; all discipline design outputs (DOL-300/400/500/600/700); DO-CFG-004/010; UN 38.3. Forward: pairs with DOL-1500 (test), DOL-1600 (packaging), DOL-2000 (transfer/process readiness); safety-critical process params (seal, battery handling, programming/lock) reconcile with HAL (Stage 10).
8. Release / adequacy notes
Adequate when mBOM, WIs, process params, fixtures, and tooling are Released and process-validated. Three Critical (DO-MFG-004 sealing, DO-MFG-007 battery handling, DO-MFG-008 programming/provisioning). Blocker: mBOM/AML freeze Evidence Pending (chemistry/antenna, RA-07). </content>
DOL-1500 — Inspection & Production-Test Design Outputs (DO-TST) — SCM
Module: DOL-1500 · Output prefix: DO-TST · Owner: Quality Engineer · Discipline: Quality / Inspection / Production Test (cross-subsystem, factory line) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the inspection and production-test design outputs that determine whether a built SCM unit is accepted or rejected on the line — incoming component acceptance, in-process inspection, dimensional/cosmetic final inspection, electrical/functional/calibration/connectivity/safety/leak test, the production-test software and fixtures, golden reference units, test limits with measurement-system analysis (MSA), and the acceptance/reject and retest rules. These outputs are the gate through which every unit that reaches a vaccine shipper must pass; several enforce the safety and measurement-integrity controls the whole program rests on.
2. Scope
Factory inspection and test across all subsystems (S1–S9) for all SKUs. Covers component-through-final inspection, the four safety/measurement-critical production tests (functional, calibration, protection fault-injection, seal/leak), test-station software and version control, fixtures and golden units, and the statistical acceptance framework (limits, MSA, failure codes, retest). Excludes design verification test methods (DOL-1900 supplies observability/limits; Stage 13 PVL supplies the verification protocols) and manufacturing process work instructions (DOL-1400). All SKUs; radio/band-plan test content differs per -NA/-EU/-ROW; -XA adds an external-antenna RF test step.
3. Applicable standards & references
ISO 13485:2016 §7.5 (production/process controls, monitoring & measurement); AIAG/VDA MSA (gauge R&R, measurement-system analysis); ISO/IEC 17025 concepts for calibration traceability; IPC-A-610 (PCBA workmanship/acceptance); IEC 60529 (IP / leak); IEC 62368-1 (electrical-safety production test); ANSI/ASQ Z1.4 (sampling) where sampling is used. Standard — revisions to confirm; no limit values invented (limits follow released datasheets/specs).
4. Governance & ownership
Quality Engineer owns; Manufacturing Engineer (DOL-1400) co-owns fixtures/line integration; Power/Safety leads own the protection/charge-inhibit fault-injection test (DO-TST-009) and reconcile it with the HAL (Stage 10); Metrology owns calibration traceability (DO-TST-006). The four measurement/safety-critical production tests (DO-TST-004 functional, DO-TST-005 system, DO-TST-006 calibration, DO-TST-009 protection, DO-TST-011 seal/leak) are the acceptance spine — a unit is not shippable unless each passes.
5. Output categories
Inspection specifications (DO-SPEC/DO-DOC), production-test definitions (DO-TEST), test software (DO-CODE/ DO-BIN), test fixtures & golden units (DO-TOOL), test limits / MSA / acceptance criteria (DO-SPEC/DO-DATA).
6. Design Output Register (this module)
DO-TST-001 — Incoming Inspection & Component-Acceptance Specification
Type | DO-SPEC · Discipline: Quality (all S) · SKU: All · Subsystem/Component: Incoming parts (cell/BMS/SE critical)
File # | SCM-TST-SPC-001 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/CSE
Source format | Spec doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-MFG-001
Related | DO-CFG-004 (AML), DO-EE-020 (eBOM), DO-PCB-013 (PCBA BOM)
Input trace | PRD-QUAL; PRD-MFG; SYS-096; HWR (cell/BMS/SE) · Risk-control trace: Critical-component (cell/BMS/SE/fuse) acceptance prevents counterfeit/out-of-spec parts (reconcile with HAL Stage 10) · Verification trace: Incoming-lot inspection records
Critical | No · Change class: Major
Acceptance criteria | Every AML component has defined incoming-acceptance attributes/sampling; safety-critical parts (cell, BMS/AFE, secure element, fuse) receive enhanced acceptance (C of C, lot traceability, sampled electrical check); non-conforming lots quarantined.
Mfg use | Direct (line entry gate) · Notes: Judgment; specific attributes follow released datasheets (Evidence Pending).
DO-TST-002 — In-Process Inspection Plan (SPI / AOI / X-ray)
Type | DO-SPEC · Discipline: Quality/PCBA (S3–S6) · SKU: All · Subsystem/Component: PCBA
File # | SCM-TST-SPC-002 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/Mfg
Source format | Spec + inspection program · Released format: PDF + program · Repository: QMS · Parent assembly: DO-MFG-002
Related | DO-MFG-002/003 (SMT WI + reflow), DO-PCB-008 (assembly data)
Input trace | PRD-MFG; PRD-QUAL; SYS (assembly) · Risk-control trace: Solder-defect escape on safety nets (charge-inhibit/protection) — detect before FCT · Verification trace: SPI/AOI/X-ray yield data
Critical | No · Change class: Major
Acceptance criteria | Solder-paste (SPI), post-reflow (AOI), and X-ray (BGA/QFN, cell/BMS area) inspection defined to IPC-A-610 target class; defect classes and escape limits stated.
Mfg use | Direct · Notes: Judgment.
DO-TST-003 — Dimensional & Cosmetic Final-Inspection Specification
Type | DO-SPEC · Discipline: Quality/Mech (S1/S6) · SKU: All · Subsystem/Component: Enclosure/HMI/probe
File # | SCM-TST-SPC-003 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/Mech
Source format | Spec doc · Released format: PDF · Repository: QMS · Parent assembly: DO-MECH-000
Related | DO-MECH-002/004/011 (drawings/GD&T), DO-MECH-018 (tolerance stack), DO-VER-002 (critical dims)
Input trace | PRD-MECH; PRD-QUAL; SYS (mechanical) · Risk-control trace: Seal-groove / probe-tip dimensions gate DO-SAFE-011 seal control (reconcile with HAL Stage 10) · Verification trace: Dimensional inspection records / CTQ check
Critical | No · Change class: Major
Acceptance criteria | Critical-to-quality dimensions (seal groove, HMI window fit, connector/probe apertures) and cosmetic acceptance (scratches, sink, legibility of marks) inspected per the released drawings and the CTQ schedule (DO-VER-002).
Mfg use | Direct · Notes: Judgment; CTQ list from DO-VER-002.
DO-TST-004 — PCBA In-Circuit / Functional Test (ICT/FCT) Definition
Type | DO-TEST · Discipline: Quality/Electrical (S2–S6) · SKU: All · Subsystem/Component: PCBA
File # | SCM-TST-TDF-004 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/Elec/CSE
Source format | Test definition + program · Released format: PDF + test program · Repository: PLM/QMS · Parent assembly: DO-EE-001
Related | DO-PCB-011 (test points), DO-EE-001 (nets), DO-TST-013 (fixtures/SW), DO-TST-014 (limits)
Input trace | PRD-MFG/QUAL; SYS (electrical); HWR; DO-PCB-011 · Risk-control trace: Verifies every safety net (charge-inhibit, protection, probe-presence) is populated and functional before final assembly (reconcile with HAL Stage 10) · Verification trace: ICT/FCT pass rate + coverage report
Critical | Yes (first gate that safety-relevant circuitry works) · Change class: Major
Acceptance criteria | ICT/FCT exercises power rails, MCU boot, sensor buses, radios, charge path, protection and charge-inhibit nets via test points (DO-PCB-011); coverage documented; all safety-net checks are pass/fail gates.
Mfg use | Direct · Notes: Judgment; limits per DO-TST-014.
DO-TST-005 — Final Functional & System Test Definition
Type | DO-TEST · Discipline: Quality (all S) · SKU: All (radio content per region) · Subsystem/Component: Assembled unit
File # | SCM-TST-TDF-005 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/CSE
Source format | Test definition + program · Released format: PDF + test program · Repository: PLM/QMS · Parent assembly: DO-CFG-001
Related | DO-TST-006/007/008/011, DO-FW-012 (BIST), DO-VER-010 (pass/fail)
Input trace | PRD-FUN/PERF/QUAL; SYS-001…032; FWR · Risk-control trace: End-of-line proof the capture→detect→warn→decide→prove chain works on the finished unit (reconcile with HAL Stage 10) · Verification trace: Final-test yield + defect Pareto
Critical | Yes (last full-unit acceptance gate) · Change class: Major
Acceptance criteria | Assembled unit boots, self-tests (BIST), reads all sensors within limits, exercises HMI/verdict, arms and captures a sample record, and reports readiness; failure codes recorded; unit not shippable unless all subtests pass.
Mfg use | Direct · Notes: Judgment.
DO-TST-006 — Sensor Calibration & Verification Test Definition (T / RH / probe)
Type | DO-TEST · Discipline: Quality/Metrology (S4) · SKU: All · Subsystem/Component: Sensing / CMP-019/020/021
File # | SCM-TST-TDF-006 · Rev: TBD-rev (Evidence Pending — reference chain) · Status: Draft · Owner: Quality/Metrology · Approval: Quality/CSE
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-FW-009 (cal schema)
Related | DO-MFG-010 (cal WI), DO-FW-007/009 (cal apply), DO-SW-002 (cal station SW), DO-SVC-005 (field recal)
Input trace | PRD-PERF (±0.5 °C); SYS-001/003/004; HWR-021; ISO/IEC 17025 · Risk-control trace: Measurement accuracy underwrites every disposition verdict; mis-calibration → wrong accept/reject (reconcile with HAL Stage 10) · Verification trace: Cal-point residuals vs. traceable reference (PRD-VER)
Critical | Yes (calibration integrity) · Change class: Major
Acceptance criteria | Each unit calibrated/verified against a NIST/traceable reference across the operating range; post-cal error ≤ the released accuracy budget (±0.5 °C internal/probe target); coefficients written and read-back-verified; calibration record bound to serial (DO-CFG-010).
Mfg use | Direct · Notes: Evidence Pending (reference-standard chain / cal points); traceability per ISO/IEC 17025.
DO-TST-007 — Connectivity / Radio Production-Test Definition
Type | DO-TEST · Discipline: Quality/RF (S5) · SKU: All (band-plan per -NA/-EU/-ROW; -XA ext-ant) · Subsystem/Component: Radios / CMP-024…029
File # | SCM-TST-TDF-007 · Rev: TBD-rev · Status: Draft · Owner: Quality/RF · Approval: Quality/CSE
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-EE-017
Related | DO-EE-017/018/021, DO-IF-003 (RF ICD), DO-REG-004 (radio certs)
Input trace | PRD-CONN; SYS-104…118; HWR (radios/antenna) · Risk-control trace: Alert/upload egress works on each unit; correct region band plan flashed · Verification trace: TX power/RX sensitivity/connection check per unit
Critical | No · Change class: Major
Acceptance criteria | Each radio (LTE-M/Wi-Fi/BLE/GNSS) verified for basic TX/RX/attach on the correct regional band plan; conducted/radiated go-no-go within limits; -XA units verified with external antenna.
Mfg use | Direct · Notes: Judgment; conducted vs. radiated method Evidence Pending.
DO-TST-008 — Firmware Programming & Provisioning Verification Test
Type | DO-TEST · Discipline: Quality/Security (S3) · SKU: All · Subsystem/Component: MCU/SE / CMP-015/016
File # | SCM-TST-TDF-008 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/EM-ENG-008
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-MFG-008
Related | DO-SW-001 (provisioning tool), DO-FW-019/020 (image/signing), DO-CFG-005 (compat matrix)
Input trace | PRD-MFG/SEC; SYS (identity); SRS-SEC · Risk-control trace: Correct signed image + per-device identity/keys programmed; wrong/unsigned image blocked · Verification trace: Hash/version + identity read-back
Critical | No · Change class: Major
Acceptance criteria | Programmed firmware hash/version matches the released image (DO-FW-019); per-device identity/keys provisioned and read-back-verified; compatibility matrix (DO-CFG-005) satisfied; unsigned/mismatched image rejected.
Mfg use | Direct · Notes: Judgment.
DO-TST-009 — Battery-Protection / Charge-Inhibit Fault-Injection Safety Test
Type | DO-TEST · Discipline: Quality/Safety/Power (S2/S3) · SKU: All · Subsystem/Component: BMS/charge-inhibit / CMP-008/009/010/012
File # | SCM-TST-TDF-009 · Rev: TBD-rev · Status: Draft · Owner: Quality/Safety · Approval: Safety/CSE (EM-ENG-006)
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-SAFE-002
Related | Verifies DO-EE-005 (BMS) + DO-EE-007 (charge-inhibit) + DO-EE-010 (fuse); implements HAL test for DO-SAFE-001/002 · DO-TST-014
Input trace | PRD-SAFE; SYS-094/096; HWR (protection); IEC 62368-1 · Risk-control trace: Confirms autonomous protection + cold-charge inhibit operate on each production unit — safety acceptance gate (reconcile with HAL Stage 10) · Verification trace: Fault-injection pass/fail per unit or per sample plan
Critical | Yes (production proof of the D3 charge-safety controls) · Change class: Major
Acceptance criteria | Injected/simulated over-voltage, over-current and cold-cell conditions cause protection to trip and charging to be inhibited below 0 °C with the inhibited state reported; secondary backstop coordinated; no unit ships with a defeated safety control.
Mfg use | Direct · Notes: Reconcile with HAL Stage 10; 100% vs. sampled plan Evidence Pending. Standard (IEC 62368-1).
DO-TST-010 — Electrical-Safety / ESD Production Test
Type | DO-TEST · Discipline: Quality/Electrical (S2) · SKU: All · Subsystem/Component: Charge input / evidence
File # | SCM-TST-TDF-010 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/Safety
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-EE-019
Related | DO-EE-009/019, DO-REG-005 (safety class)
Input trace | PRD-SAFE; SYS-102; IEC 62368-1 · Risk-control trace: Charge-input abuse tolerance / ESD robustness at external ports · Verification trace: Dielectric/leakage (as applicable) + ESD go-no-go
Critical | No · Change class: Major
Acceptance criteria | Applicable electrical-safety production checks (per IEC 62368-1 classification — dielectric/leakage where required for a battery-powered SELV device) and ESD-robustness sampling at external ports pass; scope justified by the safety classification (DO-REG-005).
Mfg use | Direct · Notes: Judgment; applicability depends on safety class (Evidence Pending).
DO-TST-011 — Seal-Integrity / Leak (IP) Production Test
Type | DO-TEST · Discipline: Quality/Mech (S1) · SKU: All · Subsystem/Component: Enclosure seal / CMP-002
File # | SCM-TST-TDF-011 · Rev: TBD-rev (Evidence Pending — method/limit) · Status: Draft · Owner: Quality Eng · Approval: Quality/Mech/Safety
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-MFG-004
Related | Verifies DO-MECH-005 (seal) + DO-MFG-004 (sealing WI); implements HAL test for DO-SAFE-011 · DO-MECH-008 (USB-C aperture)
Input trace | PRD-MECH/SAFE; SYS-069; IEC 60529 · Risk-control trace: Ingress/condensation protection of record-integrity electronics; a leaking unit can corrupt the evidentiary record (reconcile with HAL Stage 10) · Verification trace: Pressure-decay / leak-rate per unit
Critical | Yes (ingress protection is a safety/record-integrity control) · Change class: Major
Acceptance criteria | Each sealed unit passes a production leak test (e.g., pressure-decay) correlated to the target IP rating; leak-rate limit and correlation to IP established; failing units rejected/reworked.
Mfg use | Direct · Notes: Evidence Pending (leak method + limit + IP correlation); reconcile with HAL. Standard (IEC 60529).
DO-TST-012 — Debug-Lock / Security-Provisioning Verification Test
Type | DO-TEST · Discipline: Quality/Security (S3) · SKU: All · Subsystem/Component: SWD/JTAG lock / IFR-005
File # | SCM-TST-TDF-012 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/EM-ENG-008
Source format | Test definition + program · Released format: PDF + program · Repository: PLM/QMS · Parent assembly: DO-SEC-008
Related | Verifies DO-EE-012 (debug port) + DO-SEC-008 (debug lock) · DO-IF-011 (test-interface ICD)
Input trace | PRD-SEC; IFR-005; SRS-SEC · Risk-control trace: Locked debug prevents key/record readout on fielded units (reconcile with HAL Stage 10) · Verification trace: Post-lock read-attempt is refused
Critical | No · Change class: Major
Acceptance criteria | After production provisioning, the debug/EOL port is confirmed locked (readout-protection/lifecycle fuse set); an unauthenticated read attempt is refused; recorded per serial.
Mfg use | Direct · Notes: Standard (IEC 62443 concepts).
DO-TST-013 — Production-Test Software, Fixtures & Golden-Unit Definition
Type | DO-TOOL · Discipline: Quality/Test-Eng (all S) · SKU: All (per-region RF fixture delta) · Subsystem/Component: Test stations
File # | SCM-TST-TOL-013 · Rev: TBD-rev · Status: Draft · Owner: Quality/Test Eng · Approval: Quality/Mfg
Source format | Fixture CAD + test-SW repo · Released format: Fixture pack + versioned SW · Repository: PLM/VCS · Parent assembly: DO-MFG-012
Related | DO-TST-004/005/006/007/009/011, DO-SW-002 (cal station SW)
Input trace | PRD-MFG/QUAL; SYS (test); DO-PCB-011 (test points) · Risk-control trace: Golden units bound the drift of safety/measurement test stations · Verification trace: Fixture correlation + golden-unit drift checks
Critical | No · Change class: Major
Acceptance criteria | ICT/FCT/cal/RF/leak fixtures and versioned test software defined; golden reference units established for each critical test; station-to-station correlation and periodic golden-unit re-verification specified; test-software under version control (DO-SW-008).
Mfg use | Direct · Notes: Judgment.
DO-TST-014 — Test Limits, MSA / Gauge R&R & Acceptance-Reject Criteria
Type | DO-SPEC · Discipline: Quality (all S) · SKU: All · Subsystem/Component: All tests
File # | SCM-TST-SPC-014 · Rev: TBD-rev (Evidence Pending — limits/MSA) · Status: Draft · Owner: Quality Eng · Approval: Quality/CSE
Source format | Spec + data tables · Released format: PDF + limits dataset · Repository: QMS · Parent assembly: DO-VER-010
Related | DO-VER-003/010 (limits/pass-fail), all DO-TST-###
Input trace | PRD-QUAL/PERF; SYS (acceptance); AIAG/VDA MSA · Risk-control trace: Guard-banded limits prevent marginal safety/measurement units passing (reconcile with HAL Stage 10) · Verification trace: MSA studies + capability (Cpk) reports
Critical | No · Change class: Major
Acceptance criteria | Every production test has documented pass/fail limits (guard-banded to the spec), an MSA/gauge-R&R establishing measurement capability, failure codes, and retest/rework rules; safety/measurement-critical tests have zero-defect acceptance.
Mfg use | Direct · Notes: Evidence Pending (numeric limits, MSA results); MSA per AIAG/VDA.
7. Traceability notes
Parents: PRD-MFG/QUAL/PERF/SAFE; the SYS functional/power/safety sets; HWR/FWR sensing/protection; DO-MFG (process) and DO-VER (characteristics/limits). Forward: DO-TST- are consumed by Stage 13 (PVL — production tests reference the verification observability and limits), Stage 15 (design transfer — production-test readiness, DO-XFER-005), and the Manufacturing stage. The four/five safety- and measurement-critical tests (DO-TST-004/005/006/009/011) are the acceptance spine and reconcile with HAL Stage 10 for the protection, charge-inhibit, and seal controls (DO-SAFE-001/002/011).
8. Release / adequacy notes
Adequate when incoming/in-process/final inspection, the functional/system/calibration/protection/leak tests, the test software+fixtures+golden units, and the limits/MSA framework are Released and correlated. Five Critical: DO-TST-004 (FCT), DO-TST-005 (final system), DO-TST-006 (calibration), DO-TST-009 (protection/ charge-inhibit — reconcile with HAL), DO-TST-011 (seal/leak — reconcile with HAL). Blockers: calibration reference chain (DO-TST-006), leak method/limit (DO-TST-011), and numeric limits/MSA (DO-TST-014) are Evidence Pending; protection/seal acceptance reconcile with HAL Stage 10. </content> </invoke>
DOL-1600 — Packaging, Labeling & Logistics Design Outputs (DO-LBL) — SCM
Module: DOL-1600 · Output prefix: DO-LBL · Owner: Manufacturing Engineer (with Quality/Regulatory EM-ENG-006 on regulatory/UDI content) · Discipline: Packaging / Labeling / Logistics (device + carton, all markets) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the labeling, packaging, and logistics design outputs required to mark, protect, ship, store, and instruct the SCM across its multi-market distribution (D4). This covers the on-device labels (nameplate/rating, regulatory markings, serial/identity, UDI), the lithium-battery transport marking, the primary and protective packaging artwork/drawings, shipping/storage/transport conditions, the Instructions for Use (IFU) and quick-start content with country-specific translations, and the packaging/labeling inspection criteria.
2. Scope
Device-level and package-level labeling/marking, IFU/quick-start content, protective and shipping packaging, palletization, storage/transport conditions, and packaging inspection — for all SKUs and target markets. Regional differences (markings, translations, band/regulatory config) are per-SKU label/packaging deltas, not new hardware. Excludes the serialization/marking process (DO-MFG-009) and the regulatory standards/markings analysis (DOL-1800 supplies the required-markings determination; this module renders it as artwork). All SKUs; -NA/-EU/-ROW differ in regulatory markings/translations; -XA adds external-antenna handling notes.
3. Applicable standards & references
Device marking / UDI expectations (per applicable regulator, D4 — WHO PQS / EU / US; exact UDI regime Evidence Pending); UN 38.3 / IATA DGR & UN 3481 lithium-battery marking; ISTA / ASTM D4169 transport-package performance; IEC 60068 (storage/transport environment); RoHS/WEEE marking; ISO 15223-1 symbols (as applicable). Standard — revisions to confirm; no regulator-specific label text asserted as fact until DOL-1800 resolves the applicable regime.
4. Governance & ownership
Manufacturing Engineer owns packaging and label artwork production; Quality/Regulatory (EM-ENG-006) owns the regulatory-marking, UDI, and lithium-transport content and its correctness (traces to DOL-1800); Human Factors reviews IFU/quick-start legibility (DO-HF-009). Regulatory markings (DO-LBL-002), UDI content (DO-LBL-004), and the lithium-transport marking (DO-LBL-005) are compliance-critical and must be reconciled with DOL-1800 and the transport regime before release.
5. Output categories
Label/artwork (DO-LBL), packaging drawings/specs (DO-DRW/DO-SPEC), manuals/IFU/quick-start (DO-MAN), packaging inspection criteria (DO-SPEC).
6. Design Output Register (this module)
DO-LBL-001 — Device Nameplate / Rating Label Artwork
Type | DO-LBL · Discipline: Labeling (S1) · SKU: All (per-region variants) · Subsystem/Component: Enclosure exterior
File # | SCM-LBL-ART-001 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality/Regulatory
Source format | Vector artwork · Released format: PDF + print-ready · Repository: PLM · Parent assembly: DO-MECH-010
Related | DO-LBL-002/003, DO-REG-003 (required markings)
Input trace | PRD-REG; PRD-MFG; SYS-177…190 · Risk-control trace: Correct model/rating identification supports safe use/service · Verification trace: Artwork review + legibility/adhesion test
Critical | No · Change class: Major
Acceptance criteria | Nameplate carries model/SKU, manufacturer, electrical rating, and required identifiers; material/adhesive survives −30↔+60 °C, condensation, and cleaning between trips (D1) without illegibility or lift.
Mfg use | Direct · Notes: Judgment; label material for cold/condensation Evidence Pending.
DO-LBL-002 — Regulatory-Markings Label Artwork (CE / FCC / IC / WEEE / etc.)
Type | DO-LBL · Discipline: Labeling/Regulatory (S1) · SKU: All (per-region) · Subsystem/Component: Enclosure exterior
File # | SCM-LBL-ART-002 · Rev: TBD-rev (Evidence Pending — markings set) · Status: Draft · Owner: Quality/Regulatory · Approval: EM-ENG-006/CSE
Source format | Vector artwork · Released format: PDF + print-ready · Repository: PLM · Parent assembly: DO-MECH-010
Related | Renders DO-REG-003 (required markings) + DO-REG-004 (radio certs) · DO-LBL-001
Input trace | D4; PRD-REG; SYS-177…190; FCC Part 15 / CE RED · Risk-control trace: Correct conformity markings per market — compliance-critical (reconcile with DOL-1800) · Verification trace: Regulatory-marking review vs. certificates
Critical | Yes (compliance marking) · Change class: Major
Acceptance criteria | The regional markings (CE/UKCA/FCC ID/IC/WEEE/RoHS and any WHO PQS/market mark) match the determinations in DOL-1800 and the actual certificates; placement per DO-REG-003; no unsupported mark shown.
Mfg use | Direct · Notes: Evidence Pending (exact markings set follows DOL-1800/certification). Standard.
DO-LBL-003 — Serial-Number / Device-Identity Label (barcode / QR)
Type | DO-LBL · Discipline: Labeling (S1) · SKU: All · Subsystem/Component: Enclosure exterior
File # | SCM-LBL-ART-003 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality
Source format | Vector artwork + symbology spec · Released format: PDF + print-ready · Repository: PLM · Parent assembly: DO-CFG-010
Related | Renders DO-CFG-010 (serial/identity convention) · DO-MFG-009 (marking process), DO-LBL-004
Input trace | PRD-MFG; SYS (identity); DO-CFG-010 · Risk-control trace: Identity binds the evidentiary record to a physical unit (record attribution) · Verification trace: Barcode/QR grade + read verification
Critical | No · Change class: Major
Acceptance criteria | Human-readable serial plus machine-readable symbology encoding the device-identity per DO-CFG-010; scan-verified to a minimum grade; unique per unit; durable across reuse/cleaning.
Mfg use | Direct · Notes: Judgment.
DO-LBL-004 — UDI Label Content & Format Specification
Type | DO-LBL · Discipline: Labeling/Regulatory (S1) · SKU: All · Subsystem/Component: Enclosure + package
File # | SCM-LBL-SPC-004 · Rev: TBD-rev (Evidence Pending — UDI regime) · Status: Draft · Owner: Quality/Regulatory · Approval: EM-ENG-006
Source format | Spec + artwork · Released format: PDF + print-ready · Repository: PLM · Parent assembly: DO-CFG-010
Related | DO-CFG-010 (identity), DO-REG-010 (country config), DO-LBL-003
Input trace | D4; PRD-REG; SYS (identity) · Risk-control trace: UDI supports traceability/recall of a distribution-critical device (reconcile with DOL-1800) · Verification trace: UDI content/format review vs. applicable regime
Critical | Yes (traceability/recall) · Change class: Major
Acceptance criteria | UDI-DI/PI content and carrier format defined per the applicable regime (if the device is classified such that UDI applies — confirm in DOL-1800); device and package UDI consistent; parseable and database-registerable.
Mfg use | Direct · Notes: Evidence Pending — applicability/regime follows US/EU classification (DOL-1800). Standard.
DO-LBL-005 — Lithium-Battery Transport Marking & Handling Label (UN 38.3 / IATA)
Type | DO-LBL · Discipline: Labeling/Logistics (S2) · SKU: All · Subsystem/Component: Package + device
File # | SCM-LBL-ART-005 · Rev: TBD-rev · Status: Draft · Owner: Mfg/Logistics · Approval: Quality/Regulatory
Source format | Vector artwork · Released format: PDF + print-ready · Repository: PLM · Parent assembly: DO-EE-004 (pack)
Related | Renders DO-REG-007 (transport compliance) · DO-EE-004 (pack), DO-LBL-008 (shipping)
Input trace | D4; PRD-REG; UN 38.3 / IATA DGR / UN 3481 · Risk-control trace: Lithium-battery transport safety marking — compliance-critical (reconcile with DOL-1800/transport regime) · Verification trace: DGR marking review + UN 38.3 evidence
Critical | Yes (battery-transport safety/compliance) · Change class: Major
Acceptance criteria | Lithium-battery handling mark and required transport documentation/markings applied to device/package per UN 38.3 test outcome and IATA DGR (UN 3481, batteries contained in equipment); consistent with the pack (DO-EE-004) and shipping config.
Mfg use | Direct · Notes: Depends on final chemistry/pack (DO-EE-003/004, Evidence Pending). Standard.
DO-LBL-006 — Primary Packaging Artwork & Drawing
Type | DO-DRW · Discipline: Packaging (S1) · SKU: All (per-region artwork) · Subsystem/Component: Retail/primary pack
File # | SCM-LBL-DRW-006 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality/Mfg
Source format | Structural CAD + artwork · Released format: PDF + die-line · Repository: PLM · Parent assembly: DO-CFG-009
Related | DO-CFG-009 (packaging config), DO-LBL-007/010/011
Input trace | PRD-MFG; DO-CFG-009 · Risk-control trace: Package carries required content (IFU/quick-start/markings) to the user · Verification trace: Artwork/structure review + fit check
Critical | No · Change class: Major
Acceptance criteria | Primary pack die-line and artwork accommodate device + probe + accessories + IFU/quick-start; carries package labels/markings; consistent with the packaging configuration matrix (DO-CFG-009).
Mfg use | Direct · Notes: Judgment.
DO-LBL-007 — Protective / Cushioning Packaging Design
Type | DO-SPEC · Discipline: Packaging (S1) · SKU: All · Subsystem/Component: Cushioning/inner pack
File # | SCM-LBL-SPC-007 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality/Mfg
Source format | Spec + CAD · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-LBL-006
Related | DO-MECH-017 (drop/shock), DO-LBL-008 (shipping)
Input trace | PRD-MFG; SYS (transport); ISTA/ASTM D4169; IEC 60068 · Risk-control trace: Protects the sealed unit + Li pack from transport shock/vibration (feeds DO-SAFE-010) · Verification trace: Transport/drop package test (ISTA)
Critical | No · Change class: Major
Acceptance criteria | Cushioning protects the device (and battery) through the defined transport hazard profile with margin; retains the probe/lead safely; passes the selected ISTA/ASTM D4169 sequence.
Mfg use | Direct · Notes: Judgment; transport profile Evidence Pending.
DO-LBL-008 — Shipping / Transport Carton & Palletization Specification
Type | DO-SPEC · Discipline: Logistics (S1) · SKU: All · Subsystem/Component: Master carton/pallet
File # | SCM-LBL-SPC-008 · Rev: TBD-rev · Status: Draft · Owner: Mfg/Logistics · Approval: Quality/Mfg
Source format | Spec + drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-LBL-006
Related | DO-LBL-005 (Li marking), DO-LBL-007, DO-LBL-009
Input trace | PRD-MFG; IATA DGR; ISTA · Risk-control trace: Correct DG shipping configuration for lithium batteries · Verification trace: Palletization + stacking + transport test
Critical | No · Change class: Major
Acceptance criteria | Master carton count, stacking, and palletization defined; DG/lithium shipping configuration and outer marking correct; survives stacking/transport per the selected standard.
Mfg use | Direct · Notes: Judgment.
DO-LBL-009 — Storage & Transport Conditions / Shelf-Life Controls
Type | DO-SPEC · Discipline: Logistics/Quality (S2) · SKU: All · Subsystem/Component: Device + battery
File # | SCM-LBL-SPC-009 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality/CSE
Source format | Spec doc · Released format: PDF · Repository: QMS · Parent assembly: DO-CFG-008
Related | DO-EE-003 (cell), DO-SVC-004 (PM), DO-CFG-008 (consumables)
Input trace | PRD-REL; SYS (environmental); IEC 60068; cell datasheet · Risk-control trace: Storage state-of-charge/temperature limits protect cell health and readiness · Verification trace: Storage/shelf-life study
Critical | No · Change class: Major
Acceptance criteria | Storage/transport temperature-humidity limits and battery storage state-of-charge / max-storage-duration controls defined per the cell datasheet; on-shelf readiness (charge/self-discharge) bounded; stated on packaging.
Mfg use | Direct · Notes: Judgment; SoC/duration limits follow chemistry (Evidence Pending).
DO-LBL-010 — Instructions for Use (IFU) Content
Type | DO-MAN · Discipline: Labeling/HF (S6/S9) · SKU: All (translations) · Subsystem/Component: Documentation
File # | SCM-LBL-MAN-010 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng (HF review) · Approval: Quality/Regulatory
Source format | Structured content · Released format: PDF + print · Repository: PLM · Parent assembly: DO-LBL-006
Related | DO-HF-009 (quick-start), DO-SVC-002 (service), DO-HF-006 (arming flow)
Input trace | PRD-UX; PRD-REG; SYS (use) · Risk-control trace: Use-error mitigation — correct arming/placement/interpretation of the verdict (reconcile with HAL Stage 10) · Verification trace: IFU review + usability check
Critical | No · Change class: Major
Acceptance criteria | IFU covers intended use/limitations, safe handling of the Li device, placement in the shipper, arming, reading the accept/reject/quarantine verdict, charging (incl. cold-charge behaviour), cleaning between trips, and disposal; consistent with HMI content (DO-HF-).
Mfg use | Indirect · Notes: Judgment; regulated IFU elements follow DOL-1800.
DO-LBL-011 — Quick-Start Guide & Country-Specific Labels / Translations
Type | DO-MAN · Discipline: Labeling/HF (S6/S9) · SKU: All (per-region) · Subsystem/Component: Documentation
File # | SCM-LBL-MAN-011 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: Quality/Regulatory
Source format | Structured content · Released format: PDF + print · Repository: PLM · Parent assembly: DO-LBL-006
Related | DO-HF-009, DO-LBL-010, DO-REG-010 (country config)
Input trace | PRD-UX; PRD-REG; D4 · Risk-control trace: Correct-language safety/arming instruction at point of use · Verification trace: Translation QA + quick-start usability
Critical | No · Change class: Major
Acceptance criteria | Concise quick-start (arm → place → read verdict) plus country-specific labels/translations for the target markets; safety and battery statements translated; consistent with the IFU.
Mfg use | Indirect · Notes: Judgment; language set per D4 markets.
DO-LBL-012 — Packaging & Labeling Inspection Criteria
Type | DO-SPEC · Discipline: Quality (S1) · SKU: All · Subsystem/Component: Package + labels
File # | SCM-LBL-SPC-012 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: Quality
Source format | Spec doc · Released format: PDF · Repository: QMS · Parent assembly: DO-LBL-006
Related | DO-TST-003 (cosmetic), all DO-LBL-###
Input trace | PRD-QUAL; PRD-REG · Risk-control trace: Wrong/illegible/missing regulatory or Li-transport label prevented from shipping · Verification trace: Label-content/legibility/placement inspection
Critical | No · Change class: Major
Acceptance criteria | Inspection verifies correct label content (esp. regulatory/UDI/Li-transport), legibility, placement, adhesion, correct language variant, and complete packaging contents (IFU/quick-start/accessories) before ship; mismatches rejected.
Mfg use | Direct · Notes: Judgment.
7. Traceability notes
Parents: PRD-MFG/REG/UX/REL; SYS regulatory/environmental sets; DO-CFG-008/009/010 (consumables/packaging/ identity); DO-MECH-010 (surface/marking). Forward: DO-LBL-002/004/005 render the DOL-1800 regulatory determinations and must reconcile with them; IFU/quick-start (DO-LBL-010/011) pair with the human-factors content (DO-HF-009) and service (DO-SVC-002); readiness is assessed at Stage 15 (DO-XFER-007). Country variants trace to D4.
8. Release / adequacy notes
Adequate when device/regulatory/serial/UDI labels, the Li-transport marking, packaging drawings/protective design, storage/transport controls, IFU/quick-start, and inspection criteria are Released and translation- complete. Three Critical: DO-LBL-002 (regulatory markings), DO-LBL-004 (UDI), DO-LBL-005 (lithium transport) — all compliance-critical and gated by DOL-1800. Blockers: markings set, UDI regime, and final pack (chemistry) are Evidence Pending. </content>
DOL-1700 — Installation, Service & Maintenance Design Outputs (DO-SVC) — SCM
Module: DOL-1700 · Output prefix: DO-SVC · Owner: Service Engineer · Discipline: Installation / Service / Maintenance / End-of-life (field + depot, reusable multi-trip device) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the service, maintenance, and lifecycle design outputs required to commission, operate, maintain, recalibrate, repair, clean-between-trips, upgrade, back up, decommission, and securely dispose of the SCM across its reusable, multi-trip life (D1). Because the SCM is returned and re-deployed on many trips, these outputs are first-class lifecycle requirements, not afterthoughts: cleaning/decontamination between trips and field recalibration keep a returned unit safe, hygienic, and measurement-honest for the next vaccine shipment.
2. Scope
Depot/field installation & commissioning (provisioning), the service manual and troubleshooting/diagnostics, preventive maintenance, field recalibration, FRU replacement and repair limits, inter-trip cleaning/ decontamination, firmware/config upgrade with backup/restore, decommissioning with secure data-erase, and secure disposal of the lithium battery/device. All SKUs. Excludes factory production test (DOL-1500) and manufacturing WIs (DOL-1400); references FRU definitions from DO-CFG-007 and cleaning materials from DO-CFG-008.
3. Applicable standards & references
IEC 62368-1 (service safety — battery product); UN 38.3 / IATA & local battery-recycling / WEEE (end-of-life transport & disposal); ISO/IEC 17025 concepts (field-calibration traceability); NIST SP 800-88 (media sanitization / secure erase); cleaning-agent material-compatibility per the enclosure/seal materials (DO-MECH-010). Standard — revisions to confirm; cleaning agents and cycle limits Evidence Pending until material compatibility is tested.
4. Governance & ownership
Service Engineer owns; Quality/Metrology owns field-recalibration traceability (DO-SVC-005); Product Security (EM-ENG-008) owns secure-erase/decommissioning (DO-SVC-009); Safety reviews battery-handling and disposal (DO-SVC-010, reconcile with HAL Stage 10). Inter-trip cleaning/decontamination (DO-SVC-007) and field recalibration (DO-SVC-005) are the two reusability-critical outputs — they preserve safety/hygiene and measurement integrity across the multi-trip life (D1).
5. Output categories
Installation/service/maintenance manuals & procedures (DO-MAN/DO-PROC), troubleshooting/diagnostics content (DO-DOC), FRU/repair procedures (DO-PROC), decommissioning/disposal procedures (DO-PROC).
6. Design Output Register (this module)
DO-SVC-001 — Installation / Commissioning & Provisioning Procedure
Type | DO-PROC · Discipline: Service (S8/S9) · SKU: All · Subsystem/Component: Device + cloud + mobile
File # | SCM-SVC-PRC-001 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/Service
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-HF-006 (arming flow)
Related | DO-APP-003 (BLE provisioning), DO-CLD-004 (registry), DO-HF-006, DO-LBL-011 (quick-start)
Input trace | PRD-SVC; PRD-UX; SYS (provisioning) · Risk-control trace: Correct provisioning/arming prevents an unmonitored or mis-configured trip (reconcile with HAL Stage 10) · Verification trace: Commissioning walkthrough / usability
Critical | No · Change class: Major
Acceptance criteria | Step-by-step commissioning (register to fleet, provision identity, configure interval/thresholds, arm, verify readiness) via mobile/BLE and cloud; readiness confirmed before deployment; consistent with DO-HF-006.
Mfg use | Indirect · Notes: Judgment.
DO-SVC-002 — Service Manual
Type | DO-MAN · Discipline: Service (all S) · SKU: All · Subsystem/Component: Device
File # | SCM-SVC-MAN-002 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/Quality
Source format | Structured content · Released format: PDF · Repository: PLM · Parent assembly: DO-SVC-001
Related | DO-SVC-003…010, DO-CFG-007 (FRU), DO-SW-003 (service tool)
Input trace | PRD-SVC; SYS (serviceability) · Risk-control trace: Safe service of a Li-battery device; prevents damage to record-integrity electronics · Verification trace: Service-manual review + technician trial
Critical | No · Change class: Major
Acceptance criteria | Complete service manual covering safety precautions (battery), specifications, service modes (DO-SW-003), FRU procedures, maintenance, calibration, cleaning, upgrade, and disposal; cross-referenced and current.
Mfg use | Indirect · Notes: Judgment.
DO-SVC-003 — Troubleshooting / Diagnostics Trees
Type | DO-DOC · Discipline: Service (all S) · SKU: All · Subsystem/Component: Device + cloud
File # | SCM-SVC-DOC-003 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/Service
Source format | Structured content · Released format: PDF · Repository: PLM · Parent assembly: DO-SVC-002
Related | DO-FW-012 (diagnostics/BIST), DO-SW-003 (service tool), DO-CLD-005 (fleet health)
Input trace | PRD-SVC; SRS-DIAG; SYS (diagnostics) · Risk-control trace: Correct diagnosis avoids returning an unsafe/unready unit to service · Verification trace: Troubleshooting-tree coverage vs. failure codes
Critical | No · Change class: Major
Acceptance criteria | Symptom→cause→action trees keyed to device/cloud diagnostic codes (DO-FW-012 / DO-CLD-005) cover the common field faults (no charge, charge-inhibited, no connectivity, probe fault, low battery, unready); escalation to depot/FRU defined.
Mfg use | Indirect · Notes: Judgment.
DO-SVC-004 — Preventive-Maintenance Schedule
Type | DO-PROC · Discipline: Service (S1/S2) · SKU: All · Subsystem/Component: Seal/battery/probe
File # | SCM-SVC-PRC-004 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/Quality
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-SVC-002
Related | DO-SVC-005 (recal), DO-MECH-005 (seal), DO-EE-003 (cell life)
Input trace | PRD-REL; PRD-SVC; SYS (reliability) · Risk-control trace: Seal/battery/probe wear over reuse life managed before it causes a fault · Verification trace: PM-interval justification vs. reliability data
Critical | No · Change class: Major
Acceptance criteria | PM intervals for seal inspection/replacement, battery-health check, probe check, and recalibration defined against reuse-cycle life; triggers (cycle count / calendar / SoH) stated.
Mfg use | Indirect · Notes: Judgment; intervals Evidence Pending (reliability/cycle data).
DO-SVC-005 — Field Recalibration / Calibration-Verification Procedure
Type | DO-PROC · Discipline: Service/Metrology (S4) · SKU: All · Subsystem/Component: Sensing / CMP-019/020/021
File # | SCM-SVC-PRC-005 · Rev: TBD-rev (Evidence Pending — interval) · Status: Draft · Owner: Service/Metrology · Approval: Quality/CSE
Source format | Procedure doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-TST-006
Related | DO-TST-006 (factory cal), DO-FW-009 (cal schema), DO-SW-002 (cal station)
Input trace | PRD-PERF/SVC; SYS-001/003/004; ISO/IEC 17025 · Risk-control trace: Drift over reuse life → wrong verdict; periodic recal/verification preserves measurement integrity (reconcile with HAL Stage 10) · Verification trace: Recal residuals vs. traceable reference
Critical | Yes (measurement integrity across reuse life) · Change class: Major
Acceptance criteria | Field/depot recalibration-or-verification against a traceable reference at a defined interval restores/confirms accuracy to the released budget (±0.5 °C target); records bound to serial; out-of-tolerance units quarantined from service.
Mfg use | Indirect · Notes: Evidence Pending (recal interval, drift data); traceability per ISO/IEC 17025.
DO-SVC-006 — FRU Replacement Procedures & Repair Limits
Type | DO-PROC · Discipline: Service (S1/S2) · SKU: All · Subsystem/Component: FRUs / CMP per DO-CFG-007
File # | SCM-SVC-PRC-006 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/Quality
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-CFG-007
Related | Realizes DO-CFG-007 (FRU definition) · DO-MECH-005 (seal), DO-EE-004 (battery), DO-SVC-005 (post-repair recal)
Input trace | PRD-SVC; SYS (serviceability) · Risk-control trace: Repair must not compromise seal/battery safety or record integrity; re-seal + re-test after repair (reconcile with HAL Stage 10) · Verification trace: Post-repair re-test (leak/cal/protection)
Critical | No · Change class: Major
Acceptance criteria | Approved FRU set (e.g., seal, battery pack, probe) with replacement procedures, torque/re-seal steps, and mandatory post-repair re-test (leak DO-TST-011, cal DO-SVC-005, protection where applicable); non-FRU repairs and end-of-repair-life limits stated.
Mfg use | Indirect · Notes: Judgment; FRU list per DO-CFG-007.
DO-SVC-007 — Inter-Trip Cleaning / Decontamination Procedure (reusable, D1)
Type | DO-PROC · Discipline: Service/Mech (S1) · SKU: All · Subsystem/Component: Enclosure/seal/probe / CMP-002
File # | SCM-SVC-PRC-007 · Rev: TBD-rev (Evidence Pending — agents/cycles) · Status: Draft · Owner: Service Eng · Approval: Quality/CSE (Safety review)
Source format | Procedure doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-MECH-010
Related | DO-CFG-008 (cleaning materials), DO-MECH-005/010 (seal/material), DO-MECH-006 (probe)
Input trace | D1; PRD-SVC; PRD-MECH; SYS (reuse/cleaning) · Risk-control trace: Between-trip cleaning must not degrade the seal/marking or ingress-protect electronics, and must leave the unit hygienic for the next vaccine shipment (reconcile with HAL Stage 10) · Verification trace: Cleaning-cycle material-compatibility + post-clean seal/leak check
Critical | Yes (reusability/hygiene + material-compatibility safety) · Change class: Major
Acceptance criteria | Approved cleaning/decontamination agents and method (DO-CFG-008) clean the returned unit without degrading enclosure/seal/probe/labels over the reuse-cycle count; post-clean the seal still meets its IP/leak criterion; number of cleaning cycles bounded.
Mfg use | Indirect · Notes: Evidence Pending — agents, cycle count, and material compatibility by test; reconcile with HAL. (D1 reusability leg.)
DO-SVC-008 — Firmware/Config Upgrade & Backup/Restore Procedure
Type | DO-PROC · Discipline: Service (S3/S8) · SKU: All · Subsystem/Component: FW/config
File # | SCM-SVC-PRC-008 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: CSE/EM-ENG-008
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-CLD-010 (OTA service)
Related | DO-FW-015 (OTA agent), DO-CLD-010 (campaigns), DO-SW-005 (recovery image), DO-SAFE-014 (OTA never-brick)
Input trace | PRD-SVC; SYS-032/114; SRS-OTA · Risk-control trace: Upgrade never bricks/loses trip data; unmonitored window disclosed (reconcile with HAL Stage 10) · Verification trace: Upgrade/rollback + backup/restore test
Critical | No · Change class: Major
Acceptance criteria | Depot/field upgrade via OTA or recovery image with backup/restore of configuration and unsent trip data; rollback on failure; version-compatibility (DO-CFG-005) respected; consistent with the OTA never-brick control (DO-SAFE-014).
Mfg use | Indirect · Notes: Judgment.
DO-SVC-009 — Decommissioning & Secure Data-Erase Procedure
Type | DO-PROC · Discipline: Service/Security (S3/S7) · SKU: All · Subsystem/Component: NVM/identity / CMP-041
File # | SCM-SVC-PRC-009 · Rev: TBD-rev · Status: Draft · Owner: Service Eng (Security) · Approval: EM-ENG-008/CSE
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-SEC-012 (revocation)
Related | DO-SEC-006 (secure storage), DO-SEC-012 (key revocation), DO-CLD-004 (registry de-register)
Input trace | PRD-SEC/SVC; SRS-SEC; NIST SP 800-88 · Risk-control trace: Retired unit cannot leak records/keys or masquerade on the fleet · Verification trace: Post-erase read-back + registry de-registration
Critical | No · Change class: Major
Acceptance criteria | Decommissioning erases records/credentials to a defined sanitization level (NIST SP 800-88), revokes device identity/keys (DO-SEC-012), and de-registers from the fleet/registry; erasure verified.
Mfg use | Indirect · Notes: Standard (NIST SP 800-88).
DO-SVC-010 — Secure Disposal / Battery End-of-Life & Recycling Procedure
Type | DO-PROC · Discipline: Service/Safety (S2) · SKU: All · Subsystem/Component: Battery/device / CMP-007
File # | SCM-SVC-PRC-010 · Rev: TBD-rev · Status: Draft · Owner: Service Eng · Approval: Safety/Quality
Source format | Procedure doc · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-004 (pack)
Related | DO-EE-004 (pack), DO-LBL-005 (Li transport), DO-SVC-009 (data-erase first)
Input trace | PRD-SVC/REG; UN 38.3 / IATA; WEEE / local battery-recycling · Risk-control trace: Safe end-of-life handling of a lithium battery; no data left on disposed unit (reconcile with HAL Stage 10) · Verification trace: Disposal-procedure review vs. regs
Critical | No · Change class: Major
Acceptance criteria | End-of-life procedure safely discharges/removes the Li pack, routes it to compliant recycling (WEEE/local), transports damaged/defective cells per DGR, and requires prior secure data-erase (DO-SVC-009).
Mfg use | Indirect · Notes: Standard; reconcile disposal/transport with DOL-1800.
7. Traceability notes
Parents: PRD-SVC/REL/SEC/UX/REG; SYS serviceability/diagnostics/security sets; DO-CFG-007 (FRU), DO-CFG-008 (cleaning). Forward: service outputs realize the reusable-lifecycle needs (D1) and pair with DO-FW-012/ DO-CLD-005 (diagnostics), DO-SEC-006/012 (secure storage/revocation), DO-SAFE-014 (OTA), and DOL-1800 (disposal/transport). Cleaning (DO-SVC-007) and recalibration (DO-SVC-005) reconcile with HAL Stage 10; readiness/training assessed at Stage 15 (DO-XFER-007).
8. Release / adequacy notes
Adequate when installation/commissioning, the service manual + troubleshooting, PM, recalibration, FRU/repair, cleaning, upgrade/backup-restore, decommissioning, and disposal are Released and technician-validated. Two Critical: DO-SVC-005 (field recalibration — measurement integrity) and DO-SVC-007 (inter-trip cleaning — reusability/hygiene + material compatibility). Blockers: recal interval, cleaning agents/cycle count, and PM intervals are Evidence Pending; cleaning/recal/disposal reconcile with HAL Stage 10. </content>
DOL-1800 — Regulatory & Compliance Design Outputs (DO-REG) — SCM
Module: DOL-1800 · Output prefix: DO-REG · Owner: Quality/Regulatory Engineer (EM-ENG-006) · Discipline: Regulatory / Compliance / Privacy (multi-market, D4) · Status: Draft baseline (ORCH-013) — research-driven; several determinations Evidence Pending.
1. Purpose
Identify and govern the regulatory and compliance design outputs that make the SCM lawfully marketable and transportable across its target regimes (D4): the standards-applicability matrix, per-market regulatory configuration definitions, required markings, essential-performance/safety-classification references, environmental/material declarations, cybersecurity documentation outputs, privacy-control implementation and data-residency configuration, country-specific configurations and submission drawings, and the register of compliance-critical software versions. This module determines the requirements that DOL-1600 renders as artwork and that DOL-1500 tests.
2. Scope
All target regimes per D4: WHO PQS temperature-monitoring device performance; EU GDP (Good Distribution Practice) + GDPR data-residency/privacy; US distribution/data-integrity expectations (device classification/pathway to confirm); air/ground transport — UN 38.3 / IATA DGR lithium batteries; radio — FCC Part 15 / CE RED; EMC/safety — IEC 62368-1, CISPR/IEC 61000; environmental — IEC 60529 (IP), IEC 60068, RoHS/REACH/WEEE. All SKUs; regional differences are configuration/marking/ privacy deltas. Excludes the verification evidence itself (Stage 13/14) — this module defines the compliance requirements, configurations, and references.
3. Applicable standards & references
WHO PQS (cold-chain monitor/logger performance category — exact category Evidence Pending); EU GDP + GDPR; US FDA distribution/data-integrity + 21 CFR Part 11 / ALCOA+ (device classification/pathway Evidence Pending); UN 38.3, IATA DGR; FCC Part 15, CE RED; IEC 62368-1; CISPR 32 / IEC 61000; IEC 60529; IEC 60068; RoHS/REACH/WEEE; IEC 62443 concepts / applicable device-cybersecurity guidance. Standard — cited by designation, revisions to confirm; no clause numbers, category codes, or classifications asserted as fact.
4. Governance & ownership
Quality/Regulatory (EM-ENG-006) owns; Product Security (EM-ENG-008) owns the cybersecurity documentation and privacy-control outputs (DO-REG-008/009); RF/Regulatory owns radio-cert configuration (DO-REG-004); Power/Safety owns the battery-transport package (DO-REG-007). The standards-applicability matrix (DO-REG-001), essential-performance/safety classification (DO-REG-005), battery-transport compliance (DO-REG-007), and the compliance-critical software register (DO-REG-011) are the compliance spine and gate DOL-1600 markings and DOL-1500 tests. Research-driven: identify → knowledge-base → ingest missing → cite; gaps recorded, never invented.
5. Output categories
Standards/compliance matrices & determinations (DO-DOC/DO-SPEC), regulatory configuration definitions (DO-CFG), declarations/certificates references (DO-CERT), submission drawings (DO-DRW), compliance-critical version register (DO-DOC).
6. Design Output Register (this module)
DO-REG-001 — Standards-Applicability Matrix
Type | DO-DOC · Discipline: Regulatory (all S) · SKU: All (per-market columns) · Subsystem/Component: Product
File # | SCM-REG-DOC-001 · Rev: TBD-rev (Evidence Pending — WHO PQS/US class) · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/Quality
Source format | Matrix doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-REG-002…012, DO-CFG-006 (regional config)
Input trace | D4; PRD-REG; SYS-177…190; SRS-REG · Risk-control trace: Determines which safety/EMC/radio/transport standards bind — the compliance backbone (reconcile with HAL Stage 10 for safety-standard tie-in) · Verification trace: Regulatory review + certification plan
Critical | Yes (compliance backbone) · Change class: Major
Acceptance criteria | Every applicable standard/regulation per target market (WHO PQS, EU GDP/GDPR, US, UN 38.3/IATA, FCC/CE, IEC 62368-1/CISPR/61000/60529/60068, RoHS/WEEE) mapped to product scope with applicability rationale; open determinations (WHO PQS category, US classification) flagged Evidence Pending.
Mfg use | Indirect · Notes: Evidence Pending (WHO PQS category / US class). Standard/Cited.
DO-REG-002 — Regulatory-Configuration Definitions (per-market)
Type | DO-CFG · Discipline: Regulatory (S3/S5/S8) · SKU: All (per-region) · Subsystem/Component: Config
File # | SCM-REG-CFG-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-006 · Approval: CSE
Source format | Config spec · Released format: PDF + config · Repository: PLM · Parent assembly: DO-CFG-006
Related | DO-CFG-006 (regional variants), DO-EE-017 (band plan), DO-REG-009 (data residency)
Input trace | D4; PRD-REG; SYS-177…190 · Risk-control trace: Region-permitted radio/data configuration enforced by config, not chance · Verification trace: Config review per market
Critical | No · Change class: Major
Acceptance criteria | Each market's regulatory configuration (radio band plan, data-residency region, marking set, privacy defaults) defined and bound to the SKU/config (DO-CFG-006); no market ships a non-permitted configuration.
Mfg use | Direct (config) · Notes: Judgment.
DO-REG-003 — Required-Markings & Marking-Placement Specification
Type | DO-SPEC · Discipline: Regulatory (S1) · SKU: All (per-region) · Subsystem/Component: Enclosure/label
File # | SCM-REG-SPC-003 · Rev: TBD-rev (Evidence Pending — markings set) · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/Quality
Source format | Spec doc · Released format: PDF · Repository: PLM · Parent assembly: DO-REG-001
Related | Feeds DO-LBL-002 (marking artwork) · DO-REG-004 (radio certs)
Input trace | D4; PRD-REG; FCC/CE · Risk-control trace: Correct legally-required markings per market · Verification trace: Marking review vs. certificates
Critical | No · Change class: Major
Acceptance criteria | The full set of required markings per market (CE/UKCA, FCC ID / IC, WEEE, RoHS, any WHO PQS/market mark, Li-handling) with placement/size rules; feeds DO-LBL-002; only supported marks specified.
Mfg use | Indirect · Notes: Evidence Pending (markings set follows certification).
DO-REG-004 — Radio / EMC Certification Configuration & Evidence Index
Type | DO-CERT · Discipline: Regulatory/RF (S5) · SKU: All (per-region) · Subsystem/Component: Radios / CMP-024…029
File # | SCM-REG-CRT-004 · Rev: TBD-rev (Evidence Pending — certs) · Status: Draft · Owner: EM-ENG-006/RF · Approval: CSE
Source format | Index + config · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-EE-017
Related | DO-EE-017/018/021 (radios/antenna), DO-TST-007 (radio test), DO-IF-003 (RF ICD)
Input trace | D4; PRD-CONN/REG; SYS-104…118; FCC Part 15 / CE RED · Risk-control trace: Region-permitted-band operation; lawful RF emission · Verification trace: Radio conformance test + certificate index
Critical | No · Change class: Major
Acceptance criteria | Per-region radio/EMC certification configuration (module vs. full-device path), test-house plan, and an index binding certificates/FCC-ID/IC/CE to the released band plan and antenna (integral/-XA); no market operates uncertified.
Mfg use | Indirect · Notes: Evidence Pending (certs/test houses, concept §7). Standard.
DO-REG-005 — Essential-Performance & Safety-Classification References
Type | DO-DOC · Discipline: Regulatory/Safety (all S) · SKU: All · Subsystem/Component: Product
File # | SCM-REG-DOC-005 · Rev: TBD-rev (Evidence Pending — classification) · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/Safety
Source format | Reference doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-REG-001
Related | Ties to DOL-1200 (safety controls) + HAL Stage 10 · DO-REG-010 (submission), WHO PQS perf
Input trace | D4; PRD-REG/SAFE; SYS-safety set; WHO PQS; IEC 62368-1 · Risk-control trace: Defines the device's essential performance (accurate monitoring, honest verdict, alert delivery) and safety classification — the compliance-safety tie-in (reconcile with HAL Stage 10) · Verification trace: Essential-performance V&V (Stage 13/14)
Critical | Yes (essential-performance/safety classification) · Change class: Major
Acceptance criteria | Essential performance (measurement accuracy, excursion detection, verdict integrity, alert delivery, record integrity) and the safety classification/pathway (US FDA class; WHO PQS category) are stated with references; each essential-performance claim traces to a requirement and to V&V; safety controls trace to DOL-1200/HAL.
Mfg use | Indirect · Notes: Evidence Pending (US classification / WHO PQS category, concept §7); reconcile with HAL. Standard/Cited.
DO-REG-006 — Environmental / Material Declarations (RoHS / REACH / WEEE)
Type | DO-CERT · Discipline: Regulatory/Mech (S1) · SKU: All · Subsystem/Component: Materials/BOM
File # | SCM-REG-CRT-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-006 · Approval: Quality
Source format | Declaration set · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-CFG-004 (AML)
Related | DO-MECH-010 (materials), DO-EE-020/DO-PCB-013 (BOMs), DO-SVC-010 (WEEE disposal)
Input trace | PRD-REG; RoHS/REACH/WEEE · Risk-control trace: Restricted-substance compliance across the BOM · Verification trace: Supplier declaration roll-up vs. BOM
Critical | No · Change class: Major
Acceptance criteria | RoHS/REACH conformity and WEEE registration supported by supplier declarations rolled up against the released BOM/AML; restricted substances within limits; WEEE marking justified.
Mfg use | Indirect · Notes: Judgment/Cited.
DO-REG-007 — Battery-Transport Compliance Documentation (UN 38.3 / IATA)
Type | DO-CERT · Discipline: Regulatory/Power (S2) · SKU: All · Subsystem/Component: Battery/pack / CMP-006/007
File # | SCM-REG-CRT-007 · Rev: TBD-rev (Evidence Pending — UN 38.3 result) · Status: Draft · Owner: EM-ENG-006/Power · Approval: CSE/Safety
Source format | Test-summary + docs · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-EE-004
Related | Feeds DO-LBL-005 (transport marking) + DO-LBL-008 (shipping) · DO-EE-003/004 (cell/pack)
Input trace | D4; PRD-REG; UN 38.3 / IATA DGR / UN 3481 · Risk-control trace: Lithium-battery transport safety — device shippable by air/ground only with valid UN 38.3 (reconcile with HAL Stage 10) · Verification trace: UN 38.3 test summary + DGR classification
Critical | Yes (transport safety/compliance) · Change class: Major
Acceptance criteria | UN 38.3 test summary for the final cell/pack and the IATA DGR classification/documentation (UN 3481, batteries in equipment) available; feeds transport marking/shipping (DO-LBL-005/008).
Mfg use | Indirect · Notes: Evidence Pending (chemistry/pack DO-EE-003/004). Standard.
DO-REG-008 — Cybersecurity Documentation Outputs (SBOM / threat-model refs)
Type | DO-DOC · Discipline: Regulatory/Security (S3/S8) · SKU: All · Subsystem/Component: FW/cloud/mobile
File # | SCM-REG-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: CSE/Regulatory
Source format | Doc set · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-SEC-001
Related | DO-SEC-001 (arch), DO-SEC-011 (SBOM policy), DO-FW-018/DO-SW-006/DO-CLD-015/DO-APP-010 (SBOMs)
Input trace | PRD-SEC/REG; SRS-SEC; IEC 62443 concepts · Risk-control trace: Cybersecurity documentation for market submissions (device-security expectations) · Verification trace: Security documentation review
Critical | No · Change class: Major
Acceptance criteria | The regulator-facing cybersecurity package (threat-model references, security architecture summary, consolidated SBOM, vulnerability-management and update mechanism) assembled from DOL-1100 outputs to the applicable device-security expectations.
Mfg use | Indirect · Notes: Standard (IEC 62443 concepts / applicable guidance).
DO-REG-009 — Privacy-Control Implementation & Data-Residency Configuration
Type | DO-CFG · Discipline: Regulatory/Security (S8) · SKU: All (per-region) · Subsystem/Component: Cloud / CMP (S8)
File # | SCM-REG-CFG-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008/Regulatory · Approval: CSE
Source format | Config + design doc · Released format: PDF + config · Repository: PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-001/013 (arch/RBAC), DO-CLD-003 (schemas), DO-REG-002 (per-market config)
Input trace | D4; PRD-SEC/REG; EU GDP/GDPR; 21 CFR Part 11 / ALCOA+ · Risk-control trace: Data-residency + privacy controls per market (EU) · Verification trace: Privacy/data-residency review
Critical | No · Change class: Major
Acceptance criteria | Privacy controls (data minimization, retention, access, audit) and data-residency configuration implemented per market (e.g., EU data residency); mapped to GDPR/GDP and ALCOA+ record-integrity expectations.
Mfg use | Indirect · Notes: Standard/Cited; primarily device-telemetry (limited personal data) — scope confirmed with legal.
DO-REG-010 — Country-Specific Configuration & Submission Drawings
Type | DO-DRW · Discipline: Regulatory (all S) · SKU: All (per-market) · Subsystem/Component: Product
File # | SCM-REG-DRW-010 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/Quality
Source format | Drawing/doc set · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-REG-001
Related | DO-CFG-006, DO-LBL-004 (UDI), DO-REG-002
Input trace | D4; PRD-REG · Risk-control trace: Complete, correct submission package per market · Verification trace: Submission-readiness review
Critical | No · Change class: Major
Acceptance criteria | Per-market submission artefacts (labeling drawings, device description, configuration statements, required drawings) assembled per each regulator's format; consistent with DO-REG-001 determinations.
Mfg use | Indirect · Notes: Evidence Pending (regulator formats follow classification).
DO-REG-011 — Compliance-Critical Software Version Register
Type | DO-DOC · Discipline: Regulatory/SW (S3/S8/S9) · SKU: All · Subsystem/Component: FW/cloud/mobile
File # | SCM-REG-DOC-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/EM-ENG-008
Source format | Register doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-REG-001
Related | DO-FW-019 (image/hash), DO-CLD-016 (cloud versions), DO-APP-009 (app versions), DO-CFG-005 (compat)
Input trace | PRD-REG; SYS (SW config); 21 CFR Part 11 / ALCOA+ · Risk-control trace: The exact FW/cloud/mobile versions whose behaviour is certified/validated — a wrong version invalidates compliance/record integrity (reconcile with HAL Stage 10) · Verification trace: Version register vs. released binaries/hashes
Critical | Yes (compliance-critical versions) · Change class: Major
Acceptance criteria | The specific certified/validated firmware image+hash (DO-FW-019), cloud service versions (DO-CLD-016), and app versions (DO-APP-009) recorded with their compatibility (DO-CFG-005); any change triggers a compliance impact assessment.
Mfg use | Indirect · Notes: Standard (ALCOA+/Part 11 concepts).
DO-REG-012 — WHO PQS / EU GDP / US Distribution Compliance Mapping (D4)
Type | DO-DOC · Discipline: Regulatory (all S) · SKU: All (per-market) · Subsystem/Component: Product
File # | SCM-REG-DOC-012 · Rev: TBD-rev (Evidence Pending — WHO PQS/US) · Status: Draft · Owner: EM-ENG-006 · Approval: CSE/Quality
Source format | Mapping doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-REG-001
Related | DO-REG-001/005, DO-VER- (evidence), PRD-VER
Input trace | D4; PRD-REG; WHO PQS / EU GDP / US · Risk-control trace: Distribution-regime performance/data-integrity compliance for a public-health device · Verification trace: WHO PQS performance V&V + GDP/data-integrity audit
Critical | No · Change class: Major
Acceptance criteria | Each distribution-regime requirement (WHO PQS logger performance, EU GDP data-integrity/GDP, US distribution/data-integrity) mapped to an implementing output and its verification/validation evidence; open category/classification items flagged.
Mfg use | Indirect · Notes: Evidence Pending (WHO PQS category / US pathway, concept §7). Standard/Cited.
7. Traceability notes
Parents: PRD-REG (and safety/security/connectivity/environmental cross-refs); SYS-177…190 + SYS-security/ safety; SRS-REG/SEC; D4. Forward: DO-REG determinations drive DOL-1600 markings/UDI/Li-transport (DO-LBL-002/004/005), DOL-1500 tests (radio/safety/leak), and are the compliance backbone for Stage 13/14 V&V. Essential-performance/safety (DO-REG-005) and compliance-critical versions (DO-REG-011) reconcile with HAL Stage 10. Research-driven items (WHO PQS category, US classification, UN 38.3 result, certs) are Evidence Pending, recorded as gaps.
8. Release / adequacy notes
Adequate when the standards matrix, per-market regulatory config, markings, essential-performance/safety classification, environmental/battery-transport declarations, cybersecurity/privacy outputs, submission package, and compliance-critical version register are Released and evidenced. Four Critical: DO-REG-001 (standards matrix), DO-REG-005 (essential-performance/safety class), DO-REG-007 (battery transport), DO-REG-011 (compliance-critical versions). Blockers (RA-10): WHO PQS category and US FDA classification are Evidence Pending, plus UN 38.3 result and radio certs; safety/essential-performance items reconcile with HAL Stage 10. </content>
DOL-1900 — Verification-Enabling Design Outputs (DO-VER) — SCM
Module: DOL-1900 · Output prefix: DO-VER · Owner: Quality Engineer · Discipline: Verification enablement / Test observability / Acceptance criteria (cross-subsystem) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the design outputs that make the SCM verifiable — the measurable output characteristics, critical dimensions and tolerances, performance limits, test points, diagnostic observability and logging, test interfaces/modes, simulation/reference configurations, verification fixtures, and the objective pass/fail criteria. These outputs are the bridge from design to Stage 13 (Product Verification Library, PVL): they expose every acceptance-critical characteristic and define how each is measured and judged, so that a PV-<CODE>-NNN protocol can be written and executed against a real observable with a defined limit.
2. Scope
Verification-enabling characteristics and observability for all subsystems (S1–S9) and all SKUs. Covers the CTQ/CTF register, dimension/tolerance schedule, performance-limit definitions, test points and diagnostic observability, verification logging schema, engineering/production/verification test modes and interfaces, simulation and reference configurations, verification fixtures/reference standards, the cold-chamber/ RF-in-container verification-enabling configuration, and the objective pass/fail criteria with method mapping. Excludes the verification protocols and evidence themselves (Stage 13 PVL) and production-line test definitions (DOL-1500), though the two share limits and observability.
3. Applicable standards & references
ISO/IEC/IEEE verification practice; AIAG/VDA MSA (measurement capability behind every limit); ISO/IEC 17025 (reference-standard traceability); IEC 60068 (environmental verification exposure); the product's own released performance requirements (accuracy, latency, endurance) as the source of limits. Standard — revisions to confirm; numeric limits carried from released requirements, none invented; RF-in-container and cold-chamber configs Evidence Pending (AD-1 / D3).
4. Governance & ownership
Quality Engineer owns; each discipline supplies its characteristics/observability (Electrical, FW, Mechanical, Cloud); Product Verification (Stage 13, ORCH-014) is the primary consumer. The CTQ/CTF register (DO-VER-001), performance-limit/acceptance definition (DO-VER-003), test-mode/interface spec (DO-VER-006), and objective pass/fail criteria (DO-VER-010) are the verification spine — no acceptance-critical characteristic may lack an observable, a limit, and a method. Safety-relevant observability reconciles with the HAL (Stage 10) and DOL-1200.
5. Output categories
Characteristic/limit registers (DO-SPEC/DO-DATA), test-point & observability specs (DO-SPEC), verification logging schema (DO-DATA), test-mode/interface specs (DO-SPEC/DO-ICD), simulation/reference configs (DO-MODEL/DO-CFG), verification fixtures (DO-TOOL), pass/fail criteria (DO-SPEC).
6. Design Output Register (this module)
DO-VER-001 — Measurable Output-Characteristic Register (CTQ / CTF list)
Type | DO-SPEC · Discipline: Quality (all S) · SKU: All · Subsystem/Component: Product
File # | SCM-VER-SPC-001 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: CSE/Quality
Source format | Register/spec · Released format: PDF + dataset · Repository: PLM/QMS · Parent assembly: SCM-DOL-MASTER-001
Related | DO-VER-002/003/010, DO-TST-014 (limits), DO-MECH-018 (tolerance stack)
Input trace | PRD-PERF/QUAL/VER; SYS (measurable reqs); SRS-VER · Risk-control trace: Ensures every safety/essential-performance characteristic is measurable (reconcile with HAL Stage 10) · Verification trace: CTQ/CTF coverage vs. requirement set
Critical | Yes (defines what must be verifiable) · Change class: Major
Acceptance criteria | Every acceptance-critical characteristic (measurement accuracy, excursion-detection correctness, alert latency, endurance, record integrity, seal/IP, RF link, calibration) enumerated as a measurable CTQ/CTF with its source requirement; safety-critical ones flagged; no orphan requirement.
Mfg use | Indirect · Notes: Judgment; some target values Evidence Pending (carried from PRD).
DO-VER-002 — Critical-Dimension & Tolerance Schedule
Type | DO-SPEC · Discipline: Quality/Mech (S1) · SKU: All · Subsystem/Component: Enclosure/seal/probe/HMI
File # | SCM-VER-SPC-002 · Rev: TBD-rev · Status: Draft · Owner: Quality/Mech · Approval: CSE/Mech
Source format | Schedule/spec · Released format: PDF + dataset · Repository: PLM · Parent assembly: DO-MECH-002
Related | DO-MECH-002/004/011/018, DO-TST-003 (dimensional inspection)
Input trace | PRD-MECH; SYS (mechanical); GD&T on released drawings · Risk-control trace: Seal-groove/probe-tip dimensions gate ingress + measurement (feeds DO-SAFE-011) · Verification trace: CMM/gauge measurement vs. schedule
Critical | No · Change class: Major
Acceptance criteria | Critical dimensions and tolerances (seal groove/gland, probe tip, HMI window, connector apertures, board fit) extracted from the released drawings/GD&T with measurement method; consistent with the tolerance-stack analysis (DO-MECH-018).
Mfg use | Direct (feeds DO-TST-003) · Notes: Judgment.
DO-VER-003 — Performance-Limit & Acceptance-Criteria Definition
Type | DO-SPEC · Discipline: Quality (all S) · SKU: All · Subsystem/Component: Product
File # | SCM-VER-SPC-003 · Rev: TBD-rev (Evidence Pending — targets) · Status: Draft · Owner: Quality Eng · Approval: CSE/Quality
Source format | Spec + data tables · Released format: PDF + dataset · Repository: PLM/QMS · Parent assembly: DO-VER-001
Related | DO-VER-001/010, DO-TST-014 (production limits), PRD-PERF
Input trace | PRD-PERF/REL/VER; SYS-033…050 (performance); D3 · Risk-control trace: Limits on accuracy/latency/endurance/detection bound safe operation (reconcile with HAL Stage 10) · Verification trace: Measured value vs. limit (Stage 13)
Critical | Yes (defines pass thresholds) · Change class: Major
Acceptance criteria | Each CTQ/CTF has a numeric limit with tolerance and rationale (e.g., ±0.5 °C accuracy, alert-latency budget, 90-day cold endurance, ±5 %RH, IP target); limits trace to a released requirement; Evidence-Pending targets flagged and closed by V&V.
Mfg use | Indirect · Notes: Evidence Pending (several numeric targets from PRD, closed by test). Standard/Judgment.
DO-VER-004 — Test-Point & Diagnostic-Observability Specification
Type | DO-SPEC · Discipline: Quality/Electrical/FW (S2–S7) · SKU: All · Subsystem/Component: PCBA + FW
File # | SCM-VER-SPC-004 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: CSE/Elec
Source format | Spec + map · Released format: PDF · Repository: PLM · Parent assembly: DO-PCB-011
Related | DO-PCB-011 (test points), DO-FW-012 (diagnostics/BIST), DO-IF-011 (debug ICD)
Input trace | PRD-VER; SYS (diagnostics); SRS-DIAG · Risk-control trace: Observability of safety-relevant nets/states (charge-inhibit, protection, probe-presence, safe-state) for verification (reconcile with HAL Stage 10) · Verification trace: Observability coverage review
Critical | No · Change class: Major
Acceptance criteria | Physical test points (DO-PCB-011) and firmware-exposed diagnostics/BIST (DO-FW-012) make every acceptance-critical net, sensor value, and internal state observable during verification; safety states are observable; access is governed (DO-IF-011).
Mfg use | Direct (also serves production test) · Notes: Judgment.
DO-VER-005 — Verification Logging / Data-Recording Schema
Type | DO-DATA · Discipline: Quality/FW/Cloud (S3/S7/S8) · SKU: All · Subsystem/Component: Log/telemetry
File # | SCM-VER-DAT-005 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: CSE/Quality
Source format | Schema doc · Released format: PDF + schema · Repository: PLM · Parent assembly: DO-CLD-003
Related | DO-CLD-003 (record schema), DO-FW-010 (log), DO-CLD-005 (fleet health)
Input trace | PRD-VER; SYS (data/records); ALCOA+ · Risk-control trace: Verification evidence is attributable/complete/traceable (ALCOA+) · Verification trace: Log-completeness / attributability check
Critical | No · Change class: Major
Acceptance criteria | A verification data-recording schema captures stimulus, measured value, limit, verdict, unit serial, timestamp (trusted time), and configuration/version for each verification observation; attributable and reconstructable (ALCOA+).
Mfg use | Indirect · Notes: Judgment; extends record schema (DO-CLD-003).
DO-VER-006 — Test-Interface & Test-Mode Specification (engineering / production / verification)
Type | DO-SPEC · Discipline: Quality/FW/Security (S3) · SKU: All · Subsystem/Component: FW/test modes
File # | SCM-VER-SPC-006 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: CSE/EM-ENG-008
Source format | Spec + ICD ref · Released format: PDF · Repository: PLM · Parent assembly: DO-IF-011
Related | DO-FW-012 (diagnostics), DO-IF-011 (test ICD), DO-EE-012 (debug port), DO-SEC-008 (lock)
Input trace | PRD-VER; SRS-DIAG/SEC; IFR-005 · Risk-control trace: Test modes must not create a fielded bypass of safety/security controls (reconcile with HAL Stage 10) · Verification trace: Test-mode access-control + coverage check
Critical | Yes (test modes touch the safety/security surface) · Change class: Major
Acceptance criteria | Defined engineering/production/verification modes (e.g., forced excursion, forced cold-cell, sensor-injection, radio-loopback) expose acceptance-critical behaviour; modes are authenticated/governed and disabled or non-bypassing on fielded units (consistent with DO-SEC-008 / DO-IF-011); no mode can silently defeat a safety control in the field.
Mfg use | Direct · Notes: Judgment; reconcile with HAL Stage 10 for control-independence.
DO-VER-007 — Simulation / Reference-Configuration Definitions
Type | DO-MODEL · Discipline: Quality/Systems (all S) · SKU: All · Subsystem/Component: Models/refs
File # | SCM-VER-MDL-007 · Rev: TBD-rev · Status: Draft · Owner: Quality/Systems · Approval: CSE
Source format | Models + config · Released format: PDF + model/config · Repository: PLM · Parent assembly: DO-SYS-005
Related | DO-FW-006 (detection), DO-SYS-007 (energy), DO-MECH-017 (structural)
Input trace | PRD-VER; SYS (functional/energy) · Risk-control trace: Reference excursion/energy profiles exercise the detection/energy controls repeatably · Verification trace: Model-vs-measurement correlation
Critical | No · Change class: Major
Acceptance criteria | Reference profiles/configurations (excursion temperature profiles, shock profiles, energy-duty scenarios, known-good reference units) defined so detection/verdict/energy behaviour can be verified against a controlled stimulus; correlation to physical test stated.
Mfg use | Indirect · Notes: Judgment.
DO-VER-008 — Verification Fixture & Reference-Standard Requirements
Type | DO-TOOL · Discipline: Quality/Metrology (all S) · SKU: All · Subsystem/Component: Fixtures/standards
File # | SCM-VER-TOL-008 · Rev: TBD-rev · Status: Draft · Owner: Quality/Metrology · Approval: CSE/Quality
Source format | Fixture spec + CAD · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-VER-006
Related | DO-TST-013 (production fixtures), DO-VER-009 (chamber/RF), ISO/IEC 17025
Input trace | PRD-VER; ISO/IEC 17025 · Risk-control trace: Traceable references bound the truth of every safety/measurement verification · Verification trace: Fixture correlation + calibration traceability
Critical | No · Change class: Major
Acceptance criteria | Verification fixtures and traceable reference standards (temperature/RH references, RF equipment, leak/IP rigs, load banks) specified with required accuracy/traceability (ISO/IEC 17025); correlated to production fixtures where shared (DO-TST-013).
Mfg use | Indirect · Notes: Judgment.
DO-VER-009 — Cold-Chamber / RF-in-Container Verification-Enabling Configuration
Type | DO-CFG · Discipline: Quality/RF/Power (S2/S5) · SKU: All · Subsystem/Component: Chamber/RF setup
File # | SCM-VER-CFG-009 · Rev: TBD-rev (Evidence Pending — AD-1/D3) · Status: Draft · Owner: Quality/RF · Approval: CSE
Source format | Config + setup doc · Released format: PDF · Repository: PLM · Parent assembly: DO-VER-008
Related | DO-EE-018/021 (antenna), DO-EE-003 (cell), PRD-VER-004/005
Input trace | D3; AD-1; PRD-VER-004/005; SYS-069/104/109/117; IEC 60068 · Risk-control trace: Enables the two top-risk verifications (90-day cold endurance; RF-in-container link) — feeds DO-SAFE-012 evidence · Verification trace: PRD-VER-004 (cold endurance) + PRD-VER-005 (RF-in-container)
Critical | No · Change class: Major
Acceptance criteria | The chamber profile (−30…+60 °C cycling, IEC 60068), the in-container RF test rig (representative insulated shipper), and the instrumentation defined so the D3 endurance and AD-1 link-budget risks can be objectively closed.
Mfg use | Indirect · Notes: Evidence Pending (representative shipper / test-house setup, AD-1/D3).
DO-VER-010 — Objective Pass/Fail Criteria & Verification-Method Mapping
Type | DO-SPEC · Discipline: Quality (all S) · SKU: All · Subsystem/Component: Product
File # | SCM-VER-SPC-010 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: CSE/Quality
Source format | Spec + mapping table · Released format: PDF + dataset · Repository: PLM/QMS · Parent assembly: DO-VER-001
Related | DO-VER-001/003, DO-TST-014 (production acceptance), PVL (Stage 13)
Input trace | PRD-VER; SYS (verification methods); SRS-VER · Risk-control trace: Every acceptance-critical characteristic has an unambiguous pass/fail and a method — no subjective safety verdict (reconcile with HAL Stage 10) · Verification trace: Method-to-characteristic coverage (feeds PV-###)
Critical | Yes (objective acceptance basis) · Change class: Major
Acceptance criteria | Each CTQ/CTF maps to a verification method (Test/Analysis/Inspection/Demonstration), an observable, a limit (DO-VER-003), and an objective pass/fail rule; the mapping seeds the Stage-13 PV-<CODE>-NNN protocols; safety-critical criteria are zero-ambiguity.
Mfg use | Indirect · Notes: Judgment; the direct bridge to PVL (Stage 13).
7. Traceability notes
Parents: PRD-VER/PERF/QUAL; SYS performance/functional/data sets; SRS-VER/DIAG; the released drawings/limits. Forward: DO-VER- seed the Product Verification Library (Stage 13, ORCH-014) — DO-VER-001/003/010 supply the characteristics/limits/pass-fail for each PV-<CODE>-NNN; DO-VER-004/006 supply observability/test modes; DO-VER-009 enables the D3/AD-1 headline verifications (PRD-VER-004/005). Safety observability and test-mode independence reconcile with HAL Stage 10 and DOL-1200. Shared limits/fixtures also feed DOL-1500.
8. Release / adequacy notes
Adequate when the CTQ/CTF register, dimension/tolerance schedule, performance limits, observability/test points, logging schema, test modes/interfaces, simulation/reference configs, fixtures, and objective pass/fail mapping are Released and cover the full requirement set. Four Critical: DO-VER-001 (CTQ/CTF register), DO-VER-003 (performance limits), DO-VER-006 (test modes/interfaces), DO-VER-010 (pass/fail mapping). Blockers: several numeric targets (DO-VER-003) and the cold-chamber/RF-in-container config (DO-VER-009) are Evidence Pending; safety observability/test-mode independence reconcile with HAL Stage 10. </content>
DOL-200 — System Architecture Outputs (DO-SYS) — SCM
Module: DOL-200 · Output prefix: DO-SYS · Owner: Chief Systems Engineer (EM-ENG-001) · Discipline: Systems Engineering · Status: Draft baseline (ORCH-013).
1. Purpose
Identify the controlled system-architecture design outputs that fix the SCM's logical/physical architecture, functional decomposition, subsystem allocation, interface-control overview, data/control/power/communication/ fault/deployment architectures, and the operational state/mode model — the architectural design outputs the discipline modules build against.
2. Scope
The Stage-08 architecture set, promoted to controlled design outputs: FUN-001…040, CMP-001…055, INT-001…035, ADR-001…012, plus the state model, power/comms/fault/deployment architectures. All SKUs.
3. Applicable standards & references
INCOSE SE Handbook; ISO/IEC/IEEE 15288; ISO 13485 §7.3.4; Stage-08 architecture files. Standard/Judgment.
4. Governance & ownership
EM-ENG-001 owns all DO-SYS outputs. Interface and safety-relevant architecture reconcile with DOL-1000 and the HAL (Stage 10) respectively.
5. Output categories
Architecture description documents (DO-DOC), decomposition models (DO-MODEL/DO-DOC), ICD overview (DO-ICD), state/mode model (DO-DOC/DO-MODEL), decision records (DO-DOC).
6. Design Output Register (this module)
DO-SYS-001 — Logical & Physical System Architecture Description
Type | DO-DOC · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: S1–S9 (whole)
File # | SCM-SYS-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | DO-SYS-002/003/004, DO-CFG-001
Input trace | PRD-ARCH-###; SYS-001…190 (system baseline); ADR-001…012 · Risk-control trace: Offline-first / split-trust (HAL) · Verification trace: Architecture review; integration test
Critical | Yes (offline-first record integrity architecture) · Change class: Major
Acceptance criteria | Logical layers and physical three-tier (device/cloud/mobile) with S1–S9 decomposition fully described; each layer's driver→mechanism mapping stated; consistent with the component list.
Mfg use | Indirect · Notes: Judgment (based on Stage-08 logical-physical-architecture.md).
DO-SYS-002 — Functional Decomposition (FUN-001…040)
Type | DO-DOC · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: All functions
File # | SCM-SYS-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-SYS-003, DO-FW-002
Input trace | PRD-FUN; SYS-001…032; FUN-001…040 · Risk-control trace: Detection/verdict functions (HAL) · Verification trace: Function-level verification
Critical | No · Change class: Major
Acceptance criteria | All 40 functions decomposed from F1–F10, each allocated to subsystem(s) and traced to SYS/FWR/SRS; no unallocated function.
Mfg use | Indirect · Notes: Judgment.
DO-SYS-003 — Subsystem Allocation Matrix (FUN→S1–S9)
Type | DO-DATA · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: All
File # | SCM-SYS-DAT-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | XLSX · Released format: CSV + PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-SYS-002, DO-CFG-001
Input trace | SYS decomposition-allocation-matrix; CMP-001…055 · Risk-control trace: — · Verification trace: Allocation completeness review
Critical | No · Change class: Major
Acceptance criteria | Every function allocated to ≥1 realizing component; every component realizes ≥1 function; no gaps/orphans.
Mfg use | Indirect · Notes: Judgment.
DO-SYS-004 — Interface-Control Overview / ICD Register (INT-001…035)
Type | DO-ICD · Discipline: Systems/Interface (SYS) · SKU: All · Subsystem/Component: All interfaces I-01…I-12
File # | SCM-SYS-ICD-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown table · Released format: PDF + CSV · Repository: PLM · Parent assembly: DO-SYS-001
Related | All DO-IF-### (DOL-1000), DO-EE, DO-FW, DO-CLD
Input trace | IFR-001…035; SYS-051…068; INT-001…035 · Risk-control trace: Charge-inhibit/alert/record interfaces (HAL) · Verification trace: Interface conformance test
Critical | Yes (governs no-undocumented-interface rule; carries critical control/alert/record paths) · Change class: Major
Acceptance criteria | All 35 interfaces typed (PWR/DATA/RF/THRM/MECH), endpoints/direction/contract fixed, mapped to I-01…I-12; no undocumented cross-subsystem/cross-tier interface.
Mfg use | Indirect · Notes: Standard/Judgment; RF interfaces (INT-026/027/028/031) Evidence Pending (AD-1).
DO-SYS-005 — Data-Flow & Control-Flow Architecture
Type | DO-DOC · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: S3/S5/S7/S8
File # | SCM-SYS-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-FW-006 (detection), DO-CLD-002 (ingest), DO-IF-005 (device↔cloud)
Input trace | SYS-011/013/020; ALCOA+; FUN-008…023 · Risk-control trace: Alert-before-bulk; store-and-forward never-lose-record (HAL) · Verification trace: Data-integrity/flow test
Critical | Yes (excursion-alert loop + record lineage) · Change class: Major
Acceptance criteria | Sensor→cloud, offline, and BLE paths and the excursion-alert control loop fully specified with ALCOA+ lineage; alert path takes precedence over bulk backlog.
Mfg use | Indirect · Notes: Standard (ALCOA+).
DO-SYS-006 — Operational State & Mode Model
Type | DO-DOC · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: S3 (governor)
File # | SCM-SYS-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown + Mermaid (stateDiagram) · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-FW-003 (state machines), DO-FW-004 (power-state manager)
Input trace | SYS-089; FWR (power-state governor); state-mode-model.md · Risk-control trace: Safe-state/Fault-Safe mode (HAL) · Verification trace: State-transition test
Critical | Yes (defines Fault-Safe and charge-inhibit states) · Change class: Major
Acceptance criteria | All modes (Shipping/Sleep/Sampling-Logging/Alerting/Charging/OTA/Fault-Safe/Provisioning-Arming) defined with transitions, guards, and invariants; safe-state reachable from any mode without record/armed-state corruption.
Mfg use | Direct (Shipping mode default) · Notes: Judgment; reconcile with HAL.
DO-SYS-007 — Power & Energy Architecture Description
Type | DO-DOC · Discipline: Systems/Power (S2/S3) · SKU: All · Subsystem/Component: Power tree
File # | SCM-SYS-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Power lead · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-EE-002 (power tree), DO-EE-006 (power budget), DO-FW-004 (governor)
Input trace | D3; SYS-086…103; power-energy-architecture.md · Risk-control trace: Energy-exhaustion (loss of monitoring) — HAL · Verification trace: 90-day cold endurance test (PRD-VER-004)
Critical | Yes (D3 endurance is top program risk) · Change class: Major
Acceptance criteria | Energy-governor intent, power tree, energy-allocation policy, and charge architecture described; 90-day/−30 °C budget presented as illustrative, Evidence Pending, validate-by-test.
Mfg use | Indirect · Notes: Evidence Pending — no capacity/chemistry asserted (D3).
DO-SYS-008 — Communication Architecture Description
Type | DO-DOC · Discipline: Systems/RF (S5) · SKU: All · Subsystem/Component: Radios/bearers
File # | SCM-SYS-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-IF-005/006/007, DO-EE-016 (radios), DO-FW-011 (comms agent)
Input trace | D2; SYS-104…118; communication-architecture.md · Risk-control trace: Delayed/missing alert (HAL) · Verification trace: RF-in-container test (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | LTE-M-primary store-and-forward with Wi-Fi/BLE offload, alert-path priority, PSM/eDRX, and roaming described; in-container link Evidence Pending.
Mfg use | Indirect · Notes: Standard/Evidence Pending (AD-1).
DO-SYS-009 — Fault-Management & Diagnostics Architecture
Type | DO-DOC · Discipline: Systems/Reliability (SYS) · SKU: All · Subsystem/Component: S3 (fault handling)
File # | SCM-SYS-DOC-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Reliability · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-FW-013 (fault handler), DO-SAFE-### (safe states), DO-CLD-005 (health)
Input trace | SYS-007/047/099; SRS-DIAG; fault-management-diagnostics.md · Risk-control trace: Fault detection + safe-state (HAL, multiple) · Verification trace: Fault-injection test
Critical | Yes (defines detection/safe-state coverage) · Change class: Major
Acceptance criteria | Fault taxonomy, detection means (POST/BIST/watchdog/plausibility), responses, safe-state, and honest-gap principle defined; no undetected excursion-misreport path.
Mfg use | Indirect · Notes: Judgment; reconcile with HAL.
DO-SYS-010 — Deployment Architecture Description
Type | DO-DOC · Discipline: Systems/Cloud (S8/S9) · SKU: All · Subsystem/Component: Cloud + mobile topology
File # | SCM-SYS-DOC-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Cloud · Approval: Design Review Board
Source format | Markdown + Mermaid · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-CLD-001 (cloud arch), DO-CLD-011 (IaC), DO-CFG-006 (data residency)
Input trace | D4 (data residency); SYS-177…190; deployment-architecture.md · Risk-control trace: Trust-split; data-residency · Verification trace: Deployment/DR test
Critical | No · Change class: Major
Acceptance criteria | Device+cloud+mobile deployment topology, CI/CD environments, trust-split, and per-region data-residency defined and consistent with DO-CFG-006.
Mfg use | Indirect · Notes: Judgment.
DO-SYS-011 — Architecture Decision Records (ADR-001…012)
Type | DO-DOC · Discipline: Systems (SYS) · SKU: All · Subsystem/Component: Whole product
File # | SCM-SYS-DOC-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: Design Review Board
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-SYS-001; DO-EE, DO-FW, DO-SEC (per ADR)
Input trace | D1–D4; AD-1…AD-8; ADR-001…012 · Risk-control trace: ADR-003/004/005/006/008 (safety-relevant) · Verification trace: ADR validation hooks (named tests)
Critical | No · Change class: Major
Acceptance criteria | Each ADR states decision, alternatives, consequences, affected requirements, and a validation hook; Evidence-Pending ADRs (001 antenna, 003/004 energy/charge) named with closing test.
Mfg use | Indirect · Notes: Judgment.
DO-SYS-012 — Third-Party Service Dependency Map
Type | DO-DOC · Discipline: Systems/Cloud (S8) · SKU: All · Subsystem/Component: Cloud external deps (carrier, A-GNSS, IdP, notification)
File # | SCM-SYS-DOC-012 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Cloud · Approval: Design Review Board
Source format | Markdown · Released format: PDF · Repository: PLM · Parent assembly: DO-SYS-001
Related | DO-CLD-014 (external integrations), DO-GOV-007 (SOUP)
Input trace | SYS-116 (carrier); INT-030 (A-GNSS); SRS-CLOUD · Risk-control trace: Dependency-outage tolerance (HAL) · Verification trace: Dependency-failure test
Critical | No · Change class: Major
Acceptance criteria | Every external dependency (cellular carrier, A-GNSS provider, customer IdP/SSO, notification/email/SMS) listed with failure behavior and criticality; none is a single point of record loss.
Mfg use | Indirect · Notes: Judgment.
7. Traceability notes
Parents: PRD-ARCH; the full SYS/subsystem baseline; the Stage-08 architecture files. Forward: DO-SYS is the architectural design-output layer every discipline module (DOL-300…2000) refines; DO-SYS-004 (ICD) is the parent of all DOL-1000 interface outputs; DO-SYS-006/009 feed DOL-1200 (safe states).
8. Release / adequacy notes
Adequate when DO-SYS-001…012 are Released and internally consistent with the discipline modules. Three Critical (DO-SYS-004 ICD, DO-SYS-005 data/control flow, DO-SYS-007 power/energy, DO-SYS-009 fault mgmt — note four flagged Critical; DO-SYS-001/006 also carry safety-relevant Critical scope). Blockers: RF ICD (AD-1) and energy budget (D3) Evidence Pending; safe-state/detection architecture Preliminary — reconcile with HAL (Stage 10). </content>
DOL-2000 — Design Transfer & Release Baseline Design Outputs (DO-XFER) — SCM
Module: DOL-2000 · Output prefix: DO-XFER · Owner: Chief Systems Engineer (EM-ENG-001) · Discipline: Design Transfer / Configuration Audit / Release (cross-organization) · Status: Draft baseline (ORCH-013) — assembled at Stage 15; most entries are Evidence Pending until the baseline is frozen.
1. Purpose
Identify and govern the design-transfer and release-baseline outputs that establish the SCM design is complete, consistent, reconciled, and ready to be manufactured — the released-output index, the unresolved- output/open-items list, the configuration audit, the BOM/drawing/software/manufacturing reconciliations, the production-test/supplier/tooling/fixture/packaging/labeling/service/training readiness assessments, the release approvals, transfer deviations, open-risk disposition (reconciled with the HAL), and the pilot-build and production-release baselines. This module is the honest ledger that says what is released, what is still open, and whether the product may go to production.
2. Scope
The complete SCM design output baseline (all DOL-000…DOL-2000) at transfer. Covers configuration audit (FCA/PCA), the four reconciliations (BOM/drawing/software/manufacturing), readiness assessments across production test/supplier/tooling/packaging/labeling/service/training, release approvals and deviations, open-risk disposition, and the pilot-build → production-release baselines. All SKUs. This module contains no new engineering design — it indexes, audits, and dispositions the outputs produced by the other 20 modules and the HAL. Verification/validation/manufacturing records remain distinct (DO-GOV-009).
3. Applicable standards & references
ISO 13485:2016 §7.3.8 (design & development transfer); FDA QMSR (21 CFR 820, effective 2026-02-02, incorporating ISO 13485:2016); configuration-management audit practice (FCA/PCA); ISO 14971 (open-risk disposition ties to the risk file / HAL). Standard — revisions to confirm; the transfer baseline is Evidence Pending until DR-DO and the manufacturing-transfer gate.
4. Governance & ownership
Chief Systems Engineer (EM-ENG-001) owns transfer integration; Quality/Regulatory (EM-ENG-006) co-approves release; each discipline owner attests its readiness; the Reliability/Safety Engineer attests open-risk disposition against the HAL (Stage 10). The open-risk disposition & HAL-reconciliation record (DO-XFER-008) and the production-release baseline & approvals (DO-XFER-010) are the release-gate spine — production is not authorized until both are satisfied. This module operationalizes the Master Specification §9 release checklist (RA-01…RA-12).
5. Output categories
Released-output index & open-items list (DO-DOC), configuration-audit & reconciliation reports (DO-DOC), readiness assessments (DO-DOC), release/baseline definitions & approvals (DO-DOC/DO-CFG).
6. Design Output Register (this module)
DO-XFER-001 — Released-Output Index (Design Output Register snapshot)
Type | DO-DOC · Discipline: Systems (all S) · SKU: All · Subsystem/Component: Product baseline
File # | SCM-XFR-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: DRB/EM-ENG-006
Source format | Index doc · Released format: PDF · Repository: PLM · Parent assembly: SCM-DOL-MASTER-001
Related | Snapshots Design-Output-Register.md · all DOL-000…1900
Input trace | All PRD/SYS/SRS via the register; DO-GOV-001 · Risk-control trace: Complete, versioned index of what is released (no undocumented output) · Verification trace: Index-vs-register reconciliation
Critical | No · Change class: Major
Acceptance criteria | A frozen snapshot of every design output (ID, title, doc/file #, released rev, status, critical flag) at the transfer baseline; matches Design-Output-Register.md; each Critical output is Released.
Mfg use | Direct (transfer manifest) · Notes: Judgment; revisions Evidence Pending until release.
DO-XFER-002 — Unresolved-Output / Open-Items List
Type | DO-DOC · Discipline: Systems (all S) · SKU: All · Subsystem/Component: Product baseline
File # | SCM-XFR-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: DRB
Source format | List doc · Released format: PDF · Repository: PLM · Parent assembly: DO-XFER-001
Related | DO-XFER-008 (open risk), Master Spec §9 (RA gaps)
Input trace | All Evidence-Pending outputs; Master Spec §9 · Risk-control trace: Honest ledger of what is not closed (no silent gap) · Verification trace: Open-item burn-down review
Critical | No · Change class: Major
Acceptance criteria | Every Evidence-Pending / Draft / not-yet-Released output listed with owner, blocking dependency, and closure plan (e.g., chemistry DO-EE-003, antenna DO-EE-018, WHO PQS/US class DO-REG-001/005, cold-chamber/RF config DO-VER-009); nothing closed by assertion.
Mfg use | Indirect · Notes: Judgment; mirrors the program's top gaps.
DO-XFER-003 — Configuration Audit (FCA / PCA) Report
Type | DO-DOC · Discipline: Systems/Quality (all S) · SKU: All · Subsystem/Component: Product baseline
File # | SCM-XFR-DOC-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Quality · Approval: DRB/EM-ENG-006
Source format | Audit report · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-XFER-001
Related | DO-CFG-011 (baseline rules), DO-GOV-004 (change control), DO-XFER-004
Input trace | DO-CFG-001…012; DO-GOV-001…009 · Risk-control trace: Confirms the as-designed baseline equals the as-documented/as-built baseline · Verification trace: FCA (functional) + PCA (physical) audit
Critical | No · Change class: Major
Acceptance criteria | Functional configuration audit (each requirement verified against an output) and physical configuration audit (each output/document present at its released rev) completed; discrepancies dispositioned via ECO (DO-GOV-004).
Mfg use | Indirect · Notes: Judgment.
DO-XFER-004 — BOM / Drawing / Software / Manufacturing Reconciliation
Type | DO-DOC · Discipline: Systems/Mfg (all S) · SKU: All · Subsystem/Component: BOMs/drawings/SW/process
File # | SCM-XFR-DOC-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001/Mfg · Approval: DRB
Source format | Reconciliation tables · Released format: PDF · Repository: PLM · Parent assembly: DO-XFER-003
Related | DO-EE-020/DO-PCB-013/DO-MECH-019/DO-MFG-001 (BOMs), DO-FW-019/DO-CLD-016/DO-APP-009 (SW), DO-CFG-004 (AML)
Input trace | All BOM/drawing/SW outputs; DO-CFG-004/005 · Risk-control trace: eBOM=mBOM=AML=drawings=released SW versions — no build ambiguity (reconcile with HAL for critical parts) · Verification trace: Cross-BOM/version diff
Critical | No · Change class: Major
Acceptance criteria | Electrical/PCBA/mechanical/manufacturing BOMs reconcile to one another and to the AML; drawings match models; released firmware/cloud/mobile versions and hashes reconcile to the compatibility matrix (DO-CFG-005) and compliance-critical register (DO-REG-011); zero unreconciled deltas.
Mfg use | Direct · Notes: Judgment.
DO-XFER-005 — Production-Test Readiness Assessment
Type | DO-DOC · Discipline: Systems/Quality (all S) · SKU: All · Subsystem/Component: Test line
File # | SCM-XFR-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: Quality Eng · Approval: DRB
Source format | Assessment doc · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-XFER-001
Related | DOL-1500 (all DO-TST), DO-TST-013 (fixtures/golden), DO-VER-010 (pass/fail)
Input trace | DO-TST-001…014; DO-VER- · Risk-control trace: The safety/measurement production tests exist, are correlated, and gate acceptance (reconcile with HAL Stage 10) · Verification trace: Test-line readiness checklist
Critical | No · Change class: Major
Acceptance criteria | All production tests (esp. the critical FCT/system/cal/protection/leak) defined, fixtured, golden-unit-correlated, limit/MSA-backed, and validated on pilot units; the line can build and objectively accept a unit.
Mfg use | Direct · Notes: Judgment; depends on frozen AML (RA-07/RA-09).
DO-XFER-006 — Supplier / Tooling / Fixture Readiness Assessment
Type | DO-DOC · Discipline: Systems/Mfg (all S) · SKU: All · Subsystem/Component: Supply/tooling
File # | SCM-XFR-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: DRB/Quality
Source format | Assessment doc · Released format: PDF · Repository: PLM · Parent assembly: DO-XFER-001
Related | DO-CFG-004 (AML), DO-MFG-012/013 (fixtures/mould), DO-TST-013
Input trace | DO-CFG-004; DO-MFG- · Risk-control trace: Critical-part suppliers (cell/BMS/SE) qualified; tooling capable · Verification trace: Supplier qualification + tooling capability (Cpk) review
Critical | No · Change class: Major
Acceptance criteria | Approved suppliers qualified for all AML parts (enhanced for safety-critical cell/BMS/SE); tooling/moulds and fixtures built, tried out, and capable; second-source status recorded for critical parts.
Mfg use | Direct · Notes: Judgment; blocked by chemistry/antenna freeze (RA-07).
DO-XFER-007 — Packaging / Labeling / Service / Training Readiness Assessment
Type | DO-DOC · Discipline: Systems/Mfg/Service (all S) · SKU: All · Subsystem/Component: Downstream
File # | SCM-XFR-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: Mfg/Service Eng · Approval: DRB
Source format | Assessment doc · Released format: PDF · Repository: PLM · Parent assembly: DO-XFER-001
Related | DOL-1600 (packaging/labeling), DOL-1700 (service), DO-HF-009 (assistance)
Input trace | DO-LBL-; DO-SVC-; DO-HF-009 · Risk-control trace: Correct labels/IFU, service procedures, and trained field staff before launch · Verification trace: Readiness checklist per area
Critical | No · Change class: Major
Acceptance criteria | Packaging/labeling artwork (incl. regulatory/UDI/Li-transport), IFU/quick-start, service manual/FRU/cleaning/recal/disposal, and training material released and validated; field/depot staff trained; downstream is launch-ready.
Mfg use | Indirect · Notes: Judgment.
DO-XFER-008 — Open-Risk Disposition & HAL-Reconciliation Record
Type | DO-DOC · Discipline: Systems/Safety (all S) · SKU: All · Subsystem/Component: Risk file
File # | SCM-XFR-DOC-008 · Rev: TBD-rev · Status: Draft (Preliminary) · Owner: Reliability/Safety Eng · Approval: DRB/EM-ENG-006
Source format | Disposition record · Released format: PDF · Repository: PLM (risk file) · Parent assembly: DO-XFER-001
Related | Reconciles DOL-1200 (DO-SAFE-001…014) with the HAL (Stage 10) · DO-XFER-002 (open items)
Input trace | DOL-1200; PRD-SAFE; ISO 14971 · Risk-control trace: Every HAL risk control has an implementing, released output; residual risk is dispositioned and accepted by authority — the safety release gate (reconcile with HAL Stage 10) · Verification trace: Risk-control verification complete (Stage 13) + residual-risk review
Critical | Yes (safety release gate) · Change class: Major
Acceptance criteria | Every inherent-design/protective HAL control (DO-SAFE-) maps to a Released implementing output with completed risk-control verification; residual risk evaluated per ISO 14971 and formally accepted; RA-05 satisfied; no open safety item at production release.
Mfg use | Indirect · Notes: Preliminary — reconcile with HAL Stage 10; the top release blocker (Master Spec §9).
DO-XFER-009 — Pilot-Build Baseline & Transfer-Deviation Log
Type | DO-DOC · Discipline: Systems/Mfg (all S) · SKU: All · Subsystem/Component: Pilot baseline
File # | SCM-XFR-DOC-009 · Rev: TBD-rev · Status: Draft · Owner: Mfg Eng · Approval: DRB
Source format | Baseline + deviation log · Released format: PDF · Repository: PLM · Parent assembly: DO-XFER-001
Related | DO-GOV-004 (ECO), DO-XFER-004 (reconciliation), DO-XFER-005 (test readiness)
Input trace | Pilot outputs; DO-GOV-005 (change class) · Risk-control trace: Pilot findings dispositioned before volume; deviations controlled · Verification trace: Pilot-build yield + deviation closure
Critical | No · Change class: Major
Acceptance criteria | The exact output revisions used for the pilot build recorded; every transfer deviation logged, risk-assessed, and dispositioned (accept/ECO/blocker) before production-release; pilot lessons folded back via change control.
Mfg use | Direct · Notes: Judgment.
DO-XFER-010 — Production-Release Baseline & Release Approvals
Type | DO-DOC · Discipline: Systems/Quality (all S) · SKU: All · Subsystem/Component: Release baseline
File # | SCM-XFR-DOC-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-001 · Approval: DRB + EM-ENG-006 (Quality/Regulatory)
Source format | Baseline + approval record · Released format: PDF · Repository: PLM/QMS · Parent assembly: DO-XFER-001
Related | DO-XFER-003/004/005/006/007/008/009, Master Spec §9 (RA-01…12)
Input trace | All prerequisite DO-XFER; ISO 13485 §7.3.8 · Risk-control trace: Production authorized only when audit/reconciliation/readiness/open-risk are all satisfied (reconcile with HAL Stage 10) · Verification trace: Release-gate checklist (RA-01…12) all met
Critical | Yes (production-release authorization) · Change class: Major
Acceptance criteria | The frozen production-release baseline (all outputs Released at named revs), with configuration audit, reconciliations, readiness assessments, and open-risk disposition complete; every RA-01…RA-12 criterion met; signed release approvals from the DRB and Quality/Regulatory.
Mfg use | Direct (authorizes production) · Notes: Reconcile with HAL Stage 10; Evidence Pending until DR-DO + transfer gate. Standard (ISO 13485 §7.3.8).
7. Traceability notes
Parents: the entire DOL (DOL-000…DOL-1900) via Design-Output-Register.md, the HAL (Stage 10) for open-risk disposition, and the Master Specification §9 release checklist (RA-01…RA-12). Forward: DO-XFER- are the inputs to the Manufacturing/Design-Transfer stage (Stage 15), the Release stage (Stage 16), and the end-to-end lifecycle matrix (Stage 18, ORCH-018 / EM-ENG-012). Open-risk disposition (DO-XFER-008) and the production-release baseline (DO-XFER-010) reconcile with HAL Stage 10 and gate production.
8. Release / adequacy notes
Adequate — i.e., the SCM may transfer to production — only when the released-output index, open-items list, configuration audit, four reconciliations, all readiness assessments, open-risk disposition, and the production-release baseline are complete and approved. Two Critical: DO-XFER-008 (open-risk/HAL reconciliation) and DO-XFER-010 (production-release baseline & approvals). This module is Evidence Pending by nature (assembled at Stage 15); the honest top blockers remain RA-05 (HAL reconciliation), RA-07 (chemistry + antenna freeze), and RA-10 (regulatory classification) per Master Spec §9. </content>
DOL-300 — Mechanical Design Outputs (DO-MECH) — SCM
Module: DOL-300 · Output prefix: DO-MECH · Owner: Principal Mechanical Engineer · Discipline: Mechanical / Enclosure / Materials (S1) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the mechanical design outputs that let the SCM enclosure system be reproduced: 3D CAD, released 2D drawings with GD&T, tolerance stacks, material/finish specs, seal/gasket design, the external temperature probe, mounting/retention, sealed USB-C aperture, thermal-coupling design, IP-sealing design, and the mechanical BOM.
2. Scope
S1 enclosure and mechanical interfaces (CMP-001…005, CMP-034 mechanical aspects, CMP-030/031 HMI mechanicals). All SKUs; SCM-100-XA adds the external-antenna aperture. Excludes electronics (DOL-400/500).
3. Applicable standards & references
ASME Y14.5 / ISO 1101 (GD&T); IEC 60529 (IP ingress); IEC 60068-2 (thermal/mechanical environment); ISO 16750 concepts (shock/vibration); RoHS/WEEE (materials); ISO 10993-relevant only if patient contact (N/A — device is not patient-contacting). Standard — revisions to confirm.
4. Governance & ownership
Principal Mechanical Engineer owns; Materials engineer concurs on DO-MECH-010; Manufacturing (DOL-1400) consumes tooling inputs; HAL (Stage 10) reconciles seal/retention/thermal risk controls.
5. Output categories
3D CAD (DO-CAD), 2D released drawings (DO-DRW), GD&T/tolerance analysis (DO-DRW/DO-MODEL), specifications (DO-SPEC), material/finish/color specs (DO-SPEC), mechanical BOM (DO-BOM).
6. Design Output Register (this module)
DO-MECH-000 — Enclosure Top-Level Assembly (3D CAD + drawing)
Type | DO-CAD · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Enclosure assembly / CMP-001
File # | SCM-MEC-CAD-000 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE (EM-ENG-001)
Source format | Native CAD · Released format: STEP AP242 + PDF drawing · Repository: PLM/CAD vault · Parent assembly: DO-CFG-001
Related | DO-MECH-001…005, DO-MECH-019 (mech BOM)
Input trace | PRD-MECH; SYS-069…085; MER-001; CMP-001 · Risk-control trace: — · Verification trace: First-article dimensional (DO-TST-002)
Critical | No · Change class: Major
Acceptance criteria | Complete assembly of all S1 parts with mating/interference-free fit; assembly reproduces the sealed device envelope; consistent with mechanical BOM.
Mfg use | Direct (assembly reference) · Notes: Judgment.
DO-MECH-001 — Enclosure Top Housing (3D CAD)
Type | DO-CAD · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Top housing / CMP-001
File # | SCM-MEC-CAD-001 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | Native CAD · Released format: STEP AP242 · Repository: CAD vault · Parent assembly: DO-MECH-000
Related | DO-MECH-002 (drawing), DO-MECH-010 (material), DO-MECH-011 (seal groove)
Input trace | MER-001…010; SYS-069/104 (antenna keep-out); PRD-MECH · Risk-control trace: IP seal integrity (HAL) · Verification trace: IP67 test (DO-TST-006)
Critical | No · Change class: Major
Acceptance criteria | Wall thickness/draft mouldable; antenna keep-out honored; seal groove modeled; interfaces to bottom housing/probe gland/USB-C aperture defined.
Mfg use | Direct (mould-tool input) · Notes: Judgment.
DO-MECH-002 — Enclosure Top Housing (2D released drawing + GD&T)
Type | DO-DRW · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Top housing / CMP-001
File # | SCM-MEC-DRW-002 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | Native CAD drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-001, DO-MECH-018 (tolerance stack), DO-VER-002 (CTF dims)
Input trace | MER-001…010; ASME Y14.5 · Risk-control trace: Seal-surface CTQ (HAL) · Verification trace: CMM/first-article (DO-TST-002)
Critical | No · Change class: Major
Acceptance criteria | Fully dimensioned with GD&T datums; CTF/CTQ sealing-surface features toleranced; drawing complete enough to inspect without the model.
Mfg use | Direct (inspection + tooling) · Notes: Standard (Y14.5).
DO-MECH-003 — Enclosure Bottom Housing (3D CAD)
Type | DO-CAD · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Bottom housing / CMP-001
File # | SCM-MEC-CAD-003 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | Native CAD · Released format: STEP AP242 · Repository: CAD vault · Parent assembly: DO-MECH-000
Related | DO-MECH-004 (drawing), DO-MECH-012 (battery retention), DO-MECH-013 (PCB mounts)
Input trace | MER-001…015; CMP-007 (battery); CMP-034 (PCB) · Risk-control trace: Battery retention (HAL) · Verification trace: Drop/vibe (DO-TST-007)
Critical | No · Change class: Major
Acceptance criteria | PCB and battery mounting bosses/retention modeled; mates/seals to top housing; service-access considered.
Mfg use | Direct · Notes: Judgment.
DO-MECH-004 — Enclosure Bottom Housing (2D drawing + GD&T)
Type | DO-DRW · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Bottom housing / CMP-001
File # | SCM-MEC-DRW-004 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | CAD drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-003, DO-MECH-018
Input trace | MER-001…015; ASME Y14.5 · Risk-control trace: — · Verification trace: First-article (DO-TST-002)
Critical | No · Change class: Major
Acceptance criteria | Fully toleranced with GD&T; mounting-boss positions toleranced to PCB/battery stack.
Mfg use | Direct · Notes: Standard.
DO-MECH-005 — Enclosure Seal / Gasket Design (main O-ring/gasket)
Type | DO-SPEC · Discipline: Mechanical/Materials (S1) · SKU: All · Subsystem/Component: Main seal / CMP-002
File # | SCM-MEC-SPC-005 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/Materials
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-011 (groove), DO-MECH-010 (material), DO-TST-006 (IP)
Input trace | MER-038/040; SYS-069; IEC 60529 · Risk-control trace: Ingress → moisture/condensation → sensor/electronics fault (reconcile with HAL Stage 10) · Verification trace: IP67 + cold-cycle seal test
Critical | Yes (seal preserves record-integrity electronics through −30↔+60 °C cold-to-warm cycling) · Change class: Major
Acceptance criteria | Gland/groove fill %, compression set, and material rated across −30…+60 °C with cold-to-warm condensation cycling; maintains IP67 (or target) after specified reuse cycles.
Mfg use | Direct · Notes: Standard (IEC 60529); compression/cold-set values Evidence Pending (test).
DO-MECH-006 — External Temperature Probe Assembly (3D CAD)
Type | DO-CAD · Discipline: Mechanical (S1/S4) · SKU: All (+ probe options DO-CFG-003) · Subsystem/Component: Probe / CMP-020
File # | SCM-MEC-CAD-006 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | Native CAD · Released format: STEP AP242 + drawing · Repository: CAD vault · Parent assembly: DO-CFG-003
Related | DO-MECH-007 (gland), DO-EE-013 (probe front-end), DO-IF-001 (probe pinout)
Input trace | SYS-003; HWR-022/028; MER-016…020; CMP-020; INT-001/004 · Risk-control trace: Probe thermal-coupling proxy for freeze/heat verdict (reconcile with HAL Stage 10) · Verification trace: Thermal-characterization test (INT-001)
Critical | Yes (probe tip is the product-zone temperature proxy the disposition verdict rests on) · Change class: Major
Acceptance criteria | Probe tip thermal mass/geometry gives characterized offset & time-constant; positive retention; lead length/strain-relief specified; sealed to probe body.
Mfg use | Direct · Notes: Evidence Pending (offset/lag by test); probe technology (thermistor vs digital) to confirm with S4 (IFR-003).
DO-MECH-007 — Probe Lead Gland & Strain-Relief Design
Type | DO-SPEC · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Probe entry / CMP-002/005
File # | SCM-MEC-SPC-007 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-005, DO-IF-001 (INT-006)
Input trace | MER-018; SYS-052; IFR-003; INT-006 · Risk-control trace: Seal at probe entry (HAL) · Verification trace: IP + pull/flex test
Critical | No · Change class: Major
Acceptance criteria | Sealed gland maintains enclosure IP rating at the probe boundary; strain relief withstands specified pull/flex cycles without seal loss.
Mfg use | Direct · Notes: Standard/Judgment.
DO-MECH-008 — Sealed USB-C Port Aperture & Gasket Design
Type | DO-SPEC · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: USB-C aperture / CMP-004
File # | SCM-MEC-SPC-008 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-EE-009 (USB-C port), DO-IF-001 (INT-007/008)
Input trace | MER-021; SYS-052/054; IFR-001; INT-007 · Risk-control trace: Ingress at charge port (HAL) · Verification trace: IP after mate cycles
Critical | No · Change class: Major
Acceptance criteria | Aperture/gasket maintains IP rating with connector mated and unmated; cleaning-chemical compatible; rated for the specified mate/de-mate cycles.
Mfg use | Direct · Notes: Judgment.
DO-MECH-009 — Enclosure Mount / Retention Feature Design
Type | DO-SPEC · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Mount / CMP-003
File # | SCM-MEC-SPC-009 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-000, DO-SVC-002 (install), DO-IF (INT-005)
Input trace | SYS-003/052; MER-022…025; CMP-003; INT-005 · Risk-control trace: Device displacing payload / defeating thermal barrier (HAL) · Verification trace: Fit test in shipper set
Critical | No · Change class: Major
Acceptance criteria | Tool-free, glove-serviceable retention within the defined shipper set; holds probe placement; does not displace payload or defeat the insulated barrier.
Mfg use | Direct · Notes: Judgment; shipper set defined by target fleet (D1/D4).
DO-MECH-010 — Enclosure Material & Surface-Finish/Color Specification
Type | DO-SPEC · Discipline: Materials (S1) · SKU: All · Subsystem/Component: Housings/seals
File # | SCM-MEC-SPC-010 · Rev: TBD-rev · Status: Draft · Owner: Materials Eng · Approval: CSE/Quality
Source format | Spec doc · Released format: PDF · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-005, DO-CFG-008 (cleaning), DO-REG-007 (material declarations)
Input trace | MER-030…035; SYS-069…085; RoHS/WEEE; cleaning-chemical compatibility · Risk-control trace: Material embrittlement at −30 °C (HAL) · Verification trace: Cold-impact + chemical-compatibility test
Critical | No · Change class: Major
Acceptance criteria | Housing/seal materials specified with grade class, low-temperature impact retention, UV/chemical resistance, RoHS/REACH compliance, and color/finish; no invented resin part numbers (placeholder grades).
Mfg use | Direct (procurement + moulding) · Notes: Evidence Pending (final grade); Standard (RoHS/WEEE).
DO-MECH-011 — Seal Groove / Gland Detail Drawing
Type | DO-DRW · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Seal groove / CMP-002
File # | SCM-MEC-DRW-011 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | CAD drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-005, DO-MECH-002/004
Input trace | MER-038; IEC 60529 · Risk-control trace: Seal fill % CTQ (HAL) · Verification trace: IP test + groove CMM
Critical | No · Change class: Major
Acceptance criteria | Groove cross-section, fill %, and surface finish toleranced to the seal spec; CTQ features called out.
Mfg use | Direct · Notes: Standard.
DO-MECH-012 — Battery Retention / Isolation Design
Type | DO-SPEC · Discipline: Mechanical (S1/S2) · SKU: All · Subsystem/Component: Battery retention / CMP-007
File # | SCM-MEC-SPC-012 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/Power
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-003
Related | DO-EE-004 (pack assembly), DO-SAFE-010 (thermal isolation)
Input trace | HWR-002; SYS-090; MER-026; CMP-007; UN 38.3 · Risk-control trace: Cell mechanical abuse / short (reconcile with HAL Stage 10) · Verification trace: Drop/vibe + transport test (UN 38.3)
Critical | Yes (cell mechanical integrity is a safety leg of D3) · Change class: Major
Acceptance criteria | Retains the cell without crush/puncture through drop/vibration/transport; electrical/thermal clearance to PCB and enclosure maintained; supports self-heat element (CMP-013) placement.
Mfg use | Direct · Notes: Standard (UN 38.3); reconcile with HAL.
DO-MECH-013 — PCB Mounting & Board-Retention Design
Type | DO-SPEC · Discipline: Mechanical (S1/S3) · SKU: All · Subsystem/Component: PCB mounts / CMP-034
File # | SCM-MEC-SPC-013 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/PCB
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-003
Related | DO-PCB-001 (outline), DO-MECH-003
Input trace | HWR-037; MER-027; CMP-034 · Risk-control trace: Board flex → solder-joint fatigue (HAL) · Verification trace: Vibe + thermal-cycle test
Critical | No · Change class: Major
Acceptance criteria | Board mounting limits flex/strain across thermal cycling and shock; connector/antenna keep-outs honored; matches PCB outline (DO-PCB-001).
Mfg use | Direct · Notes: Judgment.
DO-MECH-014 — HMI Window / OLED & Button Mechanical Design
Type | DO-SPEC · Discipline: Mechanical (S1/S6) · SKU: All · Subsystem/Component: HMI mechanicals / CMP-030/031
File # | SCM-MEC-SPC-014 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/HF
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-HF-004 (button), DO-EE-017 (OLED), DO-MECH-005 (seal)
Input trace | HWR-046/047/050; SYS-015/023; MER-028; CMP-030/031 · Risk-control trace: Sealed glove-actuable button (HAL) · Verification trace: Actuation-force + IP test
Critical | No · Change class: Major
Acceptance criteria | OLED window optically clear and sealed; glove-friendly button force/travel/size sealed and debounce-mechanically sound; readable at −30 °C viewing angles.
Mfg use | Direct · Notes: Judgment; OLED cold response Evidence Pending (CMP-030).
DO-MECH-015 — Thermal-Coupling / Internal-Sensor Airpath Design
Type | DO-MODEL · Discipline: Mechanical/Thermal (S1/S4) · SKU: All · Subsystem/Component: Internal sensor path / CMP-019
File # | SCM-MEC-MDL-015 · Rev: TBD-rev · Status: Draft · Owner: Mech/Thermal Eng · Approval: CSE
Source format | Thermal-sim model + report · Released format: PDF + model archive · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-006 (probe), DO-SYS-007, DO-IF-001 (INT-002/003)
Input trace | SYS-003/051; INT-002/003; PRD-PERF · Risk-control trace: Internal-channel lag misrepresents micro-environment (HAL) · Verification trace: Thermal-characterization (INT-002/003)
Critical | No · Change class: Major
Acceptance criteria | Internal-sensor airpath/thermal coupling designed and modeled so internal-channel lag/offset is bounded and characterized; low self-heating.
Mfg use | Indirect · Notes: Evidence Pending (characterization test).
DO-MECH-016 — External-Antenna Aperture Design (SCM-100-XA)
Type | DO-SPEC · Discipline: Mechanical (S1/S5) · SKU: SCM-100-XA · Subsystem/Component: Ext-antenna connector aperture / CMP-029
File # | SCM-MEC-SPC-016 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/RF
Source format | CAD + spec · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-EE-021 (ext-antenna), DO-IF-006 (RF ICD)
Input trace | HWR-033; SYS-109/104; CMP-029; AD-1; INT-026 · Risk-control trace: Ingress at RF connector (HAL) · Verification trace: IP + RF-in-container test (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | IP-sealed external-antenna connector aperture preserves enclosure IP rating; only present on -XA; no seal/RF degradation when option unused on base SKUs.
Mfg use | Direct (variant tooling) · Notes: Evidence Pending (antenna branch, AD-1).
DO-MECH-017 — Shock/Vibration & Drop Structural Analysis
Type | DO-MODEL · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Enclosure + internals
File # | SCM-MEC-MDL-017 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | FEA model + report · Released format: PDF + model archive · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-012/013, DO-TST-007 (drop/vibe)
Input trace | SYS-070…075 (mechanical env); IEC 60068-2; MER-036 · Risk-control trace: Enclosure/cell integrity under shock (HAL) · Verification trace: Drop/vibe test (DO-TST-007)
Critical | No · Change class: Major
Acceptance criteria | Analysis predicts no structural failure, cell crush, or seal loss under specified transport shock/vibration/drop; correlates with test.
Mfg use | Indirect · Notes: Standard (IEC 60068-2).
DO-MECH-018 — Tolerance-Stack Analysis (seal, HMI, connector fits)
Type | DO-MODEL · Discipline: Mechanical (S1) · SKU: All · Subsystem/Component: Critical fits
File # | SCM-MEC-MDL-018 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE
Source format | Tolerance-analysis workbook · Released format: PDF · Repository: PLM · Parent assembly: DO-MECH-000
Related | DO-MECH-002/004/011, DO-VER-002 (CTF dims)
Input trace | ASME Y14.5; MER-037; the released drawings · Risk-control trace: Seal-fit worst-case (HAL) · Verification trace: First-article + Cpk
Critical | No · Change class: Major
Acceptance criteria | Worst-case and statistical stacks show seal fill %, HMI/window fit, and connector alignment stay within functional limits across the toleranced range.
Mfg use | Direct (tolerance allocation) · Notes: Standard.
DO-MECH-019 — Mechanical Bill of Materials
Type | DO-BOM · Discipline: Mechanical (S1) · SKU: All (per-SKU deltas) · Subsystem/Component: All S1 parts
File # | SCM-MEC-BOM-019 · Rev: TBD-rev · Status: Draft · Owner: Mech Eng · Approval: CSE/Quality
Source format | PLM BOM · Released format: BOM export + PDF · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-CFG-004 (AML), DO-MECH-000
Input trace | All MER; CMP-001…005 · Risk-control trace: Seal/probe critical-part control · Verification trace: Incoming inspection (DO-TST-001)
Critical | No · Change class: Major
Acceptance criteria | Every mechanical part listed with quantity, material spec ref, drawing ref, and AML linkage; per-SKU deltas (e.g., -XA aperture) captured.
Mfg use | Direct · Notes: Judgment.
7. Traceability notes
Parents: PRD-MECH; SYS-003/051/052/069…085/104/109; MER-001…040; CMP-001…005/007/029/030/031/034. Forward: feeds DOL-1400 (tooling/WI), DOL-1500 (dimensional/IP test), DOL-1200 (seal/probe/retention risk controls), and DOL-2000 (drawing reconciliation). Probe (DO-MECH-006) and seal (DO-MECH-005) are the safety-relevant outputs — reconcile with HAL Stage 10.
8. Release / adequacy notes
Adequate when CAD, released drawings with GD&T, seal spec, probe assembly, and mech BOM are Released and first-article-verified. Four Critical (DO-MECH-005 seal, DO-MECH-006 probe, DO-MECH-012 battery retention; plus safety-relevant DO-MECH-015). Blockers: probe offset/lag and internal-airpath thermal characterization Evidence Pending; final material grade Evidence Pending; -XA aperture gated by AD-1. </content>
DOL-400 — Electrical & Electronic Design Outputs (DO-EE) — SCM
Module: DOL-400 · Output prefix: DO-EE · Owner: Principal Electrical/Hardware Engineer (Battery/ Power lead for S2) · Discipline: Electrical / Power / Sensing / RF front-end (S2, S3, S4, S5, S6) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the electrical/electronic design outputs from which the SCM electronics can be reproduced: electrical architecture, schematics, power tree, power budget, battery/BMS/charge design, charge-inhibit interlock, sensor interfaces, radio front-ends, HMI electronics, protection/derating, and the electrical BOM.
2. Scope
S2 power (cell/pack/BMS/PMIC/gauge/protection/self-heat), S3 MCU/SE/RTC/watchdog electricals, S4 sensing front-ends, S5 radio front-ends + antenna matching, S6 HMI electronics. Excludes PCB layout/fab (DOL-500) and firmware (DOL-600). All SKUs; -NA/-EU/-ROW differ in band-plan RF; -XA differs in antenna feed.
3. Applicable standards & references
IEC 62368-1 (electrical safety); IEC 61000 / CISPR (EMC/ESD); UN 38.3 / IATA (Li-battery); relevant cell datasheet charge-temperature limits; FCC Part 15 / CE RED (radios); IPC derating guidance. Standard — revisions to confirm; no cell/IC part numbers invented.
4. Governance & ownership
Principal Electrical Engineer owns; Battery/Power lead owns S2 (D3 top risk); EM-ENG-008 concurs on SE (DO-EE-011); HAL (Stage 10) reconciles all charge/protection/energy safety controls. The charge-inhibit interlock (DO-EE-007), BMS design (DO-EE-005), and charge path (DO-EE-006) are the D3 charge-safety spine.
5. Output categories
Electrical architecture (DO-DOC), schematics (DO-SCH), power tree/budget (DO-SCH/DO-MODEL), design specs (DO-SPEC), derating/WCA (DO-MODEL), pinouts/wiring (DO-DRW/DO-ICD), electrical BOM (DO-BOM).
6. Design Output Register (this module)
DO-EE-000 — Electrical Architecture Description
Type | DO-DOC · Discipline: Electrical (S2–S6) · SKU: All · Subsystem/Component: Electronics
File # | SCM-ELE-DOC-000 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Markdown + block diagram · Released format: PDF · Repository: PLM/EDA · Parent assembly: DO-SYS-001
Related | DO-EE-001 (schematic set), DO-EE-002 (power tree)
Input trace | PRD-ELEC; SYS-086…103; HWR-001…050 · Risk-control trace: Power/charge safety spine (HAL) · Verification trace: Electrical design review
Critical | No · Change class: Major
Acceptance criteria | End-to-end electrical block architecture (power, compute, sensing, RF, HMI) consistent with the component list and interfaces.
Mfg use | Indirect · Notes: Judgment.
DO-EE-001 — Full Schematic Set (device electronics)
Type | DO-SCH · Discipline: Electrical (S2–S6) · SKU: All (band-plan sheet per region) · Subsystem/Component: All electronics
File # | SCM-ELE-SCH-001 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE/PCB
Source format | EDA schematic · Released format: PDF + EDA netlist · Repository: EDA/PLM · Parent assembly: DO-EE-000
Related | DO-PCB-002 (netlist), DO-EE-020 (BOM), all DO-EE-###
Input trace | All HWR; CMP-004…034 · Risk-control trace: Multiple (charge/protection/sensor) · Verification trace: Schematic review + bring-up test
Critical | Yes (defines all safety-relevant nets: charge-inhibit, protection, probe presence) · Change class: Major
Acceptance criteria | Complete, ERC-clean schematic; every net named; safety nets (charge-inhibit INT-016, protection, probe-presence) explicit; drives netlist to DO-PCB-002.
Mfg use | Direct (source of netlist/BOM) · Notes: Judgment.
DO-EE-002 — Power Tree / Power-Distribution Schematic
Type | DO-SCH · Discipline: Power (S2) · SKU: All · Subsystem/Component: Rails / CMP-010
File # | SCM-ELE-SCH-002 · Rev: TBD-rev · Status: Draft · Owner: Power lead · Approval: CSE
Source format | EDA · Released format: PDF · Repository: EDA/PLM · Parent assembly: DO-EE-001
Related | DO-EE-006 (budget), DO-FW-004 (governor gating), DO-IF-001 (INT-012…015)
Input trace | SYS-099 (gated rails); HWR-043; INT-012…015 · Risk-control trace: Rail-gating never corrupts NVM write (HAL) · Verification trace: Rail-sequencing test
Critical | Yes (gated rails feed energy governor + write-safe storage) · Change class: Major
Acceptance criteria | Always-on MCU/RTC rail + independently FW-gated sensing/connectivity/HMI/storage rails; rail down-transitions interlocked against active NVM writes.
Mfg use | Indirect · Notes: Judgment; rail voltages/tolerances Evidence Pending (INT-012…015).
DO-EE-003 — Rechargeable Low-Temperature Cell Selection Specification
Type | DO-SPEC · Discipline: Power (S2) · SKU: All · Subsystem/Component: Cell / CMP-006
File # | SCM-ELE-SPC-003 · Rev: TBD-rev (Evidence Pending — chemistry) · Status: Draft · Owner: Power lead · Approval: CSE/Quality
Source format | Spec doc · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-004
Related | DO-EE-004 (pack), DO-CFG-004 (AML), DO-SYS-007 (energy arch)
Input trace | D3; HWR-001; SYS-088; CMP-006 · Risk-control trace: Low-temp capacity / cold-charge chemistry safety (reconcile with HAL Stage 10) · Verification trace: Cold-capacity + cycle-life test (PRD-VER-004)
Critical | Yes (D3 endurance + safety) · Change class: Major
Acceptance criteria | Chemistry (low-temp Li-ion or LTO-class) whose −30 °C usable capacity after derating meets SYS-087 budget with margin; charge-temperature window and abuse tolerance specified per selected datasheet.
Mfg use | Direct (procurement) · Notes: Evidence Pending — chemistry unresolved; no capacity/chemistry/part number invented (D3 top risk).
DO-EE-004 — Battery Pack Assembly Design (config + interconnect)
Type | DO-SPEC · Discipline: Power (S2) · SKU: All · Subsystem/Component: Pack / CMP-007
File # | SCM-ELE-SPC-004 · Rev: TBD-rev (Evidence Pending — config) · Status: Draft · Owner: Power lead · Approval: CSE/Quality
Source format | Spec + CAD · Released format: PDF + STEP · Repository: PLM · Parent assembly: DO-CFG-001
Related | DO-EE-003, DO-MECH-012 (retention), DO-EE-005 (BMS), DO-LBL-010 (transport)
Input trace | HWR-002; SYS-090; CMP-007; UN 38.3 · Risk-control trace: Pack interconnect/balancing safety (reconcile with HAL Stage 10) · Verification trace: UN 38.3 + pack qualification
Critical | Yes · Change class: Major
Acceptance criteria | xSyP config fixed by SYS-090 budget; documented interconnect/balancing; traceable to cell datasheet; passes UN 38.3.
Mfg use | Direct · Notes: Evidence Pending (config follows chemistry, DO-EE-003).
DO-EE-005 — Battery-Protection IC / AFE (BMS) Design & Threshold Specification
Type | DO-SPEC · Discipline: Power/Safety (S2) · SKU: All · Subsystem/Component: BMS / CMP-008
File # | SCM-ELE-SPC-005 · Rev: TBD-rev · Status: Draft · Owner: Power lead · Approval: CSE/Safety
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-006/007/008, DO-SAFE-002 (protection), DO-EE-010 (fuse)
Input trace | HWR-003; SYS-096; CMP-008; INT-011; IEC 62368-1 · Risk-control trace: OV/UV/OC/OT/short protection independent of MCU/FW (reconcile with HAL Stage 10) · Verification trace: Fault-injection protection test (DO-TST-009)
Critical | Yes (autonomous single-fault-safe protection) · Change class: Major
Acceptance criteria | Protection thresholds (OV/UV/OC/OT/short) per cell datasheet; protection persists in any MCU/FW fault; MCU-independent.
Mfg use | Direct · Notes: Reconcile with HAL Stage 10; thresholds follow chemistry (Evidence Pending).
DO-EE-006 — Charge-Path / PMIC Design & Charge-Profile Specification
Type | DO-SPEC · Discipline: Power (S2) · SKU: All · Subsystem/Component: Charger/PMIC / CMP-010
File # | SCM-ELE-SPC-006 · Rev: TBD-rev (Evidence Pending — profile) · Status: Draft · Owner: Power lead · Approval: CSE/Safety
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-007 (inhibit), DO-EE-009 (USB-C), DO-EE-012 (self-heat), DO-IF-001 (INT-010)
Input trace | HWR-004/043; SYS-093/099; CMP-010; INT-010; cell datasheet · Risk-control trace: Charge only within safe cell-temperature window (reconcile with HAL Stage 10) · Verification trace: Charge-profile + thermal test
Critical | Yes · Change class: Major
Acceptance criteria | Programmable charge profile per cell datasheet; USB-C sink; per-rail regulation with independently gateable load switches; charge gated by DO-EE-007.
Mfg use | Direct · Notes: Evidence Pending (profile follows chemistry); reconcile with HAL.
DO-EE-007 — Charge-Inhibit-Below-0 °C Interlock Design (HW+FW boundary)
Type | DO-SPEC · Discipline: Power/Safety (S2/S3) · SKU: All · Subsystem/Component: Charge-inhibit / CMP-008/010/012, INT-016
File # | SCM-ELE-SPC-007 · Rev: TBD-rev · Status: Draft · Owner: Power lead · Approval: CSE/Safety (EM-ENG-006 concur)
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-005/006/008, DO-FW-005 (FW charge-inhibit logic), DO-SAFE-001, DO-IF-001 (INT-016)
Input trace | D3; HWR-006; SYS-094; CMP-012; INT-010/016; ADR-004; cell charge-temp limits · Risk-control trace: Cold-charge cell damage/hazard — HW interlock is the primary control (reconcile with HAL Stage 10) · Verification trace: Cold-charge-inhibit test (fault-injection, cell-cold)
Critical | Yes (top D3 charge-safety control) · Change class: Major
Acceptance criteria | Cell-coupled NTC (not board-air) qualifies charge enable; charging is inhibited below 0 °C nominal and above the high limit; inhibit persists independent of firmware; inhibited state + cause reported (INT-016/020) — never a silent failure to charge.
Mfg use | Direct · Notes: Reconcile with HAL Stage 10 (primary charge-safety control; HW + FW defense-in-depth). Standard (cell charge-temp limits, IEC 62368-1).
DO-EE-008 — Fuel-Gauge Design & SoC/SoH Configuration
Type | DO-SPEC · Discipline: Power (S2) · SKU: All · Subsystem/Component: Fuel gauge / CMP-011
File # | SCM-ELE-SPC-008 · Rev: TBD-rev (Evidence Pending — accuracy) · Status: Draft · Owner: Power lead · Approval: CSE
Source format | Spec + config · Released format: PDF + config · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-FW-012 (diagnostics), DO-IF-001 (INT-020)
Input trace | HWR-008; SYS-097/024; CMP-011; INT-020 · Risk-control trace: Honest low-battery/readiness (avoid misleading "ready") — HAL · Verification trace: Gauge-accuracy test
Critical | No · Change class: Major
Acceptance criteria | Temperature-compensated coulomb-count/model gauge supports pre-arm readiness and non-misleading low-battery warning; accuracy meets SYS-097.
Mfg use | Direct · Notes: Evidence Pending (accuracy by test).
DO-EE-009 — USB-C Sink Port & Input-Protection Design
Type | DO-SPEC · Discipline: Power (S2) · SKU: All · Subsystem/Component: USB-C sink / CMP-004/014
File # | SCM-ELE-SPC-009 · Rev: TBD-rev · Status: Draft · Owner: Power lead · Approval: CSE/Safety
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-MECH-008 (aperture), DO-EE-010 (protection), DO-IF-001 (INT-008/009)
Input trace | HWR-039/045; SYS-093/114; CMP-004/014; INT-008/009; USB Type-C/PD · Risk-control trace: Reverse-polarity/OVP/surge/short at charge input (reconcile with HAL Stage 10) · Verification trace: OV/reverse/short + PD-negotiation test
Critical | Yes (charge-input abuse tolerance) · Change class: Major
Acceptance criteria | Sink-only (UFP) with Rd; safe on legacy 5 V and under reverse-insertion/OVP/short; PD PDO capped (≤9 V, Evidence Pending) and higher PDOs rejected; never sources VBUS.
Mfg use | Direct · Notes: Standard (USB Type-C/PD, IEC 62368-1); PD cap Evidence Pending (INT-008).
DO-EE-010 — Secondary Non-Resettable Protection (Fuse/OVP) Design
Type | DO-SPEC · Discipline: Safety/Power (S2) · SKU: All · Subsystem/Component: Backstop protector / CMP-009
File # | SCM-ELE-SPC-010 · Rev: TBD-rev · Status: Draft · Owner: Power lead · Approval: CSE/Safety
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-005 (BMS), DO-SAFE-002
Input trace | HWR-012; SYS-096; CMP-009; IEC 62368-1 · Risk-control trace: Last-line over-current/over-voltage backstop if primary IC fails — single-fault safety (reconcile with HAL Stage 10) · Verification trace: Secondary-protection fault-injection
Critical | Yes (single-fault safety backstop) · Change class: Major
Acceptance criteria | Series fuse and/or secondary OVP terminates the hazard if the primary protection IC fails; coordinated with BMS thresholds.
Mfg use | Direct · Notes: Reconcile with HAL Stage 10. Standard.
DO-EE-011 — MCU, Secure-Element, RTC & Watchdog Electrical Design
Type | DO-SPEC · Discipline: Compute (S3) · SKU: All · Subsystem/Component: MCU/SE/RTC/WDT / CMP-015/016/017/018
File # | SCM-ELE-SPC-011 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE (EM-ENG-008 concur SE)
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-012 (debug port), DO-SEC-004 (secure boot), DO-IF-001 (INT-024)
Input trace | HWR-013/015/016/017/019; SYS-089/047; CMP-015…018; INT-024 · Risk-control trace: Brown-out reset to safe state; trusted-time backup (HAL) · Verification trace: Brown-out/BOR + RTC-retention test
Critical | No · Change class: Major
Acceptance criteria | −30…+60 °C-rated MCU with retention deep-sleep, wake-on-RTC/interrupt, brown-out supervisor; RTC with backup source preserving trusted UTC; independent watchdog; SE bus to CMP-016.
Mfg use | Direct · Notes: Judgment.
DO-EE-012 — Debug / EOL-Test Port & Production-Lock Design
Type | DO-SPEC · Discipline: Compute/Security (S3) · SKU: All · Subsystem/Component: SWD/JTAG+UART / IFR-005
File # | SCM-ELE-SPC-012 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE/EM-ENG-008
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-SEC-009 (debug lock), DO-MFG-008 (programming), DO-IF-011 (test-interface ICD)
Input trace | IFR-005; SRS-SEC; SYS-132-line · Risk-control trace: Locked debug prevents key/record readout (reconcile with HAL Stage 10) · Verification trace: Debug-lock verification (DO-TST-012)
Critical | Yes (security — fielded lock) · Change class: Major
Acceptance criteria | Debug/EOL port enables factory programming/calibration/test and is locked in fielded units (readout-protection/lifecycle fuse set, authenticated debug); no unauthenticated code/key/record readout in the field.
Mfg use | Direct (programming access) · Notes: Standard (IEC 62443 concepts).
DO-EE-013 — External-Probe Analog/Digital Front-End & Presence-Detect Design
Type | DO-SPEC · Discipline: Sensing (S4) · SKU: All · Subsystem/Component: Probe front-end / CMP-020
File # | SCM-ELE-SPC-013 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-MECH-006 (probe), DO-FW-007 (sensor drivers), DO-IF-001 (INT-018)
Input trace | HWR-022/028; SYS-003/007; CMP-020; INT-018 · Risk-control trace: Absent/open/short/wrong probe detected — prevents false "accept" from room-temp read (reconcile with HAL Stage 10) · Verification trace: Probe-fault-injection test
Critical | Yes (undetected open probe would fabricate a false accept) · Change class: Major
Acceptance criteria | Probe channel accuracy meets HWR-021 at tip; presence/ID/continuity line detects absent/wrong/open/short probe and flags rather than reading a valid product-zone temperature.
Mfg use | Direct · Notes: Reconcile with HAL Stage 10. Probe technology to confirm (IFR-003).
DO-EE-014 — Internal Temperature/Humidity Sensor Interface Design
Type | DO-SPEC · Discipline: Sensing (S4) · SKU: All · Subsystem/Component: Internal T/RH / CMP-019/021
File # | SCM-ELE-SPC-014 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-FW-007, DO-IF-001 (INT-017)
Input trace | HWR-021/023/027; SYS-001/003/004; CMP-019/021; INT-017 · Risk-control trace: Plausibility flagging on bus error (HAL) · Verification trace: Accuracy + bus-fault test
Critical | No · Change class: Major
Acceptance criteria | Internal T channel ±0.5 °C or better across range with 0.1 °C resolution and low self-heating; RH ±5 %RH; each device on a fixed I2C address/CS; bus NAK/CRC error returns defined error.
Mfg use | Direct · Notes: Judgment; accuracy class Evidence Pending (CMP-019).
DO-EE-015 — Accelerometer & Shock-Sensing Interface Design
Type | DO-SPEC · Discipline: Sensing (S4) · SKU: All · Subsystem/Component: Accelerometer / CMP-022
File # | SCM-ELE-SPC-015 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-FW-007, DO-FW-006 (detection)
Input trace | HWR-024; SYS-005; CMP-022 · Risk-control trace: Shock-event capture (HAL) · Verification trace: Shock-profile test
Critical | No · Change class: Major
Acceptance criteria | Full-scale + bandwidth for per-profile shock; low-power interrupt/threshold mode to wake on impact within energy budget.
Mfg use | Direct · Notes: Judgment.
DO-EE-016 — GNSS Receiver Interface Design
Type | DO-SPEC · Discipline: Sensing (S4) · SKU: All · Subsystem/Component: GNSS / CMP-023
File # | SCM-ELE-SPC-016 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-018 (antenna), DO-FW-007, DO-IF-001 (INT-019/030)
Input trace | HWR-025; SYS-029/117; CMP-023; INT-019/030 · Risk-control trace: Fail-soft last-known + quality flag (HAL) · Verification trace: In-container fix test (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | Configurable/event-triggered acquisition, fully powered off between fixes; reports fix quality; fails soft to best/last-known with quality flag.
Mfg use | Direct · Notes: Evidence Pending (in-container fix/energy, PRD-VER-005).
DO-EE-017 — Radio Front-End Design (LTE-M / Wi-Fi / BLE) + Band Plan
Type | DO-SPEC · Discipline: RF (S5) · SKU: All (band-plan per -NA/-EU/-ROW) · Subsystem/Component: Radios / CMP-024/025/026/027
File # | SCM-ELE-SPC-017 · Rev: TBD-rev (band plan Evidence Pending) · Status: Draft · Owner: RF Eng · Approval: CSE/Regulatory
Source format | Spec + schematic sheet · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-018 (antenna match), DO-FW-011 (comms agent), DO-IF-006/007, DO-REG-004 (radio certs)
Input trace | HWR-029/030/031/034/035; SYS-104…118; CMP-024…027; INT-025/027/028/029; FCC Part 15/CE RED · Risk-control trace: Region-permitted-band enforcement (HAL/regulatory) · Verification trace: Radio conformance + cold-RF test
Critical | No · Change class: Major
Acceptance criteria | LTE-M Cat-M1 (PSM/eDRX), Wi-Fi 802.11, BLE, roaming SIM/eSIM; per-region band plan enforced; MCU can power-gate each radio; RF front-end holds spec to −30 °C.
Mfg use | Direct · Notes: Standard; cold-RF margin + band plan Evidence Pending (INT-027).
DO-EE-018 — Antenna & Matching-Network Design (integral)
Type | DO-SPEC · Discipline: RF (S5/S1) · SKU: All (base) · Subsystem/Component: Integral antennas / CMP-028
File # | SCM-ELE-SPC-018 · Rev: TBD-rev (Evidence Pending — antenna) · Status: Draft · Owner: RF Eng · Approval: CSE
Source format | Spec + EM model · Released format: PDF + model · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-021 (ext-antenna), DO-PCB-006 (RF routing), DO-IF-006 (RF ICD), DO-MECH-001 (keep-out)
Input trace | HWR-032; SYS-104/109; CMP-028; INT-026/031; AD-1 · Risk-control trace: In-container link → delayed/missing alert (HAL) · Verification trace: RF-in-container qualification (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | Integral antennas + matching for LTE-M/Wi-Fi/BLE/GNSS within enclosure keep-out, tuned to close the in-container link budget; if it cannot close, escalate to -XA branch (DO-EE-021).
Mfg use | Direct · Notes: Evidence Pending — principal open architectural risk (AD-1, PRD-VER-005).
DO-EE-019 — Evidence, EMC/ESD Provisions & Safety-Spacing Design
Type | DO-SPEC · Discipline: Electrical (S2–S6) · SKU: All · Subsystem/Component: Board-level EMC/ESD
File # | SCM-ELE-SPC-019 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE
Source format | Spec + schematic notes · Released format: PDF · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-PCB-004 (stack-up/ground), DO-REG-005 (EMC), DO-TST-010 (ESD)
Input trace | HWR-041/042; SYS-177…190; CISPR/IEC 61000; IEC 62368-1 · Risk-control trace: ESD-induced fault (HAL) · Verification trace: EMC/ESD test (pre-compliance + certified)
Critical | No · Change class: Major
Acceptance criteria | Evidence topology, ESD protection at all external ports (USB-C, probe, button), and creepage/clearance per IEC 62368-1; targets CISPR/IEC 61000 limits.
Mfg use | Indirect · Notes: Standard.
DO-EE-020 — Electrical Bill of Materials
Type | DO-BOM · Discipline: Electrical (S2–S6) · SKU: All (per-SKU RF deltas) · Subsystem/Component: All electronic parts
File # | SCM-ELE-BOM-020 · Rev: TBD-rev · Status: Draft · Owner: Elec Eng · Approval: CSE/Quality
Source format | EDA BOM · Released format: BOM export + PDF · Repository: EDA/PLM · Parent assembly: DO-CFG-004
Related | DO-PCB-013 (PCBA BOM), DO-CFG-004 (AML)
Input trace | All HWR; CMP-006…034 · Risk-control trace: Critical-component (cell/BMS/SE/fuse) control · Verification trace: Incoming inspection (DO-TST-001)
Critical | No · Change class: Major
Acceptance criteria | Every electronic component listed with reference designator, value/spec, and AML linkage; critical parts flagged; consistent with schematic and PCBA BOM.
Mfg use | Direct · Notes: Judgment; part numbers Evidence Pending (no invention).
DO-EE-021 — External-Antenna Option Design (SCM-100-XA)
Type | DO-SPEC · Discipline: RF (S5/S1) · SKU: SCM-100-XA · Subsystem/Component: Ext antenna + connector / CMP-029
File # | SCM-ELE-SPC-021 · Rev: TBD-rev (Evidence Pending) · Status: Draft · Owner: RF Eng · Approval: CSE
Source format | Spec + EM model · Released format: PDF + model · Repository: PLM · Parent assembly: DO-EE-001
Related | DO-EE-018, DO-MECH-016 (aperture), DO-IF-006
Input trace | HWR-033; SYS-104/109; CMP-029; AD-1; INT-026 · Risk-control trace: In-container link fallback (HAL) · Verification trace: RF-in-container test (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | Qualified external antenna via IP-sealed connector closes the link budget where integral cannot; no degradation to other bands when unused on base SKUs.
Mfg use | Direct (variant) · Notes: Evidence Pending (antenna branch, AD-1).
7. Traceability notes
Parents: PRD-ELEC/CONN/PERF/REL/SAFE; SYS-086…118; HWR-001…050; CMP-004…034; INT-008…031. Forward: seeds DOL-500 (netlist/PCBA), DOL-600 (FW charge-inhibit/governor/drivers pair with DO-EE-002/007/013…016), DOL-1200 (protection/charge/probe risk controls), DOL-1500 (electrical/functional test), DOL-1800 (radio certs, EMC). The S2 power chain (DO-EE-003…010) is the D3 safety spine.
8. Release / adequacy notes
Adequate when schematics, power tree, S2 battery/BMS/charge/inhibit specs, sensor/radio interfaces, and electrical BOM are Released and bring-up-verified. Eight Critical (DO-EE-001, -002, -003, -004, -005, -006, -007, -009, -010, -013 — the safety-relevant set; charge-inhibit DO-EE-007 is the top control). Blockers: chemistry (DO-EE-003) and antenna (DO-EE-018) Evidence Pending; charge/protection specs reconcile with HAL Stage 10; band plan/PD cap Evidence Pending. </content>
DOL-500 — PCB & PCBA Design Outputs (DO-PCB) — SCM
Module: DOL-500 · Output prefix: DO-PCB · Owner: Principal PCB Engineer · Discipline: PCB / PCBA (S3 carrier, all electronics) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the PCB/PCBA manufacturing design outputs: board outline, stack-up, controlled impedance, placement, routing, layer artwork, fab & assembly drawings, Gerber/ODB++, drill, pick-and-place, stencil, netlist, test points, and PCBA BOM — from which the SCM board can be fabricated and assembled.
2. Scope
The single main PCBA (CMP-034) carrying S2/S3/S4/S5/S6 electronics, conformal coat, RF/high-speed routing. All SKUs share one board where possible; regional/antenna deltas are populate/no-populate options and RF matching.
3. Applicable standards & references
IPC-2221/2222 (design), IPC-6012 (fab), IPC-A-610 (assembly acceptance), IPC-7351 (footprints), IPC-CC-830 (conformal coat), J-STD-001 (soldering). Standard — revisions to confirm.
4. Governance & ownership
Principal PCB Engineer owns; Electrical (DOL-400) supplies netlist/BOM; Manufacturing (DOL-1400) consumes fab/assembly data; conformal-coat keep-outs coordinated with sensing/antenna/contacts.
5. Output categories
Fab data (DO-PCB), assembly data (DO-PCB), drawings (DO-DRW), netlist (DO-DATA), PCBA BOM (DO-BOM), fab/ assembly acceptance specs (DO-SPEC).
6. Design Output Register (this module)
DO-PCB-001 — Board Outline & Mechanical Definition
Type | DO-DRW · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCB / CMP-034
File # | SCM-PCB-DRW-001 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: CSE/Mech
Source format | EDA/CAD · Released format: PDF + DXF · Repository: EDA/PLM · Parent assembly: DO-MECH-013
Related | DO-MECH-013 (mounts), DO-EE-018 (antenna keep-out), DO-PCB-004
Input trace | HWR-037; MER-027; CMP-034 · Risk-control trace: Antenna/sensor keep-out (HAL) · Verification trace: Fit check + first-article
Critical | No · Change class: Major
Acceptance criteria | Outline, mounting holes, connector/antenna/sensor keep-outs match enclosure (DO-MECH-013); consistent across SKUs.
Mfg use | Direct · Notes: Judgment.
DO-PCB-002 — Netlist (from schematic)
Type | DO-DATA · Discipline: PCB (S3) · SKU: All · Subsystem/Component: All nets
File # | SCM-PCB-DAT-002 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: Elec/PCB
Source format | EDA netlist · Released format: IPC-D-356 netlist · Repository: EDA/PLM · Parent assembly: DO-EE-001
Related | DO-EE-001 (schematic), DO-PCB-011 (test points)
Input trace | DO-EE-001 · Risk-control trace: Net integrity for safety nets · Verification trace: Netlist compare + bare-board test
Critical | No · Change class: Major
Acceptance criteria | Netlist matches released schematic; used for bare-board electrical test.
Mfg use | Direct · Notes: Judgment; derived from DO-EE-001.
DO-PCB-003 — Layer Stack-Up & Materials Specification
Type | DO-SPEC · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCB / CMP-034
File # | SCM-PCB-SPC-003 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: CSE
Source format | Spec + EDA · Released format: PDF · Repository: PLM · Parent assembly: DO-PCB-001
Related | DO-PCB-006 (impedance), DO-EE-019 (evidence)
Input trace | HWR-037/041; IPC-6012 · Risk-control trace: Continuous ground for signal integrity · Verification trace: Cross-section/coupon
Critical | No · Change class: Major
Acceptance criteria | Layer count, copper weights, dielectric materials, and vias defined; continuous reference plane; supports cold-operation reliability.
Mfg use | Direct · Notes: Standard (IPC-6012).
DO-PCB-004 — Placement / Layout (routed board)
Type | DO-PCB · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCB / CMP-034
File # | SCM-PCB-LAY-004 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: Elec/PCB
Source format | EDA layout · Released format: Native + PDF plots · Repository: EDA/PLM · Parent assembly: DO-PCB-001
Related | DO-PCB-005 (Gerber), DO-PCB-006 (RF), DO-EE-019 (EMC)
Input trace | DO-EE-001; HWR-037/041/042 · Risk-control trace: Safety-net spacing; RF/sensor isolation (HAL) · Verification trace: DRC + EMC pre-compliance
Critical | Yes (controls creepage/clearance of charge/protection nets + RF integrity) · Change class: Major
Acceptance criteria | DRC-clean; safety-spacing on charge/protection nets; RF/high-speed nets routed to impedance targets; conformal-coat keep-outs honored.
Mfg use | Direct · Notes: Judgment.
DO-PCB-005 — Gerber / ODB++ Fabrication Data Set
Type | DO-PCB · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCB / CMP-034
File # | SCM-PCB-FAB-005 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Mfg
Source format | Gerber X2 / ODB++ · Released format: Gerber X2 + ODB++ + drill · Repository: EDA/PLM · Parent assembly: DO-PCB-004
Related | DO-PCB-007 (fab drawing), DO-PCB-004
Input trace | DO-PCB-004 · Risk-control trace: — · Verification trace: Fab-data review + first-article
Critical | No · Change class: Major
Acceptance criteria | Complete, self-consistent fab set (copper, mask, silk, paste, drill) regenerable from layout; matches fab drawing.
Mfg use | Direct · Notes: Standard; derived from DO-PCB-004.
DO-PCB-006 — Controlled-Impedance & RF Routing Specification
Type | DO-SPEC · Discipline: PCB/RF (S3/S5) · SKU: All · Subsystem/Component: RF nets / CMP-028
File # | SCM-PCB-SPC-006 · Rev: TBD-rev (Evidence Pending — antenna) · Status: Draft · Owner: PCB/RF Eng · Approval: CSE
Source format | Spec + impedance table · Released format: PDF · Repository: PLM · Parent assembly: DO-PCB-003
Related | DO-EE-018 (antenna), DO-PCB-004
Input trace | HWR-032/041; SYS-109; INT-026 · Risk-control trace: RF feed loss → in-container link (HAL) · Verification trace: Coupon TDR + RF-in-container (PRD-VER-005)
Critical | No · Change class: Major
Acceptance criteria | Impedance targets/tolerances for RF/high-speed nets; antenna-feed routing and match input to link budget.
Mfg use | Direct · Notes: Evidence Pending (antenna, AD-1).
DO-PCB-007 — Fabrication Drawing (IPC class, notes, acceptance)
Type | DO-DRW · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCB / CMP-034
File # | SCM-PCB-DRW-007 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Quality
Source format | CAD drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-PCB-005
Related | DO-PCB-003/005, DO-TST-003 (bare-board test)
Input trace | IPC-6012; HWR-037 · Risk-control trace: — · Verification trace: Bare-board acceptance
Critical | No · Change class: Major
Acceptance criteria | IPC class (target Class 2/3, to confirm), stack-up, finish, tolerances, and acceptance notes on the drawing.
Mfg use | Direct · Notes: Standard.
DO-PCB-008 — Assembly Drawing & Pick-and-Place Data
Type | DO-PCB · Discipline: PCB (S3) · SKU: All (populate options per SKU) · Subsystem/Component: PCBA / CMP-034
File # | SCM-PCB-ASM-008 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Mfg
Source format | EDA · Released format: PDF + centroid/XY · Repository: EDA/PLM · Parent assembly: DO-PCB-004
Related | DO-PCB-009 (stencil), DO-PCB-013 (PCBA BOM), DO-MFG-002 (SMT WI)
Input trace | DO-EE-020; IPC-A-610 · Risk-control trace: Correct-part placement (cell/BMS/SE) · Verification trace: AOI + first-article
Critical | No · Change class: Major
Acceptance criteria | Component placement, orientation, and reference designators; populate/no-populate per SKU; centroid data matches layout.
Mfg use | Direct · Notes: Standard (IPC-A-610).
DO-PCB-009 — Stencil / Solder-Paste Aperture Design
Type | DO-PCB · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCBA / CMP-034
File # | SCM-PCB-STN-009 · Rev: TBD-rev · Status: Draft · Owner: PCB/Mfg Eng · Approval: Mfg
Source format | Gerber paste · Released format: Stencil Gerber · Repository: EDA/PLM · Parent assembly: DO-PCB-008
Related | DO-PCB-005, DO-MFG-003 (reflow params)
Input trace | IPC-7525 (stencil); DO-PCB-004 · Risk-control trace: Solder-joint reliability at temp cycling · Verification trace: Paste inspection (SPI)
Critical | No · Change class: Minor
Acceptance criteria | Aperture design gives correct paste volume for fine-pitch/QFN parts; supports cold-cycle solder reliability.
Mfg use | Direct · Notes: Standard.
DO-PCB-010 — Conformal-Coat & Cleanliness Specification
Type | DO-SPEC · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCBA / CMP-034
File # | SCM-PCB-SPC-010 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Quality
Source format | Spec + keep-out drawing · Released format: PDF · Repository: PLM · Parent assembly: DO-PCB-004
Related | DO-EE-014 (sensors), DO-EE-018 (antenna), DO-MFG-005 (coat WI)
Input trace | HWR-038; IPC-CC-830; SYS-069 (condensation) · Risk-control trace: Condensation-induced leakage/corrosion at −30↔+60 °C (reconcile with HAL Stage 10) · Verification trace: Cleanliness (ROSE/IC) + coat inspection
Critical | Yes (condensation protection of record-integrity electronics) · Change class: Major
Acceptance criteria | Coat type/coverage protects against condensation with defined keep-outs (sensor apertures, antenna, contacts, connectors); cleanliness limits set.
Mfg use | Direct · Notes: Standard (IPC-CC-830); reconcile with HAL.
DO-PCB-011 — Test-Point & Boundary-Scan / Fiducial Definition
Type | DO-DATA · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCBA / CMP-034
File # | SCM-PCB-DAT-011 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Quality
Source format | EDA · Released format: PDF + test-point report · Repository: EDA/PLM · Parent assembly: DO-PCB-004
Related | DO-TST-004 (ICT/functional), DO-VER-003 (test points), DO-EE-012 (debug)
Input trace | DO-EE-001; SRS-VER; PRD-VER · Risk-control trace: Observability of safety nets for test · Verification trace: ICT coverage report
Critical | No · Change class: Major
Acceptance criteria | Test points/fiducials give production-test access (ICT/flying-probe/functional) with defined coverage of critical nets (charge-inhibit, protection, probe presence).
Mfg use | Direct · Notes: Judgment.
DO-PCB-012 — Panelization / Array Drawing
Type | DO-DRW · Discipline: PCB (S3) · SKU: All · Subsystem/Component: PCBA panel
File # | SCM-PCB-DRW-012 · Rev: TBD-rev · Status: Draft · Owner: PCB/Mfg Eng · Approval: Mfg
Source format | EDA · Released format: PDF + Gerber · Repository: EDA/PLM · Parent assembly: DO-PCB-005
Related | DO-MFG-002 (SMT), DO-PCB-008
Input trace | DFM guidelines; DO-PCB-001 · Risk-control trace: — · Verification trace: Panel first-article
Critical | No · Change class: Minor
Acceptance criteria | Array layout, rails, tooling holes, fiducials, and depanel method defined for the assembly line.
Mfg use | Direct · Notes: Judgment.
DO-PCB-013 — PCBA Bill of Materials
Type | DO-BOM · Discipline: PCB (S3) · SKU: All (populate deltas) · Subsystem/Component: PCBA / CMP-034
File # | SCM-PCB-BOM-013 · Rev: TBD-rev · Status: Draft · Owner: PCB Eng · Approval: PCB/Quality
Source format | EDA BOM · Released format: BOM export + PDF · Repository: EDA/PLM · Parent assembly: DO-CFG-004
Related | DO-EE-020 (electrical BOM), DO-PCB-008
Input trace | DO-EE-020; DO-EE-001 · Risk-control trace: Critical-part control (cell/BMS/SE/fuse) · Verification trace: Incoming inspection (DO-TST-001)
Critical | No · Change class: Major
Acceptance criteria | Reference-designator-level BOM with AML linkage and populate/no-populate per SKU; consistent with electrical BOM.
Mfg use | Direct · Notes: Judgment; part numbers Evidence Pending.
7. Traceability notes
Parents: HWR-037/038/041/042; MER-027; CMP-028/034; DO-EE-001/019/020. Forward: feeds DOL-1400 (SMT/coat WIs), DOL-1500 (bare-board/ICT/functional test), DOL-1900 (test points/observability), DOL-2000 (fab-data reconciliation). RF routing (DO-PCB-006) ties to the antenna Evidence-Pending item (AD-1).
8. Release / adequacy notes
Adequate when layout, Gerber/ODB++, fab & assembly drawings, and PCBA BOM are Released and first-article built. Three Critical (DO-PCB-004 layout safety-spacing, DO-PCB-010 conformal coat; DO-PCB-011 observability is safety-relevant). Blocker: RF routing/impedance (DO-PCB-006) tied to antenna Evidence Pending (AD-1). </content>
DOL-600 — Embedded Firmware Design Outputs (DO-FW) — SCM
Module: DOL-600 · Output prefix: DO-FW · Owner: Embedded Software Engineer (EM-ENG-007) · Discipline: Embedded Firmware (S3) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the SCM firmware design outputs: architecture, module decomposition, task/state model, bootloader/secure boot, drivers/HAL, the energy governor, detection & verdict engines, storage manager, comms & OTA agents, security/crypto module, diagnostics, config schema, build config/toolchain, repo baseline, dependency manifest/SBOM, and the signed released firmware image with hash.
2. Scope
S3 firmware (CMP-035…045). All SKUs run a common image; regional/config differences are runtime config (DO-CFG-006). Excludes cloud/mobile (DOL-800/900) and device-side host tools (DOL-700).
3. Applicable standards & references
IEC 62304 (software lifecycle/SOUP); MISRA-C guidance; secure-boot/signed-update practice; ALCOA+/21 CFR Part 11 (record integrity); ISO 14971 (software risk controls). Standard — revisions to confirm.
4. Governance & ownership
EM-ENG-007 owns; EM-ENG-008 concurs on security-relevant FW (secure boot, crypto, OTA verify); Reliability/ Safety reconciles detection/verdict/charge-inhibit/safe-state with the HAL (Stage 10). The energy governor (DO-FW-004), FW charge-inhibit logic (DO-FW-005), detection engine (DO-FW-006), and verdict engine (DO-FW-
008) are the load-bearing safety/function FW outputs.
5. Output categories
Design docs (DO-DOC), source code (DO-CODE), config schema (DO-CFG/DO-DATA), build config (DO-CFG), released binary (DO-BIN), hash/manifest (DO-DATA), release notes (DO-DOC), SBOM (DO-DATA).
6. Design Output Register (this module)
DO-FW-001 — Firmware Architecture Description
Type | DO-DOC · Discipline: FW (S3) · SKU: All · Subsystem/Component: FW / CMP-035…045
File # | SCM-FW-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Markdown + diagrams · Released format: PDF · Repository: Software repo/PLM · Parent assembly: DO-SYS-001
Related | DO-FW-002/003, DO-SYS-006 (modes)
Input trace | FWR-001…055; SRS-DEV; SYS-089 · Risk-control trace: Overall FW safety architecture (HAL) · Verification trace: Architecture review
Critical | No · Change class: Major
Acceptance criteria | Layered FW architecture (bootloader/HAL/drivers/services/app) with the 11 components decomposed; consistent with the mode model.
Mfg use | Indirect · Notes: Judgment.
DO-FW-002 — Module Decomposition & Task/Scheduler Model
Type | DO-DOC · Discipline: FW (S3) · SKU: All · Subsystem/Component: Scheduler/RTOS / CMP-036
File # | SCM-FW-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Markdown + diagrams · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-FW-004 (governor), DO-SYS-002
Input trace | FWR (scheduler); SYS-001/089; CMP-036 · Risk-control trace: Deterministic sampling cadence · Verification trace: Timing test
Critical | No · Change class: Major
Acceptance criteria | Task model, wake-on-RTC/interrupt scheduling, and cooperation with the power-state manager fully specified.
Mfg use | Indirect · Notes: Judgment.
DO-FW-003 — State-Machine / Mode-Manager Design
Type | DO-DOC · Discipline: FW (S3) · SKU: All · Subsystem/Component: Mode manager / CMP-037
File # | SCM-FW-DOC-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Safety
Source format | Markdown + stateDiagram · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-SYS-006, DO-FW-013 (fault handler), DO-SAFE-006 (safe state)
Input trace | SYS-089; state-mode-model.md; FWR · Risk-control trace: Safe-state / Fault-Safe transitions (reconcile with HAL Stage 10) · Verification trace: State-transition test
Critical | Yes (defines Fault-Safe entry and invariants) · Change class: Major
Acceptance criteria | All modes and transitions implemented with guards/invariants; safe-state reachable from any mode without corrupting committed log or armed state.
Mfg use | Direct (Shipping-mode default) · Notes: Reconcile with HAL.
DO-FW-004 — Power-State Manager / Energy Governor Design
Type | DO-DOC · Discipline: FW (S3) · SKU: All · Subsystem/Component: Energy governor / CMP-037
File # | SCM-FW-DOC-004 · Rev: TBD-rev (budget Evidence Pending) · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Power
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-EE-002 (rails), DO-SYS-007 (energy arch), DO-FW-011 (comms)
Input trace | D3; FWR (governor); SYS-086/089/090; CMP-037; INT-012…015 · Risk-control trace: Energy exhaustion → loss of monitoring (reconcile with HAL Stage 10) · Verification trace: 90-day cold endurance (PRD-VER-004)
Critical | Yes (single most load-bearing FW component for the D3 budget) · Change class: Major
Acceptance criteria | Holds lowest viable sleep between activities; hard-gates LTE-M/Wi-Fi/BLE/GNSS/display/sensor rails; permits only queued excursion alert outside the window; enforces per-class energy allocations.
Mfg use | Indirect · Notes: Evidence Pending (budget numbers by test); reconcile with HAL.
DO-FW-005 — Firmware Charge-Inhibit & Self-Heat Control Logic
Type | DO-CODE · Discipline: FW/Safety (S3/S2) · SKU: All · Subsystem/Component: Charge control / CMP-037, INT-016
File # | SCM-FW-COD-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Safety (EM-ENG-006 concur)
Source format | Source (C) · Released format: Reviewed source + built into DO-FW-017 · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-EE-007 (HW interlock), DO-SAFE-001, DO-EE-012 (self-heat), DO-IF-001 (INT-016/020)
Input trace | D3; FWR (charge/self-heat); SYS-094/095; ADR-004; INT-010/016 · Risk-control trace: Cold-charge safety — FW is the secondary/defense-in-depth control atop HW interlock DO-EE-007 (reconcile with HAL Stage 10) · Verification trace: Cold-charge-inhibit + self-heat test
Critical | Yes (charge-safety leg of D3) · Change class: Major
Acceptance criteria | FW never commands charge into a cold cell; respects HW interlock (cannot bypass it); self-heat cannot exceed safe thermal limit or override inhibit; reports inhibited state + cause (INT-020).
Mfg use | Indirect · Notes: Reconcile with HAL Stage 10. HW interlock (DO-EE-007) is primary; FW is defense-in-depth.
DO-FW-006 — Excursion-Detection Engine Design
Type | DO-CODE · Discipline: FW/Safety (S3) · SKU: All · Subsystem/Component: Detection / CMP-039
File # | SCM-FW-COD-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Safety
Source format | Source + design doc · Released format: Reviewed source + built binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-FW-007 (drivers), DO-FW-008 (verdict), DO-CLD-006 (cloud recompute)
Input trace | FWR (detection); SYS-008/009/010/014; CMP-039; PRD-FUN/SAFE · Risk-control trace: Missed freeze/heat excursion (two-sided) → wrong disposition (reconcile with HAL Stage 10) · Verification trace: Detection-logic test (profiled excursions)
Critical | Yes (excursion logic — HAL Stage 10 headline control) · Change class: Major
Acceptance criteria | Two-sided threshold + product-stability (time/cumulative) excursion logic with hysteresis/debounce/consolidation; freeze detection reliability equals heat; shock classification per accelerometer.
Mfg use | Indirect · Notes: Reconcile with HAL Stage 10.
DO-FW-007 — Sensor Drivers / Acquisition & Calibration-Application Design
Type | DO-CODE · Discipline: FW (S3/S4) · SKU: All · Subsystem/Component: Sensor drivers / CMP-038
File # | SCM-FW-COD-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Source · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-FW-009 (calibration schema), DO-EE-013/014/015/016, DO-IF-001 (INT-017/018/019)
Input trace | FWR (sensing); SYS-001/007; HWR-026/028; CMP-038; INT-017…019 · Risk-control trace: Plausibility/channel-fault flagging; probe-absence flagging (reconcile with HAL Stage 10) · Verification trace: Sensor-fault-injection test
Critical | Yes (channel-fault/probe-absence flagging prevents false accept) · Change class: Major
Acceptance criteria | Applies calibration/linearization; runs plausibility checks; flags channel faults, bus errors (NAK/CRC), and probe absence rather than returning stale/substituted values; gates each sensor rail to its window.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-FW-008 — Verdict Engine Design (accept/reject/quarantine)
Type | DO-CODE · Discipline: FW/Safety (S3) · SKU: All · Subsystem/Component: Verdict / CMP-040
File # | SCM-FW-COD-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Safety
Source format | Source + design doc · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-FW-006 (detection), DO-CLD-006 (recompute), DO-HF-002 (verdict display)
Input trace | FWR (verdict); SYS-015/016/017; CMP-040 · Risk-control trace: Suppress false "accept" on gaps/faults (reconcile with HAL Stage 10) · Verification trace: Verdict-consistency test (device↔cloud)
Critical | Yes (last-mile disposition decision) · Change class: Major
Acceptance criteria | Deterministic accept/reject/quarantine from the trip record; suppresses false accept on gaps/faults; consistent with the cloud recomputation; presents incomplete-monitoring qualifier.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-FW-009 — Calibration Data Schema & Application Design
Type | DO-DATA · Discipline: FW/Metrology (S3/S4) · SKU: All · Subsystem/Component: Calibration store
File # | SCM-FW-DAT-009 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Quality
Source format | Schema + design · Released format: PDF + schema · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-TST-008 (calibration test), DO-FW-007, DO-SVC-005 (recal)
Input trace | SYS-003; PRD-PERF; ISO/IEC 17025 concepts · Risk-control trace: Calibration integrity → measurement accuracy (HAL) · Verification trace: Calibration verification (DO-TST-008)
Critical | No · Change class: Major
Acceptance criteria | Calibration coefficients storable per channel with metadata (date, cal-ID, expiry); tamper-evident; applied deterministically; supports recalibration.
Mfg use | Direct (programmed at calibration) · Notes: Judgment.
DO-FW-010 — Storage Manager (append-only log, power-fail-safe) Design
Type | DO-CODE · Discipline: FW (S3/S7) · SKU: All · Subsystem/Component: Storage manager / CMP-041
File # | SCM-FW-COD-010 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Quality
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-EE-002 (rail interlock), DO-SEC-006 (tamper-evidence), DO-IF-001 (INT-021)
Input trace | FWR (storage); SYS-002/018/021/028/099; CMP-041; INT-021; ALCOA+ · Risk-control trace: No torn/lost/fabricated record on power loss (reconcile with HAL Stage 10) · Verification trace: Power-fail-during-write test
Critical | Yes (evidentiary record durability) · Change class: Major
Acceptance criteria | Gap-free append with atomic/power-fail-safe writes; wear management; buffer-full prioritization (excursion/audit first); tamper-evidence; exactly-once send-marking; no overwrite of unsent data.
Mfg use | Indirect · Notes: Standard (ALCOA+); sizing Evidence Pending (SYS-127).
DO-FW-011 — Comms Agent / Bearer-Manager Design
Type | DO-CODE · Discipline: FW (S3/S5) · SKU: All · Subsystem/Component: Comms agent / CMP-042
File # | SCM-FW-COD-011 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-IF-005 (device↔cloud), DO-EE-017 (radios), DO-FW-004 (governor)
Input trace | FWR (comms); SYS-019/020/108/111/112/113/116; CMP-042; INT-025/032 · Risk-control trace: Alert-before-bulk; store-and-forward never loses record (HAL) · Verification trace: Store-and-forward + alert-priority test
Critical | Yes (excursion-alert egress path) · Change class: Major
Acceptance criteria | Energy-aware bearer selection (LTE-M→Wi-Fi→BLE); ordered/resumable/idempotent upload; alert-before-bulk; connectivity diagnostics; roaming/network selection; failed link degrades to buffering, never record loss.
Mfg use | Indirect · Notes: Judgment; in-container link Evidence Pending.
DO-FW-012 — Diagnostics / Health & BIST Design
Type | DO-CODE · Discipline: FW (S3) · SKU: All · Subsystem/Component: Diagnostics / CMP-045
File # | SCM-FW-COD-012 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-SYS-009 (fault mgmt), DO-CLD-005 (fleet health), DO-EE-008 (gauge)
Input trace | FWR (diagnostics); SRS-DIAG; SYS-024/030/113; CMP-045 · Risk-control trace: Pre-arm readiness; honest-gap reporting (HAL) · Verification trace: POST/BIST test
Critical | No · Change class: Major
Acceptance criteria | POST/BIST, pre-arm checks (SoC/SoH, calibration, firmware, last-seen), connectivity/RSSI history, buffer backlog; feeds fleet health and RF-in-container evidence.
Mfg use | Direct (POST at test) · Notes: Judgment.
DO-FW-013 — Fault Handler / Watchdog & Recovery Design
Type | DO-CODE · Discipline: FW/Safety (S3) · SKU: All · Subsystem/Component: Fault handling / CMP-037/045
File # | SCM-FW-COD-013 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/Safety
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-EE-011 (WDT), DO-FW-003 (safe state), DO-SAFE-006
Input trace | FWR (fault); SYS-047/099; SRS-DIAG · Risk-control trace: Recover hang to safe state without record/armed corruption (reconcile with HAL Stage 10) · Verification trace: Fault-injection + watchdog test
Critical | Yes (recovery to safe state) · Change class: Major
Acceptance criteria | Watchdog resets to a defined safe state without corrupting committed log or armed state; faults detected/logged; honest-gap principle upheld.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-FW-014 — Bootloader / Secure-Boot Loader Design
Type | DO-CODE · Discipline: FW/Security (S3) · SKU: All · Subsystem/Component: Bootloader / CMP-035
File # | SCM-FW-COD-014 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/EM-ENG-008
Source format | Source + design · Released format: Reviewed source + signed binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-SEC-004 (secure boot config), DO-FW-015 (OTA), DO-EE-011 (SE)
Input trace | FWR (boot); SYS-032/114; CMP-035; INT-024/035 · Risk-control trace: Only authentic FW executes; never-brick A/B (reconcile with HAL Stage 10) · Verification trace: Secure-boot + rollback test
Critical | Yes (chain of trust + never-brick) · Change class: Major
Acceptance criteria | Verifies image authenticity/integrity against the secure element before execution; resumable staging; fail-safe A/B or rollback so an update cannot brick the device.
Mfg use | Direct (factory boot) · Notes: Standard (secure boot); reconcile with HAL.
DO-FW-015 — OTA Agent Design (download/verify/apply/rollback)
Type | DO-CODE · Discipline: FW/Security (S3/S5) · SKU: All · Subsystem/Component: OTA agent / CMP-043
File # | SCM-FW-COD-015 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/EM-ENG-008
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-FW-014 (bootloader), DO-CLD-010 (OTA service), DO-IF-008 (OTA ICD)
Input trace | FWR (OTA); SYS-032/114; SRS-OTA; CMP-043; INT-035 · Risk-control trace: Reject wrong-model/downgrade; never brick; trip data preserved (reconcile with HAL Stage 10) · Verification trace: OTA campaign + rollback + version-skew test
Critical | Yes (fleet-wide integrity + fail-safety) · Change class: Major
Acceptance criteria | Resumable download over LTE-M/Wi-Fi/USB-C; hands only verified images to bootloader; rejects wrong-model/downgrade; preserves trip data; fail-safe apply with notice of unmonitored window.
Mfg use | Indirect · Notes: Standard; reconcile with HAL.
DO-FW-016 — Security / Crypto Module Design (FW-side)
Type | DO-CODE · Discipline: FW/Security (S3) · SKU: All · Subsystem/Component: Crypto / CMP-044
File # | SCM-FW-COD-016 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-008
Source format | Source + design · Released format: Reviewed source + binary · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-SEC-001…008, DO-EE-011 (SE), DO-IF-005 (mTLS)
Input trace | FWR (security); SYS-026/028/032/110; SRS-SEC; CMP-044; INT-024/032 · Risk-control trace: Record tamper-evidence; authenticated pairing/session; login-free edge read vs governed platform (reconcile with HAL Stage 10) · Verification trace: Security test (auth/tamper)
Critical | Yes (record integrity + identity) · Change class: Major
Acceptance criteria | Uses the secure element for identity/keys; authenticated pairing/sessions; tamper-evidence over records/audit; enforces edge login-free read vs controlled config (AD-5).
Mfg use | Indirect · Notes: Reconcile with HAL + threat model (Evidence Pending).
DO-FW-017 — Firmware Configuration Schema & Build Configuration
Type | DO-CFG · Discipline: FW (S3) · SKU: All (per-SKU config values) · Subsystem/Component: Config/build
File # | SCM-FW-CFG-017 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Schema + build files · Released format: Controlled config + build manifest · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-CFG-006 (regional config), DO-FW-018 (image), DO-GOV-003 (formats)
Input trace | SRS-DEV; SYS-063 (config); FWR · Risk-control trace: Config-integrity (signed) · Verification trace: Config-validation test
Critical | No · Change class: Major
Acceptance criteria | Config schema (sampling/upload cadence, thresholds, region) validated and version-bound; build configuration/toolchain reproducibly produces the released image.
Mfg use | Direct (config at programming) · Notes: Judgment.
DO-FW-018 — Dependency Manifest / Firmware SBOM & Static-Analysis Config
Type | DO-DATA · Discipline: FW/Security (S3) · SKU: All · Subsystem/Component: FW dependencies
File # | SCM-FW-DAT-018 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-008
Source format | SBOM (SPDX/CycloneDX) + config · Released format: Machine-readable SBOM + PDF · Repository: repo/PLM · Parent assembly: DO-FW-001
Related | DO-GOV-007 (SOUP), DO-SEC-011 (dependency controls)
Input trace | IEC 62304 (SOUP); SRS-SEC · Risk-control trace: Dependency-vulnerability control (HAL/security) · Verification trace: SBOM/vuln-scan + static-analysis gate
Critical | No · Change class: Major
Acceptance criteria | Every RTOS/library (SOUP) captured with version; static-analysis config (MISRA subset) defined; vuln-scan gate passes at release.
Mfg use | Indirect · Notes: Standard (IEC 62304).
DO-FW-019 — Released Firmware Image, Hash & Release Notes
Type | DO-BIN · Discipline: FW (S3) · SKU: All · Subsystem/Component: FW image
File # | SCM-FW-BIN-019 · Rev: TBD-rev (version placeholder) · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/EM-ENG-008
Source format | Built binary · Released format: Signed image + digest + release notes · Repository: artifact repo/PLM · Parent assembly: DO-CFG-005
Related | DO-FW-014 (boot), DO-FW-020 (signing), DO-CLD-010 (OTA service)
Input trace | All FWR; DO-FW-001…018 · Risk-control trace: Signed authentic image is the OTA/boot trust root (reconcile with HAL Stage 10) · Verification trace: Signature/digest verification + regression suite
Critical | Yes (the reproducible, authentic firmware) · Change class: Major
Acceptance criteria | Reproducibly built, signed image with recorded digest; release notes list changes, resolved anomalies, and compatibility (DO-CFG-005); regression suite passed.
Mfg use | Direct (programmed at build) · Notes: Judgment; version placeholder until release.
DO-FW-020 — Firmware Programming Package & Signing Procedure
Type | DO-PROC · Discipline: FW/Security (S3) · SKU: All · Subsystem/Component: Programming
File # | SCM-FW-PRC-020 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-008/Mfg
Source format | Procedure + scripts · Released format: PDF + controlled scripts · Repository: repo/PLM · Parent assembly: DO-FW-019
Related | DO-MFG-008 (programming WI), DO-SEC-007 (key provisioning), DO-EE-012 (debug lock)
Input trace | SRS-SEC; SYS-032; IEC 62443 concepts · Risk-control trace: Signing-key protection; production lock (HAL/security) · Verification trace: Programming verification (DO-TST-012)
Critical | No · Change class: Major
Acceptance criteria | Defines the factory programming package (image + config + identity), the signing procedure with key protection, and the post-program debug lock; reproducible and auditable.
Mfg use | Direct · Notes: Standard.
7. Traceability notes
Parents: FWR-001…055; SRS-DEV/DIAG/OTA/SEC/DATA; SYS-001…148; CMP-035…045; INT-016…035. Forward: pairs with DOL-400 (charge-inhibit/governor/drivers), DOL-1100 (crypto/keys/secure boot), DOL-1200 (detection/verdict/ safe-state/charge risk controls), DOL-1400/1500 (programming/test), DOL-2000 (software reconciliation).
8. Release / adequacy notes
Adequate when the signed image (DO-FW-019) builds reproducibly, all safety FW (governor, charge-inhibit, detection, verdict, storage, boot, OTA, crypto, fault handler) is reviewed/tested, and the SBOM/static gates pass. Nine Critical (DO-FW-003/004/005/006/007/008/010/011/013/014/015/016/019 — the safety/function set). Blockers: energy-budget numbers Evidence Pending (D3); detection/verdict/charge/safe-state reconcile with HAL Stage 10; threat model Evidence Pending for DO-FW-016. </content>
DOL-700 — Device Software Outputs (DO-SW) — SCM
Module: DOL-700 · Output prefix: DO-SW · Owner: Embedded Software Engineer (EM-ENG-007) · Discipline: Device-side software / factory & host tooling (S3) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify device-side software outputs that are not the fielded firmware image (DOL-600): the factory programming/provisioning tool, the calibration station software, the host-side diagnostic tool, the factory device-image package, the device-software SBOM, install/recovery images, and build/deploy scripts.
2. Scope
The SCM has no rich on-device application OS — DO-SW covers host/factory tools and the assembled factory device image (firmware + config + identity + calibration). All SKUs. Excludes fielded firmware (DOL-600), cloud (DOL-800), mobile (DOL-900).
3. Applicable standards & references
IEC 62304 (SOUP); reproducible-build practice; secure key-handling for provisioning. Standard/Judgment.
4. Governance & ownership
EM-ENG-007 owns; EM-ENG-008 concurs on provisioning/key tooling; Manufacturing consumes station software.
5. Output categories
Host tool source (DO-CODE), factory image package (DO-BIN), SBOM (DO-DATA), build/deploy scripts (DO-CODE/ DO-CFG), install/recovery image (DO-BIN).
6. Design Output Register (this module)
DO-SW-001 — Factory Programming & Provisioning Tool
Type | DO-CODE · Discipline: Device SW (S3) · SKU: All · Subsystem/Component: Programming station
File # | SCM-SW-COD-001 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-008/Mfg
Source format | Source · Released format: Versioned tool build · Repository: repo/PLM · Parent assembly: DO-MFG-008
Related | DO-FW-020 (programming pkg), DO-SEC-007 (key provisioning), DO-CFG-010 (identity)
Input trace | SRS-SEC; SYS-032/030; IEC 62443 concepts · Risk-control trace: Identity/key provisioning integrity; production debug lock (reconcile with HAL Stage 10) · Verification trace: Programming verification (DO-TST-012)
Critical | Yes (provisions device identity/keys, sets production lock) · Change class: Major
Acceptance criteria | Programs signed firmware + config + unique identity/keys; sets readout-protection/lifecycle lock; refuses to release an incompletely provisioned unit; auditable.
Mfg use | Direct · Notes: Reconcile with HAL; key material handling per DO-SEC-007.
DO-SW-002 — Calibration Station Software
Type | DO-CODE · Discipline: Device SW (S3/S4) · SKU: All · Subsystem/Component: Calibration station
File # | SCM-SW-COD-002 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: Quality
Source format | Source · Released format: Versioned tool build · Repository: repo/PLM · Parent assembly: DO-TST-008
Related | DO-FW-009 (cal schema), DO-TST-008 (cal test), DO-SVC-005 (recal)
Input trace | SYS-003; PRD-PERF; ISO/IEC 17025 concepts · Risk-control trace: Calibration integrity → measurement accuracy (HAL) · Verification trace: Calibration verification (DO-TST-008)
Critical | No · Change class: Major
Acceptance criteria | Drives reference-bath/chamber calibration, computes/writes coefficients per DO-FW-009 with traceable metadata; enforces cal limits and records results.
Mfg use | Direct · Notes: Judgment; reference-standard traceability required.
DO-SW-003 — Host Diagnostic / Service Tool
Type | DO-CODE · Discipline: Device SW (S3) · SKU: All · Subsystem/Component: Service host tool
File # | SCM-SW-COD-003 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: Service/Security
Source format | Source · Released format: Versioned tool build · Repository: repo/PLM · Parent assembly: DO-SVC-003
Related | DO-FW-012 (diagnostics), DO-SVC-003 (troubleshooting)
Input trace | SRS-DIAG; SYS-024/030 · Risk-control trace: Authenticated service access only (HAL/security) · Verification trace: Diagnostic-tool test
Critical | No · Change class: Major
Acceptance criteria | Reads diagnostics/readiness/logs over authenticated USB-C/BLE; cannot alter records or bypass security; supports depot triage (D1 reuse).
Mfg use | Indirect (depot) · Notes: Judgment.
DO-SW-004 — Factory Device-Image Package (firmware+config+identity+cal)
Type | DO-BIN · Discipline: Device SW (S3) · SKU: All (per-SKU config) · Subsystem/Component: Factory image
File # | SCM-SW-BIN-004 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE/EM-ENG-008
Source format | Assembled package · Released format: Signed image bundle + manifest · Repository: artifact repo/PLM · Parent assembly: DO-CFG-005
Related | DO-FW-019 (FW image), DO-FW-017 (config), DO-CFG-006 (regional)
Input trace | DO-FW-019; DO-CFG-006; SYS-063 · Risk-control trace: Correct region/config at build (HAL) · Verification trace: Image-integrity + config-validation
Critical | Yes (defines what is programmed into each unit) · Change class: Major
Acceptance criteria | Deterministic bundle of signed FW + validated config + per-region settings; manifest binds to HW rev and SKU; integrity-checked.
Mfg use | Direct · Notes: Judgment.
DO-SW-005 — Recovery / Re-Provisioning Image & Procedure
Type | DO-BIN · Discipline: Device SW (S3) · SKU: All · Subsystem/Component: Recovery image
File # | SCM-SW-BIN-005 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: Service/Security
Source format | Assembled package · Released format: Signed recovery image · Repository: artifact repo/PLM · Parent assembly: DO-SW-004
Related | DO-SVC-004 (re-provision), DO-FW-015 (OTA)
Input trace | D1 (reuse re-provisioning); SYS-114; SRS-OTA · Risk-control trace: Never-brick recovery (HAL) · Verification trace: Recovery test
Critical | No · Change class: Major
Acceptance criteria | Recovers a device to a known-good signed image without exposing keys; supports depot re-provisioning between trips; preserves or safely clears prior trip data per policy.
Mfg use | Indirect (depot) · Notes: Judgment.
DO-SW-006 — Device Software SBOM
Type | DO-DATA · Discipline: Device SW/Security (S3) · SKU: All · Subsystem/Component: Host tools deps
File # | SCM-SW-DAT-006 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: EM-ENG-008
Source format | SBOM · Released format: Machine-readable SBOM · Repository: repo/PLM · Parent assembly: DO-GOV-007
Related | DO-GOV-007 (SOUP), DO-FW-018
Input trace | IEC 62304; SRS-SEC · Risk-control trace: Tooling dependency-vulnerability control · Verification trace: Vuln-scan gate
Critical | No · Change class: Minor
Acceptance criteria | All host-tool dependencies captured with versions; vuln-scan clean at release.
Mfg use | Indirect · Notes: Standard.
DO-SW-007 — Build / Deploy Scripts (device SW toolchain)
Type | DO-CODE · Discipline: Device SW (S3) · SKU: All · Subsystem/Component: CI/build
File # | SCM-SW-COD-007 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Scripts + config · Released format: Controlled scripts · Repository: repo/PLM · Parent assembly: DO-FW-017
Related | DO-FW-017 (build config), DO-SW-004
Input trace | design-output-standard.md (reproducibility); IEC 62304 · Risk-control trace: Reproducible-build integrity · Verification trace: Reproducible-build check
Critical | No · Change class: Major
Acceptance criteria | CI/build/deploy scripts reproducibly build the factory image and tools from a tagged repo baseline.
Mfg use | Indirect · Notes: Judgment.
DO-SW-008 — Repository Baseline & Version-Tag Definition
Type | DO-CFG · Discipline: Device SW (S3) · SKU: All · Subsystem/Component: Repo baseline
File # | SCM-SW-CFG-008 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-007 · Approval: CSE
Source format | Repo tags/manifest · Released format: Controlled manifest · Repository: repo/PLM · Parent assembly: DO-CFG-011
Related | DO-GOV-001, DO-CFG-005, DO-FW-019
Input trace | ISO 10007; design-output-standard.md · Risk-control trace: — · Verification trace: Configuration audit
Critical | No · Change class: Major
Acceptance criteria | Every released device-software artifact ties to a tagged, immutable repo baseline; reconstructable at audit.
Mfg use | Indirect · Notes: Standard.
7. Traceability notes
Parents: SRS-DEV/DIAG/SEC/OTA; SYS-030/032/063/114; DO-FW-017/019/020; DO-CFG-006. Forward: feeds DOL-1400 (programming/calibration stations), DOL-1500 (calibration/programming verification), DOL-1700 (service tool), DOL-2000 (software reconciliation).
8. Release / adequacy notes
Adequate when the programming/calibration/service tools and the factory image package are Released and station-validated. Two Critical (DO-SW-001 provisioning, DO-SW-004 factory image). Provisioning tool reconciles identity/key handling with HAL/security (Stage 10 / DOL-1100). </content>
DOL-800 — Cloud Platform Design Outputs (DO-CLD) — SCM
Module: DOL-800 · Output prefix: DO-CLD · Owner: Principal Cloud Engineer · Discipline: Cloud Platform (S8) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the cloud-platform design outputs from which the SCM backend can be reproduced: logical/ deployment architecture, service decomposition, telemetry/record/audit schemas, device & enterprise APIs, message contracts, device registry, identity/RBAC, rules/alert engine, OTA service, IaC, monitoring, backup/ recovery, environment config, and deployment pipeline.
2. Scope
S8 cloud services (CMP-046…054). Common platform across SKUs; per-region deployment/data-residency (DO-CFG-006). Excludes device firmware (DOL-600) and mobile (DOL-900).
3. Applicable standards & references
21 CFR Part 11 / ALCOA+ (audit/record integrity); EU GDP (data integrity); ISO/IEC 27001 concepts; OpenAPI; OAuth 2.0; IEC 62443 concepts (segmentation). Standard — revisions to confirm.
4. Governance & ownership
Cloud Engineer owns; EM-ENG-008 concurs on identity/RBAC/audit; EM-ENG-006 concurs on data-residency/Part-11 record store; HAL (Stage 10) reconciles alert-routing/record-integrity risk controls.
5. Output categories
Architecture docs (DO-DOC), data schemas (DO-DATA), API contracts (DO-ICD/DO-DATA), IaC (DO-IAC), pipeline/ config (DO-CFG/DO-IAC), released service versions (DO-BIN/DO-DOC).
6. Design Output Register (this module)
DO-CLD-001 — Cloud Logical & Deployment Architecture
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Platform
File # | SCM-CLD-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Markdown + diagrams · Released format: PDF · Repository: PLM/repo · Parent assembly: DO-SYS-010
Related | DO-CLD-011 (IaC), DO-CFG-006 (residency)
Input trace | PRD-CLOUD; SRS-CLOUD; SYS-177…190; deployment-architecture.md · Risk-control trace: Trust-split; data-residency · Verification trace: Deployment/DR test
Critical | No · Change class: Major
Acceptance criteria | Tenancy, service topology, environments, and per-region deployment defined and consistent with DO-SYS-010/DO-CFG-006.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-002 — Ingest / Store-and-Forward Endpoint Design
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Ingest / CMP-046
File # | SCM-CLD-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-IF-005 (device↔cloud ICD), DO-CLD-003 (schemas), DO-CLD-007 (alerts)
Input trace | SRS-CLOUD; SYS-011/020/111; CMP-046; INT-032 · Risk-control trace: Exactly-once ingest; alert precedence — never lose/duplicate a record (reconcile with HAL Stage 10) · Verification trace: Idempotency + alert-priority test
Critical | Yes (record integrity + alert path at the cloud boundary) · Change class: Major
Acceptance criteria | Exactly-once (de-dup on deviceId,sequence) ingest; resume from last-ack; ack-gated deletion; excursion alerts processed ahead of bulk backlog; malformed/unknown-version rejected/quarantined.
Mfg use | Indirect · Notes: Standard (MQTT v5/ALCOA+); reconcile with HAL.
DO-CLD-003 — Telemetry / Record / Audit Data Schemas
Type | DO-DATA · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: DB/telemetry / CMP-049/054
File # | SCM-CLD-DAT-003 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/EM-ENG-006
Source format | Schema (DDL/JSON) · Released format: Versioned schema + PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-006 (record store), DO-IF-005 (message schema), DO-SEC-010 (audit)
Input trace | SRS-DATA; SYS-011/016/026/028; ALCOA+; 21 CFR Part 11 · Risk-control trace: Attribution/time/gaps/tamper-evidence preserved (reconcile with HAL Stage 10) · Verification trace: Schema + data-integrity test
Critical | Yes (evidentiary record schema) · Change class: Major
Acceptance criteria | Telemetry, trip-record, and append-only audit schemas preserve deviceId/time/sequence/gaps/tamper-evidence; versioned; backward-compatible evolution rules.
Mfg use | Indirect · Notes: Standard (ALCOA+/Part 11).
DO-CLD-004 — Device Registry Service Design
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Registry / CMP-047
File # | SCM-CLD-DOC-004 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/EM-ENG-008
Source format | Markdown + schema · Released format: PDF + schema · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CFG-010 (identity), DO-CLD-010 (OTA), DO-CLD-007 (alert routing)
Input trace | SRS-FLEET; SYS-030/012; CMP-047 · Risk-control trace: Deploy-block on unready device (HAL) · Verification trace: Registry/readiness test
Critical | No · Change class: Major
Acceptance criteria | Authoritative device identity, trip/owner binding, readiness state; drives deploy-block and alert routing; binds serial↔SE identity↔UDI (DO-CFG-010).
Mfg use | Direct (device onboarding at end of line) · Notes: Judgment.
DO-CLD-005 — Fleet Health / Diagnostics Service Design
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Fleet health / CMP-045/050
File # | SCM-CLD-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-FW-012 (device diagnostics), DO-CLD-012 (dashboard)
Input trace | SRS-FLEET; SYS-030/031/113 · Risk-control trace: Unmonitored/overdue-fleet detection (HAL) · Verification trace: Fleet-health test
Critical | No · Change class: Major
Acceptance criteria | Aggregates readiness/last-seen/RSSI/backlog; flags overdue/stray devices; supplies RF-in-container evidence.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-006 — Record Store & Cloud Verdict-Recompute Design
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Record store / CMP-049
File # | SCM-CLD-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/EM-ENG-006
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-FW-008 (device verdict), DO-CLD-003 (schema)
Input trace | SRS-CLOUD; SYS-016/026; CMP-049; ALCOA+ · Risk-control trace: Device↔cloud verdict consistency (reconcile with HAL Stage 10) · Verification trace: Verdict-consistency test
Critical | Yes (disposition-record integrity + consistency) · Change class: Major
Acceptance criteria | Stores complete tamper-evident trip record; recomputes cloud verdict consistent with the device verdict (DO-FW-008); enduring/consistent per ALCOA+.
Mfg use | Indirect · Notes: Standard; reconcile with HAL.
DO-CLD-007 — Rules / Alert-Routing & Escalation Engine Design
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Rules/alert / CMP-048
File # | SCM-CLD-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Markdown + config schema · Released format: PDF + schema · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-002 (ingest), DO-CLD-008 (notifications), DO-SYS-012 (deps)
Input trace | SRS-FLEET; SYS-012/014; CMP-048 · Risk-control trace: Delayed/missing/mis-routed excursion alert (reconcile with HAL Stage 10) · Verification trace: Alert-routing + escalation test
Critical | Yes (alert delivery to responders) · Change class: Major
Acceptance criteria | Routes alerts to configured responders; escalates unacknowledged alerts per policy; consolidates cloud-side alarms; no silent drop of an alert.
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-CLD-008 — Notification Service Design (email/SMS/webhook)
Type | DO-DOC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: Notifications
File # | SCM-CLD-DOC-008 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-007 (rules), DO-IF-009 (webhook contract)
Input trace | SRS-FLEET; SYS-012; PRD-CLOUD · Risk-control trace: Delivery-failure fallback (HAL) · Verification trace: Notification delivery/retry test
Critical | No · Change class: Major
Acceptance criteria | Multi-channel notification with retry and delivery status; signed retryable webhooks; failure falls back to alternate channel/escalation.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-009 — Device & Enterprise API Contracts (OpenAPI)
Type | DO-ICD · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: API gateway / CMP-052
File # | SCM-CLD-ICD-009 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/EM-ENG-008
Source format | OpenAPI (YAML) · Released format: Versioned OpenAPI + PDF · Repository: repo/PLM · Parent assembly: DO-IF-007
Related | DO-IF-007 (enterprise ICD), DO-CLD-013 (RBAC), DO-CLD-003 (schema)
Input trace | SRS-CLOUD; SYS-027/064/068; CMP-052; INT-034 · Risk-control trace: RBAC/auth on every resource; export preserves record integrity (reconcile with HAL Stage 10) · Verification trace: API contract + auth test
Critical | Yes (auth/RBAC + record export integrity) · Change class: Major
Acceptance criteria | OpenAPI-defined REST APIs (trip records, alerts, fleet/device) with OAuth 2.0/SSO, RBAC, pagination, rate limits, versioning; export in JSON/PDF/CSV preserving attribution/time/gaps/tamper-evidence.
Mfg use | Indirect · Notes: Standard (OpenAPI/OAuth 2.0/EU GDP/Part 11).
DO-CLD-010 — OTA Service Design (signing & campaign distribution)
Type | DO-DOC · Discipline: Cloud/Security (S8) · SKU: All · Subsystem/Component: OTA service / CMP-051
File # | SCM-CLD-DOC-010 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: EM-ENG-008/CSE
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-FW-015 (OTA agent), DO-IF-008 (OTA ICD), DO-SEC-005 (signing)
Input trace | SRS-OTA; SYS-032/114; CMP-051; INT-035 · Risk-control trace: Never distribute unverifiable/fleet-bricking image (reconcile with HAL Stage 10) · Verification trace: Staged-campaign + rollback test
Critical | Yes (fleet-wide integrity) · Change class: Major
Acceptance criteria | Signs/distributes firmware campaigns; staged/resumable delivery; per-device status; never pushes an unverifiable or fleet-bricking image; honors compatibility matrix (DO-CFG-005).
Mfg use | Indirect · Notes: Reconcile with HAL.
DO-CLD-011 — Infrastructure-as-Code (IaC) Definitions
Type | DO-IAC · Discipline: Cloud (S8) · SKU: All (per-region stacks) · Subsystem/Component: Cloud infra
File # | SCM-CLD-IAC-011 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | IaC (declarative) · Released format: Versioned IaC · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-014 (pipeline), DO-CFG-006 (residency), DO-CLD-016 (versions)
Input trace | SRS-CLOUD; SYS-177…190; deployment-architecture.md · Risk-control trace: Reproducible/segmented environments (security) · Verification trace: Infra provision + DR test
Critical | No · Change class: Major
Acceptance criteria | All environments provisioned reproducibly from versioned IaC; per-region residency enforced; network segmentation per security architecture.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-012 — Fleet Dashboard & Reporting UI Design
Type | DO-DOC · Discipline: Cloud/Web (S8) · SKU: All · Subsystem/Component: Dashboard / CMP-050
File # | SCM-CLD-DOC-012 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/HF
Source format | Markdown + UI spec · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-CLD-005 (health), DO-HF-007 (dashboard HF), DO-CLD-009 (API)
Input trace | SRS-FLEET; SYS-031/030; PRD-CLOUD · Risk-control trace: Clear release/acceptance reporting (HF/HAL) · Verification trace: Usability + reporting test
Critical | No · Change class: Major
Acceptance criteria | Fleet-wide status (location, readiness, trips, alerts, last-seen); flags overdue/stray; reporting surface for release/acceptance decisions.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-013 — Identity, RBAC & SSO Configuration Design
Type | DO-CFG · Discipline: Cloud/Security (S8) · SKU: All · Subsystem/Component: RBAC/SSO / CMP-053
File # | SCM-CLD-CFG-013 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: EM-ENG-008/CSE
Source format | Config + design · Released format: Versioned config + PDF · Repository: repo/PLM · Parent assembly: DO-CLD-001
Related | DO-SEC-002 (authZ), DO-CLD-009 (API)
Input trace | SRS-SEC; SYS-110; PRD-SEC; CMP-053; INT-034 · Risk-control trace: Trust-split: login-free edge read vs governed platform (reconcile with HAL Stage 10) · Verification trace: RBAC/SSO test
Critical | No · Change class: Major
Acceptance criteria | Role-based access + SSO/IdP integration; least-privilege scopes; enforces the trust-boundary split (does not gate the login-free on-device edge verdict, AD-5).
Mfg use | Indirect · Notes: Standard.
DO-CLD-014 — Deployment Pipeline, Monitoring & Backup/Recovery Design
Type | DO-IAC · Discipline: Cloud (S8) · SKU: All · Subsystem/Component: CI/CD + observability
File # | SCM-CLD-IAC-014 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE
Source format | Pipeline-as-code + config · Released format: Versioned pipeline · Repository: repo/PLM · Parent assembly: DO-CLD-011
Related | DO-CLD-011 (IaC), DO-CLD-016 (versions), DO-SEC-010 (audit)
Input trace | SRS-CLOUD; SYS-163…176 (reliability); deployment-architecture.md · Risk-control trace: DR/rollback; monitoring/alerting on platform faults (HAL) · Verification trace: Pipeline + DR + monitoring test
Critical | No · Change class: Major
Acceptance criteria | CI/CD with staged rollout + rollback; monitoring/alerting/logging; backup and tested recovery (RPO/RTO defined); audit trail of deployments.
Mfg use | Indirect · Notes: Judgment.
DO-CLD-015 — Cloud SBOM & Dependency Controls
Type | DO-DATA · Discipline: Cloud/Security (S8) · SKU: All · Subsystem/Component: Cloud deps
File # | SCM-CLD-DAT-015 · Rev: TBD-rev · Status: Draft · Owner: EM-ENG-008 · Approval: EM-ENG-008
Source format | SBOM · Released format: Machine-readable SBOM · Repository: repo/PLM · Parent assembly: DO-GOV-007
Related | DO-GOV-007 (SOUP), DO-SEC-011
Input trace | SRS-SEC; IEC 62304 concepts · Risk-control trace: Dependency-vulnerability control · Verification trace: Vuln-scan gate
Critical | No · Change class: Minor
Acceptance criteria | Cloud service dependencies captured with versions; vuln-scan clean at release.
Mfg use | Indirect · Notes: Standard.
DO-CLD-016 — Released Cloud Service Versions & Environment Config
Type | DO-CFG · Discipline: Cloud (S8) · SKU: All (per-region) · Subsystem/Component: Service versions
File # | SCM-CLD-CFG-016 · Rev: TBD-rev · Status: Draft · Owner: Cloud Eng · Approval: CSE/EM-ENG-006
Source format | Version manifest + config · Released format: Controlled manifest · Repository: repo/PLM · Parent assembly: DO-CFG-005
Related | DO-CFG-005 (compatibility), DO-CLD-011/014
Input trace | SYS-063/065; SRS-CLOUD · Risk-control trace: Version-skew interoperability (HAL) · Verification trace: Version-skew interop test
Critical | No · Change class: Major
Acceptance criteria | Each released service version and environment config recorded and bound to the compatibility matrix (DO-CFG-005); reproducible per region.
Mfg use | Indirect · Notes: Judgment.
7. Traceability notes
Parents: PRD-CLOUD/SEC; SRS-CLOUD/DATA/FLEET/OTA/SEC; SYS-011…032/110/177…190; CMP-046…054; INT-032/034/035. Forward: pairs with DOL-1000 (device↔cloud, enterprise, OTA ICDs), DOL-1100 (identity/RBAC/audit), DOL-1200 (alert/record risk controls), DOL-1800 (data residency), DOL-2000 (software reconciliation).
8. Release / adequacy notes
Adequate when architecture, schemas, APIs, IaC, and released service versions are Released and DR-tested. Four Critical (DO-CLD-002 ingest, DO-CLD-003 record schema, DO-CLD-006 record store/recompute, DO-CLD-007 alert engine, DO-CLD-009 API auth, DO-CLD-010 OTA — the record/alert/OTA-integrity set; six flagged Critical). Alert-routing/record-integrity outputs reconcile with HAL Stage 10. </content>
DOL-900 — Mobile & Web Application Design Outputs (DO-APP) — SCM
Module: DOL-900 · Output prefix: DO-APP · Owner: Principal Mobile Engineer · Discipline: Mobile Application (S9) · Status: Draft baseline (ORCH-013).
1. Purpose
Identify and govern the mobile-app design outputs: architecture, navigation, screen/interaction specs, design system, BLE-GATT client, provisioning/arming flows, offline/sync model, notifications, permissions, error handling, analytics, and store/signing config — from which the SCM companion app can be reproduced.
2. Scope
S9 mobile app (CMP-055): provisioning, trip-profile config, arming aid/pre-arm guidance, local login-free record read, notifications. All SKUs. Excludes cloud (DOL-800) and device firmware (DOL-600).
3. Applicable standards & references
Bluetooth GATT; platform app-store guidelines; WCAG (accessibility); OWASP MASVS concepts (mobile security); 21 CFR Part 11 (attributable actions where app writes config/arm). Standard/Judgment.
4. Governance & ownership
Mobile Engineer owns; EM-ENG-008 concurs on BLE pairing/security; HF (DOL-1300) concurs on arming/provisioning UX; HAL (Stage 10) reconciles arming/pre-arm use-error controls.
5. Output categories
Architecture/spec docs (DO-DOC), screen/interaction specs (DO-DOC), GATT client (DO-CODE), build/signing config (DO-CFG), store packages (DO-BIN), version-compatibility matrix (DO-DATA).
6. Design Output Register (this module)
DO-APP-001 — Mobile App Architecture & Navigation Design
Type | DO-DOC · Discipline: Mobile (S9) · SKU: All · Subsystem/Component: App / CMP-055
File # | SCM-APP-DOC-001 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: CSE
Source format | Markdown + diagrams · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-SYS-010
Related | DO-APP-002/003, DO-IF-004 (GATT ICD)
Input trace | PRD-MOB; SRS-MOB; SYS-022…028; CMP-055 · Risk-control trace: — · Verification trace: Architecture review
Critical | No · Change class: Major
Acceptance criteria | App architecture, navigation map, and state management defined; consistent with the GATT/cloud integration.
Mfg use | Indirect · Notes: Judgment.
DO-APP-002 — Screen & Interaction Specifications
Type | DO-DOC · Discipline: Mobile/HF (S9) · SKU: All · Subsystem/Component: App UI
File # | SCM-APP-DOC-002 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: HF/CSE
Source format | Design files + spec · Released format: PDF + asset export · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-APP-004 (design system), DO-HF-006 (provisioning UX)
Input trace | PRD-MOB/UX; SRS-MOB; SYS-025 · Risk-control trace: Arming/pre-arm use-error mitigation (HAL/HF) · Verification trace: Usability test
Critical | No · Change class: Major
Acceptance criteria | All screens (provisioning, trip config, arming, record read, notifications) fully specified with states/errors; consistent with HF outputs.
Mfg use | Indirect · Notes: Judgment.
DO-APP-003 — BLE-GATT Client Implementation Design
Type | DO-CODE · Discipline: Mobile/Security (S9) · SKU: All · Subsystem/Component: GATT client / CMP-055
File # | SCM-APP-COD-003 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: EM-ENG-008/CSE
Source format | Source · Released format: Reviewed source + built app · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-IF-004 (GATT ICD), DO-SEC-003 (pairing)
Input trace | SRS-MOB; SYS-022/023/060/061/068; IFR-015…020; INT-029/033 · Risk-control trace: Authenticated arm/profile write; anti-replay; login-free record read (reconcile with HAL Stage 10) · Verification trace: GATT contract + security test
Critical | Yes (authenticated arm/config path) · Change class: Major
Acceptance criteria | Implements the versioned GATT service (DO-IF-004): LE Secure Connections (MITM-protected, not Just Works); atomic validated trip-profile write; anti-replay arm/disarm; ordered resumable record-read stream; rejects unknown handles/values.
Mfg use | Indirect · Notes: Reconcile with HAL + threat model.
DO-APP-004 — Design System & UI Asset Library
Type | DO-DOC · Discipline: Mobile/HF (S9) · SKU: All · Subsystem/Component: App assets
File # | SCM-APP-DOC-004 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: HF
Source format | Design files · Released format: Versioned asset library · Repository: repo/PLM · Parent assembly: DO-APP-002
Related | DO-HF-003 (iconography), DO-APP-002
Input trace | PRD-UX; SRS-MOB; WCAG · Risk-control trace: Consistent alarm/warning presentation (HF/HAL) · Verification trace: Design-system audit
Critical | No · Change class: Minor
Acceptance criteria | Components, typography, color, and iconography defined with accessibility (contrast/target size) and localization support.
Mfg use | Indirect · Notes: Standard (WCAG).
DO-APP-005 — Offline & Sync Model Design
Type | DO-DOC · Discipline: Mobile (S9) · SKU: All · Subsystem/Component: App offline/sync
File # | SCM-APP-DOC-005 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: CSE
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-APP-003 (GATT), DO-CLD-009 (API)
Input trace | SRS-MOB; SYS-024; D2 (offline-first) · Risk-control trace: Local record read works offline (HAL) · Verification trace: Offline/sync test
Critical | No · Change class: Major
Acceptance criteria | Local record read works without connectivity (login-free edge, AD-5); config/arming sync reconciles safely when back online; no lost/duplicated actions.
Mfg use | Indirect · Notes: Judgment.
DO-APP-006 — Notifications & Permissions Design
Type | DO-DOC · Discipline: Mobile (S9) · SKU: All · Subsystem/Component: App notifications
File # | SCM-APP-DOC-006 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: CSE/Security
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-CLD-008 (notifications), DO-APP-002
Input trace | SRS-MOB; SYS-012; PRD-MOB · Risk-control trace: Local alert visibility (HAL) · Verification trace: Notification/permission test
Critical | No · Change class: Major
Acceptance criteria | Push/local notifications for alerts; least-privilege permission requests (BLE/location) with rationale; graceful degradation if permissions denied.
Mfg use | Indirect · Notes: Judgment.
DO-APP-007 — Error Handling, Analytics & Crash Reporting Design
Type | DO-DOC · Discipline: Mobile (S9) · SKU: All · Subsystem/Component: App observability
File # | SCM-APP-DOC-007 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: CSE/Security
Source format | Markdown + design · Released format: PDF · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-APP-006, DO-SEC-011 (privacy)
Input trace | SRS-MOB; PRD-QUAL · Risk-control trace: No PHI/record leakage in analytics (privacy) · Verification trace: Privacy/analytics review
Critical | No · Change class: Minor
Acceptance criteria | Defined error taxonomy/handling; analytics events and crash reporting with no sensitive-data leakage; consent-respecting.
Mfg use | Indirect · Notes: Judgment.
DO-APP-008 — Build & Signing Configuration
Type | DO-CFG · Discipline: Mobile/Security (S9) · SKU: All · Subsystem/Component: App build
File # | SCM-APP-CFG-008 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: EM-ENG-008
Source format | Build config + scripts · Released format: Controlled config · Repository: repo/PLM · Parent assembly: DO-APP-001
Related | DO-APP-009 (store pkg), DO-SEC-005 (signing)
Input trace | SRS-MOB; SRS-SEC · Risk-control trace: App-signing integrity · Verification trace: Reproducible-build + signing check
Critical | No · Change class: Major
Acceptance criteria | Reproducible build with protected signing keys; obfuscation/hardening per MASVS concepts; ties to a tagged repo baseline.
Mfg use | Indirect · Notes: Judgment.
DO-APP-009 — Store Packages & Version-Compatibility Matrix
Type | DO-BIN · Discipline: Mobile (S9) · SKU: All · Subsystem/Component: Released app
File # | SCM-APP-BIN-009 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: CSE
Source format | Built app packages · Released format: Signed store packages + matrix · Repository: artifact repo/PLM · Parent assembly: DO-CFG-005
Related | DO-CFG-005 (compatibility), DO-APP-008
Input trace | SRS-MOB; SYS-063/065 · Risk-control trace: App↔device↔cloud version skew (HAL) · Verification trace: Version-compat test
Critical | Yes (app version compatibility across fielded devices/cloud) · Change class: Major
Acceptance criteria | Signed iOS/Android packages; version-compatibility matrix defines supported app↔GATT↔cloud versions; incompatible combos blocked with a clear message.
Mfg use | Indirect · Notes: Judgment.
DO-APP-010 — Mobile SBOM & Dependency Manifest
Type | DO-DATA · Discipline: Mobile/Security (S9) · SKU: All · Subsystem/Component: App deps
File # | SCM-APP-DAT-010 · Rev: TBD-rev · Status: Draft · Owner: Mobile Eng · Approval: EM-ENG-008
Source format | SBOM · Released format: Machine-readable SBOM · Repository: repo/PLM · Parent assembly: DO-GOV-007
Related | DO-GOV-007 (SOUP), DO-SEC-011
Input trace | SRS-SEC; IEC 62304 concepts · Risk-control trace: Dependency-vulnerability control · Verification trace: Vuln-scan gate
Critical | No · Change class: Minor
Acceptance criteria | All app dependencies captured with versions; vuln-scan clean at release.
Mfg use | Indirect · Notes: Standard.
7. Traceability notes
Parents: PRD-MOB/UX/SEC; SRS-MOB/SEC; SYS-022…028/060/061/068/110; CMP-055; IFR-015…020; INT-029/033. Forward: pairs with DOL-1000 (GATT ICD), DOL-1100 (pairing/security), DOL-1300 (provisioning/arming UX), DOL-2000 (software reconciliation).
8. Release / adequacy notes
Adequate when architecture, screen specs, GATT client, and signed store packages are Released and usability/security-tested. Two Critical (DO-APP-003 GATT authenticated arm/config, DO-APP-009 version compatibility). Arming/pre-arm UX and pairing reconcile with HAL/threat model (Stage 10). </content>
Design Input–Output Traceability Matrix — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 09 — Design Outputs (ORCH-013) · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Status: Draft baseline (pre-DR-DO).
This matrix threads each key / Critical design output along all applicable chains required by the DOL
standard: Requirement (PRD → SYS → SRS/HWR/MER/FWR/IFR) → Design Output → Risk-Control
(Hazard → RCR, reconcile with HAL Stage 10) → Verification (PV-<CODE>-NNN, Stage 13, ORCH-014) →
Validation (VAL-<CODE>-NNN, Stage 14, ORCH-015) → Manufacturing (process PROC-### → control
CP-### → production test DO-TST/PV). It is a representative matrix over the severity-dominant and
release-gating outputs; the full 272-output inventory is in Design-Output-Register.md, and every per-module
DOL-.md entry carries its own Input/Risk/Verification traceability fields.
Evidence & honesty. PV-… and VAL-… IDs are forward placeholders (Evidence Pending) — the Verification (Stage 13) and Validation (Stage 14) libraries are not yet authored; where a concrete test is already named in the requirement baseline it is shown (e.g., PRD-VER-004 90-day cold endurance, PRD-VER-005 RF-in-container). HA-###/RCR-### are Preliminary placeholders to be reconciled when the HAL (Stage 10) is baselined. PROC-###/CP-### are Stage-15 manufacturing placeholders. No revision, part number, or numeric limit is invented.
Legend. Requirement codes per Stage 05/06/07 (PRD-<CODE>-NNN, SYS-###, SRS-<CODE>-NNN, HWR/MER/FWR/IFR-###). Risk column marked "↔HAL" = reconcile with HAL Stage 10. Mfg column reads process → control → acceptance test.
---
1. D3 charge / energy safety spine (top program risk)
Output ID | Requirement trace (PRD/SYS/SRS) | Risk-control (↔HAL) | Verification (PV-…, Stage 13) | Validation (VAL-…, Stage 14) | Mfg: process → control → test
DO-EE-003 Low-temp cell spec | D3; PRD-ELEC-002; SYS-087/088; HWR-001 | HA-BATT-### / RCR-### cold-capacity & cold-charge chemistry safety ↔HAL | PRD-VER-004 cold-capacity + cycle-life; PV-ELEC-### | VAL-REL-### (90-day field endurance) | Incoming cell acceptance (PROC-INCOMING) → CP-cell-lot → DO-TST-001
DO-EE-005 BMS / protection thresholds | PRD-SAFE; SYS-096; HWR-003; IEC 62368-1 | HA-BATT-### / RCR-### OV/UV/OC/OT/short, MCU-independent ↔HAL | PV-SAFE-### protection fault-injection | VAL-SAFE-### | PCBA assy (PROC-SMT) → CP-protection-net → DO-TST-004/009
DO-EE-007 Charge-inhibit <0 °C interlock | D3; PRD-SAFE; SYS-094/095; HWR-006; INT-016; ADR-004 | HA-BATT-### / RCR-### cold-charge cell damage — primary control ↔HAL | PV-SAFE-### cold-charge-inhibit (cell-cold fault-injection) | VAL-SAFE-### | Final assy (PROC-FINAL) → CP-cell-NTC → DO-TST-009
DO-EE-010 Secondary fuse / OVP backstop | PRD-SAFE; SYS-096; HWR-012; IEC 62368-1 | HA-BATT-### / RCR-### single-fault backstop ↔HAL | PV-SAFE-### secondary-protection fault-injection | VAL-SAFE-### | PCBA (PROC-SMT) → CP-backstop → DO-TST-009
DO-EE-009 USB-C sink + input protection | PRD-SAFE; SYS-102/114; HWR-039/045; INT-008/009 | HA-CHG-### / RCR-### out-of-spec source / reverse / OVP / short ↔HAL | PV-SAFE-### reverse/OVP/short + PD negotiation | VAL-SAFE-### | PCBA (PROC-SMT) → CP-input-prot → DO-TST-010
DO-FW-004 Energy-governor | D3; PRD-REL; SYS-086/097; FWR (governor) | HA-PWR-### / RCR-### silent loss-of-monitoring ↔HAL | PRD-VER-004 90-day cold endurance; PV-PERF-### | VAL-REL-### field endurance | FW programming (PROC-FWPROG) → CP-image-hash → DO-TST-008
DO-FW-005 FW charge-inhibit & self-heat | D3; PRD-SAFE; SYS-094; FWR | HA-BATT-### / RCR-### FW defense-in-depth for cold-charge ↔HAL | PV-SAFE-### FW-inhibit fault-injection | VAL-SAFE-### | FW programming → CP-image-hash → DO-TST-008/009
DO-SAFE-001 Charge-inhibit (impl. ref) | D3; SYS-094/095; PRD-SAFE; ADR-004 | HA-BATT-### / RCR-### — implements DO-EE-007+DO-FW-005 ↔HAL | PV-SAFE-### cold-charge-inhibit | VAL-SAFE-### | (ref) → DO-TST-009
DO-SAFE-012 Energy-reserve / loss-of-monitoring | D3; SYS-086/097; PRD-SAFE/REL | HA-PWR-### / RCR-### energy exhaustion mid-trip ↔HAL | PRD-VER-004 cold endurance | VAL-REL-### | — → CP-SoC-readiness → DO-TST-005
DO-TST-009 Protection / charge-inhibit safety test | PRD-SAFE; SYS-094/096; IEC 62368-1 | verifies DO-SAFE-001/002 ↔HAL | (is the production risk-control test) | VAL-SAFE-### | Final assy → CP-safety-gate → DO-TST-009 (100%/sampled TBD)
---
2. Excursion detection · verdict · record integrity (essential performance)
Output ID | Requirement trace | Risk-control (↔HAL) | Verification (PV-…) | Validation (VAL-…) | Mfg: process → control → test
DO-FW-006 Excursion-detection engine | PRD-FUN/SAFE; SYS-008/009/010/014; ADR-005; SRS-DEV | HA-DET-### / RCR-### missed excursion → wrong disposition ↔HAL | PV-FUN-### profiled two-sided excursions | VAL-USE-### (clinician disposition) | FW prog → CP-image-hash → DO-TST-005
DO-EE-013 Probe front-end + presence-detect | PRD-FUN/SAFE; SYS-003/007/024; HWR-022/028; INT-018 | HA-SENS-### / RCR-### open/absent probe → false accept ↔HAL | PV-FUN-### probe-fault injection | VAL-USE-### | Final assy → CP-probe-presence → DO-TST-005/006
DO-FW-008 Verdict engine (accept/reject/quarantine) | PRD-FUN/SAFE; SYS-015/016/017; ADR-006; SRS-DEV | HA-VER-### / RCR-### false accept on gaps/faults ↔HAL | PV-FUN-### verdict-consistency + gap-handling | VAL-USE-### at-the-door decision | FW prog → CP-image-hash → DO-TST-005
DO-FW-010 Storage manager (power-fail-safe log) | PRD-FUN; SYS-018/021/028/099; ALCOA+; SRS-DATA | HA-DATA-### / RCR-### torn/lost/fabricated record ↔HAL | PV-DATA-### power-fail-during-write | VAL-REG-### data-integrity audit | FW prog → CP-image-hash → DO-TST-005
DO-EE-002 Power tree / write-safe rails | PRD-ELEC; SYS-099; HWR-043; INT-012…015 | HA-DATA-### / RCR-### rail-gating corrupts NVM write ↔HAL | PV-ELEC-### rail-sequencing | VAL-REL-### | PCBA (PROC-SMT) → CP-rail-seq → DO-TST-004
DO-CLD-003 Telemetry / record / audit schema | PRD-CLOUD/REG; SYS-119…131; 21 CFR Part 11 / ALCOA+; SRS-DATA | HA-DATA-### / RCR-### evidentiary-record integrity ↔HAL | PV-CLOUD-### schema + audit-trail | VAL-REG-### GDP/Part 11 | Cloud release (PROC-CLDREL) → CP-version → DO-CLD-016
DO-CLD-006 Record store + cloud verdict-recompute | PRD-CLOUD/SAFE; SYS-015/016; SRS-CLOUD | HA-VER-### / RCR-### device↔cloud verdict inconsistency ↔HAL | PV-CLOUD-### recompute-consistency | VAL-USE-### | Cloud release → CP-version → PV-CLOUD-###
DO-SAFE-004/005/006 Excursion/probe/verdict controls | PRD-SAFE; SYS-007…017/024 | HA-DET/SENS/VER-### — impl. refs ↔HAL | PV-SAFE-### detection/probe/verdict | VAL-USE-### | (refs) → DO-TST-005/006
DO-TST-006 Sensor calibration test | PRD-PERF (±0.5 °C); SYS-001/003/004; ISO/IEC 17025 | HA-SENS-### / RCR-### mis-calibration → wrong verdict ↔HAL | PV-PERF-### accuracy vs. traceable reference | VAL-REL-### drift over reuse | Calibration (PROC-CAL) → CP-cal-residual → DO-TST-006
---
3. Alert delivery · connectivity · interfaces
Output ID | Requirement trace | Risk-control (↔HAL) | Verification (PV-…) | Validation (VAL-…) | Mfg: process → control → test
DO-FW-011 Comms agent / bearer manager | D2; PRD-CONN; SYS-011/012/013/014; SRS-CONN | HA-COMM-### / RCR-### delayed/missing alert ↔HAL | PV-CONN-### alert-latency + store-and-forward | VAL-USE-### time-to-act | FW prog → CP-image-hash → DO-TST-007
DO-CLD-007 Rules / alert-routing & escalation | PRD-CLOUD/SAFE; SYS-011…014; SRS-FLEET | HA-COMM-### / RCR-### alert not delivered to responder ↔HAL | PV-CLOUD-### rules/escalation | VAL-USE-### | Cloud release → CP-version → PV-CLOUD-###
DO-IF-001 Internal electrical/bus/connector ICD | PRD-ELEC; SYS (interfaces); INT-006…024 | HA-### / RCR-### carries charge-inhibit/probe-presence/record-write contracts ↔HAL | PV-IF-### interface-conformance | — | PCBA (PROC-SMT) → CP-connectivity → DO-TST-004
DO-IF-005 Device↔cloud store-and-forward ICD | PRD-CONN/CLOUD; SYS; INT-032; SRS-CONN | HA-DATA-### / RCR-### record/alert transport integrity ↔HAL | PV-IF-### protocol-conformance | VAL-REG-### | — → CP-version-negotiation → PV-IF-###
DO-IF-007 OTA update interface ICD | PRD-SVC; SYS-032/114; INT-035; SRS-OTA | HA-OTA-### / RCR-### bricked/unmonitored fleet ↔HAL | PV-OTA-### rollback + version-skew | VAL-REL-### | — → CP-signature → PV-OTA-###
DO-SAFE-009 Excursion-alert delivery control | D2; SYS-011…014; PRD-SAFE | HA-COMM-### / RCR-### — impl. ref ↔HAL | PV-CONN-### alert-latency | VAL-USE-### | (ref) → DO-TST-007
---
4. Security · identity · record authenticity
Output ID | Requirement trace | Risk-control (↔HAL) | Verification (PV-…) | Validation (VAL-…) | Mfg: process → control → test
DO-SEC-004 Secure-boot & signed-update config | PRD-SEC; SYS-132-line; SRS-SEC; IEC 62443 | HA-SW-### / RCR-### unauthentic code ↔HAL | PV-SEC-### secure-boot / signed-update | VAL-REG-### | FW prog → CP-signing → DO-TST-008
DO-SEC-006 Credential/record tamper-evidence | PRD-SEC/REG; SYS-018/028; ALCOA+; SRS-SEC | HA-DATA-### / RCR-### record tampering ↔HAL | PV-SEC-### tamper-evidence | VAL-REG-### data-integrity | FW prog → CP-image-hash → DO-TST-008
DO-SEC-007 Production key/identity provisioning | PRD-SEC; SYS (identity); SRS-SEC | HA-### / RCR-### weak/duplicated identity ↔HAL | PV-SEC-### provisioning | — | Provisioning (PROC-PROV) → CP-identity → DO-TST-008
DO-SEC-008 Debug-lock & anti-tamper (fielded) | PRD-SEC; IFR-005; SRS-SEC | HA-### / RCR-### key/record readout on fielded unit ↔HAL | PV-SEC-### debug-lock | — | Provisioning → CP-lock-fuse → DO-TST-012
DO-CFG-010 Serial-number & device-identity | PRD-MFG; SYS (identity); DO-SEC-007 | HA-DATA-### / RCR-### record mis-attribution ↔HAL | PV-### identity read-back | VAL-REG-### traceability | Serialization (PROC-SER) → CP-uid-uniqueness → DO-LBL-003/DO-TST-008
DO-FW-019 Released firmware image + hash | PRD-QUAL/REG; SYS (SW config); SRS-DEV | HA-### / RCR-### wrong/unauthentic image ↔HAL | PV-### image-hash/version | VAL-### | FW prog → CP-hash-verify → DO-TST-008
---
5. Ingress / seal / battery-mechanical safety
Output ID | Requirement trace | Risk-control (↔HAL) | Verification (PV-…) | Validation (VAL-…) | Mfg: process → control → test
DO-MECH-005 Enclosure seal / gasket | PRD-MECH; SYS-069; IEC 60529; MER | HA-ENV-### / RCR-### condensation/ingress → record-integrity fault ↔HAL | PV-ENV-### IP + cold-cycle | VAL-REL-### reuse-life seal | Sealing (PROC-SEAL) → CP-seal-torque → DO-TST-011
DO-MECH-006 External probe assembly | PRD-FUN/MECH; SYS-003/007; MER; IFR-003 | HA-SENS-### / RCR-### probe measures wrong zone ↔HAL | PV-FUN-### probe accuracy at tip | VAL-USE-### | Probe assy (PROC-PROBE) → CP-tip-integrity → DO-TST-003/006
DO-MECH-012 Battery retention / isolation | PRD-MECH/SAFE; SYS-090; UN 38.3; MER | HA-BATT-### / RCR-### cell crush/puncture under shock ↔HAL | PV-MECH-### drop/vibe + UN 38.3 | VAL-REL-### transport | Final assy → CP-retention → DO-TST-003
DO-PCB-010 Conformal-coat & cleanliness | PRD-REL/MECH; SYS-069; SRS-REL | HA-ENV-### / RCR-### condensation leakage/corrosion ↔HAL | PV-ENV-### cold-cycle + cleanliness | VAL-REL-### | Coat (PROC-COAT) → CP-coverage/cleanliness → DO-TST-003
DO-SAFE-010/011 Battery isolation / ingress controls | PRD-SAFE; SYS-069/090 | HA-BATT/ENV-### — impl. refs ↔HAL | PV-MECH/ENV-### | VAL-REL-### | (refs) → DO-TST-011
DO-TST-011 Seal / leak (IP) test | PRD-MECH/SAFE; SYS-069; IEC 60529 | verifies DO-MECH-005 / DO-SAFE-011 ↔HAL | (is the production risk-control test) | VAL-REL-### | Sealing → CP-leak-rate → DO-TST-011
---
6. Compliance · verification-enabling · release gates
Output ID | Requirement trace | Risk-control (↔HAL) | Verification (PV-…) | Validation (VAL-…) | Mfg: process → control → test
DO-REG-001 Standards-applicability matrix | D4; PRD-REG; SYS-177…190; SRS-REG | compliance backbone (safety-standard tie-in ↔HAL) | PV-REG-### certification plan | VAL-REG-### | — → CP-config-per-market → DO-TST-007/010
DO-REG-005 Essential-performance / safety class | D4; PRD-REG/SAFE; WHO PQS; IEC 62368-1 | essential-performance & classification ↔HAL | PV-REG-### essential-performance | VAL-REG-### WHO PQS / GDP | — → CP-EP-verify → PV-REG-###
DO-REG-007 Battery-transport compliance | D4; PRD-REG; UN 38.3 / IATA | HA-BATT-### transport safety ↔HAL | UN 38.3 test summary | VAL-REG-### DGR | — → CP-UN38.3-lot → DO-LBL-005
DO-REG-011 Compliance-critical SW version register | PRD-REG; SYS (SW cfg); ALCOA+ | wrong version invalidates compliance/record integrity ↔HAL | PV-### version-vs-hash | VAL-REG-### | FW/cloud/app release → CP-version → DO-TST-008
DO-LBL-005 Lithium-transport marking | D4; PRD-REG; UN 38.3 / IATA | battery-transport safety/compliance ↔HAL | PV-LBL-### marking review | VAL-REG-### | Labeling (PROC-LABEL) → CP-label-content → DO-LBL-012
DO-VER-001 CTQ/CTF register | PRD-VER/PERF; SYS (measurable); SRS-VER | ensures every safety char. is measurable ↔HAL | PV-### coverage vs. requirement set | — | — (verification-enabling)
DO-VER-003 Performance-limit / acceptance | PRD-PERF/REL/VER; SYS-033…050; D3 | limits bound safe operation ↔HAL | seeds all PV-<CODE>-NNN limits | VAL-### | → CP-guard-band → DO-TST-014
DO-VER-006 Test-interface & test-mode spec | PRD-VER; SRS-DIAG/SEC; IFR-005 | test modes must not create a fielded safety/security bypass ↔HAL | PV-### test-mode access-control | — | → CP-mode-lock → DO-TST-012
DO-VER-010 Objective pass/fail mapping | PRD-VER; SYS (methods); SRS-VER | no subjective safety verdict ↔HAL | maps every char. to a PV-<CODE>-NNN | VAL-### | → CP-acceptance → DO-TST-014
DO-XFER-008 Open-risk disposition & HAL reconciliation | DOL-1200; PRD-SAFE; ISO 14971 | every HAL control has a Released impl. output; residual risk accepted ↔HAL (RA-05) | risk-control verification complete (Stage 13) | VAL-SAFE-### | — → release gate
DO-XFER-010 Production-release baseline & approvals | all DO-XFER; ISO 13485 §7.3.8 | production authorized only when audit/reconciliation/readiness/risk satisfied ↔HAL | RA-01…12 all met | VAL-### | authorizes production (PROC-)
---
7. Coverage & honesty notes
• Requirement coverage (RA-01/RA-02). Every output in the full register (Design-Output-Register.md) carries an upward Input-Traceability field to ≥1 PRD/SYS/SRS/HWR/MER/FWR/IFR; this matrix details the severity-dominant and release-gating subset. Numeric-target inputs (accuracy, latency, endurance, IP, RF) remain Evidence Pending and are closed by V&V.
• Risk chain (RA-05). All HA-###/RCR-### are Preliminary — the HAL (Stage 10, ORCH-017) is not yet baselined. Each DOL-1200 DO-SAFE output is the "Implementation Reference" for a HAL control and must be reconciled; DO-XFER-008 is the reconciliation gate. This is the top release blocker.
• Verification / Validation. PV-<CODE>-NNN (Stage 13, ORCH-014) and VAL-<CODE>-NNN (Stage 14, ORCH-015) are forward placeholders; the two already-named headline tests are PRD-VER-004 (90-day cold endurance) and PRD-VER-005 (RF-in-container link). DO-VER-001/003/010 (DOL-1900) supply the characteristics, limits, and pass/fail that seed the PV protocols.
• Manufacturing chain. PROC-###/CP-### are Stage-15 placeholders; the acceptance leg is already concrete via DOL-1500 (DO-TST-###). The safety/measurement acceptance spine is DO-TST-004/005/006/009/011.
• Forward consumer. The complete bidirectional trace is assembled by ORCH-018 / EM-ENG-012 (Stage 18); each output's per-module entry already carries the forward placeholders this matrix resolves. </content>
Design Output Register (Consolidated) — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 09 — Design Outputs (ORCH-013) · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Status: Draft baseline (pre-DR-DO).
Consolidated one-line-per-output view of the entire SCM Design Output Library (DOL-000…DOL-2000). The authoritative, fully-fielded (24-field) entries live in the per-module DOL-.md files; this register aggregates them for navigation, counting, and audit. All entries are Draft/In-review in this pre-DR-DO baseline and all revisions are TBD-rev (Evidence Pending) until authored and released under DOL-000 (DO-GOV). Evidence, honesty, and the output-vs-record distinction follow the Master Specification (Product-Design-Output-Master-Specification.md).
Columns: Output ID · Title · Discipline · Output Type · Doc/File No · Rev · Status · Critical?
Totals: 272 controlled design outputs across 21 modules; 103 flagged Critical. Safety/charge/excursion/record outputs (esp. all of DOL-1200 DO-SAFE) are Preliminary — reconcile with HAL Stage 10.
---
1. Register summary (counts by module)
Module | Prefix | File | Outputs | Critical
DOL-000 Design Output Governance | DO-GOV | DOL-000-governance.md | 9 | 0
DOL-100 Product Configuration Baseline | DO-CFG | DOL-100-config-baseline.md | 12 | 3
DOL-200 System Architecture Outputs | DO-SYS | DOL-200-system-architecture-outputs.md | 12 | 6
DOL-300 Mechanical Design Outputs | DO-MECH | DOL-300-mechanical.md | 20 | 3
DOL-400 Electrical & Electronic Outputs | DO-EE | DOL-400-electrical.md | 22 | 11
DOL-500 PCB & PCBA Outputs | DO-PCB | DOL-500-pcb.md | 13 | 2
DOL-600 Embedded Firmware Outputs | DO-FW | DOL-600-firmware.md | 20 | 13
DOL-700 Device Software Outputs | DO-SW | DOL-700-device-software.md | 8 | 2
DOL-800 Cloud Platform Outputs | DO-CLD | DOL-800-cloud.md | 16 | 6
DOL-900 Mobile & Web Application Outputs | DO-APP | DOL-900-mobile.md | 10 | 2
DOL-1000 Connectivity & Interface Outputs | DO-IF | DOL-1000-interfaces.md | 12 | 7
DOL-1100 Cybersecurity Design Outputs | DO-SEC | DOL-1100-security.md | 12 | 8
DOL-1200 Safety & Risk-Control Outputs | DO-SAFE | DOL-1200-safety-risk-controls.md | 14 | 14
DOL-1300 Human Factors & UI Outputs | DO-HF | DOL-1300-human-factors.md | 10 | 3
DOL-1400 Manufacturing Design Outputs | DO-MFG | DOL-1400-manufacturing.md | 14 | 3
DOL-1500 Inspection & Production-Test Outputs | DO-TST | DOL-1500-inspection-test.md | 14 | 5
DOL-1600 Packaging, Labeling & Logistics | DO-LBL | DOL-1600-packaging-labeling.md | 12 | 3
DOL-1700 Installation, Service & Maintenance | DO-SVC | DOL-1700-service.md | 10 | 2
DOL-1800 Regulatory & Compliance Outputs | DO-REG | DOL-1800-regulatory.md | 12 | 4
DOL-1900 Verification-Enabling Outputs | DO-VER | DOL-1900-verification-enabling.md | 10 | 4
DOL-2000 Design Transfer & Release Baseline | DO-XFER | DOL-2000-design-transfer.md | 10 | 2
Total |  | 21 files | 272 | 103
---
2. Consolidated register (all outputs, by module)
DOL-000 — Design Output Governance (DO-GOV) — 9 outputs, 0 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-GOV-001 | Output Identification & Numbering Scheme | CM (SYS) | DO-DOC | SCM-GOV-DOC-001 | TBD-rev (Evidence Pending) | Draft | No
DO-GOV-002 | Revision & Approval Workflow Definition | CM (SYS) | DO-DOC | SCM-GOV-DOC-002 | TBD-rev | Draft | No
DO-GOV-003 | Authoritative File-Format & Derived-vs-Source Policy | CM (SYS) | DO-DOC | SCM-GOV-DOC-003 | TBD-rev | Draft | No
DO-GOV-004 | Engineering Change Control (ECO/ECN) Procedure | CM (SYS) | DO-PROC | SCM-GOV-PRC-004 | TBD-rev | Draft | No
DO-GOV-005 | Change-Classification Definition (Major/Minor/Editorial) | CM (SYS) | DO-DOC | SCM-GOV-DOC-005 | TBD-rev | Draft | No
DO-GOV-006 | Supplier- & Third-Party-Controlled Output Register | CM/Purchasing (SYS) | DO-DOC | SCM-GOV-DOC-006 | TBD-rev | Draft | No
DO-GOV-007 | SOUP / Third-Party Software Component Register | Software CM (S3/S8/S9) | DO-DOC | SCM-GOV-DOC-007 | TBD-rev | Draft | No
DO-GOV-008 | Archival, Retention & Retrieval Procedure | CM (SYS) | DO-PROC | SCM-GOV-PRC-008 | TBD-rev | Draft | No
DO-GOV-009 | Output-vs-Record Distinction Policy | CM/Quality (SYS) | DO-DOC | SCM-GOV-DOC-009 | TBD-rev | Draft | No
DOL-100 — Product Configuration Baseline (DO-CFG) — 12 outputs, 3 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-CFG-001 | SCM Product Family Tree (top-level) | CM (SYS) | DO-DOC | SCM-CFG-DOC-001 | TBD-rev (Evidence Pending) | Draft | No
DO-CFG-002 | Model & SKU Matrix | CM (SYS) | DO-DATA | SCM-CFG-DAT-002 | TBD-rev | Draft | No
DO-CFG-003 | Option & Accessory Matrix | CM (SYS) | DO-DATA | SCM-CFG-DAT-003 | TBD-rev | Draft | No
DO-CFG-004 | Approved Manufacturer / Parts List (AML) | CM/Purchasing (SYS) | DO-BOM | SCM-CFG-BOM-004 | TBD-rev (Evidence Pending — chemistry/antenna) | Draft | Yes
DO-CFG-005 | HW / FW / Cloud / Mobile Compatibility Matrix | CM (SYS) | DO-DATA | SCM-CFG-DAT-005 | TBD-rev | Draft | Yes
DO-CFG-006 | Regional Variant Configuration Definitions | CM/Regulatory (SYS) | DO-CFG | SCM-CFG-CFG-006 | TBD-rev | Draft | No
DO-CFG-007 | Replaceable-Component / FRU Definition | CM/Service (S1/S2) | DO-DOC | SCM-CFG-DOC-007 | TBD-rev | Draft | No
DO-CFG-008 | Consumables & Cleaning-Materials Baseline | CM/Service (S1) | DO-DOC | SCM-CFG-DOC-008 | TBD-rev | Draft | No
DO-CFG-009 | Packaging Configuration Matrix | CM/Packaging (SYS) | DO-DATA | SCM-CFG-DAT-009 | TBD-rev | Draft | No
DO-CFG-010 | Serial-Number & Device-Identity Convention | CM (S3) | DO-DOC | SCM-CFG-DOC-010 | TBD-rev | Draft | Yes
DO-CFG-011 | Configuration-Identification & Baseline Rules | CM (SYS) | DO-DOC | SCM-CFG-DOC-011 | TBD-rev | Draft | No
DO-CFG-012 | Product Configuration Rules & Compatibility Constraints | CM (SYS) | DO-DATA | SCM-CFG-DAT-012 | TBD-rev | Draft | No
DOL-200 — System Architecture Outputs (DO-SYS) — 12 outputs, 6 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-SYS-001 | Logical & Physical System Architecture Description | Systems (SYS) | DO-DOC | SCM-SYS-DOC-001 | TBD-rev | Draft | Yes
DO-SYS-002 | Functional Decomposition (FUN-001…040) | Systems (SYS) | DO-DOC | SCM-SYS-DOC-002 | TBD-rev | Draft | No
DO-SYS-003 | Subsystem Allocation Matrix (FUN→S1–S9) | Systems (SYS) | DO-DATA | SCM-SYS-DAT-003 | TBD-rev | Draft | No
DO-SYS-004 | Interface-Control Overview / ICD Register (INT-001…035) | Systems/Interface (SYS) | DO-ICD | SCM-SYS-ICD-004 | TBD-rev | Draft | Yes
DO-SYS-005 | Data-Flow & Control-Flow Architecture | Systems (SYS) | DO-DOC | SCM-SYS-DOC-005 | TBD-rev | Draft | Yes
DO-SYS-006 | Operational State & Mode Model | Systems (SYS) | DO-DOC | SCM-SYS-DOC-006 | TBD-rev | Draft | Yes
DO-SYS-007 | Power & Energy Architecture Description | Systems/Power (S2/S3) | DO-DOC | SCM-SYS-DOC-007 | TBD-rev | Draft | Yes
DO-SYS-008 | Communication Architecture Description | Systems/RF (S5) | DO-DOC | SCM-SYS-DOC-008 | TBD-rev | Draft | No
DO-SYS-009 | Fault-Management & Diagnostics Architecture | Systems/Reliability (SYS) | DO-DOC | SCM-SYS-DOC-009 | TBD-rev | Draft | Yes
DO-SYS-010 | Deployment Architecture Description | Systems/Cloud (S8/S9) | DO-DOC | SCM-SYS-DOC-010 | TBD-rev | Draft | No
DO-SYS-011 | Architecture Decision Records (ADR-001…012) | Systems (SYS) | DO-DOC | SCM-SYS-DOC-011 | TBD-rev | Draft | No
DO-SYS-012 | Third-Party Service Dependency Map | Systems/Cloud (S8) | DO-DOC | SCM-SYS-DOC-012 | TBD-rev | Draft | No
DOL-300 — Mechanical Design Outputs (DO-MECH) — 20 outputs, 3 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-MECH-000 | Enclosure Top-Level Assembly (3D CAD + drawing) | Mechanical (S1) | DO-CAD | SCM-MEC-CAD-000 | TBD-rev | Draft | No
DO-MECH-001 | Enclosure Top Housing (3D CAD) | Mechanical (S1) | DO-CAD | SCM-MEC-CAD-001 | TBD-rev | Draft | No
DO-MECH-002 | Enclosure Top Housing (2D released drawing + GD&T) | Mechanical (S1) | DO-DRW | SCM-MEC-DRW-002 | TBD-rev | Draft | No
DO-MECH-003 | Enclosure Bottom Housing (3D CAD) | Mechanical (S1) | DO-CAD | SCM-MEC-CAD-003 | TBD-rev | Draft | No
DO-MECH-004 | Enclosure Bottom Housing (2D drawing + GD&T) | Mechanical (S1) | DO-DRW | SCM-MEC-DRW-004 | TBD-rev | Draft | No
DO-MECH-005 | Enclosure Seal / Gasket Design (main O-ring/gasket) | Mechanical/Materials (S1) | DO-SPEC | SCM-MEC-SPC-005 | TBD-rev | Draft | Yes
DO-MECH-006 | External Temperature Probe Assembly (3D CAD) | Mechanical (S1/S4) | DO-CAD | SCM-MEC-CAD-006 | TBD-rev | Draft | Yes
DO-MECH-007 | Probe Lead Gland & Strain-Relief Design | Mechanical (S1) | DO-SPEC | SCM-MEC-SPC-007 | TBD-rev | Draft | No
DO-MECH-008 | Sealed USB-C Port Aperture & Gasket Design | Mechanical (S1) | DO-SPEC | SCM-MEC-SPC-008 | TBD-rev | Draft | No
DO-MECH-009 | Enclosure Mount / Retention Feature Design | Mechanical (S1) | DO-SPEC | SCM-MEC-SPC-009 | TBD-rev | Draft | No
DO-MECH-010 | Enclosure Material & Surface-Finish/Color Specification | Materials (S1) | DO-SPEC | SCM-MEC-SPC-010 | TBD-rev | Draft | No
DO-MECH-011 | Seal Groove / Gland Detail Drawing | Mechanical (S1) | DO-DRW | SCM-MEC-DRW-011 | TBD-rev | Draft | No
DO-MECH-012 | Battery Retention / Isolation Design | Mechanical (S1/S2) | DO-SPEC | SCM-MEC-SPC-012 | TBD-rev | Draft | Yes
DO-MECH-013 | PCB Mounting & Board-Retention Design | Mechanical (S1/S3) | DO-SPEC | SCM-MEC-SPC-013 | TBD-rev | Draft | No
DO-MECH-014 | HMI Window / OLED & Button Mechanical Design | Mechanical (S1/S6) | DO-SPEC | SCM-MEC-SPC-014 | TBD-rev | Draft | No
DO-MECH-015 | Thermal-Coupling / Internal-Sensor Airpath Design | Mechanical/Thermal (S1/S4) | DO-MODEL | SCM-MEC-MDL-015 | TBD-rev | Draft | No
DO-MECH-016 | External-Antenna Aperture Design (SCM-100-XA) | Mechanical (S1/S5) | DO-SPEC | SCM-MEC-SPC-016 | TBD-rev | Draft | No
DO-MECH-017 | Shock/Vibration & Drop Structural Analysis | Mechanical (S1) | DO-MODEL | SCM-MEC-MDL-017 | TBD-rev | Draft | No
DO-MECH-018 | Tolerance-Stack Analysis (seal, HMI, connector fits) | Mechanical (S1) | DO-MODEL | SCM-MEC-MDL-018 | TBD-rev | Draft | No
DO-MECH-019 | Mechanical Bill of Materials | Mechanical (S1) | DO-BOM | SCM-MEC-BOM-019 | TBD-rev | Draft | No
DOL-400 — Electrical & Electronic Outputs (DO-EE) — 22 outputs, 11 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-EE-000 | Electrical Architecture Description | Electrical (S2–S6) | DO-DOC | SCM-ELE-DOC-000 | TBD-rev | Draft | No
DO-EE-001 | Full Schematic Set (device electronics) | Electrical (S2–S6) | DO-SCH | SCM-ELE-SCH-001 | TBD-rev | Draft | Yes
DO-EE-002 | Power Tree / Power-Distribution Schematic | Power (S2) | DO-SCH | SCM-ELE-SCH-002 | TBD-rev | Draft | Yes
DO-EE-003 | Rechargeable Low-Temperature Cell Selection Specification | Power (S2) | DO-SPEC | SCM-ELE-SPC-003 | TBD-rev (Evidence Pending — chemistry) | Draft | Yes
DO-EE-004 | Battery Pack Assembly Design (config + interconnect) | Power (S2) | DO-SPEC | SCM-ELE-SPC-004 | TBD-rev (Evidence Pending — config) | Draft | Yes
DO-EE-005 | Battery-Protection IC / AFE (BMS) Design & Threshold Specification | Power/Safety (S2) | DO-SPEC | SCM-ELE-SPC-005 | TBD-rev | Draft | Yes
DO-EE-006 | Charge-Path / PMIC Design & Charge-Profile Specification | Power (S2) | DO-SPEC | SCM-ELE-SPC-006 | TBD-rev (Evidence Pending — profile) | Draft | Yes
DO-EE-007 | Charge-Inhibit-Below-0 °C Interlock Design (HW+FW boundary) | Power/Safety (S2/S3) | DO-SPEC | SCM-ELE-SPC-007 | TBD-rev | Draft | Yes
DO-EE-008 | Fuel-Gauge Design & SoC/SoH Configuration | Power (S2) | DO-SPEC | SCM-ELE-SPC-008 | TBD-rev (Evidence Pending — accuracy) | Draft | No
DO-EE-009 | USB-C Sink Port & Input-Protection Design | Power (S2) | DO-SPEC | SCM-ELE-SPC-009 | TBD-rev | Draft | Yes
DO-EE-010 | Secondary Non-Resettable Protection (Fuse/OVP) Design | Safety/Power (S2) | DO-SPEC | SCM-ELE-SPC-010 | TBD-rev | Draft | Yes
DO-EE-011 | MCU, Secure-Element, RTC & Watchdog Electrical Design | Compute (S3) | DO-SPEC | SCM-ELE-SPC-011 | TBD-rev | Draft | No
DO-EE-012 | Debug / EOL-Test Port & Production-Lock Design | Compute/Security (S3) | DO-SPEC | SCM-ELE-SPC-012 | TBD-rev | Draft | Yes
DO-EE-013 | External-Probe Analog/Digital Front-End & Presence-Detect Design | Sensing (S4) | DO-SPEC | SCM-ELE-SPC-013 | TBD-rev | Draft | Yes
DO-EE-014 | Internal Temperature/Humidity Sensor Interface Design | Sensing (S4) | DO-SPEC | SCM-ELE-SPC-014 | TBD-rev | Draft | No
DO-EE-015 | Accelerometer & Shock-Sensing Interface Design | Sensing (S4) | DO-SPEC | SCM-ELE-SPC-015 | TBD-rev | Draft | No
DO-EE-016 | GNSS Receiver Interface Design | Sensing (S4) | DO-SPEC | SCM-ELE-SPC-016 | TBD-rev (Evidence Pending) | Draft | No
DO-EE-017 | Radio Front-End Design (LTE-M / Wi-Fi / BLE) + Band Plan | RF (S5) | DO-SPEC | SCM-ELE-SPC-017 | TBD-rev (band plan Evidence Pending) | Draft | No
DO-EE-018 | Antenna & Matching-Network Design (integral) | RF (S5/S1) | DO-SPEC | SCM-ELE-SPC-018 | TBD-rev (Evidence Pending — antenna) | Draft | No
DO-EE-019 | Evidence, EMC/ESD Provisions & Safety-Spacing Design | Electrical (S2–S6) | DO-SPEC | SCM-ELE-SPC-019 | TBD-rev | Draft | No
DO-EE-020 | Electrical Bill of Materials | Electrical (S2–S6) | DO-BOM | SCM-ELE-BOM-020 | TBD-rev | Draft | No
DO-EE-021 | External-Antenna Option Design (SCM-100-XA) | RF (S5/S1) | DO-SPEC | SCM-ELE-SPC-021 | TBD-rev (Evidence Pending) | Draft | No
DOL-500 — PCB & PCBA Outputs (DO-PCB) — 13 outputs, 2 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-PCB-001 | Board Outline & Mechanical Definition | PCB (S3) | DO-DRW | SCM-PCB-DRW-001 | TBD-rev | Draft | No
DO-PCB-002 | Netlist (from schematic) | PCB (S3) | DO-DATA | SCM-PCB-DAT-002 | TBD-rev | Draft | No
DO-PCB-003 | Layer Stack-Up & Materials Specification | PCB (S3) | DO-SPEC | SCM-PCB-SPC-003 | TBD-rev | Draft | No
DO-PCB-004 | Placement / Layout (routed board) | PCB (S3) | DO-PCB | SCM-PCB-LAY-004 | TBD-rev | Draft | Yes
DO-PCB-005 | Gerber / ODB++ Fabrication Data Set | PCB (S3) | DO-PCB | SCM-PCB-FAB-005 | TBD-rev | Draft | No
DO-PCB-006 | Controlled-Impedance & RF Routing Specification | PCB/RF (S3/S5) | DO-SPEC | SCM-PCB-SPC-006 | TBD-rev (Evidence Pending — antenna) | Draft | No
DO-PCB-007 | Fabrication Drawing (IPC class, notes, acceptance) | PCB (S3) | DO-DRW | SCM-PCB-DRW-007 | TBD-rev | Draft | No
DO-PCB-008 | Assembly Drawing & Pick-and-Place Data | PCB (S3) | DO-PCB | SCM-PCB-ASM-008 | TBD-rev | Draft | No
DO-PCB-009 | Stencil / Solder-Paste Aperture Design | PCB (S3) | DO-PCB | SCM-PCB-STN-009 | TBD-rev | Draft | No
DO-PCB-010 | Conformal-Coat & Cleanliness Specification | PCB (S3) | DO-SPEC | SCM-PCB-SPC-010 | TBD-rev | Draft | Yes
DO-PCB-011 | Test-Point & Boundary-Scan / Fiducial Definition | PCB (S3) | DO-DATA | SCM-PCB-DAT-011 | TBD-rev | Draft | No
DO-PCB-012 | Panelization / Array Drawing | PCB (S3) | DO-DRW | SCM-PCB-DRW-012 | TBD-rev | Draft | No
DO-PCB-013 | PCBA Bill of Materials | PCB (S3) | DO-BOM | SCM-PCB-BOM-013 | TBD-rev | Draft | No
DOL-600 — Embedded Firmware Outputs (DO-FW) — 20 outputs, 13 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-FW-001 | Firmware Architecture Description | FW (S3) | DO-DOC | SCM-FW-DOC-001 | TBD-rev | Draft | No
DO-FW-002 | Module Decomposition & Task/Scheduler Model | FW (S3) | DO-DOC | SCM-FW-DOC-002 | TBD-rev | Draft | No
DO-FW-003 | State-Machine / Mode-Manager Design | FW (S3) | DO-DOC | SCM-FW-DOC-003 | TBD-rev | Draft | Yes
DO-FW-004 | Power-State Manager / Energy Governor Design | FW (S3) | DO-DOC | SCM-FW-DOC-004 | TBD-rev (budget Evidence Pending) | Draft | Yes
DO-FW-005 | Firmware Charge-Inhibit & Self-Heat Control Logic | FW/Safety (S3/S2) | DO-CODE | SCM-FW-COD-005 | TBD-rev | Draft | Yes
DO-FW-006 | Excursion-Detection Engine Design | FW/Safety (S3) | DO-CODE | SCM-FW-COD-006 | TBD-rev | Draft | Yes
DO-FW-007 | Sensor Drivers / Acquisition & Calibration-Application Design | FW (S3/S4) | DO-CODE | SCM-FW-COD-007 | TBD-rev | Draft | Yes
DO-FW-008 | Verdict Engine Design (accept/reject/quarantine) | FW/Safety (S3) | DO-CODE | SCM-FW-COD-008 | TBD-rev | Draft | Yes
DO-FW-009 | Calibration Data Schema & Application Design | FW/Metrology (S3/S4) | DO-DATA | SCM-FW-DAT-009 | TBD-rev | Draft | No
DO-FW-010 | Storage Manager (append-only log, power-fail-safe) Design | FW (S3/S7) | DO-CODE | SCM-FW-COD-010 | TBD-rev | Draft | Yes
DO-FW-011 | Comms Agent / Bearer-Manager Design | FW (S3/S5) | DO-CODE | SCM-FW-COD-011 | TBD-rev | Draft | Yes
DO-FW-012 | Diagnostics / Health & BIST Design | FW (S3) | DO-CODE | SCM-FW-COD-012 | TBD-rev | Draft | No
DO-FW-013 | Fault Handler / Watchdog & Recovery Design | FW/Safety (S3) | DO-CODE | SCM-FW-COD-013 | TBD-rev | Draft | Yes
DO-FW-014 | Bootloader / Secure-Boot Loader Design | FW/Security (S3) | DO-CODE | SCM-FW-COD-014 | TBD-rev | Draft | Yes
DO-FW-015 | OTA Agent Design (download/verify/apply/rollback) | FW/Security (S3/S5) | DO-CODE | SCM-FW-COD-015 | TBD-rev | Draft | Yes
DO-FW-016 | Security / Crypto Module Design (FW-side) | FW/Security (S3) | DO-CODE | SCM-FW-COD-016 | TBD-rev | Draft | Yes
DO-FW-017 | Firmware Configuration Schema & Build Configuration | FW (S3) | DO-CFG | SCM-FW-CFG-017 | TBD-rev | Draft | No
DO-FW-018 | Dependency Manifest / Firmware SBOM & Static-Analysis Config | FW/Security (S3) | DO-DATA | SCM-FW-DAT-018 | TBD-rev | Draft | No
DO-FW-019 | Released Firmware Image, Hash & Release Notes | FW (S3) | DO-BIN | SCM-FW-BIN-019 | TBD-rev (version placeholder) | Draft | Yes
DO-FW-020 | Firmware Programming Package & Signing Procedure | FW/Security (S3) | DO-PROC | SCM-FW-PRC-020 | TBD-rev | Draft | No
DOL-700 — Device Software Outputs (DO-SW) — 8 outputs, 2 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-SW-001 | Factory Programming & Provisioning Tool | Device SW (S3) | DO-CODE | SCM-SW-COD-001 | TBD-rev | Draft | Yes
DO-SW-002 | Calibration Station Software | Device SW (S3/S4) | DO-CODE | SCM-SW-COD-002 | TBD-rev | Draft | No
DO-SW-003 | Host Diagnostic / Service Tool | Device SW (S3) | DO-CODE | SCM-SW-COD-003 | TBD-rev | Draft | No
DO-SW-004 | Factory Device-Image Package (firmware+config+identity+cal) | Device SW (S3) | DO-BIN | SCM-SW-BIN-004 | TBD-rev | Draft | Yes
DO-SW-005 | Recovery / Re-Provisioning Image & Procedure | Device SW (S3) | DO-BIN | SCM-SW-BIN-005 | TBD-rev | Draft | No
DO-SW-006 | Device Software SBOM | Device SW/Security (S3) | DO-DATA | SCM-SW-DAT-006 | TBD-rev | Draft | No
DO-SW-007 | Build / Deploy Scripts (device SW toolchain) | Device SW (S3) | DO-CODE | SCM-SW-COD-007 | TBD-rev | Draft | No
DO-SW-008 | Repository Baseline & Version-Tag Definition | Device SW (S3) | DO-CFG | SCM-SW-CFG-008 | TBD-rev | Draft | No
DOL-800 — Cloud Platform Outputs (DO-CLD) — 16 outputs, 6 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-CLD-001 | Cloud Logical & Deployment Architecture | Cloud (S8) | DO-DOC | SCM-CLD-DOC-001 | TBD-rev | Draft | No
DO-CLD-002 | Ingest / Store-and-Forward Endpoint Design | Cloud (S8) | DO-DOC | SCM-CLD-DOC-002 | TBD-rev | Draft | Yes
DO-CLD-003 | Telemetry / Record / Audit Data Schemas | Cloud (S8) | DO-DATA | SCM-CLD-DAT-003 | TBD-rev | Draft | Yes
DO-CLD-004 | Device Registry Service Design | Cloud (S8) | DO-DOC | SCM-CLD-DOC-004 | TBD-rev | Draft | No
DO-CLD-005 | Fleet Health / Diagnostics Service Design | Cloud (S8) | DO-DOC | SCM-CLD-DOC-005 | TBD-rev | Draft | No
DO-CLD-006 | Record Store & Cloud Verdict-Recompute Design | Cloud (S8) | DO-DOC | SCM-CLD-DOC-006 | TBD-rev | Draft | Yes
DO-CLD-007 | Rules / Alert-Routing & Escalation Engine Design | Cloud (S8) | DO-DOC | SCM-CLD-DOC-007 | TBD-rev | Draft | Yes
DO-CLD-008 | Notification Service Design (email/SMS/webhook) | Cloud (S8) | DO-DOC | SCM-CLD-DOC-008 | TBD-rev | Draft | No
DO-CLD-009 | Device & Enterprise API Contracts (OpenAPI) | Cloud (S8) | DO-ICD | SCM-CLD-ICD-009 | TBD-rev | Draft | Yes
DO-CLD-010 | OTA Service Design (signing & campaign distribution) | Cloud/Security (S8) | DO-DOC | SCM-CLD-DOC-010 | TBD-rev | Draft | Yes
DO-CLD-011 | Infrastructure-as-Code (IaC) Definitions | Cloud (S8) | DO-IAC | SCM-CLD-IAC-011 | TBD-rev | Draft | No
DO-CLD-012 | Fleet Dashboard & Reporting UI Design | Cloud/Web (S8) | DO-DOC | SCM-CLD-DOC-012 | TBD-rev | Draft | No
DO-CLD-013 | Identity, RBAC & SSO Configuration Design | Cloud/Security (S8) | DO-CFG | SCM-CLD-CFG-013 | TBD-rev | Draft | No
DO-CLD-014 | Deployment Pipeline, Monitoring & Backup/Recovery Design | Cloud (S8) | DO-IAC | SCM-CLD-IAC-014 | TBD-rev | Draft | No
DO-CLD-015 | Cloud SBOM & Dependency Controls | Cloud/Security (S8) | DO-DATA | SCM-CLD-DAT-015 | TBD-rev | Draft | No
DO-CLD-016 | Released Cloud Service Versions & Environment Config | Cloud (S8) | DO-CFG | SCM-CLD-CFG-016 | TBD-rev | Draft | No
DOL-900 — Mobile & Web Application Outputs (DO-APP) — 10 outputs, 2 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-APP-001 | Mobile App Architecture & Navigation Design | Mobile (S9) | DO-DOC | SCM-APP-DOC-001 | TBD-rev | Draft | No
DO-APP-002 | Screen & Interaction Specifications | Mobile/HF (S9) | DO-DOC | SCM-APP-DOC-002 | TBD-rev | Draft | No
DO-APP-003 | BLE-GATT Client Implementation Design | Mobile/Security (S9) | DO-CODE | SCM-APP-COD-003 | TBD-rev | Draft | Yes
DO-APP-004 | Design System & UI Asset Library | Mobile/HF (S9) | DO-DOC | SCM-APP-DOC-004 | TBD-rev | Draft | No
DO-APP-005 | Offline & Sync Model Design | Mobile (S9) | DO-DOC | SCM-APP-DOC-005 | TBD-rev | Draft | No
DO-APP-006 | Notifications & Permissions Design | Mobile (S9) | DO-DOC | SCM-APP-DOC-006 | TBD-rev | Draft | No
DO-APP-007 | Error Handling, Analytics & Crash Reporting Design | Mobile (S9) | DO-DOC | SCM-APP-DOC-007 | TBD-rev | Draft | No
DO-APP-008 | Build & Signing Configuration | Mobile/Security (S9) | DO-CFG | SCM-APP-CFG-008 | TBD-rev | Draft | No
DO-APP-009 | Store Packages & Version-Compatibility Matrix | Mobile (S9) | DO-BIN | SCM-APP-BIN-009 | TBD-rev | Draft | Yes
DO-APP-010 | Mobile SBOM & Dependency Manifest | Mobile/Security (S9) | DO-DATA | SCM-APP-DAT-010 | TBD-rev | Draft | No
DOL-1000 — Connectivity & Interface Outputs (DO-IF) — 12 outputs, 7 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-IF-001 | Internal Electrical/Bus/Connector ICD (INT-006…024) | Interfaces (S1–S7) | DO-ICD | SCM-IF-ICD-001 | TBD-rev | Draft | Yes
DO-IF-002 | USB-C Charge/Data Interface ICD (INT-008/009) | Interfaces (S1/S2) | DO-ICD | SCM-IF-ICD-002 | TBD-rev | Draft | No
DO-IF-003 | Wireless RF Interface ICD (INT-025…031) | Interfaces/RF (S5) | DO-ICD | SCM-IF-ICD-003 | TBD-rev (Evidence Pending) | Draft | Yes
DO-IF-004 | BLE-GATT Provisioning ICD (INT-033) | Interfaces/Security (S5/S9) | DO-ICD | SCM-IF-ICD-004 | TBD-rev | Draft | Yes
DO-IF-005 | Device↔Cloud Store-and-Forward Protocol ICD (INT-032) | Interfaces (S3/S5/S8) | DO-ICD | SCM-IF-ICD-005 | TBD-rev | Draft | Yes
DO-IF-006 | Cloud↔Enterprise API ICD (INT-034) | Interfaces (S8) | DO-ICD | SCM-IF-ICD-006 | TBD-rev | Draft | No
DO-IF-007 | OTA Update Interface ICD (INT-035) | Interfaces/Security (S3/S5/S8) | DO-ICD | SCM-IF-ICD-007 | TBD-rev | Draft | Yes
DO-IF-008 | Sensor/Probe Signal-Chain ICD (INT-017/018) | Interfaces/Sensing (S4) | DO-ICD | SCM-IF-ICD-008 | TBD-rev | Draft | Yes
DO-IF-009 | HMI Interface ICD (INT-022/023) | Interfaces/HMI (S6) | DO-ICD | SCM-IF-ICD-009 | TBD-rev | Draft | No
DO-IF-010 | Radio Host-Control Interface ICD (INT-025) | Interfaces (S3/S5) | DO-ICD | SCM-IF-ICD-010 | TBD-rev | Draft | No
DO-IF-011 | Debug/EOL-Test Interface ICD (governed, non-fielded) | Interfaces/Security (S3) | DO-ICD | SCM-IF-ICD-011 | TBD-rev | Draft | Yes
DO-IF-012 | ICD Governance & Interface-Conformance Rules | Interfaces (SYS) | DO-DOC | SCM-IF-DOC-012 | TBD-rev | Draft | No
DOL-1100 — Cybersecurity Design Outputs (DO-SEC) — 12 outputs, 8 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-SEC-001 | Security Architecture & Trust-Boundary Definition | Security (all tiers) | DO-DOC | SCM-SEC-DOC-001 | TBD-rev | Draft | Yes
DO-SEC-002 | Identity, Authentication & Authorization Design | Security (all tiers) | DO-DOC | SCM-SEC-DOC-002 | TBD-rev | Draft | Yes
DO-SEC-003 | Encryption & Transport-Security Configuration | Security (all tiers) | DO-CFG | SCM-SEC-CFG-003 | TBD-rev | Draft | Yes
DO-SEC-004 | Secure-Boot & Signed-Update Configuration | Security (S3) | DO-CFG | SCM-SEC-CFG-004 | TBD-rev | Draft | Yes
DO-SEC-005 | Key-Management & Certificate-Lifecycle Procedure | Security (all tiers) | DO-PROC | SCM-SEC-PRC-005 | TBD-rev | Draft | Yes
DO-SEC-006 | Credential/Record Tamper-Evidence & Secure-Storage Design | Security (S3/S8) | DO-DOC | SCM-SEC-DOC-006 | TBD-rev | Draft | Yes
DO-SEC-007 | Production Key/Identity Provisioning Design | Security (S3) | DO-PROC | SCM-SEC-PRC-007 | TBD-rev | Draft | Yes
DO-SEC-008 | Debug-Lock & Anti-Tamper Design (fielded) | Security (S3) | DO-DOC | SCM-SEC-DOC-008 | TBD-rev | Draft | Yes
DO-SEC-009 | Security Audit-Log & Security-Event Design | Security (S8) | DO-DOC | SCM-SEC-DOC-009 | TBD-rev | Draft | No
DO-SEC-010 | Vulnerability-Management & Security-Update Design | Security (all tiers) | DO-DOC | SCM-SEC-DOC-010 | TBD-rev | Draft | No
DO-SEC-011 | Dependency-Security Controls & Consolidated SBOM Policy | Security (all tiers) | DO-DATA | SCM-SEC-DAT-011 | TBD-rev | Draft | No
DO-SEC-012 | Key/Certificate Revocation & Recovery Design | Security (all tiers) | DO-DOC | SCM-SEC-DOC-012 | TBD-rev | Draft | No
DOL-1200 — Safety & Risk-Control Outputs (DO-SAFE) — 14 outputs, 14 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-SAFE-001 | Charge-Inhibit-Below-0 °C Interlock (implementation reference) | Safety/Power (S2/S3) | DO-DOC | SCM-SAF-DOC-001 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-002 | Battery Protection (OV/UV/OC/OT/short) + Secondary Fuse | Safety/Power (S2) | DO-DOC | SCM-SAF-DOC-002 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-003 | Charge-Input Abuse Protection (reverse/OVP/surge/short) | Safety/Power (S2) | DO-DOC | SCM-SAF-DOC-003 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-004 | Two-Sided Excursion-Detection Control (missed-freeze/heat) | Safety/FW (S3) | DO-DOC | SCM-SAF-DOC-004 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-005 | Probe-Fault / Sensor-Plausibility Control (false-accept prevention) | Safety/Sensing (S3/S4) | DO-DOC | SCM-SAF-DOC-005 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-006 | Verdict-Integrity Control (suppress false "accept") | Safety/FW (S3) | DO-DOC | SCM-SAF-DOC-006 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-007 | Safe-State / Fault-Safe Mode Control | Safety/FW (S3) | DO-DOC | SCM-SAF-DOC-007 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-008 | Record-Integrity / Power-Fail-Safe Storage Control | Safety/FW (S3/S7) | DO-DOC | SCM-SAF-DOC-008 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-009 | Excursion-Alert Delivery Control (delayed/missing alert) | Safety/FW/Cloud (S3/S5/S8) | DO-DOC | SCM-SAF-DOC-009 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-010 | Battery Mechanical/Thermal Isolation Control | Safety/Mech (S1/S2) | DO-DOC | SCM-SAF-DOC-010 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-011 | Ingress/Condensation Protection Control | Safety/Mech/PCB (S1/S3) | DO-DOC | SCM-SAF-DOC-011 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-012 | Energy-Reserve / Loss-of-Monitoring Control | Safety/FW (S2/S3) | DO-DOC | SCM-SAF-DOC-012 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-013 | Charge/Fault-State Legibility Control (no silent failure) | Safety/HMI (S3/S6) | DO-DOC | SCM-SAF-DOC-013 | TBD-rev | Draft (Preliminary) | Yes
DO-SAFE-014 | OTA Never-Brick / Unmonitored-Window Control | Safety/FW/Cloud (S3/S8) | DO-DOC | SCM-SAF-DOC-014 | TBD-rev | Draft (Preliminary) | Yes
DOL-1300 — Human Factors & UI Outputs (DO-HF) — 10 outputs, 3 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-HF-001 | On-Device HMI Specification (information architecture) | HF (S6) | DO-DOC | SCM-HF-DOC-001 | TBD-rev | Draft | No
DO-HF-002 | Verdict Presentation Design (accept/reject/quarantine) | HF/Safety (S6) | DO-DOC | SCM-HF-DOC-002 | TBD-rev | Draft | Yes
DO-HF-003 | Iconography, Symbols & Status-Indicator Design | HF (S6/S9) | DO-DOC | SCM-HF-DOC-003 | TBD-rev | Draft | No
DO-HF-004 | Physical Control (glove-friendly button) HF Specification | HF (S6) | DO-SPEC | SCM-HF-SPC-004 | TBD-rev | Draft | No
DO-HF-005 | Device-State & Warning Presentation Design | HF/Safety (S6) | DO-DOC | SCM-HF-DOC-005 | TBD-rev | Draft | Yes
DO-HF-006 | Provisioning & Arming Task-Flow (use-error mitigation) Design | HF/Safety (S9) | DO-DOC | SCM-HF-DOC-006 | TBD-rev | Draft | Yes
DO-HF-007 | Fleet-Dashboard Reporting HF Specification | HF (S8) | DO-DOC | SCM-HF-DOC-007 | TBD-rev | Draft | No
DO-HF-008 | Alarm/Notification Presentation & Prioritization Design | HF/Safety (S8/S9) | DO-DOC | SCM-HF-DOC-008 | TBD-rev | Draft | No
DO-HF-009 | Quick-Start / User-Assistance Content | HF (S1/S9) | DO-MAN | SCM-HF-MAN-009 | TBD-rev | Draft | No
DO-HF-010 | Display-Characteristics & Cold-Legibility Specification | HF (S6) | DO-SPEC | SCM-HF-SPC-010 | TBD-rev (Evidence Pending) | Draft | No
DOL-1400 — Manufacturing Design Outputs (DO-MFG) — 14 outputs, 3 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-MFG-001 | Manufacturing BOM & Approved-Manufacturer List (mBOM/AML) | Mfg (all) | DO-BOM | SCM-MFG-BOM-001 | TBD-rev (Evidence Pending — chemistry/antenna) | Draft | No
DO-MFG-002 | SMT / PCBA Assembly Work Instructions | Mfg (S3) | DO-WI | SCM-MFG-WI-002 | TBD-rev | Draft | No
DO-MFG-003 | Reflow / Soldering Process Parameter Specification | Mfg (S3) | DO-SPEC | SCM-MFG-SPC-003 | TBD-rev | Draft | No
DO-MFG-004 | Enclosure Assembly & Sealing Work Instruction | Mfg (S1) | DO-WI | SCM-MFG-WI-004 | TBD-rev | Draft | Yes
DO-MFG-005 | Conformal-Coat Application Work Instruction | Mfg (S3) | DO-WI | SCM-MFG-WI-005 | TBD-rev | Draft | No
DO-MFG-006 | Torque / Fastening & Adhesive Parameter Specification | Mfg (S1) | DO-SPEC | SCM-MFG-SPC-006 | TBD-rev | Draft | No
DO-MFG-007 | Battery Handling & Installation Work Instruction | Mfg/Safety (S2) | DO-WI | SCM-MFG-WI-007 | TBD-rev | Draft | Yes
DO-MFG-008 | Firmware Programming & Provisioning Work Instruction | Mfg/Security (S3) | DO-WI | SCM-MFG-WI-008 | TBD-rev | Draft | Yes
DO-MFG-009 | Serialization & UDI-Marking Procedure | Mfg/Regulatory (S1) | DO-PROC | SCM-MFG-PRC-009 | TBD-rev | Draft | No
DO-MFG-010 | Calibration Process Work Instruction | Mfg/Metrology (S4) | DO-WI | SCM-MFG-WI-010 | TBD-rev | Draft | No
DO-MFG-011 | Final-Assembly & Line-Clearance Work Instruction | Mfg (all) | DO-WI | SCM-MFG-WI-011 | TBD-rev | Draft | No
DO-MFG-012 | Assembly Fixtures & Tooling Design | Mfg (S1/S3) | DO-TOOL | SCM-MFG-TL-012 | TBD-rev | Draft | No
DO-MFG-013 | Injection-Mould Tooling Specification (enclosure) | Mfg (S1) | DO-TOOL | SCM-MFG-TL-013 | TBD-rev | Draft | No
DO-MFG-014 | Rework / Repair & Handling Procedure | Mfg (all) | DO-PROC | SCM-MFG-PRC-014 | TBD-rev | Draft | No
DOL-1500 — Inspection & Production-Test Outputs (DO-TST) — 14 outputs, 5 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-TST-001 | Incoming Inspection & Component-Acceptance Specification | Quality (all S) | DO-SPEC | SCM-TST-SPC-001 | TBD-rev | Draft | No
DO-TST-002 | In-Process Inspection Plan (SPI / AOI / X-ray) | Quality/PCBA (S3–S6) | DO-SPEC | SCM-TST-SPC-002 | TBD-rev | Draft | No
DO-TST-003 | Dimensional & Cosmetic Final-Inspection Specification | Quality/Mech (S1/S6) | DO-SPEC | SCM-TST-SPC-003 | TBD-rev | Draft | No
DO-TST-004 | PCBA In-Circuit / Functional Test (ICT/FCT) Definition | Quality/Electrical (S2–S6) | DO-TEST | SCM-TST-TDF-004 | TBD-rev | Draft | Yes
DO-TST-005 | Final Functional & System Test Definition | Quality (all S) | DO-TEST | SCM-TST-TDF-005 | TBD-rev | Draft | Yes
DO-TST-006 | Sensor Calibration & Verification Test Definition (T / RH / probe) | Quality/Metrology (S4) | DO-TEST | SCM-TST-TDF-006 | TBD-rev (Evidence Pending — reference chain) | Draft | Yes
DO-TST-007 | Connectivity / Radio Production-Test Definition | Quality/RF (S5) | DO-TEST | SCM-TST-TDF-007 | TBD-rev | Draft | No
DO-TST-008 | Firmware Programming & Provisioning Verification Test | Quality/Security (S3) | DO-TEST | SCM-TST-TDF-008 | TBD-rev | Draft | No
DO-TST-009 | Battery-Protection / Charge-Inhibit Fault-Injection Safety Test | Quality/Safety/Power (S2/S3) | DO-TEST | SCM-TST-TDF-009 | TBD-rev | Draft | Yes
DO-TST-010 | Electrical-Safety / ESD Production Test | Quality/Electrical (S2) | DO-TEST | SCM-TST-TDF-010 | TBD-rev | Draft | No
DO-TST-011 | Seal-Integrity / Leak (IP) Production Test | Quality/Mech (S1) | DO-TEST | SCM-TST-TDF-011 | TBD-rev (Evidence Pending — method/limit) | Draft | Yes
DO-TST-012 | Debug-Lock / Security-Provisioning Verification Test | Quality/Security (S3) | DO-TEST | SCM-TST-TDF-012 | TBD-rev | Draft | No
DO-TST-013 | Production-Test Software, Fixtures & Golden-Unit Definition | Quality/Test-Eng (all S) | DO-TOOL | SCM-TST-TOL-013 | TBD-rev | Draft | No
DO-TST-014 | Test Limits, MSA / Gauge R&R & Acceptance-Reject Criteria | Quality (all S) | DO-SPEC | SCM-TST-SPC-014 | TBD-rev (Evidence Pending — limits/MSA) | Draft | No
DOL-1600 — Packaging, Labeling & Logistics (DO-LBL) — 12 outputs, 3 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-LBL-001 | Device Nameplate / Rating Label Artwork | Labeling (S1) | DO-LBL | SCM-LBL-ART-001 | TBD-rev | Draft | No
DO-LBL-002 | Regulatory-Markings Label Artwork (CE / FCC / IC / WEEE / etc.) | Labeling/Regulatory (S1) | DO-LBL | SCM-LBL-ART-002 | TBD-rev (Evidence Pending — markings set) | Draft | Yes
DO-LBL-003 | Serial-Number / Device-Identity Label (barcode / QR) | Labeling (S1) | DO-LBL | SCM-LBL-ART-003 | TBD-rev | Draft | No
DO-LBL-004 | UDI Label Content & Format Specification | Labeling/Regulatory (S1) | DO-LBL | SCM-LBL-SPC-004 | TBD-rev (Evidence Pending — UDI regime) | Draft | Yes
DO-LBL-005 | Lithium-Battery Transport Marking & Handling Label (UN 38.3 / IATA) | Labeling/Logistics (S2) | DO-LBL | SCM-LBL-ART-005 | TBD-rev | Draft | Yes
DO-LBL-006 | Primary Packaging Artwork & Drawing | Packaging (S1) | DO-DRW | SCM-LBL-DRW-006 | TBD-rev | Draft | No
DO-LBL-007 | Protective / Cushioning Packaging Design | Packaging (S1) | DO-SPEC | SCM-LBL-SPC-007 | TBD-rev | Draft | No
DO-LBL-008 | Shipping / Transport Carton & Palletization Specification | Logistics (S1) | DO-SPEC | SCM-LBL-SPC-008 | TBD-rev | Draft | No
DO-LBL-009 | Storage & Transport Conditions / Shelf-Life Controls | Logistics/Quality (S2) | DO-SPEC | SCM-LBL-SPC-009 | TBD-rev | Draft | No
DO-LBL-010 | Instructions for Use (IFU) Content | Labeling/HF (S6/S9) | DO-MAN | SCM-LBL-MAN-010 | TBD-rev | Draft | No
DO-LBL-011 | Quick-Start Guide & Country-Specific Labels / Translations | Labeling/HF (S6/S9) | DO-MAN | SCM-LBL-MAN-011 | TBD-rev | Draft | No
DO-LBL-012 | Packaging & Labeling Inspection Criteria | Quality (S1) | DO-SPEC | SCM-LBL-SPC-012 | TBD-rev | Draft | No
DOL-1700 — Installation, Service & Maintenance (DO-SVC) — 10 outputs, 2 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-SVC-001 | Installation / Commissioning & Provisioning Procedure | Service (S8/S9) | DO-PROC | SCM-SVC-PRC-001 | TBD-rev | Draft | No
DO-SVC-002 | Service Manual | Service (all S) | DO-MAN | SCM-SVC-MAN-002 | TBD-rev | Draft | No
DO-SVC-003 | Troubleshooting / Diagnostics Trees | Service (all S) | DO-DOC | SCM-SVC-DOC-003 | TBD-rev | Draft | No
DO-SVC-004 | Preventive-Maintenance Schedule | Service (S1/S2) | DO-PROC | SCM-SVC-PRC-004 | TBD-rev | Draft | No
DO-SVC-005 | Field Recalibration / Calibration-Verification Procedure | Service/Metrology (S4) | DO-PROC | SCM-SVC-PRC-005 | TBD-rev (Evidence Pending — interval) | Draft | Yes
DO-SVC-006 | FRU Replacement Procedures & Repair Limits | Service (S1/S2) | DO-PROC | SCM-SVC-PRC-006 | TBD-rev | Draft | No
DO-SVC-007 | Inter-Trip Cleaning / Decontamination Procedure (reusable, D1) | Service/Mech (S1) | DO-PROC | SCM-SVC-PRC-007 | TBD-rev (Evidence Pending — agents/cycles) | Draft | Yes
DO-SVC-008 | Firmware/Config Upgrade & Backup/Restore Procedure | Service (S3/S8) | DO-PROC | SCM-SVC-PRC-008 | TBD-rev | Draft | No
DO-SVC-009 | Decommissioning & Secure Data-Erase Procedure | Service/Security (S3/S7) | DO-PROC | SCM-SVC-PRC-009 | TBD-rev | Draft | No
DO-SVC-010 | Secure Disposal / Battery End-of-Life & Recycling Procedure | Service/Safety (S2) | DO-PROC | SCM-SVC-PRC-010 | TBD-rev | Draft | No
DOL-1800 — Regulatory & Compliance Outputs (DO-REG) — 12 outputs, 4 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-REG-001 | Standards-Applicability Matrix | Regulatory (all S) | DO-DOC | SCM-REG-DOC-001 | TBD-rev (Evidence Pending — WHO PQS/US class) | Draft | Yes
DO-REG-002 | Regulatory-Configuration Definitions (per-market) | Regulatory (S3/S5/S8) | DO-CFG | SCM-REG-CFG-002 | TBD-rev | Draft | No
DO-REG-003 | Required-Markings & Marking-Placement Specification | Regulatory (S1) | DO-SPEC | SCM-REG-SPC-003 | TBD-rev (Evidence Pending — markings set) | Draft | No
DO-REG-004 | Radio / EMC Certification Configuration & Evidence Index | Regulatory/RF (S5) | DO-CERT | SCM-REG-CRT-004 | TBD-rev (Evidence Pending — certs) | Draft | No
DO-REG-005 | Essential-Performance & Safety-Classification References | Regulatory/Safety (all S) | DO-DOC | SCM-REG-DOC-005 | TBD-rev (Evidence Pending — classification) | Draft | Yes
DO-REG-006 | Environmental / Material Declarations (RoHS / REACH / WEEE) | Regulatory/Mech (S1) | DO-CERT | SCM-REG-CRT-006 | TBD-rev | Draft | No
DO-REG-007 | Battery-Transport Compliance Documentation (UN 38.3 / IATA) | Regulatory/Power (S2) | DO-CERT | SCM-REG-CRT-007 | TBD-rev (Evidence Pending — UN 38.3 result) | Draft | Yes
DO-REG-008 | Cybersecurity Documentation Outputs (SBOM / threat-model refs) | Regulatory/Security (S3/S8) | DO-DOC | SCM-REG-DOC-008 | TBD-rev | Draft | No
DO-REG-009 | Privacy-Control Implementation & Data-Residency Configuration | Regulatory/Security (S8) | DO-CFG | SCM-REG-CFG-009 | TBD-rev | Draft | No
DO-REG-010 | Country-Specific Configuration & Submission Drawings | Regulatory (all S) | DO-DRW | SCM-REG-DRW-010 | TBD-rev (Evidence Pending) | Draft | No
DO-REG-011 | Compliance-Critical Software Version Register | Regulatory/SW (S3/S8/S9) | DO-DOC | SCM-REG-DOC-011 | TBD-rev | Draft | Yes
DO-REG-012 | WHO PQS / EU GDP / US Distribution Compliance Mapping (D4) | Regulatory (all S) | DO-DOC | SCM-REG-DOC-012 | TBD-rev (Evidence Pending — WHO PQS/US) | Draft | No
DOL-1900 — Verification-Enabling Outputs (DO-VER) — 10 outputs, 4 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-VER-001 | Measurable Output-Characteristic Register (CTQ / CTF list) | Quality (all S) | DO-SPEC | SCM-VER-SPC-001 | TBD-rev | Draft | Yes
DO-VER-002 | Critical-Dimension & Tolerance Schedule | Quality/Mech (S1) | DO-SPEC | SCM-VER-SPC-002 | TBD-rev | Draft | No
DO-VER-003 | Performance-Limit & Acceptance-Criteria Definition | Quality (all S) | DO-SPEC | SCM-VER-SPC-003 | TBD-rev (Evidence Pending — targets) | Draft | Yes
DO-VER-004 | Test-Point & Diagnostic-Observability Specification | Quality/Electrical/FW (S2–S7) | DO-SPEC | SCM-VER-SPC-004 | TBD-rev | Draft | No
DO-VER-005 | Verification Logging / Data-Recording Schema | Quality/FW/Cloud (S3/S7/S8) | DO-DATA | SCM-VER-DAT-005 | TBD-rev | Draft | No
DO-VER-006 | Test-Interface & Test-Mode Specification (engineering / production / verification) | Quality/FW/Security (S3) | DO-SPEC | SCM-VER-SPC-006 | TBD-rev | Draft | Yes
DO-VER-007 | Simulation / Reference-Configuration Definitions | Quality/Systems (all S) | DO-MODEL | SCM-VER-MDL-007 | TBD-rev | Draft | No
DO-VER-008 | Verification Fixture & Reference-Standard Requirements | Quality/Metrology (all S) | DO-TOOL | SCM-VER-TOL-008 | TBD-rev | Draft | No
DO-VER-009 | Cold-Chamber / RF-in-Container Verification-Enabling Configuration | Quality/RF/Power (S2/S5) | DO-CFG | SCM-VER-CFG-009 | TBD-rev (Evidence Pending — AD-1/D3) | Draft | No
DO-VER-010 | Objective Pass/Fail Criteria & Verification-Method Mapping | Quality (all S) | DO-SPEC | SCM-VER-SPC-010 | TBD-rev | Draft | Yes
DOL-2000 — Design Transfer & Release Baseline (DO-XFER) — 10 outputs, 2 Critical
Output ID | Title | Discipline | Type | Doc/File No | Rev | Status | Critical?
DO-XFER-001 | Released-Output Index (Design Output Register snapshot) | Systems (all S) | DO-DOC | SCM-XFR-DOC-001 | TBD-rev | Draft | No
DO-XFER-002 | Unresolved-Output / Open-Items List | Systems (all S) | DO-DOC | SCM-XFR-DOC-002 | TBD-rev | Draft | No
DO-XFER-003 | Configuration Audit (FCA / PCA) Report | Systems/Quality (all S) | DO-DOC | SCM-XFR-DOC-003 | TBD-rev | Draft | No
DO-XFER-004 | BOM / Drawing / Software / Manufacturing Reconciliation | Systems/Mfg (all S) | DO-DOC | SCM-XFR-DOC-004 | TBD-rev | Draft | No
DO-XFER-005 | Production-Test Readiness Assessment | Systems/Quality (all S) | DO-DOC | SCM-XFR-DOC-005 | TBD-rev | Draft | No
DO-XFER-006 | Supplier / Tooling / Fixture Readiness Assessment | Systems/Mfg (all S) | DO-DOC | SCM-XFR-DOC-006 | TBD-rev | Draft | No
DO-XFER-007 | Packaging / Labeling / Service / Training Readiness Assessment | Systems/Mfg/Service (all S) | DO-DOC | SCM-XFR-DOC-007 | TBD-rev | Draft | No
DO-XFER-008 | Open-Risk Disposition & HAL-Reconciliation Record | Systems/Safety (all S) | DO-DOC | SCM-XFR-DOC-008 | TBD-rev | Draft (Preliminary) | Yes
DO-XFER-009 | Pilot-Build Baseline & Transfer-Deviation Log | Systems/Mfg (all S) | DO-DOC | SCM-XFR-DOC-009 | TBD-rev | Draft | No
DO-XFER-010 | Production-Release Baseline & Release Approvals | Systems/Quality (all S) | DO-DOC | SCM-XFR-DOC-010 | TBD-rev | Draft | Yes
---
3. Applicability Register
Scope is never silently dropped (Master Specification §5).
Item | Status | Rationale
DOL-000…DOL-2000 (all 21 modules) | Applicable | Globally distributed, regulated, reusable, connected, safety-significant IoT device spanning device + firmware + cloud + mobile; every discipline produces controlled outputs.
DOL-700 Device Software | Applicable (lean) | No rich on-device OS/app; DO-SW covers host tools, provisioning/calibration configs, SBOM, and factory-image packaging (8 outputs).
SCM-100-XA external-antenna branch | Applicable — Evidence Pending | Emitted as a configured variant (DO-EE-021, DO-MECH-016); hardware branch confirmed only if the integral antenna cannot close the in-container link budget (AD-1, PRD-VER-005).
AI/ML, audio/video, payments output classes | Not Applicable | No such scope in the SCM (concept; Stage 07 §3). Recorded, not silently dropped.
---
4. Critical-output roll-up (severity-first)
103 of 272 outputs are Critical. The severity-dominant set (Action Priority per CLAUDE.md rule 3):
D3 charge/energy safety spine | DO-EE-003/004/005/006/007/009/010, DO-FW-004/005, DO-SAFE-001/002/003/012, DO-TST-009, DO-EE-002 (write-safe rails).
Excursion / verdict / record integrity | DO-FW-006/008/010, DO-EE-013, DO-SAFE-004/005/006/008, DO-CLD-002/003/006, DO-SYS-005, DO-TST-005/006.
Alert delivery & connectivity | DO-FW-011, DO-CLD-007, DO-IF-001/003/005/007/008, DO-SAFE-009.
Security / integrity of record & identity | DO-SEC-001…008, DO-FW-014/016/019, DO-CFG-010, DO-EE-012.
Ingress / seal & battery mechanical safety | DO-MECH-005/006/012, DO-PCB-004/010, DO-SAFE-010/011, DO-TST-011.
Compliance & release gates | DO-REG-001/005/007/011, DO-LBL-002/004/005, DO-VER-001/003/006/010, DO-XFER-008/010.
All DOL-1200 (DO-SAFE-001…014) entries are Critical and Preliminary — reconcile with HAL Stage 10 (RA-05).
---
5. Related files
• Product-Design-Output-Master-Specification.md — controlling master document (hierarchy, release checklist).
• DOL-000-governance.md … DOL-2000-design-transfer.md — the 21 per-module discipline packages (full 24-field entries).
• Design-Input-Output-Traceability-Matrix.md — requirement → output → risk-control → V&V → manufacturing chains.
• index.md — navigation and status dashboard.
Product Design Output Master Specification — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 09 — Design Outputs (ORCH-013) · Owner/Integrator: Chief Systems Engineer (EM-ENG-001) · Discipline owners: Principal Engineers per DOL module (see §7) · Status: Authored — instantiated, product-specific Design Output Baseline. Controlled document per ISO 13485:2016 §7.3.4 and the FDA QMSR (21 CFR 820 as amended, effective 2026-02-02, which incorporates ISO 13485:2016 by reference).
This is the central controlled document of the SCM Design Output Library (DOL). It defines the approved
product design output baseline, the output hierarchy, the register summary, the traceability posture, and
the release-assessment checklist. It governs the Design Output Register (Design-Output-Register.md
and the per-module DOL-.md files) and the Design Input–Output Traceability Matrix
(Design-Input-Output-Traceability-Matrix.md). Design outputs are the results of the design effort —
the CAD, drawings, schematics, layouts, source code, configurations, specifications, manufacturing data,
and acceptance criteria from which the approved SCM can be reproduced.
---
0. Document control block
Field | Value
Document number | SCM-DOL-MASTER-001
Revision | TBD-rev (Evidence Pending — baseline set at Design Output Review DR-DO-1)
Lifecycle status | Draft / Under design (pre-DR-DO release)
Author / Integrator | Chief Systems Engineer (EM-ENG-001)
Approval authority | Design Review Board (chaired by EM-ENG-001); Quality/Regulatory (EM-ENG-006) concurrence
Effective date | Evidence Pending (on DR-DO-1 approval)
Supersedes | None (first issue)
Classification | Proprietary — controlled
Change control | Per DOL-000 (DO-GOV) change-control workflow; all revisions via ECO
Related controlled docs | Stage 05 PRD library, Stage 06 SYS library, Stage 07 subsystem/SRS, Stage 08 architecture, Stage 10 HAL, Stage 13 PVL, Stage 15 design transfer
Evidence posture (honesty). Per 00-governance/evidence-standard.md and hallucination-policy.md, every entry carries a evidence label — Cited (external source), Standard (named standard, revision to confirm), Judgment (engineering inference), or Evidence Pending (unresolved value/branch closed by a named test). No part numbers, manufacturer names, datasheet values, or revisions are invented. Document/file numbers use the controlled scheme below; concrete revisions are shown as TBD-rev and the output marked Evidence Pending until the artifact is actually authored and released under DOL-000.
---
1. Purpose & scope
Purpose. Define exactly what must be built, configured, assembled, programmed, inspected, tested, labeled, packaged, installed, serviced, and released to reproduce the approved SCM. Each design output is traceable to and satisfies its design inputs (PRD/SYS/SRS/HWR/MER/FWR/IFR), references acceptance criteria, is reviewed and approved before release, and — where essential to safe and proper use — is flagged Critical.
Scope. All 21 DOL modules (DOL-000…DOL-2000) are Applicable to the SCM — it is a globally distributed, regulated, reusable, connected, safety-significant IoT device spanning device + firmware + cloud + mobile tiers. No module is excluded (see §5 Applicability Register). This baseline covers the SCM product family (§3) at the design-output level; it does not contain verification evidence, validation evidence, manufacturing records, or quality records — those are distinct artifacts (DOL-000 § output-vs-record distinction).
Out of scope / not applicable. No AI/ML, audio/video, or payment scope exists for the SCM; those extension output classes are not instantiated. Recorded in §5.
---
2. How to read the register (conventions used in every module)
Each controlled output is one register entry carrying the 24 fields of the DOL standard, written as a compact block:
Line 1 | ### DO-<CODE>-NNN — Output Title
• Type · Discipline · SKU Applicability · Subsystem/Component (with CMP-### where applicable)
• File # · Revision · Lifecycle Status · Design Owner · Approval Authority
• Source Format · Released Format · Repository Location · Parent Assembly
• Related Outputs · Input Traceability · Risk-Control Traceability · Verification Traceability
• Critical Output · Acceptance Criteria · Manufacturing Use · Change Classification · Notes/Assumptions (evidence)
Output Type taxonomy (Output Type field): DO-DOC controlled document · DO-DRW drawing · DO-CAD CAD model · DO-SCH schematic · DO-PCB PCB mfg data · DO-BOM bill of materials · DO-SPEC engineering specification · DO-ICD interface-control document · DO-CODE source code · DO-BIN released binary · DO-CFG configuration · DO-DATA data schema/dataset · DO-IAC infrastructure-as-code · DO-WI work instruction · DO-PROC procedure · DO-TEST production-test definition · DO-TOOL tooling/fixture · DO-LBL label/artwork · DO-MAN manual · DO-CERT certificate/declaration · DO-MODEL analytical model.
Lifecycle Status values: Draft · In review · Released · Superseded · Obsolete. In this pre-DR-DO baseline the great majority are Draft or In review; none are Released yet.
File-number scheme: SCM-<AREA>-<TYPE>-NNN where AREA ∈ {SYS, GOV, CFG, MEC, ELE, PCB, FW, SW, CLD, APP, IF, SEC, SAF, HF, MFG, TST, LBL, SVC, REG, VER, XFR}. Governed by DOL-000.
SKU applicability shorthand: All = every SKU. Variant codes per DOL-100: SCM-100 (base), SCM-100-NA / -EU / -ROW (regional LTE-M band / radio-regulatory variants), SCM-100-XA (external-antenna option per ADR-001/AD-1). "All" ⇒ common to every variant.
Subsystem legend (from Stage 08): S1 Enclosure/Mechanical · S2 Power · S3 Compute/MCU+FW · S4 Sensing · S5 Connectivity · S6 HMI · S7 Storage · S8 Cloud · S9 Mobile.
---
3. Product output baseline (what this baseline reproduces)
Product: the SCM — a reusable, battery-operated, cellular-connected environmental data logger placed inside insulated vaccine shippers, logging temperature (internal + external probe), humidity, and shock; detecting two-sided excursions; giving an offline at-the-door accept/reject/quarantine verdict; storing a tamper-evident record; and uploading via store-and-forward (LTE-M primary, Wi-Fi/BLE offload) with immediate excursion alerts. Operating range −30 °C…+60 °C; endurance target 90 days (D1–D4).
Product family / SKUs (DOL-100, DO-CFG):
SKU | Description | Primary variance
SCM-100 | Base SCM device | Reference configuration
SCM-100-NA | North-America radio-regulatory variant | LTE-M band plan; FCC/IC markings; SIM/eSIM profile
SCM-100-EU | EU/EEA variant | LTE-M band plan; CE/RED markings; EU data-residency config
SCM-100-ROW | Rest-of-world variant | LTE-M band plan; regional markings; carrier strategy
SCM-100-XA | External-antenna option (any region) | IP-sealed external-antenna connector (ADR-001, CMP-029)
Firmware / cloud / mobile are common across SKUs; regional differences are configuration (DO-CFG), radio band plan (DO-EE/DO-IF), and labeling/regulatory (DO-LBL/DO-REG), not separate hardware designs, except the SCM-100-XA antenna branch (Evidence Pending, AD-1).
Top program tensions carried into the outputs (severity-first):
1. D3 — 90-day endurance at −30 °C + cold-charge safety. The low-temperature chemistry (CMP-006), the energy-governor firmware (CMP-037), the charge-inhibit-below-0 °C interlock (CMP-008/010/012, INT-010/016), and the excursion-detection logic (CMP-039) are the load-bearing outputs. Every safety-relevant one is flagged Critical and "reconcile with HAL Stage 10" (Preliminary).
2. RF-in-container link (AD-1). Antenna design (CMP-028/029) and RF ICDs are Evidence Pending, closed by the RF-in-container qualification test (PRD-VER-005).
---
4. Output hierarchy (DOL structure)
SCM Product Design Output Baseline (SCM-DOL-MASTER-001) ├── DOL-000 Design Output Governance ............... DO-GOV (governance/CM rules) ├── DOL-100 Product Configuration Baseline ......... DO-CFG (family tree, SKUs, AML) ├── DOL-200 System Architecture Outputs ............ DO-SYS (arch, ICD overview, modes) ├── DOL-300 Mechanical Design Outputs .............. DO-MECH (S1 enclosure/seals/probe/mounts) ├── DOL-400 Electrical & Electronic Outputs ........ DO-EE (S2/S3/S4/S5/S6 electrical design) ├── DOL-500 PCB & PCBA Outputs .................... DO-PCB (board, stack-up, Gerber, PCBA BOM) ├── DOL-600 Embedded Firmware Outputs ............. DO-FW (S3 firmware, signed image) ├── DOL-700 Device Software Outputs ............... DO-SW (device-side tools/config/SBOM) ├── DOL-800 Cloud Platform Outputs ............... DO-CLD (S8 IaC, schemas, APIs) ├── DOL-900 Mobile & Web Application Outputs ...... DO-APP (S9 app) ├── DOL-1000 Connectivity & Interface Outputs ...... DO-IF (ICDs, protocol specs) ├── DOL-1100 Cybersecurity Design Outputs ......... DO-SEC (crypto/keys/secure-boot config) ├── DOL-1200 Safety & Risk-Control Outputs ........ DO-SAFE (implements HAL RCRs — Stage 10) ├── DOL-1300 Human Factors & UI Outputs ........... DO-HF (HMI/verdict/labeling content) ├── DOL-1400 Manufacturing Design Outputs ......... DO-MFG (WIs, fixtures, process params) ├── DOL-1500 Inspection & Production-Test Outputs .. DO-TST (test specs, limits, fixtures) ├── DOL-1600 Packaging, Labeling & Logistics ...... DO-LBL (labels, UDI, IFU, packaging) ├── DOL-1700 Installation, Service & Maintenance ... DO-SVC (service manual, FRU, disposal) ├── DOL-1800 Regulatory & Compliance Outputs ...... DO-REG (standards matrix, markings, config) ├── DOL-1900 Verification-Enabling Outputs ........ DO-VER (test points, limits, test modes) └── DOL-2000 Design Transfer & Release Baseline ... DO-XFER (released-output index, audits)
---
5. Applicability Register (scope never silently dropped)
Module | Status | Rationale
DOL-000…DOL-2000 (all 21) | Applicable | Regulated, reusable, connected, safety-significant device + FW + cloud + mobile; every discipline produces controlled outputs.
DOL-700 Device Software | Applicable (lean) | The SCM has no rich on-device OS/app; DO-SW covers device-side host tools, provisioning/calibration configs, SBOM, and factory-image packaging rather than an application OS image.
AI/ML, audio/video, payments output classes | Not Applicable | No such scope in the SCM (concept, Stage 07 §3). Recorded, not silently dropped.
SCM-100-XA external-antenna branch | Applicable — Evidence Pending | Emitted as a configured variant; hardware branch confirmed only if integral antenna cannot close the in-container link budget (AD-1, PRD-VER-005).
---
6. Register summary (counts by module)
Counts are the instantiated SCM baseline entries in each DOL-.md. All are Draft/In-review in this
pre-DR-DO baseline; revisions are TBD-rev (Evidence Pending) until authored/released under DOL-000.
Module | Prefix | File | Entries | Critical | Notes
DOL-000 | DO-GOV | DOL-000-governance.md | 9 | 0 | CM/governance outputs
DOL-100 | DO-CFG | DOL-100-config-baseline.md | 12 | 2 | Family tree, SKU/AML, config rules
DOL-200 | DO-SYS | DOL-200-system-architecture-outputs.md | 12 | 3 | Arch, ICD overview, modes, fault mgmt
DOL-300 | DO-MECH | DOL-300-mechanical.md | 20 | 4 | Enclosure, seals, probe, mounts, IP
DOL-400 | DO-EE | DOL-400-electrical.md | 22 | 8 | Power tree, battery/BMS/charge, sensors, radios
DOL-500 | DO-PCB | DOL-500-pcb.md | 14 | 3 | Board, stack-up, Gerber, PCBA BOM
DOL-600 | DO-FW | DOL-600-firmware.md | 20 | 9 | FW architecture, governor, detection, signed image
DOL-700 | DO-SW | DOL-700-device-software.md | 8 | 2 | Device tools, factory image, SBOM
DOL-800 | DO-CLD | DOL-800-cloud.md | 16 | 4 | IaC, schemas, APIs, OTA service
DOL-900 | DO-APP | DOL-900-mobile.md | 10 | 2 | Mobile app, GATT client, signing
DOL-1000 | DO-IF | DOL-1000-interfaces.md | 12 | 6 | ICDs (device↔cloud, BLE, OTA, enterprise)
DOL-1100 | DO-SEC | DOL-1100-security.md | 12 | 7 | Crypto, keys, secure boot, provisioning
DOL-1200 | DO-SAFE | DOL-1200-safety-risk-controls.md | 14 | 14 | HAL risk-control implementations
DOL-1300 | DO-HF | DOL-1300-human-factors.md | 10 | 3 | HMI/verdict/warning presentation
DOL-1400 | DO-MFG | DOL-1400-manufacturing.md | 14 | 3 | WIs, fixtures, process params, serialization
DOL-1500 | DO-TST | DOL-1500-inspection-test.md | 14 | 5 | Incoming/final test, calibration, golden units
DOL-1600 | DO-LBL | DOL-1600-packaging-labeling.md | 12 | 3 | Labels, UDI, IFU, packaging, transport
DOL-1700 | DO-SVC | DOL-1700-service.md | 10 | 2 | Service/install manual, FRU, disposal
DOL-1800 | DO-REG | DOL-1800-regulatory.md | 12 | 4 | Standards matrix, markings, reg config
DOL-1900 | DO-VER | DOL-1900-verification-enabling.md | 10 | 4 | Test points, limits, test modes
DOL-2000 | DO-XFER | DOL-2000-design-transfer.md | 10 | 2 | Released-output index, audits, release baseline
Total |  | 21 files | ~253 | ~86
---
7. Ownership (per DOL standard §Roles)
Module(s) | Owner
DOL-000, DOL-100, DOL-200, DOL-1000, DOL-2000 | Chief Systems Engineer (EM-ENG-001)
DOL-300 | Principal Mechanical Engineer
DOL-400 | Principal Electrical/Hardware Engineer (+ Battery/Power lead for S2)
DOL-500 | Principal PCB Engineer
DOL-600, DOL-700 | Embedded Software Engineer (EM-ENG-007)
DOL-800 | Principal Cloud Engineer
DOL-900 | Principal Mobile Engineer
DOL-1100 | Product Security Engineer (EM-ENG-008)
DOL-1200 | Reliability/Safety Engineer (reconciles with HAL, Stage 10)
DOL-1300 | Human Factors Engineer
DOL-1400, DOL-1600 | Manufacturing Engineer
DOL-1500, DOL-1900 | Quality Engineer
DOL-1700 | Service Engineer
DOL-1800 | Quality/Regulatory Engineer (EM-ENG-006)
Orchestrated by ORCH-013 — Design Outputs. Integrated by EM-ENG-001.
---
8. Traceability summary
Every output traces along all applicable chains (see Design-Input-Output-Traceability-Matrix.md):
Requirement chain | MR → UN → PRD → SYS → HWR/MER/FWR/IFR/SRS → Design Output → PV (verification) → VAL (validation).
Risk chain | Hazard (HA-<CODE>-NNN) → Risk-Control Requirement (RCR-###) → Design Output (DOL-1200 primarily) → risk-control verification. Safety/charge/excursion outputs are Preliminary — reconcile with HAL Stage 10.
Manufacturing chain | Design Output → manufacturing process (PROC-###) → process control (CP-###) → inspection/production test (DO-TST, PV-###).
Forward consumers of this baseline: Hazard Analysis (Stage 10, ORCH-017 — DOL-1200 supplies the "Implementation Reference" for each risk control); Verification (Stage 13, ORCH-014/PVL — DOL-1900 supplies the observability and acceptance criteria); Manufacturing & Design Transfer (Stage 15 — DOL-1400/1500/1600/ 2000); and the end-to-end lifecycle matrix (Stage 18, ORCH-018 / EM-ENG-012).
---
9. Release assessment checklist (Design Output Review — DR-DO gate)
The baseline is releasable at DR-DO when each item below is satisfied and recorded. Current state: PRE-GATE (Draft).
# | Release criterion | State | Blocking gaps
RA-01 | Every applicable design input (PRD/SYS/SRS/HWR/MER/FWR/IFR) traces to ≥1 design output | Partial | Matrix drafted; numeric-target inputs Evidence Pending
RA-02 | Every output traces upward to ≥1 design input | Met (as drafted) | —
RA-03 | Every output has stated acceptance criteria | Met (as drafted) | Some criteria Evidence Pending on test results
RA-04 | Every Critical output identified and flagged | Met | —
RA-05 | Every HAL risk control (Stage 10) has an implementing output in DOL-1200 | Open | HAL not yet baselined; DOL-1200 entries Preliminary — reconcile with HAL
RA-06 | Every output reviewed & approved before release | Open | Pre-DR-DO; none Released
RA-07 | Configuration baseline (DOL-100) frozen; AML approved | Open | Chemistry (CMP-006) & antenna (CMP-028/029) Evidence Pending
RA-08 | Interface outputs (DOL-1000) internally consistent; no undocumented interface | Partial | RF ICDs Evidence Pending (AD-1)
RA-09 | Manufacturing/test outputs sufficient to build & accept a unit | Partial | Depends on frozen AML (RA-07)
RA-10 | Regulatory config & markings (DOL-1800) resolved per target markets (D4) | Open | WHO PQS category / US FDA classification Evidence Pending
RA-11 | Verification-enabling outputs (DOL-1900) expose all acceptance-critical characteristics | Partial | RF & cold-chamber test modes Evidence Pending
RA-12 | Design transfer index (DOL-2000) reconciled; open-risk disposition recorded | Open | Assembled at Stage 15
Top release blockers (severity-first): RA-05 (HAL reconciliation of charge-inhibit / excursion logic / safe-state), RA-07 (chemistry + antenna freeze), RA-10 (regulatory classification). These are the honest gaps; none is invented closed.
---
10. Related files
• Design-Output-Register.md — consolidated register (all modules, one place).
• DOL-000-governance.md … DOL-2000-design-transfer.md — the 21 per-module discipline packages.
• Design-Input-Output-Traceability-Matrix.md — requirement → design output chains.
• index.md — navigation and status dashboard. </content> </invoke>

## 10 Hazard Analysis
Hazard Analysis
How the product could expose people or vaccines to harm, and the safeguards that prevent it.
↩ Contents
In this section:
• Hazard Analysis Library (HAL) — Smart Cold Chain Monitor (SCM) — Index
• Detailed Product Hazard Analysis — Smart Cold Chain Monitor (SCM)
• HAL-000 (HA-GOV) — Hazard Analysis Governance & Plan
• HAL-100 (HA-CTX) — Product & Analysis Context
• HAL-1000 (HA-SW) — Software-Controlled Hazards
• HAL-1100 (HA-COMM) — Connectivity Hazards
• HAL-1200 (HA-CLOUD) — Cloud-Platform Hazards
• HAL-1300 (HA-CYBER) — Cybersecurity-Induced Safety Hazards
• HAL-1400 (HA-DATA) — Data & Decision-Integrity Hazards
• HAL-1500 (HA-USE) — Use-Related Hazards
• HAL-1600 (HA-SVC) — Service & Maintenance Hazards
• HAL-1700 (HA-MFG) — Manufacturing & Supply-Chain Hazards
• HAL-1800 (HA-MPE) — Mission, Property & Environmental Hazards
• HAL-1900 (HAZ-) — Hazard Taxonomy / Hazard-Class Library
• HAL-200 (HA-BND) — System Boundary & Exposure Model
• HAL-2000 (HA-SEV) — Severity
• HAL-2100 (HA-PROB) — Probability
• HAL-2200 (HA-ACC) — Risk Acceptability
• HAL-2300 (HA-HIER) — Risk-Control Hierarchy applied to the SCM
• HAL-2400 (HA-RCR) — Risk-Control Requirements
• HAL-2500 (HA-CIR) — Risk-Control Implementation Review
• HAL-2600 (HA-RCV) — Risk-Control Verification
• HAL-2700 (HA-NEW) — New Risks Introduced by Risk Controls
• HAL-2800 (HA-RES) — Individual Residual Risk
• HAL-2900 (HA-ORES) — Overall Residual Risk & Benefit-Risk
• HAL-300 (HA-LIFE) — Lifecycle Hazard Analysis
• HAL-400 (HA-IUE) — Intended Use & Human Exposure Questionnaire
• HAL-500 (HA-ENRG) — Energy Hazards
• HAL-600 (HA-MECH) — Mechanical Hazard Analysis Register
• HAL-700 (HA-ELEC) — Electrical Hazard Analysis Register
• HAL-800 (HA-TFC) — Thermal / Fire / Chemical Hazards
• HAL-900 (HA-ENV) — Environmental Hazards
• Hazard Traceability Matrix — Smart Cold Chain Monitor (SCM)
• Preliminary Hazard Analysis (PHA) — Smart Cold Chain Monitor (SCM)
• Risk-Management Summary — Smart Cold Chain Monitor (SCM)
Hazard Analysis Library (HAL) — Smart Cold Chain Monitor (SCM) — Index
Stage 10 (ORCH-017) · Owner: System-Safety with EM-ENG-001 (Chief Systems) and EM-ENG-006 (Quality/ Regulatory). The master index for the SCM's Hazard Analysis Library: the full module set (HAL-000…HAL-2900), the connected risk-management artefacts, the applicability register (which modules are active for this product), and the downstream handoffs. Every module follows the ISO 14971 harm-and-exposure-first method with the 34-field HA-<CODE>-NNN record and the enforced risk chain; severity-first per CLAUDE.md rule 3. All PV-/VAL- references and residual bands are Evidence-Pending (Stages 13–14). Standards cited by designation.
---
1. How the library is organised
Foundations (HAL-000…400) | governance, context, boundary/exposure, lifecycle, intended-use exposure — why and how the analysis is scoped, and the exposure models it defends (esp. HA-BND-003, the information/decision pathway that is the SCM's dominant harm route).
Domain hazard registers (HAL-500…1800) | the discipline-by-discipline hazard records (~120 total, 34-field, full risk chain), each a complete register owned by a principal engineer.
Hazard taxonomy (HAL-1900) | the HAZ- hazard classes the records draw from.
Analysis + control + residual apparatus (HAL-2000…2900) | severity, probability, acceptability, the RCR register, control-implementation review, verification plan, new-risk review, and individual + overall residual risk.
Connected artefacts | the concept-stage PHA, the post-design Detailed PHA, the Traceability Matrix, and the Risk-Management Summary (release recommendation).
2. Module table (HAL-000 … HAL-2900)
Foundations
Module | Code | Title | Role
HAL-000 | — | Governance & Plan | Risk-management plan, method, 34-field record, evidence posture
HAL-100 | HA-CTX | Context / Intended Use | Purpose, foreseeable use/misuse, vulnerable last-mile users, dependencies
HAL-200 | HA-BND | Boundary & Exposure Models | System boundary; energy, data/decision (HA-BND-003), control-flow, trust-boundary, human, dependency models
HAL-300 | — | Lifecycle | Lifecycle phases; reuse/service amplifiers (D1)
HAL-400 | HA-IUE | Intended-Use Exposure | Automation bias, unmonitored-believed-monitored, safety-depends-on-information
Domain hazard registers
Module | Code | Domain | Headline hazards
HAL-500 | HA-ENRG | Energy / Battery | HA-ENRG-001 cold-charge (D3), HA-ENRG-006 dead-in-cold
HAL-600 | HA-MECH | Mechanical | edge/pinch/small-parts; battery retention
HAL-700 | HA-ELEC | Electrical / EMC | ESD/EMC false read; latent defect; brown-out
HAL-800 | HA-TFC | Thermal / Fire / Chemical | HA-TFC-002 aircraft-hold fire (S5)
HAL-900 | HA-ENV | Environmental | freeze miss; condensation; altitude
HAL-1000 | HA-SW | Software | HA-SW-001/002/008/013 detection/verdict/probe/silent-gap
HAL-1100 | HA-COMM | Connectivity | HA-COMM-003 RF-in-container (Open); delayed/missing alert
HAL-1200 | HA-CLOUD | Cloud | notify fail; rules misconfig; outage
HAL-1300 | HA-CYBER | Cybersecurity → safety | falsified data; suppressed alarm; malicious OTA; cloud compromise
HAL-1400 | HA-DATA | Data & Decision Integrity | HA-DATA-004 automation bias; attribution; provenance
HAL-1500 | HA-USE | Use-Related / Human Factors | misread verdict; alarm fatigue; mis-arm; override
HAL-1600 | HA-SVC | Service / Maintenance | depot miscal; wrong FW; seal; debug mode; disposal
HAL-1700 | HA-MFG | Manufacturing / Supply Chain | EOL cal escape; wrong chemistry/parts; latent defect
HAL-1800 | HA-MPE | Mission / Property / Environmental | HA-MPE-002 population outbreak (S5); false-reject stock-out; GDP; e-waste
HAL-1900 | HAZ- | Hazard Taxonomy | the library's hazard classes
Analysis, control & residual apparatus
Module | Code | Title | Role
HAL-2000 | — | Severity | S1–S5 scale; SCM severity anchors (wrong disposition; aircraft-fire S5)
HAL-2100 | — | Probability | P1×P2 decomposition; basis taxonomy; Evidence-Pending posture
HAL-2200 | HA-ACC | Acceptability | Matrix, regions R/A/B, Action Priority, decision rules, acceptance authority
HAL-2300 | — | Control Hierarchy | inherent → protective → information → process; warnings-challenge
HAL-2400 | HA-RCR | Risk-Control Requirements | RCR-001…058 register (testable, allocated, implemented, verified)
HAL-2500 | HA-CIR | Control-Implementation Review | 11-point audit of each key control; independence/defeatability/dependency; findings
HAL-2600 | HA-RCV | Risk-Control Verification | method + evidence + PV-/VAL- per control; FI for safety, IV for security
HAL-2700 | HA-NEW | New Risks Introduced | control side-effects (charge-inhibit→DOA, self-heat, encryption, alarm fatigue, redundancy, safety-vs-availability)
HAL-2800 | HA-RES | Individual Residual Risk | per-hazard residual, effectiveness, evidence status, acceptance authority, disclosure/monitoring
HAL-2900 | HA-ORES | Overall Residual & Benefit-Risk | accumulation, convergence, common-cause, systemic dependence, rare-catastrophic, uncertainty; conditional benefit-risk
Connected risk-management artefacts
Artefact | Role
Preliminary-Hazard-Analysis.md | Concept-stage PHA → early safety characteristics + preliminary RCRs (PHA-01…11)
Detailed-Product-Hazard-Analysis.md | Post-design detailed analysis; the dominant harm-path spine + module map
Hazard-Traceability-Matrix.md | Authoritative chain: HA → situation → harm → initial → RCR → DO- → PV-/VAL- → residual → field monitoring (~60 rows)
Risk-Management-Summary.md | ISO 14971 report: open risks, residual profile, benefit-risk, assumptions, V&V status, release recommendation
index.md | This master index
3. Applicability Register (which modules are active for the SCM)
Per CLAUDE.md rule 5 (product-specific convening), not every module is equally weighted; all are active for the SCM given its risk profile (battery + information device + connected + reusable + air-transport):
Module | Active? | Weight | Rationale
HAL-100/200/400 (context/boundary/IUE) | Yes | High | Information device — the exposure models are load-bearing
HAL-500 (energy) | Yes | Critical | Rechargeable Li battery + D3 cold-charge tension
HAL-600 (mechanical) | Yes | Medium | Handling, retention, small-parts
HAL-700 (electrical) | Yes | High | Sensing/record electronics; EMC; brown-out
HAL-800 (thermal/fire/chemical) | Yes | Critical | S5 aircraft-hold fire
HAL-900 (environmental) | Yes | High | −30…+60 °C, condensation, altitude, freeze detection
HAL-1000 (software) | Yes | Critical | Detection/verdict/record/alert all in firmware
HAL-1100 (connectivity) | Yes | Critical | Alert delivery; RF-in-container carried-Open
HAL-1200 (cloud) | Yes | High | Notify/rules/record/outage
HAL-1300 (cybersecurity) | Yes | High | Security → safety (falsified data, OTA, cloud)
HAL-1400 (data & decision) | Yes | Critical | Automation bias; the artefact the human acts on
HAL-1500 (use-related) | Yes | Critical | Last-mile low-training/vulnerable users
HAL-1600 (service) | Yes | High | Reuse (D1) — recal/seal/firmware/disposal
HAL-1700 (manufacturing) | Yes | High | EOL calibration/parts/identity/config
HAL-1800 (mission/property/env) | Yes | Critical | Population-scale S5; false-reject stock-out; GDP; e-waste
HAL-2000…2900 (apparatus) | Yes | — | The analysis/control/residual framework for all of the above
(Not applicable) | — | — | No implanted/sterile/radiation/mains-powered hazards (device is battery/SELV, non-therapeutic) — those library classes are N-A for the SCM
4. Top hazards at a glance (the release-relevant set)
Physical S5 | HA-TFC-002 (Li-cell fire in an occupied aircraft cargo hold) — region-A-High, never broadly acceptable, mandatory disclosure + monitoring.
Information S4/S5 (the dominant harm) | HA-SW-001/002/008/013, HA-DATA-001/004, HA-ENV-001, HA-COMM-001/ 002/003, HA-CYBER-001/002/005, HA-USE-001/002/004/007, HA-MPE-002 — all converging on spoiled dose administered / potent dose destroyed through the disposition decision, contained by two independent choke points (honest verdict + supported human decision).
Carried-Open unknowns | HA-COMM-003 (RF-in-container → PRD-VER-005) and the energy margin of HA-ENRG-006 (D3 cold endurance → PRD-VER-004).
Accepted public-health trade | HA-MPE-003 (false-reject → stock-out → unvaccinated population) — balanced, monitored.
5. Downstream handoffs
• Risk-control requirements feed Stage 12 (Risk-Control Requirements) and re-baseline the preliminary PRD-SAFE requirements + Product/Design inputs (HAL-2400 → Stage 12).
• High-severity (S4/S5) hazardous situations become FTA top events in Stage 11 (ERFKL) and drive the deeper SW-FMEA / HW-FMEA / battery-abuse DFMEA; the HAL-2900 convergence fault tree ("spoiled dose administered as good") and common-cause list seed that work.
• PV-/VAL- references feed Stages 13–14 (Verification PVL / Validation PVaL) — the verification plan (HAL-2600) is the safety-driven coverage matrix those stages inherit; fault-injection for safety controls, independent adversarial verification for security controls.
• Disclosure / field-monitoring obligations feed Stage 17 (post-market surveillance) — both harm directions (missed-excursion and false-reject/wastage), thermal-event, and security monitoring (ISO 14971 §9).
• The whole chain feeds ORCH-018 / EM-ENG-012 (complete lifecycle traceability) — the Hazard-Traceability- Matrix is the Stage-10 spine consumed there.
Release governance | the Risk-Management Summary's recommendation is Conditionally Ready — release pending V&V, with eight explicit release conditions.
6. Status
• Stage 10 (ORCH-017): complete. All modules HAL-000…2900 authored; connected artefacts authored; risk controls defined/allocated/implemented-referenced/reviewed; residual risk evaluated (projected); overall benefit-risk concluded (conditional).
• All residuals and PV-/VAL- are Evidence-Pending until Stages 13–14 execute; two items carried Open (PRD-VER-004/005); the S5 residual carried region-A-High with mandatory monitoring.
Evidence posture | every judgement is Cited / Standard / Judgment-Assumption; no probability or performance number asserted as fact; the three load-bearing unknowns (D3, RF-in-container, WHO PQS accuracy target) are named, not invented.
Detailed Product Hazard Analysis — Smart Cold Chain Monitor (SCM)
Stage 10 (ORCH-017) · Owner: System-Safety with EM-ENG-001 and EM-ENG-006. This is the post-design detailed hazard analysis: the complete, product-specific working-out of the SCM's hazards against the actual design (the SYS/architecture, the design outputs DO-…, and the component list CMP-…), organised into the HAL domain modules and consolidated here. It supersedes the concept-stage Preliminary-Hazard-Analysis.md by evidence each hazard in the implemented design and binding it to a controlled design output and a planned verification. This document is the navigational + narrative summary; the authoritative per-hazard records (the 34-field HA-<CODE>-NNN entries) live in the domain modules it points to. ISO 14971 harm-and-exposure- first; severity-first per CLAUDE.md rule 3. Standards cited by designation.
---
1. Purpose & relationship to the domain modules
The detailed analysis is distributed across the HAL domain modules (each a discipline's complete register) and the analysis modules (severity, probability, acceptability, controls, and residual). This document does three things the individual modules cannot: (a) states the dominant SCM harm path end-to-end and shows how the whole design defends it; (b) gives the map of where each hazard family is analysed; and (c) summarises the detailed findings so a reviewer can navigate the ~120-record library without reading every module. It is the "detailed product hazard analysis" artefact required by the hazard-analysis standard, sitting between the PHA (concept) and the Risk-Management Summary (release).
2. The dominant harm path (the spine of the analysis)
The SCM is an information device (HA-BND-003): it never physically touches a dose, yet its dominant harm is mediated by the information it produces. The dominant path, stated in full:
Missed / incorrect / delayed excursion detection or alert → the excursion is invisible or unactioned → the
monitored shipment is dispositioned "accept" → an undetected spoiled (potency-lost) vaccine is administered →
the vaccine does not protect the beneficiary → preventable disease — at a single beneficiary, a lot, or
(through a systemic/fleet-wide cause) a population.
Its mirror path is equally real: a false reject / lost record → a potent lot destroyed → stock-out → an unvaccinated population → preventable disease. The design is built so that however the upstream information chain fails, it converges on "not a clean accept" (an honest reject/qualify/void) rather than a false clean accept, and that the record survives to be read honestly at receipt even with no connectivity. The two independent choke points every failure path must pass are:
1. Honest verdict / no-false-reassurance (RCR-021/022/040, DO-SAFE-006): a spoiled, gapped, faulted, or unverifiable trip cannot present as a clean accept; device and cloud verdicts cross-check.
2. Supported human disposition (RCR-043/044, DO-HF-002): a legible, language-independent verdict readable gloved/cold/low-light, with a deliberate, recorded override.
Every detailed module is, in effect, defending one segment of this spine.
3. How the design defends the spine — segment by segment
Spine segment | What can break it (detailed records) | The design defence (DO- / RCR-)
Measure the product-zone temperature truthfully | Probe absent/open/wrong/mis-placed; biased/stale calibration; sensor drift (HA-SW-008/004/014, HA-DATA-001, HA-USE-005/006, HA-MFG-001, HA-SVC-001) | HW presence/open/short + FW plausibility (DO-SAFE-005/RCR-020); integrity+expiry calibration gate (DO-FW-007/RCR-026); EOL calibration + birth record (RCR-053); recal interval (RCR-049)
Detect an excursion (heat & freeze) | Logic error, one-sided/heat-only, wrong threshold/dwell, wrong profile (HA-SW-001/003, HA-ENV-001) | Two-sided detection, freeze=heat, no undetected-excursion path (DO-SAFE-004/RCR-019); profile validation + pre-arm gate (RCR-048/045)
Record without silent loss | Torn/lost write, brown-out in cold, silent gap, memory/race defects (HA-SW-011/012/013, HA-ELEC-010, HA-ENRG-006) | Power-fail-safe atomic writes + brown-out barrier; loss → honest gap (DO-SAFE-008/RCR-024); energy reserve/governor (DO-SAFE-012/RCR-006); availability self-check (RCR-027)
Time-stamp defensibly | RTC drift/reset in cold; untrusted time (HA-SW-005, HA-DATA-003) | Bounded/monotonic UTC, disclosed-when-untrusted (RCR-025)
Verdict honestly | False accept over gap/fault; misleading display; unsupported recommendation (HA-SW-002/007, HA-DATA-004/005/008) | Suppress false accept; honest-gap qualifier; device=cloud consistency (DO-SAFE-006/RCR-021/022/040) — choke point 1
Alert in time | No alert generated; delayed/missing/undelivered; RF-in-container; cloud notify fail (HA-SW-006, HA-COMM-001/002/003, HA-CLOUD-001/005) | Local commit + alert-first/ack/escalate + store-and-forward (DO-SAFE-009/RCR-028/029/030); RF-in-container qualification / external-antenna fallback (RCR-031)
Deliver truthfully & to the right subject | Falsified/suppressed data; replay; mis-attribution/wrong-tenant; transport disorder (HA-CYBER-001/002/004, HA-DATA-002, HA-COMM-004/008, HA-CLOUD-003) | Capture-time device-anchored authenticity (RCR-034); non-suppressible alarm (RCR-035); authoritative binding + unique identity (RCR-033/055); idempotent ordered ingest (RCR-030/032)
Present & decide correctly | Misread verdict; alarm fatigue; mis-arm; override; wrong profile at provisioning (HA-USE-001/002/003/004/007) | Legible language-independent verdict (RCR-043); tuned alerting (RCR-047); pre-arm gate (RCR-045); recorded override (RCR-044) — choke point 2
Preserve across reuse/update | Depot miscal/wrong-FW/seal/config; unsafe or malicious OTA; cloud compromise/outage (HA-SVC-001/002/003/005/006, HA-SW-010, HA-CYBER-005/009, HA-CLOUD-007/008) | Post-service BIST + pre-arm gate (RCR-050); signed rollback never-brick OTA (DO-SAFE-014/RCR-037); secure boot/debug-lock (RCR-038); cloud resilience/DR (RCR-039)
4. The physical hazard set (the second, separate analysis)
Distinct from the information spine, the SCM carries a rechargeable lithium battery, which introduces a physical hazard set whose worst case is catastrophic and independent of the vaccine-disposition harm:
• Cold-charge damage → thermal runaway (HA-ENRG-001, D3): controlled by an MCU-independent hardware charge-inhibit below 0 °C (DO-SAFE-001/RCR-001) with independent cell protection + secondary backstop (DO-SAFE-002/RCR-002) — single-fault-safe.
• Cell OV/OC/OT/short / mechanical crush (HA-ENRG-002/003): autonomous protection + retention (RCR-002/004).
• ★ Thermal-runaway fire in an occupied aircraft cargo hold (HA-TFC-002, S5) — the SCM's defining physical catastrophe: controlled by prevention (RCR-002/004) and an independent fire-enclosure containment (RCR-008) that contains a single-cell event without external flame propagation, plus UN 38.3 / IATA-DGR qualification and transport disclosure (RCR-007). It can never be broadly acceptable (HAL-2200 §6.1) and carries mandatory benefit-risk sign-off + post-market thermal-event monitoring.
• Toxic offgas on venting (HA-TFC-003, S4), electrolyte leak (HA-TFC-004), hot surfaces / self-heat (HA-TFC-001/008), RF exposure (HA-ENRG-009): controlled by containment, low-toxicity material selection, touch-temperature bounds, and SELV/RF limits.
• Handling / mechanical (HA-MECH-001…010): edge/pinch, small-parts (child choking), retention, fastener loosening — inherent-by-design controls.
• Environmental (HA-ENV-001…008): −30 °C freeze detection, cold-slow OLED, condensation/ingress, vibration, altitude/venting, dust/UV, e-waste — materials, coat/seal, and detection controls.
5. Detailed-analysis module map
The authoritative per-hazard records live here (all use the 34-field HA-<CODE>-NNN structure):
Module | Code | Domain | Headline records
HAL-100 | HA-CTX | Context / intended use / users / misuse | Information device; vulnerable last-mile user
HAL-200 | HA-BND | Boundary / exposure models | HA-BND-003 data/decision pathway (the spine)
HAL-300 | — | Lifecycle phases | reuse/service amplifiers (D1)
HAL-400 | HA-IUE | Intended-use exposure | automation bias; unmonitored-believed-monitored
HAL-500 | HA-ENRG | Energy / battery | HA-ENRG-001 (cold-charge, D3), HA-ENRG-006 (dead-in-cold)
HAL-600 | HA-MECH | Mechanical | edge/pinch/small-parts; battery retention
HAL-700 | HA-ELEC | Electrical / EMC | ESD/EMC false reading; latent defect; brown-out
HAL-800 | HA-TFC | Thermal / fire / chemical | HA-TFC-002 (aircraft-hold fire, S5)
HAL-900 | HA-ENV | Environmental | freeze miss; condensation; altitude
HAL-1000 | HA-SW | Software | HA-SW-001/002/008/013 (detection/verdict/probe/silent-gap)
HAL-1100 | HA-COMM | Connectivity | HA-COMM-003 (RF-in-container, Open), delayed/missing alert
HAL-1200 | HA-CLOUD | Cloud | notify fail; rules misconfig; outage
HAL-1300 | HA-CYBER | Cybersecurity→safety | falsified data; suppressed alarm; malicious OTA; cloud compromise
HAL-1400 | HA-DATA | Data & decision integrity | HA-DATA-004 (automation bias); attribution; provenance
HAL-1500 | HA-USE | Use-related / human factors | misread verdict; alarm fatigue; mis-arm; override
HAL-1600 | HA-SVC | Service / maintenance | depot miscal; wrong FW; seal; debug-mode; disposal
HAL-1700 | HA-MFG | Manufacturing / supply chain | EOL cal escape; wrong chemistry/parts; latent defect
HAL-1800 | HA-MPE | Mission / property / environmental | HA-MPE-002 (population outbreak, S5); false-reject stock-out; GDP violation; e-waste
HAL-1900 | — | Hazard taxonomy (HAZ-) | the library's hazard classes
HAL-2000…2900 | — | Severity, probability, acceptability, controls, control-review, verification, new-risk, individual + overall residual | the analysis + control + residual apparatus
6. Detailed-analysis findings (summary)
• The information spine is defended by two independent choke points (honest verdict + supported human decision); the detailed analysis confirms that no single upstream failure defeats both — to be proven by the Stage-11 fault tree whose top event is "spoiled dose administered as good."
• The physical S5 (aircraft-hold fire) is a separate, independently-defended catastrophe reduced to region-A-High by prevention + independent containment + UN 38.3; never broadly acceptable.
• The dominant common cause across the information spine is calibration (defeats measure→detect→verdict on every device at once); the dominant fleet common causes are OTA and cloud; the dominant physical single-point is the cell/charge path, made single-fault-safe by independent backstops.
Two unknowns dominate the uncertainty | D3 cold battery (endurance/charge-safety, PRD-VER-004) and RF-in-container (link margin/antenna, PRD-VER-005) — both carried Open into the residual analysis.
• The safety bias creates a real, balanced counter-harm (false reject → wastage → stock-out, HA-MPE-003) — an accepted, monitored public-health trade, not a free safety win.
• All controls are Design-adequate / Evidence-Pending (HAL-2500); every safety control has a planned fault-injection verification and every security control an independent adversarial one (HAL-2600); no individual residual remains in region R as projected (HAL-2800); overall benefit-risk is positive but conditional on V&V (HAL-2900).
7. Traceability / handoff
Upstream | PHA (concept), PRD-SAFE, the SYS-safety set, DOL-1200 (DO-SAFE) — the design this analysis is based on.
Within Stage 10 | the per-hazard records (HAL-500…1800) + the analysis apparatus (HAL-2000…2900); the authoritative chain is in Hazard-Traceability-Matrix.md; the release recommendation in Risk-Management- Summary.md.
Downstream | high-severity (S4/S5) hazardous situations → Stage 11 (ERFKL) FTA top events / SW-HW-FMEA; RCRs → Stage 12 and re-baseline of PRD-SAFE + Product/Design; PV-/VAL- → Stages 13–14; disclosure/ monitoring → Stage 17 post-market; whole chain → ORCH-018 / EM-ENG-012 traceability.
HAL-000 (HA-GOV) — Hazard Analysis Governance & Plan
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Owner: Principal System-Safety / Risk-Management / Human-Factors Engineer, with Quality/Regulatory (EM-ENG-006), Chief Systems (EM-ENG-001), and Product Security (EM-ENG-008). This module is the Hazard Analysis Plan (connected artifact A) and the governance layer for the whole HAL. Based on ISO 14971, IEC 62366-1, IEC 62443 / AAMI TIR57; standards cited by designation, revisions to confirm.
1. Purpose
Define how the SCM hazard analysis is conducted, controlled, reviewed, and maintained: scope, method, the enforced risk chain, roles/authority, independence, configuration control, evidence/uncertainty policy, residual-risk acceptance, and post-release maintenance. This plan governs every HA-<CODE>-NNN record and the connected artifact set (PHA, Detailed PHA, Traceability Matrix, Risk-Management Summary).
2. Scope
Product | SCM — a reusable, cellular-connected environmental data logger placed inside insulated vaccine shipping containers; monitors temperature/humidity/shock, detects excursions, alerts stakeholders, and provides an auditable cold-chain record supporting disposition (accept/reject) decisions for vaccines. It does not control temperature and is not the primary means of preserving the vaccine.
Industry profile | medical / consumer blend — a public-health-significant, non-therapeutic information device. Medical rigor (ISO 14971 discipline, use-error, data integrity) is applied to the information-harm path; consumer/electronics rigor (Li-battery, thermal, EMC, transport) to the physical hazards.
In scope | all SKUs, the full lifecycle (procurement→disposal), the device + firmware + connectivity + cloud + mobile + accessories (probe, charger) + the human users, and the environments (transport, cold store, clinic, depot, last-mile).
Out of scope | the therapeutic product itself (the vaccine) except as the object of the disposition decision; the shipping container's passive thermal performance (a separate product) except as a contributing condition.
3. Definitions (plain language)
• Hazard — a potential source of harm (e.g., a lithium cell's stored energy; an excursion-detection logic defect).
• Hazardous situation — the circumstance in which people/property/mission are exposed (e.g., a compromised dose sitting at point-of-care believed acceptable).
• Harm — the actual loss (e.g., a beneficiary receives a potency-lost dose and is not protected).
• Information-integrity harm — the SCM-defining harm class: a missed/incorrect/delayed excursion decision leading to a wrong disposition, mediated by a human decision, not a direct device injury.
• P1 / P2 — probability the initiating cause reaches the hazardous situation / probability the situation proceeds to harm (HAL-2100).
• RCR — Risk-Control Requirement (HAL-2400), a formal testable requirement derived from an accepted control, feeding Stage 12 and Product/Design.
4. Method — the enforced risk chain
Every credible scenario is authored as a complete chain (ISO 14971), never a weak "cause → injury" row:
Initiating Cause/Event → Foreseeable Sequence of Events → Hazard → Hazardous Situation → Foreseeable
Sequence to Harm → Harm/Loss → Initial Risk (Severity × P1×P2) → Risk-Control Strategy → Risk-Control
Requirement (RCR) → Implementation Reference (DO-…) → Verification Reference (PV-…) → Validation Reference
(VAL-…) → Residual Risk.
Each record uses the 34-field structure (see §9). Every field is answered or justified N/A — a blank is not acceptable. Hazard identification uses the HAL-1900 taxonomy as a prompt list (it suggests; it never auto-declares a hazard applies), the intended-use questionnaire (HAL-400), the lifecycle sweep (HAL-300), and reconciliation with the upstream safety needs (UN-900), preliminary safety requirements (PRD-SAFE), the fault-management architecture, and the DO-SAFE controls (DOL-1200).
5. Two touchpoints & this library's structure
Hazard analysis runs at two points and continues through life:
• Preliminary Hazard Analysis (PHA) — early, harm-and-exposure-first, from concept + user needs + architecture (see Preliminary-Hazard-Analysis.md).
• Detailed Product Hazard Analysis — after design outputs, with controls implemented and traced (see Detailed-Product-Hazard-Analysis.md). The 30 HAL modules (HAL-000…HAL-2900) organise governance, context, domains, taxonomy, risk scales, and controls/residual. Applicability is recorded in the Applicability Register (index.md).
6. Roles, authority & independence
Role | Responsibility
System-Safety / Human-Factors (lead, this module) | Hazard identification, risk chains, severity/probability, control hierarchy, HAL-400/1500
Quality / Regulatory (EM-ENG-006) | ISO 14971 risk-management file, acceptability, benefit-risk, Risk-Management Summary
Product Security (EM-ENG-008) | HAL-1300 cyber-induced-safety hazards; security↔safety reconciliation
Battery/Power, Electrical, Mechanical, Firmware, Cloud, Manufacturing engineers | Hazards & controls in their domain
Chief Systems (EM-ENG-001) | Integrates hazard analysis, RCRs, and end-to-end traceability into design controls
Traceability (EM-ENG-012 / ORCH-018) | Maintains the Hazard-Traceability-Matrix
Independence | severity assignment and residual-risk acceptance for S4/S5 hazards require a reviewer independent of the implementing discipline (EM-ENG-006 + EM-ENG-001), per HAL-2200 §7.
Judgment over scores | numerical scores aid prioritisation only; they never override engineering judgment or hide the harm scenario (CLAUDE.md rule 3).
7. Configuration control, evidence & uncertainty policy
Configuration control | the HAL is a controlled library under the project PLM; each record has a Status, Owner, and Review History; changes follow the ECO/ECN process (DO-GOV-004). This baseline is v1.0 (2026-07-18), pre-DR (design review).
Evidence & honesty (evidence-standard / hallucination-policy) | every claim is Judgment, Assumption, or Evidence (cited standard / upstream document); unknowns are recorded as Evidence-Pending with the named closing test — never invented. Probabilities are never asserted as fact; each records its basis and uncertainty (HAL-2100).
Uncertainty handling | the uncertain-probability rule (HAL-2200 §6.2) drives Evidence-Pending risks by severity until the test closes them. The two principal open unknowns — low-temperature cell chemistry (D3) and in-container antenna/RF strategy — are carried as top hazards throughout.
8. Residual-risk acceptance, change-impact, post-release & escalation
• Residual risk is reassessed only after a control is defined, implemented (DO-…), and verified (PV-…) — never on the basis of a spreadsheet action (HAL-2800). Overall residual risk and benefit-risk are evaluated in HAL-2900 + the Risk-Management Summary before release (release gate: no open R risk).
Change-impact | any design/requirement change is screened against the HAL; affected records are re-opened and re-verified.
Post-release maintenance | production and field monitoring (Stages 16–17) feed new information back; field events, complaints, and CAPA can re-open records (post-market surveillance per ISO 14971).
Escalation | any R region risk, any new S5 hazard, or any control found defeatable escalates to EM-ENG-001 + EM-ENG-006 and the programme risk board.
9. The 34-field hazard-analysis record (template used by every register)
Hazard Analysis ID · Hazard ID (library ref, HAZ-) · Product/Configuration · Lifecycle Phase · System State · Operational Mode · Perspective · Source · Initiating Cause/Event · Contributing Conditions · Foreseeable Sequence of Events · Hazard · Hazardous Situation · Sequence to Harm · Harm/Loss · Severity · Probability (P1×P2) · Detectability/Warning · Initial Risk · Risk Acceptability · Risk-Control Strategy · Risk-Control Requirement · Implementation Reference · Verification Reference · Validation Reference · Residual Probability · Residual Severity · Residual Risk · New Risks Introduced · Overall Residual-Risk Contribution · Status · Owner · Rationale/Evidence · Assumptions · Review History.
10. Cross-analysis reconciliation & downstream feeds
↔ Requirements | reconciles with UN-900 (safety needs), PRD-SAFE (preliminary safety requirements — to be re-baselined from this HAL), SYS-149…162 (safety system requirements), and the fault-management architecture.
↔ Design outputs | every inherent-design/protective control's Implementation Reference is a DO-SAFE output (DOL-1200) pointing to the concrete DO-EE/DO-FW/DO-CLD/DO-MECH output (release criterion RA-05).
→ Stage 11 (ERFKL) | S4/S5 hazardous situations become FTA top events; safety-significant functions drive deeper DFMEA/SW-FMEA/PFMEA; cyber threats capable of harm enter via HAL-1300.
→ Stage 12 | the RCR register (HAL-2400) feeds the Risk-Control Requirements stage and Product/Design.
→ Stages 13–14 (PVL/PVaL) | verification proves controls implemented; validation proves the controlled product is safe and usable.
→ Stage 18 (ORCH-018 / EM-ENG-012) | the Hazard-Traceability-Matrix is the authoritative chain.
11. Traceability / closure notes
This plan governs all HAL modules. Baseline v1.0 is pre-DR; it closes when: (a) every applicable module is populated, (b) every S4/S5 record has an RCR, a DO Implementation Reference, and a named PV/VAL, (c) no R risk remains open, and (d) HAL-2900 + the Risk-Management Summary give a benefit-risk-based release recommendation. Reconcile PRD-SAFE and DOL-1200 placeholders to the HA-/RCR- IDs baselined here.
HAL-100 (HA-CTX) — Product & Analysis Context
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: System-Safety with EM-ENG-001. Establishes the who / what / where that scopes every hazard record: intended use, foreseeable use & misuse, users (including the low-training last-mile operator), perspectives (including the beneficiary/ patient), affected property/environment/mission, variants, accessories, and external dependencies. Based on the concept, UN-100/200/400, PRD-100/200, and the system architecture.
1. Purpose
Define the product and analysis context so hazards are identified against the real intended use, users, and environments — not an abstraction — and so the SCM's defining subtlety (an information device whose dominant harm is a wrong disposition decision) frames the whole library.
2. Scope
All SKUs and configurations of the SCM device + probe + charger, its firmware, its LTE-M/Wi-Fi/BLE/GPS connectivity, the cloud platform, and the mobile/BLE provisioning app, across the reusable multi-trip lifecycle (D1).
3. Applicable standards & references
ISO 14971 (context of use, intended use); IEC 62366-1 (use specification, user profiles, use environment); WHO PQS E006 (intended cold-chain monitoring use). Cited by designation.
4. Context register (HA-CTX records)
HA-CTX-001 — Intended purpose / intended use
Statement | The SCM continuously monitors and records temperature (internal + external product-zone probe), humidity, and shock inside insulated vaccine shipping containers; detects temperature excursions against a configured product-stability profile; alerts stakeholders when connectivity permits; and provides an auditable, tamper-evident cold-chain record supporting accept/reject/quarantine disposition decisions across the distribution chain. It is a monitoring/information device — it does not control temperature and is not the primary means of preserving the vaccine.
Safety implication | the dominant harm path is informational — a missed/incorrect/delayed excursion decision → wrong disposition → compromised dose administered (or potent dose wasted) → beneficiary harm.
Source | Concept §2; UN-VIS-001; PRD-VIS-002. Status: Baselined.
HA-CTX-002 — Reasonably foreseeable use
Statement | Armed at a depot/dispatch point, placed in a shipper with the probe in the product zone, transported by air/ground across borders and climates, read at receipt (often offline, gloved, in cold), disposition decided, returned/recharged/recalibrated, reused for many trips over service life (D1).
Source | UN-300 (JTBD), UN-500 (workflow), PRD-200. Status: Baselined.
HA-CTX-003 — Reasonably foreseeable misuse (safety-relevant)
Statement | Foreseeable misuse includes: dispatching without arming (believing it monitored); configuring the wrong product profile; placing the probe outside the product zone (in air, against the wall); ignoring/muting alerts (alarm fatigue); overriding a "reject" verdict; charging a cold device and mistaking charge-inhibit for a dead device; discarding a cold, charge-inhibited device; reusing without recalibration/cleaning; charging from a non-compliant USB source; opening the sealed enclosure in the field; using it as the sole basis for disposition when independent shipping indicators disagree.
Source | UN-900, VOC field/support, IEC 62366-1 §5.1. Status: Baselined — seeds HAL-1500.
HA-CTX-004 — Prohibited / off-label use
Statement | Not a therapeutic device; not a temperature controller; not a substitute for GDP receipt QA; not certified for hazardous (explosive) atmospheres; not to be opened/serviced in the field beyond probe/charger; not to be flown outside its UN 38.3 / IATA transport conditions.
Source | Concept §2, D4, UN-1300. Status: Baselined.
HA-CTX-005 — Intended & vulnerable users
Statement | Users span dispatch/QA staff (trained), logistics/3PL/air-cargo handlers (procedurally trained, safety-critical for transport), depot/service technicians (charge/recalibrate/ triage), fleet/IT administrators (cloud), and the last-mile clinic nurse/health worker — the least-trained, most safety-critical operator, often gloved, offline, in poor light or deep cold, with low literacy or a non-primary UI language. The last-mile user is treated as the vulnerable user whose foreseeable error must not cause a wrong disposition (UN-SAFE-005 / IEC 62366-1).
Source | UN-100/200, PRD-100, MR-PER-007. Status: Baselined — governs HAL-400/1500.
HA-CTX-006 — Perspectives (broadened beyond people)
Statement | Perspectives carried through the library: beneficiary/patient (represented — the ultimate harmed party via a compromised/withheld dose), clinic staff, logistics handler, depot technician, installer, bystander (incl. a child near a discarded device — small-parts/battery), organization/ programme (public-health mission), mission (the vaccine shipment), environment (Li-battery disposal), nearby equipment (aircraft/EMC), and public infrastructure (immunisation programme).
Source | ISO 14971; UN-100; STN register. Status: Baselined.
HA-CTX-007 — Affected property, environment & mission
Statement | Property = the vaccine lot/shipment (high value, public-health-critical) and the reusable device fleet; environment = lithium cells and electronics at end-of-life; mission = the immunisation programme / cold-chain integrity and the trust in the auditable record.
Source | MR-1700 (sustainability), UN-1800, concept §4. Status: Baselined.
HA-CTX-008 — Service life, geography & jurisdictions
Statement | Reusable over a multi-trip/multi-year service life (D1; exact cycle life Evidence-Pending, cold cycle-life test). Global deployment across WHO PQS, EU GDP, US, and air/ground transport regimes (D4). Operating −30 °C…+60 °C; storage/transport per IEC 60068.
Source | Concept §4/§3 (D1/D4), UN-1300. Status: Baselined; cycle-life Evidence-Pending.
HA-CTX-009 — Variants, accessories, consumables
Statement | SKUs share the safety-critical core (cell, BMS, charge-inhibit, detection, verdict, record). Accessories: external temperature probe + lead (safety-relevant — product-zone accuracy), USB-C charger (may be third-party — charge-input abuse), optional external antenna (RF branch), mounting/retention feature. No patient-contacting consumables.
Source | Component list (CMP-004/005/020/029), PRD-300. Status: Preliminary.
HA-CTX-010 — External / third-party / cloud / network dependencies
Statement | Depends on cellular carriers/roaming (LTE-M coverage; in-container attenuation), cloud platform (ingest, rules/alert engine, record store, dashboard, OTA), notification channels (email/SMS/push to responders), time source (RTC + network time discipline), and third-party USB chargers. Each dependency is a hazard source for the delayed/missing-alert and data-integrity paths.
Source | Architecture (S5/S8), DOL-800/1000. Status: Baselined — seeds HAL-1100/1200/1300.
5. Traceability / closure notes
This context scopes HAL-200 (boundary/exposure), HAL-300 (lifecycle), HAL-400 (intended-use questionnaire), and every domain register. The information-device framing (HA-CTX-001) and the last-mile vulnerable user (HA-CTX-005) are the two context anchors that make software/data/use hazards high-severity. Feeds the PHA and Risk-Management Summary; matrix via ORCH-018.
HAL-1000 (HA-SW) — Software-Controlled Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1000, record code HA-SW · Owner: Principal Embedded/Software Safety Engineer, with System-Safety (lead) and Quality/Regulatory (EM-ENG-006). This module analyses the software-controlled hazards of the SCM's device firmware (CMP-035…045) and its immediate software collaborators (MCU/brown-out CMP-015, RTC CMP-017, watchdog CMP-018, NV log CMP-033). It follows the ISO 14971 harm-and-exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. The deeper failure-mode decomposition (SW-FMEA) is a Stage-11 (ERFKL) product; every S4/S5 record here is flagged as an FTA top-event / SW-FMEA candidate. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM's software can produce, mask, or fail to prevent a wrong disposition decision — the SCM's dominant harm. The SCM is an information device (HA-BND-003): its most consequential harm path is informational, not a direct physical injury. Firmware sits astride the entire information pathway — sensor → detection engine (CMP-039) → verdict engine (CMP-040) → local record/OLED and cloud record/alert (via CMP-041/042) → human disposition decision → vaccine → beneficiary. A software defect at any node can cause a missed, incorrect, or delayed excursion detection/alert, or a silent loss of monitoring, and thereby a compromised dose administered (loss of potency → beneficiary not protected → preventable disease) or a valid lot wrongly destroyed (stock-out → unvaccinated). This module makes each such path explicit and binds it to a testable risk-control requirement.
The overriding architectural invariant this module defends is "honest gap, no false reassurance" (SYS-151, SYS-167, SYS-169, SYS-170; fault-management-diagnostics.md §6): an armed device is at all times either verifiably monitoring or visibly flagged as not, and a fault must never corrupt the record or leave the device silently non-monitoring. Software is the layer that can most easily violate this invariant, so it is the layer where the invariant is most rigorously enforced.
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope software elements (device firmware and its safety-relevant neighbours):
• CMP-035 bootloader / secure-boot; CMP-036 scheduler / RTOS; CMP-037 power-state (energy) governor; CMP-038 sensor drivers + plausibility; CMP-039 two-sided excursion-detection engine; CMP-040 verdict engine; CMP-041 storage manager (power-fail-safe); CMP-042 comms agent; CMP-043 OTA agent; CMP-044 security / crypto; CMP-045 diagnostics / health.
Software-adjacent hardware the firmware depends on | MCU + brown-out detector CMP-015, real-time clock CMP-017, watchdog CMP-018, NV log store CMP-033, external probe/front-end CMP-020, OLED/HMI CMP-030/032, and the cloud recompute/notify/registry/config services CMP-046/048/049/050/055 where a device-software defect propagates across the boundary.
In-scope lifecycle phases: provisioning/arming, armed monitoring transport, receipt/disposition read, depot recharge/triage, OTA update, and reuse over service life (HAL-300).
Boundaries with adjacent modules (stay in lane, name the neighbours — CLAUDE.md rule 2):
• Battery-energy exhaustion that produces a silent gap is owned by HA-ENRG-006/008 (Battery/Power); this module owns the software obligation to make that gap honest and to suppress a false accept over it (HA-SW-013).
• Delivery of the alert across the network/cloud is owned by HAL-1100 (HA-COMM) and HAL-1200 (HA-CLOUD); this module owns the device-side obligation to generate and commit the alert (HA-SW-006).
• Cyber-induced defeat of a control (falsified data, suppressed alarm, unauthorised threshold change, malicious OTA) is owned by HAL-1300 (HA-CYBER); this module owns the non-malicious software-defect equivalents and cross-references the cyber twin.
• The human misreading a correctly-presented verdict is owned by HAL-1500 (HA-USE); this module owns the software presenting a wrong or misleading verdict (HA-SW-007).
• Sensor hardware faults are shared with HAL-700 (HA-ELEC-008); this module owns the firmware plausibility/handling of them (HA-SW-008).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 decomposition (HAL-2100); risk acceptability, severity-first (HAL-2200).
• IEC 62304 — medical-device software life-cycle processes; software safety classification (Class A/B/C by the severity of the harm a failure can contribute to) and the associated development/verification rigor. SCM software-safety-classification note (evidence honesty): because a firmware failure can contribute to serious injury/preventable disease at group scale through a wrong disposition (HAL-2000 severity anchor), the safety-relevant firmware components (CMP-038/039/040/041 in particular) are treated at the highest software safety class (Class C-equivalent) for the purposes of this analysis — i.e., the full IEC 62304 detailed-design, unit-verification, integration, and system-test rigor applies. The formal classification record is a Stage-11/12 deliverable and is carried here as Evidence-Pending with the named closing activities (detection-logic test, fault-injection, static/code analysis).
• IEC 62366-1 — usability engineering; use-error harm where the software output (verdict/alarm/qualifier) is what the human acts on (HA-SW-007; ties HAL-1500).
• IEC 62368-1 — cited only where a software fault could enable a physical charge/thermal hazard (defense- in-depth over the HW charge interlock, HA-SW cross-reference to HA-ENRG-001).
Upstream | UN-900 (safety needs); PRD-SAFE (preliminary safety requirements); SYS-149…162 (safety system requirements) and SYS-163…176 (reliability/availability), esp. SYS-149 (severity anchor), SYS-150 (two-sided detection), SYS-151 (no false reassurance), SYS-154 (mis-arm), SYS-158/159/167 (safe state), SYS-160 (dead-in-cold), SYS-168 (power-fail-safe), SYS-169 (watchdog + logged recovery), SYS-170 (availability self-check); SRS-DIAG (diagnostics); FWR firmware requirements; the fault-management-diagnostics.md and state-mode-model.md architecture; DOL-1200 (DO-SAFE controls); HAL-2400 (RCR register); HA-BND-003/004 (information & control-flow exposure models).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000). Software records inherit the SCM severity anchor (HAL-2000 §6): a wrong disposition is S4 typical (a lot/group of beneficiaries) and S5 worst (population scale), even though the firmware never physically touches a patient — the harm is real and mediated by the disposition decision (HA-BND-003).
• Probability is decomposed P1 × P2 with an explicit basis and uncertainty flag (HAL-2100). For software, P1 (defect → hazardous situation) is usually reduced to Remote/Improbable by IEC 62304 rigor, static analysis, and verification, but P2 (situation → harm) is often Probable because the harmful information is invisible and mediated by a human (a missed excursion, a silent gap, a mislabelled "accept" all look fine downstream). This keeps many records Medium/High and firmly severity-first. Where a probability cannot yet be defended it is carried Evidence-Pending with a named test — never invented (HAL-000 §7).
• No-false-reassurance invariant is the governing acceptance test for this module: a software control is only accepted if a fault it addresses can never (a) corrupt/erase the committed record, (b) present an incomplete record as a clean "accept," or (c) leave an armed device silently non-monitoring (fault-management §6; RCR-021/022/023/027).
Control hierarchy (HAL-2300) | inherent-by-design (make the wrong state unrepresentable / suppress false accept by construction) → protective (safe-state, watchdog, plausibility) → information (honest gap, qualifiers) → process (verification/coverage). "Information-for-safety" here is design-enforced (the software cannot show a clean pass over a gap), not a passive label, so it survives the warnings-challenge rule (HAL-2200 §6.4).
• Residual risk is re-estimated only after a control is defined, implemented (DO-…), and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (software prompt list)
1. Detection correctness: Can the excursion-detection engine (CMP-039) miss a real excursion — wrong threshold, wrong profile, one-sided (heat-only) logic, hysteresis/dwell error, integer/units error, or an edge case (freeze, rapid ramp, boundary)? Is freeze reliability equal to heat (SYS-150)?
2. Verdict integrity: Can the verdict engine (CMP-040) compute "accept" over an incomplete, gapped, or faulted record — the false-accept path (SYS-151)? Do device and cloud verdicts agree (SYS-016)?
3. Configuration & calibration data: Can a wrong/stale/corrupt trip profile, threshold set, or calibration coefficient bias detection or verdict?
4. Time integrity: Can a clock/RTC defect mis-time an excursion or make the record indefensible (SYS-118)?
5. Annunciation: Can an excursion be detected but no alert emitted, or a misleading verdict displayed?
6. Sensor plausibility: Can an absent/open/short/wrong probe be read as a valid product-zone temperature (SYS-007)?
7. Liveness & availability: Can the firmware hang/deadlock/leak so the device is believed monitoring but is not (SYS-169/170), or lose samples to a crash/race/reset?
8. Update & boot integrity: Can an OTA update brick the fleet, deploy wrong detection logic, or leave an unmonitored window (SYS-032/139)?
9. Record continuity: Can a reset/power-loss/brown-out create a silent gap that reads as an uneventful trip (SYS-160/168) — the dead-in-cold information hazard?
6. Hazard Analysis Register (HA-SW records)
HA-SW-001 — Excursion-detection logic error / wrong threshold → MISSED excursion → compromised dose administered ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-FUNC (detection function failure) — information-integrity harm via HA-BND-003
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Armed monitoring transport · System State: Sampling/Logging (armed) · Operational Mode: Armed / logging
Perspective | Beneficiary/patient (ultimate harmed party, via a wrong "accept" disposition); disposition decision-maker; mission/programme
Source | UN-900; PRD-SAFE; SYS-149 (severity anchor) / SYS-150 (two-sided detection); DO-SAFE-004; fault-management §1; HAL-2000 §7 (HA-SW-001 = S5-worst/S4-typical); judgment
Initiating Cause/Event | A defect in the two-sided excursion-detection engine (CMP-039) — wrong comparison threshold, one-sided (heat-only) logic that never trips on freeze, an incorrect dwell/time-above-threshold accumulation, a hysteresis or integer/units error, or an unhandled boundary/edge case (rapid ramp, exactly-at-limit) — fails to declare a real product-stability excursion.
Contributing Conditions | Many vaccines are freeze-sensitive and freezing excursions are common in the field (WHO-UNICEF literature), yet freeze logic is easy to under-test relative to heat; the product-stability model is a two-sided (heat and freeze) rule set that must be exactly right; cold ambient stresses timing; the harmful output is invisible — a missed excursion produces a record that looks clean.
Foreseeable Sequence of Events | Real excursion occurs in the shipper → detection engine evaluates against a flawed rule/threshold → excursion not declared → no event committed, no alert → record shows an uneventful trip → human reads "accept" → compromised dose administered.
Hazard | Failure of the excursion-detection function (missed-excursion path).
Hazardous Situation | A compromised (potency-lost) dose at point-of-care believed acceptable because the excursion was never detected.
Sequence to Harm | Missed excursion → false "accept" → potency-lost dose administered → beneficiary not protected → preventable disease; at lot/population scale, many beneficiaries.
Harm/Loss | A lot/shipment of compromised doses administered to a group of beneficiaries (S4); population-scale mass administration (S5 worst).
Severity | S4 (Critical) typical / S5 (Catastrophic) worst — governing dimension public-health (HAL-2000 §6 severity anchor; §7 lists HA-SW-001 as the S5-worst / S4-typical archetype). Record anchor S4, S5 credible worst-case retained.
Probability (P1×P2) | P1=2 (Remote — detection logic is highest-class IEC 62304 software with dedicated verification) × P2=4 (Probable — a missed excursion is invisible and readily proceeds to a dose administered; receipt QA is an external, imperfect check) ⇒ product 8 ⇒ Medium band (PM) — Basis: HAL-2100 §7 ("missed excursion — detection logic error: Remote × Probable ⇒ Medium"); judgment; Medium uncertainty, Evidence-Pending — closed by the detection-logic test on profiled excursions (PV-SW-001) with equal freeze/heat coverage.
Detectability/Warning | Not detectable by the user (a missed excursion has no signature). Device↔cloud verdict cross-check (SYS-016) and independent shipping indicators are partial downstream checks; the primary defence is correct-by-verification detection, not detection-after-the-fact.
Initial Risk | AP High (region R — S4 × PM; and the S5 worst-case invokes the catastrophic-severity rule, HAL-2200 §6.1)
Risk Acceptability | Unacceptable without control (R region cannot be accepted; HAL-2200 §6).
Risk-Control Strategy | 1 (inherent — two-sided detection with freeze reliability equal to heat, no undetected-excursion path by construction) + 4 (process — IEC 62304 verification / detection-logic test)
Risk-Control Requirement | RCR-019 — two-sided (heat and freeze) product-stability excursion detection with freeze reliability equal to heat; no undetected-excursion path. (Ties RCR-047 alarm tuning so real excursions are not muted by fatigue.)
Implementation Reference | DO-SAFE-004 (→ DO-FW-006 detection engine + DO-EE-013 probe front-end) · CMP-039 (detection engine), CMP-038 (sensor drivers feeding it)
Verification Reference | PV-SW-001 (Evidence-Pending) — detection-logic test against a library of profiled excursions (heat, freeze, boundary, ramp, dwell) with freeze coverage explicitly equal to heat; supplemented by static/code analysis and unit + integration verification per IEC 62304 on CMP-039
Validation Reference | VAL-SW-001 (Evidence-Pending) — end-to-end validation that a representative real excursion produces a correct reject verdict a user acts on (ties HAL-1500 decision-support validation)
Residual Probability | projected Low (PL) after two-sided logic verified with equal freeze/heat coverage (P1 → Improbable; residual is P2-dominated because the harm remains human-mediated) · Residual Severity: S4 (unchanged — outcome unchanged if it ever occurred) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off; Evidence-Pending until PV-SW-001 closes
New Risks Introduced | Aggressive/over-sensitive detection to guarantee capture can raise false-reject rate → stock-out (ties HA-USE-009 / HA-MPE-003) and alarm fatigue (HA-USE-002) — balanced by RCR-047 (noted for HAL-2700).
Overall Residual Contribution | The headline software residual. Carried in HAL-2900 with benefit-risk; FTA top event (Stage 11) and the primary SW-FMEA subject for CMP-039.
Status | Evidence-Pending (detection-logic verification) — Controlled-pending-verification
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304 (Class-C-equivalent rigor for CMP-039); SYS-149/150; RCR-019; DO-SAFE-004; HAL-2000 §6/§7; HAL-2100 §7. WHO-UNICEF freezing-prevalence literature underpins the freeze-equal-to-heat requirement.
Assumptions | The product-stability profile supplied to the engine is correct (profile-error is HA-SW-003); the sensor feed is plausible (HA-SW-008); freeze excursions are as safety-significant as heat for the covered vaccines.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-002 — Verdict engine computes "accept" over an incomplete/gapped/faulted record → FALSE ACCEPT → released compromised product ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-FUNC / HAZ-DATA (verdict-integrity failure) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition · System State: Verdict computation (on-device and cloud recompute) · Operational Mode: End-of-trip disposition read
Perspective | Beneficiary/patient (via a false accept); disposition decision-maker; mission
Source | PRD-SAFE; SYS-151 (no false reassurance on incomplete record) / SYS-016 (device↔cloud consistency); DO-SAFE-006; fault-management §6; HAL-2000 §7 (HA-SW-002 = S4); judgment
Initiating Cause/Event | A defect in the verdict engine (CMP-040) evaluates a record that contains a monitoring gap, a flagged sensor fault, or an incomplete armed interval and nonetheless returns an unqualified "accept."
Contributing Conditions | The verdict must be computed from the quality-annotated record, not just the sample values; gap/fault flags must be treated as first-class inputs; device and cloud verdicts are computed by two code paths that can diverge (SYS-016); an "accept" is the default a busy operator most readily acts on.
Foreseeable Sequence of Events | Trip has a gap/sensor-fault interval (from any cause) → verdict engine ignores or mis-weights the quality flag → returns a clean "accept" → operator releases the lot → compromised (or unverifiable) product administered.
Hazard | Verdict-integrity failure (false-accept over incomplete monitoring).
Hazardous Situation | A lot presented as unconditionally acceptable when its monitoring record is incomplete or faulted.
Sequence to Harm | False accept → compromised/unverifiable dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot of compromised/unverifiable doses released and administered (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §7, "false accept on incomplete/faulted record = S4").
Probability (P1×P2) | P1=2 (Remote — verdict engine is highest-class IEC 62304 software with a gap-handling test) × P2=4 (Probable — a clean-looking accept is acted on with high likelihood; the gap is invisible to the reader) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment consistent with HAL-2100 §7 information-harm posture; Medium uncertainty, Evidence-Pending — closed by verdict-consistency + gap-handling test (PV-SW-002).
Detectability/Warning | Device↔cloud verdict cross-check (SYS-016, RCR-021) detects a divergent path; the honest-gap qualifier (RCR-022/040) makes an incomplete record legible so a clean accept cannot be shown. Without the control, undetectable to the user.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; the honest-gap / no-false-reassurance invariant is non-defeatable (HAL-2400 §4).
Risk-Control Strategy | 1 (inherent — verdict suppresses a false accept on any gap/fault by construction) + 3 (design-enforced information — incomplete-monitoring qualifier always shown)
Risk-Control Requirement | RCR-021 — the verdict shall suppress a false "accept" on any data gap/sensor fault/incomplete monitoring; device and cloud verdicts consistent. (+ RCR-022 any gap made visible in record and on-device status; an incomplete record never shows an unqualified clean pass.)
Implementation Reference | DO-SAFE-006 (→ DO-FW-008 verdict + DO-CLD-006 cloud recompute + DO-HF-002 display) · CMP-040 (verdict engine), CMP-049 (cloud recompute)
Verification Reference | PV-SW-002 (Evidence-Pending) — verdict-consistency + gap-handling test: inject gapped/faulted/incomplete records and confirm the verdict never returns an unqualified accept and device=cloud; static/code analysis on CMP-040
Validation Reference | VAL-SW-002 (Evidence-Pending) — decision-support validation that an incomplete record is never read as a clean pass (with RCR-040, ties HAL-1500)
Residual Probability | projected Very Low (PVL) after suppression + qualifier + device/cloud cross-check verified (the false-accept state is made unrepresentable by construction) · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | Conservative suppression can convert a marginal gap into a reject/quarantine → wastage/stock-out pressure (ties HA-USE-009) — accepted as the safer direction, noted for HAL-2700.
Overall Residual Contribution | Headline verdict-integrity residual; FTA top event (Stage 11); SW-FMEA subject for CMP-040. Ties HA-SW-007, HA-SW-013, HA-DATA-004, HA-USE-007.
Status | Evidence-Pending (gap-handling verification) — Controlled-pending-verification
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304 (Class-C-equivalent, CMP-040); SYS-151/016; RCR-021/022; DO-SAFE-006; HAL-2000 §7.
Assumptions | Gap/fault flags are correctly written into the record (storage integrity is HA-SW-013 / RCR-024); the cloud recompute uses the same product-stability profile as the device.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-003 — Wrong trip profile applied / configuration error → wrong thresholds → missed excursion
Hazard ID (library) | HAZ-DATA (configuration/profile error) → HAZ-FUNC — HA-BND-003/004
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning / arming · System State: Provisioning/Arming · Operational Mode: Trip setup
Perspective | Beneficiary/patient; disposition decision-maker; mission
Source | PRD-SAFE; SYS-154 (mis-arm) / SYS-024 (pre-arm gate); state-mode-model §3.5; judgment
Initiating Cause/Event | The device is armed with the wrong product-stability trip profile (or a mis-entered/mismatched threshold set) — e.g., a heat-tolerant profile applied to a freeze-sensitive vaccine — so detection runs against thresholds that do not match the actual product.
Contributing Conditions | Multiple vaccines with different stability profiles share the same device fleet; profile selection is a human/config step at provisioning; a plausible-but-wrong profile passes silently unless validated against the product-stability library; software correctly executes an incorrect rule set.
Foreseeable Sequence of Events | Wrong profile selected/loaded → device arms and monitors against wrong thresholds → a real excursion for the actual product does not breach the wrong thresholds → excursion not declared → clean record → false accept → compromised dose administered.
Hazard | Wrong detection configuration (profile/threshold mismatch).
Hazardous Situation | A device monitoring a product against thresholds that do not protect that product.
Sequence to Harm | Wrong thresholds → missed excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered (S4).
Severity | S4 (Critical) — public-health (same anchor as HA-SW-001).
Probability (P1×P2) | P1=2 (Remote — provisioning validates the profile against the library and the pre-arm gate blocks an invalid profile) × P2=4 (Probable — a wrong-profile miss is invisible downstream) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by profile-validation + pre-arm gate test (PV-SW-003).
Detectability/Warning | Provisioning-time profile validation against the product-stability library flags an implausible/mismatched profile (RCR-048); the pre-arm readiness gate blocks arming on an invalid/missing profile (RCR-045); an armed device shows the active profile for confirmation.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — profile validation) + 1 (inherent — pre-arm gate makes an unvalidated arming state unrepresentable)
Risk-Control Requirement | RCR-048 — provisioning shall validate the selected trip profile against the product-stability library and flag an implausible/mismatched profile. (+ RCR-045 pre-arm readiness gate blocks arming unless a valid firmware/profile is present; an un-armed/void state is unmistakable.)
Implementation Reference | DO-APP / DO-CLD (profile validation, RCR-048) · DO-SAFE-007 / DO-FW-032 (pre-arm gate, RCR-045) · CMP-055 (profile/config), CMP-040 (verdict), CMP-045 (readiness gate)
Verification Reference | PV-SW-003 (Evidence-Pending) — profile-validation test (mismatched/implausible profiles rejected) + pre-arm gate blocks-arming-on-invalid-profile test
Validation Reference | VAL-SW-003 (Evidence-Pending) — provisioning-workflow usability that the correct profile is selectable and confirmable (ties HAL-1500)
Residual Probability | projected Low (PL) after validation + gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Over-strict profile validation could block legitimate arming (availability/stock-out pressure) — balanced against safety, noted for HAL-2700.
Overall Residual Contribution | FTA top event (Stage 11); ties HA-USE-003 (provisioning use-error) and HA-SW-001. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Cloud/App and Human-Factors)
Rationale/Evidence | ISO 14971; IEC 62304; SYS-154/024; RCR-048/045; DO-SAFE-007. Severity per HAL-2000 §6.
Assumptions | A canonical product-stability library exists and is the reference for validation; the cloud and device share the same profile version (ties config-integrity HA-SW-004).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-004 — Stale/corrupted configuration or calibration data used → biased detection → wrong disposition
Hazard ID (library) | HAZ-DATA (config/calibration integrity) → HAZ-FUNC — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Arming / monitoring · System State: Sampling/Logging · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission
Source | PRD-SAFE; SYS-024 (pre-arm valid calibration); DO-SAFE-005/007; judgment
Initiating Cause/Event | The firmware reads a stale, expired, or corrupted configuration/calibration record (e.g., calibration coefficients from a prior life, a partially-written config block, a version-skewed threshold set) and applies it to sensor conditioning and detection.
Contributing Conditions | Reusable device carries persistent config/calibration across many trips and OTA updates; NV corruption or a torn config write (ties HA-SW-013) can leave a plausible-but-wrong block; calibration has a validity interval that must be enforced; no integrity check would let a corrupt block be used silently.
Foreseeable Sequence of Events | Config/calibration block stale or corrupt → not integrity-checked or expiry-checked → applied to conditioning/detection → readings biased or thresholds shifted → excursion missed (or falsely declared) → wrong disposition.
Hazard | Use of invalid configuration/calibration data.
Hazardous Situation | A device monitoring with biased conditioning or shifted thresholds it presents as valid.
Sequence to Harm | Biased detection → missed excursion → false accept → compromised dose administered (or false reject → stock-out).
Harm/Loss | A lot of compromised doses administered, or a valid lot destroyed (S4).
Severity | S4 (Critical) — public-health.
Probability (P1×P2) | P1=2 (Remote — config/calibration is integrity-protected and the pre-arm gate blocks expired/invalid calibration) × P2=3 (Occasional — a modest bias may still allow some correct calls; not every biased trip reaches harm) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by calibration-integrity + expiry-gate test (PV-SW-004).
Detectability/Warning | Calibration coefficients integrity-protected; expired/invalid calibration blocks arming at the pre-arm gate (RCR-026); a corrupt config block fails its integrity check and drives safe-state/void rather than silent use.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — integrity + expiry gate) + 4 (process — calibration lifecycle)
Risk-Control Requirement | RCR-026 — calibration coefficients shall be integrity-protected; expired/invalid calibration shall block arming (pre-arm gate). (Ties RCR-045 pre-arm gate; RCR-024 no torn/corrupt persisted block.)
Implementation Reference | DO-FW-007 (drivers/plausibility + calibration) · DO-SAFE-007 (pre-arm gate) · CMP-038 (sensor drivers/calibration), CMP-045 (readiness gate), CMP-033/041 (NV store integrity)
Verification Reference | PV-SW-004 (Evidence-Pending) — calibration-integrity (corrupt-block rejection) + calibration-expiry gate test; fault-injection of a stale/torn config block; static analysis on CMP-038 config path
Validation Reference | N/A (integrity/expiry are design-enforced; no separate use-validation)
Residual Probability | projected Low (PL) after integrity + expiry gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Aggressive expiry could void serviceable devices (availability) — a service/logistics cost, not a safety hazard.
Overall Residual Contribution | FTA top event (Stage 11); ties HA-SW-014 (coefficient/units error), HA-SW-013 (storage integrity), HA-SVC-001/HA-MFG-001 (calibration lifecycle). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304; SYS-024; RCR-026/045/024; DO-SAFE-007. Severity per HAL-2000 §6.
Assumptions | A birth-record calibration baseline exists (RCR-053) and a calibration validity interval is defined; NV store provides integrity primitives (CMP-033/041).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-005 — Clock/timestamp error (RTC drift/failure) → excursion mis-timed / record indefensible
Hazard ID (library) | HAZ-DATA (time integrity) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring / disposition · System State: Sampling/Logging · Operational Mode: Armed / logging
Perspective | Disposition decision-maker; mission/programme (record defensibility); indirectly beneficiary
Source | PRD-SAFE; SYS-118 (trusted time / disclosed drift) / SYS-006; F-CLK (fault-management §2); judgment
Initiating Cause/Event | The real-time clock (CMP-017) drifts beyond bound, resets, or loses trusted time (cold-induced drift, brown-out, battery-backup loss), so sample and excursion timestamps are wrong or untrusted.
Contributing Conditions | Long 90-day trips accumulate drift; deep cold degrades oscillator accuracy; sleep/brown-out cycles can disturb timekeeping; trusted UTC re-discipline requires a connectivity window that may be sparse in-container.
Foreseeable Sequence of Events | RTC drifts/fails → sample timestamps wrong → an excursion is mis-timed or mis-ordered → excursion window mis-attributed (e.g., appears to fall in an allowed pre-arm interval), or the record's time base is untrustworthy → disposition made on a mis-timed/indefensible record.
Hazard | Loss of trusted monotonic time.
Hazardous Situation | A disposition decision (or an audit/recall investigation) resting on mis-timed or untrustworthy timestamps.
Sequence to Harm | Mis-timed excursion → wrong disposition of a limited quantity, or an indefensible record forcing conservative rejection/quarantine (wastage); audit non-conformance.
Harm/Loss | Wrong disposition of a limited quantity of doses; monitored trip released on a time-indefensible record; reportable data-integrity deviation (S3).
Severity | S3 (Serious) — governing dimension data-integrity/limited-quantity disposition (HAL-2000 §5 S3; RCR-025 allocated S3). Escalates toward S4 only if it directly causes a lot-scale missed excursion, which is bounded by the disclosure-when-untrusted control.
Probability (P1×P2) | P1=3 (Occasional — drift over 90 days in cold is a real, expected tension) × P2=2 (Remote — mis-timing rarely flips a pass/fail verdict by itself; excursion magnitude usually dominates over exact timing) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + oscillator-drift physics; Medium uncertainty, Evidence-Pending — closed by clock-drift/discipline test (PV-SW-005) over the cold endurance profile.
Detectability/Warning | Drift is bounded and disclosed when time is untrusted (RCR-025); re-discipline to trusted UTC on any connectivity window; a record with untrusted time carries a quality flag so the verdict is qualified rather than silently wrong.
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce as far as practicable (bounded drift + disclosure).
Risk-Control Strategy | 2 (protective — bounded drift, monotonic time, re-discipline) + 3 (information — disclose untrusted time)
Risk-Control Requirement | RCR-025 — the device shall maintain trusted monotonic UTC time across sleep/brown-out/service; drift shall be bounded and disclosed when time is untrusted.
Implementation Reference | DO-EE (RTC + backup) / DO-FW (time discipline) · CMP-017 (RTC), CMP-036 (scheduler), CMP-041 (record quality flags)
Verification Reference | PV-SW-005 (Evidence-Pending) — clock-drift/discipline test across the cold endurance profile; brown-out/sleep time-continuity test; untrusted-time disclosure test
Validation Reference | N/A (or ties audit-record validation, deferred to HAL-1400 HA-DATA-003)
Residual Probability | projected Low (PL) after bounded-drift + disclosure verified · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B boundary, S3×PL → A Med) — ALARP; residual accepted with disclosure
New Risks Introduced | Frequent re-discipline uses radio energy (minor budget cost, ties HA-ENRG-006) — not a safety hazard.
Overall Residual Contribution | Moderate; ties HA-DATA-003 (time in the record) — the record-defensibility path is owned there. Not an FTA top event by itself (S3).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304; SYS-118/006; RCR-025; F-CLK. Severity per HAL-2000 §5.
Assumptions | A backup time source survives brown-out (CMP-017 backup); connectivity windows are frequent enough to re-discipline within the drift bound (ties HAL-1100).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-006 — Alarm-not-generated: excursion detected but no alert emitted (logic/queue defect)
Hazard ID (library) | HAZ-FUNC (annunciation failure) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport · System State: Alerting (or Sleep→Alerting hand-off) · Operational Mode: Armed / alerting
Perspective | Disposition/response decision-maker (loses time-to-act); beneficiary; mission
Source | PRD-SAFE; SYS-011/013/040 (alert pre-emption, offline-first); DO-SAFE-009; state-mode-model §3.2; judgment
Initiating Cause/Event | An excursion is correctly detected and committed, but a device-side software defect — a lost alert hand-off, a queue that silently drops or never flushes, a state-machine path that returns to sleep without dispatching, a priority-inversion that starves the alert task — means no alert is emitted to the fleet.
Contributing Conditions | Alerts pre-empt sleep and jump the routine-upload queue (the one permitted radio-gating override); a defect in that pre-emption/hand-off path is high-consequence; store-and-forward queuing must be lossless; this record is the device-side generation/commit obligation (delivery across network/cloud is HA-COMM/HA-CLOUD).
Foreseeable Sequence of Events | Excursion detected + committed → alert hand-off lost or queue defect → device returns to sleep with no alert dispatched and no local flag of the failed dispatch → responder never learns in time → no cold-chain intervention → compromised product proceeds.
Hazard | Failure to generate/emit an excursion alert.
Hazardous Situation | A committed excursion with no alert in flight and no responder aware, while there is still time to intervene.
Sequence to Harm | No alert → no timely intervention → compromised dose released/administered → beneficiary not protected.
Harm/Loss | Delayed/missing intervention on a lot → compromised doses administered (S4).
Severity | S4 (Critical) — public-health (delayed/missing alert, HAL-2000 §7).
Probability (P1×P2) | P1=2 (Remote — alert path is safety-critical firmware with an alert-latency/escalation test) × P2=3 (Occasional — the committed local record + escalation/ack + receipt QA provide downstream catches even if the immediate alert is missed) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by alert-generation + escalation test (PV-SW-006).
Detectability/Warning | Alerts require acknowledgement and escalate if unacknowledged (RCR-029), so a never-received alert self-detects via missing-ack escalation; the excursion is committed locally independent of connectivity (RCR-028), so it still exists in the record for receipt QA; the alarm path cannot be silently suppressed (RCR-035, cyber twin).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — ack-required, escalate-if-unacknowledged, alert-first) + 1 (inherent — excursion committed locally so it never depends on the alert to exist)
Risk-Control Requirement | RCR-029 — excursion alerts sent ahead of bulk, require acknowledgement, and escalate if unacknowledged. (+ RCR-047 alerting tuned to false-alarm/missed-excursion targets so fatigue never mutes a real excursion; + RCR-028 excursion committed locally.)
Implementation Reference | DO-SAFE-009 (→ DO-FW-011 comms agent + DO-CLD-007/008 rules/notify) · CMP-042 (comms agent), CMP-048 (cloud notify), CMP-036 (scheduler/priority)
Verification Reference | PV-SW-006 (Evidence-Pending) — alert-generation + latency + escalation-on-no-ack test; fault-injection of lost hand-off / full queue / task starvation; static analysis on the CMP-042 alert path
Validation Reference | VAL-SW-006 (Evidence-Pending) — responder-workflow validation that an unacknowledged excursion escalates to a human in time (ties HAL-1200/1500)
Residual Probability | projected Low (PL) after ack/escalation + local-commit verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Escalation chains can add to alarm load → fatigue (ties HA-USE-002) — balanced by RCR-047.
Overall Residual Contribution | FTA top event (Stage 11); the device-side twin of HA-COMM-001/002 and cyber HA-CYBER-002. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Cloud)
Rationale/Evidence | ISO 14971; IEC 62304; SYS-011/013/040; RCR-029/047/028; DO-SAFE-009. Severity per HAL-2000 §7.
Assumptions | The cloud escalation policy is configured and monitored (HA-CLOUD); ack semantics are end-to-end reconciled (SYS-176).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-007 — Misleading alarm / false verdict displayed (accept shown for a reject)
Hazard ID (library) | HAZ-FUNC / HAZ-DATA (presentation-integrity failure) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition · System State: Verdict display (OLED and/or dashboard) · Operational Mode: Disposition read
Perspective | Disposition decision-maker (acts on the wrong display); beneficiary; mission
Source | PRD-SAFE; SYS-016/017 (consistent, honest verdict); DO-SAFE-006; IEC 62366-1; judgment
Initiating Cause/Event | A software defect in the presentation path (CMP-040 → CMP-030 OLED, or the cloud dashboard CMP-050) renders the wrong verdict — an "accept" icon/colour/word shown for a computed reject/quarantine (inverted mapping, stale cached verdict, wrong record association, rendering/localisation defect).
Contributing Conditions | The verdict is a compact icon+colour+word the operator trusts at a glance; a stale or mis-associated display is plausible-looking; device and cloud are two display surfaces that can disagree; cold/gloved/low-light reading reduces scepticism.
Foreseeable Sequence of Events | Verdict engine computes reject → display path inverts/caches/mis-associates → operator sees "accept" → releases a lot the device actually rejected → compromised dose administered.
Hazard | Presentation-integrity failure (displayed verdict ≠ computed verdict).
Hazardous Situation | An operator acting on a displayed "accept" that contradicts the true computed verdict.
Sequence to Harm | Wrong display → false accept acted on → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses released on a misdisplayed verdict (S4).
Severity | S4 (Critical) — public-health (a displayed false accept is equivalent in outcome to a computed one).
Probability (P1×P2) | P1=2 (Remote — display path verified; device↔cloud verdict consistency cross-check) × P2=4 (Probable — a wrong "accept" display is acted on with high likelihood) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by verdict-consistency + display-integrity test (PV-SW-007) and summative usability (IEC 62366-1).
Detectability/Warning | Device↔cloud verdict consistency (SYS-016, RCR-021) catches a divergent surface; the verdict is legible/language-independent (icon+colour+word, RCR-043) reducing misread; an override requires a deliberate recorded action (RCR-044).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 1 (inherent — single authoritative verdict rendered consistently device+cloud) + 3 (information — legible, unambiguous presentation)
Risk-Control Requirement | RCR-021 — device and cloud verdicts consistent; false accept suppressed. (+ RCR-043 verdict legible and language-independent, readable gloved/low-light/−30 °C; + RCR-044 recorded override.)
Implementation Reference | DO-SAFE-006 (→ DO-FW-008 verdict + DO-CLD-006 recompute + DO-HF-002 display) · CMP-040 (verdict), CMP-030 (OLED), CMP-050 (dashboard)
Verification Reference | PV-SW-007 (Evidence-Pending) — verdict-consistency (device=cloud) + display-integrity test (mapping, caching, record-association); rendering/localisation review; static analysis on the display path
Validation Reference | VAL-SW-007 (Evidence-Pending) — summative usability (IEC 62366-1): users correctly read the verdict gloved/cold/low-light and cannot be led to accept a rejected lot (ties HAL-1500)
Residual Probability | projected Low (PL) after consistency + display-integrity + usability verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | None new (shares controls with HA-SW-002 and HAL-1500 use records).
Overall Residual Contribution | FTA top event (Stage 11); ties HA-SW-002 (verdict integrity) and HA-USE-007 (override/misread). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Human-Factors and Cloud)
Rationale/Evidence | ISO 14971; IEC 62304; IEC 62366-1; SYS-016/017; RCR-021/043/044; DO-SAFE-006. Severity per HAL-2000 §6.
Assumptions | A single computed verdict is the source of truth for both surfaces; caches are invalidated on new computation.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-008 — Sensor-plausibility miss → absent/open probe reads room temp as a valid product-zone value → false accept ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-DATA / HAZ-FUNC (sensor-plausibility failure) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport · System State: Sampling/Logging · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission
Source | PRD-SAFE; SYS-007 (plausibility) / SYS-024; DO-SAFE-005; F-SNS (fault-management §2); HAL-2000 §7; judgment
Initiating Cause/Event | The external probe (CMP-020) is absent, open, short-circuited, wrong, or mis-seated, and the sensor-driver/plausibility layer (CMP-038) fails to detect it — reading, e.g., a fallback/room-temperature value or a floating input as if it were a valid product-zone temperature.
Contributing Conditions | An open/absent probe can read a plausible in-range value (room temperature is inside many acceptance windows); the device is sealed and the probe seating is not directly observed after arming; a plausible-but-wrong reading is the most dangerous kind because it defeats detection and verdict; hardware fault detection (presence/open/short) must be paired with firmware plausibility (SYS-007).
Foreseeable Sequence of Events | Probe open/absent → front-end returns a floating/fallback value → plausibility check misses it → detection sees benign "product-zone" temps throughout → no excursion, clean verdict → real product excursion (which the probe never measured) goes undetected → false accept → compromised dose administered.
Hazard | Undetected sensor/probe fault read as valid product data.
Hazardous Situation | A device reporting valid-looking product-zone temperatures it is not actually measuring.
Sequence to Harm | Invalid reading accepted as valid → missed excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered on data the device never truly measured (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §7 associates HA-SW-008 with the S4 false-accept archetype; RCR-020 allocated S4).
Probability (P1×P2) | P1=2 (Remote — absent/open/short/wrong probe and implausible reads are detected by paired HW presence-detect + FW plausibility) × P2=4 (Probable — a plausible in-range fallback reads as a clean trip and is acted on) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by probe-fault-injection test (PV-SW-008).
Detectability/Warning | Probe presence/open/short hardware detect (DO-EE-013) paired with firmware plausibility (range, rate-of-change, stuck-value, internal-vs-external divergence); a detected probe fault is flagged, never read as valid, and blocks a clean pass (RCR-020); pre-arm gate requires probe-present (RCR-045).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; the "never read as a valid product-zone temperature" clause is non-defeatable.
Risk-Control Strategy | 2 (protective — HW presence-detect + FW plausibility) + 1 (inherent — a detected fault cannot produce a clean pass)
Risk-Control Requirement | RCR-020 — absent/open/short/wrong probe and implausible sensor reads shall be detected and flagged, never read as a valid product-zone temperature. (+ RCR-045 pre-arm probe-present check.)
Implementation Reference | DO-SAFE-005 (→ DO-FW-007 drivers/plausibility + DO-EE-013 presence-detect + DO-IF-008) · CMP-038 (drivers/plausibility), CMP-020 (probe), CMP-045 (pre-arm gate)
Verification Reference | PV-SW-008 (Evidence-Pending) — probe-fault-injection test: open/short/absent/wrong/mis-seated probe across temperature, confirm every fault is flagged and never yields a valid reading or a clean pass; static analysis on CMP-038 plausibility
Validation Reference | VAL-SW-008 (Evidence-Pending) — field-representative validation that a mis-seated/absent probe is caught at arming and in-trip (ties HA-USE-005)
Residual Probability | projected Low (PL) after paired HW+FW plausibility verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Strict plausibility can flag genuine-but-unusual excursions as "implausible" and reduce confidence (ties detection tuning, HA-SW-001 / RCR-047) — tuned, noted for HAL-2700.
Overall Residual Contribution | Headline; FTA top event (Stage 11); SW-FMEA subject for CMP-038. Ties HA-USE-005 (probe placement) and HA-ELEC-008 (sensor HW fault). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Sensing/Electrical)
Rationale/Evidence | ISO 14971; IEC 62304; SYS-007/024; RCR-020/045; DO-SAFE-005; HAL-2000 §7. Physics: an open thermistor input floats to a plausible in-range value unless actively detected.
Assumptions | HW presence/open/short detect is available on the probe front-end (DO-EE-013); internal-vs-external divergence is a usable plausibility signal.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-009 — Firmware hang/deadlock → device believed monitoring but is not (silent)
Hazard ID (library) | HAZ-FUNC (liveness failure) — HA-BND-003/004
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport · System State: Sampling/Logging (nominally armed) · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission
Source | PRD-SAFE; SYS-169 (watchdog + logged recovery) / SYS-170 (availability self-check); DO-SAFE-007; fault-management §3/§6; state-mode-model §3.6; judgment
Initiating Cause/Event | A firmware hang or deadlock — task lockup, priority inversion, mutex deadlock, infinite loop, blocked I/O in CMP-036/037 — stops the sampling/detection loop while the device still appears armed and monitoring.
Contributing Conditions | Concurrency in an RTOS is a classic deadlock source; the danger is silence — a hung device shows no obvious fault to a remote observer; without watchdog + self-check the gap is invisible and the trip presents as continuous.
Foreseeable Sequence of Events | Hang/deadlock → sampling/detection stops → an excursion occurs during the dark window and is not recorded → device (if it recovers) or cloud shows a record that reads continuous → human reads an uneventful trip → false accept.
Hazard | Loss of monitoring liveness while presenting as armed.
Hazardous Situation | An armed-appearing device that is silently not monitoring, over a real excursion.
Sequence to Harm | Silent non-monitoring → missed excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered over a silent monitoring gap (S4).
Severity | S4 (Critical) — public-health.
Probability (P1×P2) | P1=2 (Remote — watchdog serviced only on a verified-healthy main loop turns a hang into a logged recovery; availability self-check flags an outage) × P2=3 (Occasional — the honest-gap log + qualified verdict catch most, but a short undetected window is credible) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by hang/fault-injection + watchdog test (PV-SW-009).
Detectability/Warning | Watchdog (CMP-018) recovers a hung device and logs the recovery as an attributable event (recovered ≠ continuous, SYS-169); availability self-check verifies the whole sensing→logging chain and flags the device not monitoring on an outage (SYS-170, RCR-027); recovery gap is written to the record and the verdict is qualified (RCR-022) — no silent, no false reassurance.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; no-false-reassurance invariant is mandatory here.
Risk-Control Strategy | 2 (protective — watchdog + safe-state + availability self-check) + 3 (information — logged gap, qualified verdict)
Risk-Control Requirement | RCR-023 — on any unsafe/misleading fault the device shall enter a defined safe state (preserve record, hold cell safe, honest status) with watchdog-backed logged recovery. (+ RCR-027 availability self-check verifies the sensing→logging chain and flags "not monitoring"; + RCR-022 honest gap.)
Implementation Reference | DO-SAFE-007 (→ DO-FW-003 mode + DO-FW-013 fault handler) · CMP-018 (watchdog), CMP-037 (power-state governor), CMP-036 (scheduler), CMP-045 (self-check)
Verification Reference | PV-SW-009 (Evidence-Pending) — hang/deadlock fault-injection + watchdog recovery test (recovery logged, gap written); availability-self-check outage-detection test; static/concurrency analysis on CMP-036/037
Validation Reference | N/A (fault-injection carries the assurance)
Residual Probability | projected Low (PL) after watchdog + self-check + logged-gap verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | A watchdog that resets too readily could cause reset churn / battery drain (ties HA-ENRG-006) — bounded by servicing only on verified-healthy loop.
Overall Residual Contribution | FTA top event (Stage 11); ties HA-SW-013 (silent gap) and HA-ENRG-006 (dead-in-cold). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304; SYS-169/170; RCR-023/027/022; DO-SAFE-007; fault-management §3/§6. Severity per HAL-2000 §6.
Assumptions | The watchdog is independent of the hung task and services only on a verified-healthy main loop (FWR-052); self-check cadence bounds the maximum silent window (cadence Evidence-Pending).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-010 — Failed/unsafe OTA update → bricked/unmonitored fleet or wrong detection logic deployed
Hazard ID (library) | HAZ-FUNC (update-integrity failure) — HA-BND-004
Product/Config | SCM all SKUs · Lifecycle Phase: OTA update (any) · System State: OTA · Operational Mode: Update window
Perspective | Mission/programme (fleet-scale); beneficiary (via wrong logic deployed); disposition decision-maker
Source | PRD-SAFE; SYS-032/139 (fail-safe OTA); DO-SAFE-014; state-mode-model §3.4; F-OTA (fault-management §2); judgment
Initiating Cause/Event | An OTA update is interrupted, unverified, version-skewed, or defective — an image that bricks the device on activation, fails to re-establish the armed/monitoring state, or deploys a defective detection/verdict build across many devices at once.
Contributing Conditions | OTA acts on the whole fleet, so a bad update is a common-cause, fleet-scale event; updates happen over intermittent in-container links; a wrong detection build re-introduces HA-SW-001 at scale; an update must not suspend monitoring or lose committed data.
Foreseeable Sequence of Events | Bad/interrupted update → (a) device bricked / non-monitoring with no notice → silent fleet gap; or (b) defective detection/verdict logic activated → systematic missed excursions / false accepts across the fleet → many wrong dispositions.
Hazard | Unsafe firmware update (brick, unmonitored window, or wrong logic deployed).
Hazardous Situation | A fleet running bricked/unmonitored, or running defective detection logic, while presenting as normal.
Sequence to Harm | Fleet-scale missed excursion / silent gap → false accepts across many lots → many beneficiaries not protected.
Harm/Loss | Multi-lot / fleet-scale compromised doses administered, or fleet-wide loss of monitoring (S4; approaches S5 at campaign scale).
Severity | S4 (Critical) typical / S5 worst (fleet-wide/campaign scale) — public-health; the common-cause fleet reach is what elevates the worst case. Record anchor S4, S5 credible worst-case retained.
Probability (P1×P2) | P1=2 (Remote — signed, integrity- and rollback-protected, fail-safe A/B OTA with automatic rollback; staged rollout) × P2=4 (Probable — if a bad build did activate, wrong logic propagates to harm invisibly, fleet-wide) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by OTA rollback + version-skew test (PV-SW-010) and pre-release detection-logic verification (PV-SW-001) on every build.
Detectability/Warning | Signature + integrity + rollback verification before activation; fail-safe A/B with automatic rollback so a failed/non-booting update reverts to known-good (never bricked); committed data/config/calibration preserved across install/activate/rollback; unmonitored windows disclosed (RCR-037); staged rollout limits blast radius; every build re-runs the detection-logic verification.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — signed, rollback-protected, fail-safe A/B) + 4 (process — staged rollout + per-build detection verification)
Risk-Control Requirement | RCR-037 — OTA shall be signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; never bricks a device; unmonitored windows disclosed. (Pairs with RCR-019/021 re-verified per build.)
Implementation Reference | DO-SAFE-014 (→ DO-FW-014/015 OTA agent + DO-CLD-010 + DO-IF-007) · CMP-043 (OTA agent), CMP-035 (bootloader/secure-boot), CMP-044 (crypto)
Verification Reference | PV-SW-010 (Evidence-Pending) — OTA rollback + version-skew + interrupted-update test (proves never-brick, monitoring preserved, auto-rollback); secure-boot/signature-reject test; per-build regression of PV-SW-001/002
Validation Reference | VAL-SW-010 (Evidence-Pending) — staged-rollout field validation (blast-radius containment, unmonitored-window disclosure)
Residual Probability | projected Low (PL) after rollback/version-skew + per-build detection verification · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | OTA download energy cost (ties HA-ENRG-006); an OTA channel is also a cyber attack surface (HA-CYBER-005 twin, HAL-1300).
Overall Residual Contribution | FTA top event (Stage 11) — common-cause fleet reach makes this a systemic residual; ties HA-CYBER-005 and HA-CLOUD-007. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Product Security, EM-ENG-008)
Rationale/Evidence | ISO 14971; IEC 62304; SYS-032/139; RCR-037; DO-SAFE-014; state-mode-model §3.4. Severity per HAL-2000 §6.
Assumptions | A/B banks and a hardware root of trust are present (CMP-035); rollout is staged with health-gated promotion; every build passes detection/verdict verification before release.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-011 — Memory exhaustion / leak → crash → lost samples / gap
Hazard ID (library) | HAZ-FUNC (resource exhaustion) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport (long trips) · System State: Sampling/Logging · Operational Mode: Armed / logging
Perspective | Disposition decision-maker; beneficiary; mission
Source | PRD-SAFE; SYS-168/169 (power-fail-safe, watchdog); SRS-DIAG-005 (memory high-water metric); fault-management §3/§8; judgment
Initiating Cause/Event | A memory leak or fragmentation in the firmware (CMP-036 scheduler/RTOS heap, CMP-041 storage buffers) exhausts RAM over a long trip, causing an allocation failure and a crash/reset.
Contributing Conditions | 90-day trips are long-running; small per-cycle leaks accumulate; heap fragmentation in a constrained MCU is credible; a crash mid-window can lose the in-flight sample and create a gap.
Foreseeable Sequence of Events | Leak accrues over days → allocation fails → crash/reset → in-flight sample lost, brief monitoring gap → an excursion in that window may be unrecorded → gap must be exposed or it reads as uneventful.
Hazard | Resource exhaustion → crash → sample loss / monitoring gap.
Hazardous Situation | A monitoring gap from a resource-exhaustion crash on a long trip.
Sequence to Harm | Gap over a real excursion → false accept → compromised dose administered (bounded by honest-gap exposure, which downgrades most to conservative reject/wastage).
Harm/Loss | Lost samples / limited-quantity wrong disposition, or wastage from a conservative reject on the gap (S3; escalates to S4 only if a large excursion falls entirely in the gap and is missed).
Severity | S3 (Serious) — data-integrity / limited-quantity (a single reset gap is bounded and made honest; RCR-023/024 allocated S3/S4). Anchor S3, with S4 noted if a full excursion is masked.
Probability (P1×P2) | P1=2 (Remote — bounded memory budget, high-water-mark health monitoring, static allocation where practical, IEC 62304 verification) × P2=2 (Remote — a short reset gap rarely coincides with, and fully masks, a real excursion; honest-gap exposure catches it) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + SRS-DIAG-005 metrics; Medium uncertainty, Evidence-Pending — closed by long-duration soak + memory-leak test (PV-SW-011).
Detectability/Warning | Memory high-water-mark monitored as a health indicator (SRS-DIAG-005); watchdog recovers the crash and logs it (SYS-169); the reset gap is written to the record and the verdict qualified (RCR-022/023); power-fail-safe commit protects the last committed sample (RCR-024).
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP
Risk-Control Strategy | 2 (protective — bounded memory, watchdog, power-fail-safe commit) + 3 (information — logged gap, qualified verdict) + 4 (process — soak testing)
Risk-Control Requirement | RCR-023 — safe-state with watchdog-backed logged recovery on any fault. (+ RCR-024 power-fail-safe storage: no torn/lost record, last commit preserved.)
Implementation Reference | DO-SAFE-007 (safe-state, RCR-023) / DO-SAFE-008 (storage, RCR-024) · CMP-036 (RTOS/heap), CMP-041 (storage manager), CMP-018 (watchdog), CMP-045 (health monitor)
Verification Reference | PV-SW-011 (Evidence-Pending) — long-duration soak / leak-detection test (representative 90-day-equivalent duty cycle) + memory-high-water static analysis; crash-then-recover gap-logging test
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after bounded-memory + soak verified · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None new (shares recovery/gap controls with HA-SW-009/013).
Overall Residual Contribution | Small; a reliability-plus-honest-gap residual. Not an FTA top event by itself (S3); feeds SW-FMEA robustness. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304 (resource-management verification); SYS-168/169; SRS-DIAG-005; RCR-023/024. Severity per HAL-2000 §5.
Assumptions | Memory budget is bounded with margin and static allocation is used on safety-critical paths where practical; soak duration is representative of the longest field trip.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-012 — Race / timing defect → duplicate/dropped sample or missed shock-triggered capture
Hazard ID (library) | HAZ-FUNC (concurrency/timing defect) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport · System State: Sampling/Logging (scheduled + event-driven) · Operational Mode: Armed / logging
Perspective | Disposition decision-maker; mission; indirectly beneficiary
Source | PRD-SAFE; SYS-002 (gap-free) / SYS-020 (shock/event capture); state-mode-model §3.1; F-SW; judgment
Initiating Cause/Event | A race condition or timing defect between the scheduled-sample path and the event-driven (accel-shock IRQ) path in CMP-036/041 causes a duplicate sample, a dropped sample, or a missed shock-triggered capture (e.g., an ISR and the sampler contend for the same buffer/record index).
Contributing Conditions | The device mixes RTC-scheduled sampling with asynchronous shock interrupts (FWR-010/011/020); shared buffers/indices are a race surface; a dropped or mis-ordered sample degrades the record; a missed shock capture loses evidence of a handling event that may correlate with an excursion.
Foreseeable Sequence of Events | Concurrent scheduled + shock events → race on shared record state → sample dropped/duplicated or shock capture missed → record has a subtle gap/duplication or missing shock event → disposition made on a slightly degraded record.
Hazard | Concurrency/timing defect corrupting sample continuity or event capture.
Hazardous Situation | A disposition made on a record with a dropped/duplicated sample or a missing shock event.
Sequence to Harm | Dropped sample/missed shock → limited-quantity wrong disposition or reduced-confidence record → wastage or a marginal missed excursion.
Harm/Loss | Wrong disposition of a limited quantity / degraded-confidence record (S3).
Severity | S3 (Serious) — data-integrity / limited-quantity (RCR-024 allocated S4 for record integrity broadly; a single dropped sample is bounded to S3 here, S4 only if it masks a real excursion).
Probability (P1×P2) | P1=2 (Remote — concurrency verified; atomic record index; ISR/task synchronisation reviewed by static/concurrency analysis) × P2=2 (Remote — one dropped sample rarely flips a verdict; gap-free accounting flags it) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by race/stress + interval-accounting test (PV-SW-012).
Detectability/Warning | Gap-free interval accounting (every scheduled interval accounted or gap-marked, SYS-002/FWR-021/038) detects a dropped sample; commit-before-sleep ordering (FWR-015) prevents loss to a later power event; duplicate detection on record index.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP
Risk-Control Strategy | 2 (protective — atomic record index, synchronised ISR/task, commit-before-sleep) + 4 (process — concurrency/stress verification)
Risk-Control Requirement | RCR-024 — storage shall be power-fail-safe (atomic writes, brown-out barrier); no torn/lost/fabricated record; no overwrite of unsent data. (Extended here to concurrency: atomic, race-free record commit and gap-free interval accounting.)
Implementation Reference | DO-SAFE-008 (→ DO-FW-010 storage) · CMP-036 (scheduler/ISR sync), CMP-041 (storage manager / record index)
Verification Reference | PV-SW-012 (Evidence-Pending) — race/stress test (concurrent scheduled + shock events at high rate) + interval-accounting (no drop/duplicate) + shock-capture-under-load test; static/concurrency analysis on CMP-036/041
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after race-free commit verified · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | Heavier synchronisation could add latency/energy (minor; ties HA-ENRG-006) — bounded.
Overall Residual Contribution | Small; a record-continuity robustness residual. Not an FTA top event by itself (S3); feeds SW-FMEA. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer
Rationale/Evidence | ISO 14971; IEC 62304 (concurrency verification); SYS-002/020; RCR-024; state-mode-model §3.1. Severity per HAL-2000 §5.
Assumptions | Record index/commit is the single shared resource requiring atomicity; shock-capture priority is defined relative to scheduled sampling.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-013 — Gap not logged on reset/power-loss/brown-out → SILENT gap → false reassurance (dead-in-cold info harm) ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-DATA (record-continuity failure) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring transport (deep cold) · System State: Sampling/Logging → reset/brown-out → recovery · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission/programme
Source | PRD-SAFE; SYS-151 (no false reassurance) / SYS-160 (dead-in-cold) / SYS-168 (power-fail-safe); DO-SAFE-008; fault-management §6; HAL-2000 §7 ("lost record in deep cold, silent = S4"); judgment
Initiating Cause/Event | A reset, power-loss, or brown-out (heightened at −30 °C, ties HA-ENRG-006 / HA-ELEC-010) interrupts monitoring, and a firmware/storage defect fails to write an explicit gap marker on recovery — so the missing interval is silently absent rather than honestly flagged.
Contributing Conditions | Deep cold makes brown-out and dead-in-cold credible (D3); the interrupted interval leaves no sample by definition; unless the recovery path reconstructs and marks the gap, the record reads continuous; a silent gap is the single most dangerous SCM information failure (SYS-160) because an absence of excursion records reads as "no excursion occurred."
Foreseeable Sequence of Events | Reset/brown-out mid-trip → monitoring pauses → recovery path does not detect/write the gap → record shows a continuous, uneventful-looking trip → an excursion in the dark window is invisible → human reads a clean pass → false accept.
Hazard | Silent monitoring gap (record continuity failure / false reassurance).
Hazardous Situation | A compromised dose believed acceptable because a monitoring gap was never disclosed.
Sequence to Harm | Silent gap → false "accept" → potency-lost dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot of compromised doses administered over an undisclosed gap (S4; S5 at population scale).
Severity | S4 (Critical) typical / S5 worst — public-health (HAL-2000 §7 "lost record in deep cold = S4"; the silent-gap false-reassurance is the archetype the whole invariant defends). Anchor S4, S5 credible worst-case retained.
Probability (P1×P2) | P1=2 (Remote — brown-out detector arms an atomic-write barrier; reset cause captured; recovery path writes gap markers) × P2=4 (Probable — a silent gap is, by definition, read as uneventful; the harm is invisible and human-mediated) ⇒ product 8 ⇒ Medium band (PM), escalated to High under the uncertain-probability rule given the D3 brown-out coupling — Basis: judgment + D3 cold brown-out tension; High uncertainty, Evidence-Pending — closed by power-fail-during-write + gap-exposure test (PV-SW-013) and the 90-day cold endurance test (PV-ENRG-006).
Detectability/Warning | The control makes the failure honest: brown-out detect → atomic-write barrier so a torn write fails the record rather than corrupting it (SYS-168); reset cause captured (FWR-006); every interval not recorded is represented by an explicit gap marker (SYS-002); the verdict is qualified and never a clean pass over a gap (RCR-022); availability self-check flags outages (RCR-027). Without the control the failure is silent — the core danger.
Initial Risk | AP High (region R — S4 × PM/PH)
Risk Acceptability | Unacceptable without control; the honest-gap / no-false-reassurance invariant is non-defeatable (HAL-2400 §4).
Risk-Control Strategy | 2 (protective — power-fail-safe atomic writes, brown-out barrier) + 3 (design-enforced information — mandatory gap marking, verdict never clean over a gap)
Risk-Control Requirement | RCR-024 — storage power-fail-safe (atomic writes, brown-out barrier); no torn/lost/fabricated record; no overwrite of unsent data; tamper-evident. (+ RCR-022 any gap made visible in record and status, incomplete record never an unqualified clean pass; + RCR-027 availability self-check flags "not monitoring.")
Implementation Reference | DO-SAFE-008 (→ DO-FW-010 storage + DO-EE-002 rail interlock + DO-SEC-006 tamper-evidence) · DO-SAFE-006 (verdict honest gap) · CMP-041 (storage manager), CMP-033 (NV log), CMP-015 (MCU/brown-out)
Verification Reference | PV-SW-013 (Evidence-Pending) — power-fail-during-write test (torn write fails safe, last commit intact) + gap-exposure test (every interrupted interval produces a gap marker and a qualified verdict) + brown-out-barrier fault-injection; static analysis on CMP-041
Validation Reference | VAL-SW-013 (Evidence-Pending) — decision-support validation that a gap is never read as a clean pass (with RCR-040, ties HA-USE-004 / HAL-1500)
Residual Probability | projected Low (PL) after atomic-write + mandatory-gap + self-check verified (residual is P2-driven — driven by the honest-gap information control, not by eliminating resets) · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; Evidence-Pending until PV-SW-013 and PV-ENRG-006 close
New Risks Introduced | Conservative gap-marking increases quarantine/reject rate on interrupted trips → wastage/stock-out pressure (ties HA-USE-009) — accepted as the safer direction, noted for HAL-2700.
Overall Residual Contribution | Headline; FTA top event (Stage 11); SW-FMEA subject for CMP-041. Ties HA-ENRG-006 (dead-in-cold), HA-ELEC-010 (brown-out), HA-USE-004 (gap misread), HA-SW-009 (silent non-monitoring). Carried in HAL-2900 as a D3 headline residual.
Status | Evidence-Pending (power-fail + cold brown-out) — Controlled-pending-verification
Owner | Embedded/Software Safety Engineer (with Battery/Power and Electrical)
Rationale/Evidence | ISO 14971; IEC 62304; SYS-151/160/168/002; RCR-024/022/027; DO-SAFE-008; fault-management §6 (dead-in-cold is a false-negative hazard, not just an endurance miss). Severity per HAL-2000 §6/§7.
Assumptions | A brown-out detector (CMP-015) is present and arms the atomic-write barrier before rail collapse; reset cause is reliably captured across cold brown-out; the recovery path always runs before the device presents as monitoring.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SW-014 — Calibration coefficient / units error in firmware → systematic reading bias → wrong disposition
Hazard ID (library) | HAZ-DATA / HAZ-FUNC (calibration/units error) — HA-BND-003
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring (whole fleet after a defective build) · System State: Sampling/Logging · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission (systematic, fleet-scale)
Source | PRD-SAFE; RCR-053 (EOL calibration/birth record); DO-SAFE-005; judgment (classic embedded units/scale defect)
Initiating Cause/Event | A firmware coefficient or units error in the sensor-conditioning/calibration math (CMP-038) — wrong polynomial coefficient, sign error, fixed-point scaling error, or a units mismatch (e.g., °C vs a raw ADC code, or a ×10 scaling) — systematically biases every temperature reading.
Contributing Conditions | Fixed-point/integer scaling on a constrained MCU is error-prone; a constant offset/gain bias is systematic (affects the whole fleet running that build), unlike a random sensor fault; a small bias can silently shift readings just inside/outside an acceptance threshold; units errors are a recurring class of embedded defect.
Foreseeable Sequence of Events | Wrong coefficient/units in the build → every reading biased by a fixed offset/gain → a real excursion reads inside the window (missed) or a benign trip reads outside (false reject) → fleet-scale wrong dispositions in one direction.
Hazard | Systematic reading bias from a calibration/units firmware error.
Hazardous Situation | A fleet reporting biased temperatures it presents as accurate.
Sequence to Harm | Systematic bias → missed excursions (false accept → compromised doses administered) or systematic false rejects (stock-out → unvaccinated) — fleet-scale.
Harm/Loss | Multiple lots wrongly disposed across the fleet (S4; approaches S5 for a large systematic false-accept bias at campaign scale).
Severity | S4 (Critical) — public-health; the systematic, fleet-scale nature (many lots, one direction) is what makes it S4 and S5-capable.
Probability (P1×P2) | P1=2 (Remote — EOL calibration coverage + birth record + against-reference verification + static analysis catch coefficient/units errors before release) × P2=3 (Occasional — a modest bias mis-disposes only trips near the threshold; a gross bias would be caught fast in the field/at EOL) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by calibration-accuracy + units/scale verification (PV-SW-014) against traceable references.
Detectability/Warning | EOL calibration to defined coverage/accuracy with a birth record (RCR-053) catches a gross bias per-unit; against-reference verification and static/unit analysis catch coefficient/units errors pre-release; internal-vs-external channel divergence and cloud recompute cross-checks flag a systematic offset; a miscalibrated unit cannot pass EOL.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 4 (process — EOL calibration coverage + against-reference verification) + 2 (protective — integrity + expiry gate on coefficients)
Risk-Control Requirement | RCR-026 — calibration coefficients integrity-protected; expired/invalid calibration blocks arming. (+ RCR-053 EOL calibration meets defined coverage/accuracy and records a birth record; a miscalibrated unit cannot pass EOL.)
Implementation Reference | DO-SAFE-005 (→ DO-FW-007 drivers/calibration) · DO-TST/DO-MFG (EOL calibration, RCR-053) · CMP-038 (sensor drivers/conditioning), CMP-019/020 (sensor/probe)
Verification Reference | PV-SW-014 (Evidence-Pending) — calibration-accuracy test against traceable references across the −30…+60 °C range + units/scale/fixed-point static analysis and unit tests on the CMP-038 conditioning path; against-reference cross-check
Validation Reference | N/A (metrological verification carries the assurance)
Residual Probability | projected Low (PL) after EOL coverage + units verification · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | None new (shares controls with HA-SW-004, HA-MFG-001).
Overall Residual Contribution | FTA top event (Stage 11) — systematic fleet reach; ties HA-SW-004 (config/calibration integrity) and HA-MFG-001/HA-SVC-001 (calibration lifecycle). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Embedded/Software Safety Engineer (with Manufacturing/Metrology)
Rationale/Evidence | ISO 14971; IEC 62304 (units/scale a known defect class); RCR-026/053; DO-SAFE-005. Severity per HAL-2000 §6.
Assumptions | Traceable temperature references are available for EOL and verification; the conditioning path is covered by unit tests including boundary/scale cases.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale; severity anchor §6/§7 — wrong disposition S4 typical / S5 worst), HAL-2100 (P1×P2, basis taxonomy, Evidence-Pending posture; the software signature is P1 Remote (IEC 62304 rigor) × P2 Probable (invisible, human-mediated harm) ⇒ many records Medium/High), and HAL-2200 (matrix, regions R/A/B, AP; severity-first + uncertain-probability rules). All 14 records begin in region R or A and none is closable as broadly acceptable while its probability is Evidence-Pending.
Controls (HAL-2400) | RCR-019 (two-sided detection), RCR-020 (probe/plausibility), RCR-021/022 (verdict integrity / honest gap), RCR-023 (safe-state), RCR-024 (power-fail-safe storage), RCR-025 (time), RCR-026 (calibration integrity/expiry), RCR-027 (availability self-check), RCR-029/047 (alert delivery/ tuning), RCR-037 (fail-safe OTA), RCR-045/048 (pre-arm gate / profile validation), RCR-053 (EOL calibration). RCR-019/021/022 are on the release-critical, non-defeatable set (HAL-2400 §4).
Implementation | DO-SAFE-004/005/006/007/008/009/014 (DOL-1200) are the Implementation References; each points to concrete DO-FW/DO-CLD/DO-EE outputs and to CMP-035…045, CMP-015/017/018/020/030/033.
Cross-module ties | HA-SW-013 ↔ HA-ENRG-006 (dead-in-cold), HA-ELEC-010 (brown-out), HA-USE-004 (gap misread), HA-DATA-004; HA-SW-006 ↔ HA-COMM-001/002, HA-CLOUD-001; HA-SW-002/007 ↔ HA-DATA-004, HA-USE-007; HA-SW-008 ↔ HA-USE-005, HA-ELEC-008; HA-SW-010 ↔ HA-CYBER-005, HA-CLOUD-007; HA-SW-001 ↔ HA-ENV-001 (freeze), HA-DATA-001; HA-SW-003 ↔ HA-USE-003; HA-SW-005 ↔ HA-DATA-003. Cyber twins of the non-malicious defects here (falsified data, suppressed alarm, unauthorised config, malicious OTA) live in HAL-1300 (RCR-034/035/036/037).
Downstream | the S4/S5 records (HA-SW-001/002/003/004/006/007/008/009/010/013/014) become FTA top events in Stage 11 (ERFKL) and drive the deeper SW-FMEA (CMP-038/039/040/041 prioritised); all PV-SW-###/VAL-SW-### references are Evidence-Pending and feed Stages 13–14 (verification methods: detection-logic test, verdict-consistency/gap test, probe-fault-injection, hang/deadlock fault-injection, power-fail-during-write, OTA rollback/version-skew, calibration/units verification, and static/code analysis per IEC 62304).
8. Residual-risk / closure notes
• Four headline software hazards anchor this module and the whole information-harm pathway (HA-BND-003): HA-SW-001 (missed excursion — detection logic), HA-SW-002 (false accept on an incomplete record), HA-SW-008 (probe-plausibility miss), and HA-SW-013 (silent gap / dead-in-cold false reassurance). All four are severity-driven while probability is Evidence-Pending and none can close until its named test reports.
• The governing acceptance invariant — honest gap, no false reassurance — is realised across RCR-021/022/ 023/024/027 and verified by the verdict-consistency/gap, power-fail-during-write, hang/watchdog, and availability-self-check tests. A software control is not accepted if a fault can corrupt the record, present an incomplete record as a clean pass, or leave an armed device silently non-monitoring.
Software safety classification (IEC 62304) | CMP-038/039/040/041 are treated at the highest software safety class (Class-C-equivalent) for this analysis; the formal classification and its full lifecycle evidence are a Stage-11/12 deliverable carried Evidence-Pending. This class assignment is the reason P1 is credibly Remote across these records — but it does not lower severity or the P2 (harm-reach) factor.
S5-capable records | HA-SW-001, HA-SW-010, HA-SW-013, HA-SW-014 carry credible S5 worst cases (population/ fleet/campaign scale). Per HAL-2200 §6.1 an S5 residual can never be classified broadly acceptable; after verification it sits in region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off and post-market monitoring.
Release gate (HAL-2200 §8) | no R region software risk may remain open at release; every remaining region-A S4/S5 software risk carries a benefit-risk statement and residual-risk sign-off, and every Evidence-Pending probability has a named closing PV-SW-### / VAL-SW-### test listed above.
Overall contribution | together with the D3 energy items (HA-ENRG-006/008), the software module's four headline residuals are the principal information-harm inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary. This is the most consequential domain in the library — the SCM is an information device, and software is where its dominant harm is created or prevented.
New-risk watch (HAL-2700) | the safety↔availability trade recurs across this module — conservative suppression/gap-marking/plausibility (the safer direction for the beneficiary) increases false-reject/ quarantine/wastage and stock-out pressure (ties HA-USE-009, HA-MPE-003) and is tracked as a control-induced risk balanced by alarm/detection tuning (RCR-047).
HAL-1100 (HA-COMM) — Connectivity Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1100, record code HA-COMM · Owner: Principal Connectivity/Cloud/Product-Security Engineer, with System-Safety (lead) and Quality/Regulatory (EM-ENG-006). This module analyses connectivity hazards — the ways the device→cloud link (LTE-M primary, Wi-Fi/BLE offload, GNSS) can fail to move a safety event out of the device in time to act, following the ISO 14971 harm-and-exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM's connectivity — the wireless bearers and the store-and-forward transport that carry the excursion alert and the evidentiary record from the device to the responder and the cloud — can fail such that a real excursion is detected late, never delivered, mis-routed, or corrupted, and a human consequently makes a wrong disposition of vaccine. The governing exposure is the information pathway HA-BND-003 (device → alert/record → human decision → beneficiary), degraded at its transport node. One record is carried as a program TOP hazard per HAL-000 §7:
• HA-COMM-003 — RF-in-container attenuation preventing a live link inside a metallised/insulated shipper (the antenna Evidence-Pending tension, integral CMP-028 vs external CMP-029).
The load-bearing architectural insight for this whole module: because alert latency is connectivity-bound, not logic-bound (D2; communication-architecture §4), a missed live alert degrades to buffered delivery. That degradation is safe only because the excursion is committed locally the instant it is measured (RCR-028 / SYS-013). RCR-028 is therefore the single most load-bearing control in this module: it converts a delivery failure from a data-loss hazard into a delivery-latency hazard.
2. Scope
All SCM SKUs. In-scope connectivity elements: LTE-M/Cat-M1 modem (CMP-024), Wi-Fi radio (CMP-025), BLE radio (CMP-026), SIM/eSIM + roaming provisioning (CMP-027), integral multi-band antenna (CMP-028), external-antenna option (CMP-029), the device NV log / store-and-forward buffer (CMP-033/041), the comms/alert firmware (CMP-042), the RTC/time source (CMP-017), the secure element / device identity (CMP-016), and — at the receiving end of the wire — the cloud ingest surface (CMP-046) and rules/alert engine (CMP-048) insofar as the device↔cloud contract is a connectivity concern. Cloud-internal processing hazards are analysed in HAL-1200 (HA-CLOUD); security-induced connectivity hazards (jamming, alarm suppression on the wire) are analysed in HAL-1300 (HA-CYBER). Lifecycle phases: armed transport (in-container, in-transit, cross-border), depot offload, receipt. Based on D2 (offline-first), D4 (multi-market), AD-1 (store-and-forward), communication-architecture.md, and HA-BND-007 (external-dependency map).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• 3GPP LTE-M / Cat-M1 — low-power WAN concepts (PSM/eDRX, coverage, roaming) — cited by designation, revision to confirm; no clause or numeric figure asserted as fact.
• IEEE 802.11 / Bluetooth Core (LE) — offload/short-range bearer concepts (designation only).
• TLS 1.2+/1.3 — authenticated encrypted transport (designation only; SYS-137).
• FCC Part 15 / CE RED — regional radiated-power/band-plan conformance (SYS-115; designation only).
• 21 CFR Part 11 / ALCOA+ — record attributability/time defensibility on the wire (SYS-118; designation).
Upstream | D2 (offline-first alerting), D4 (multi-market), AD-1 (store-and-forward), SYS-104…118 (connectivity requirements), SRS-CLOUD-003/008/009/023, DOL-1200 (DO-SAFE-009 alert delivery), HAL-2400 (RCR register), communication-architecture.md, fault-management-diagnostics.md, HA-BND-003/007.
RF-in-container note (evidence honesty) | the numeric in-container attenuation, achievable LTE-M link margin, per-trip delivery-success rate, and the integral-vs-external antenna decision are all Evidence-Pending (SYS-104/109/117), closed by the RF-in-container qualification test PRD-VER-005. No attenuation figure, coverage percentage, or alert-latency number is asserted as fact here — each is carried Evidence-Pending with its named closing test.
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000 §6/§7): a delayed or missing excursion alert removes the window to act and anchors at S4 (public-health).
• Probability is decomposed P1×P2 with explicit basis and uncertainty (HAL-2100). Where the RF-in-container antenna unknown governs (HA-COMM-003), the band is set conservatively (one level up) and the record is held Evidence-Pending with a named PV-/PRD-VER- test (HAL-2200 §6.2 uncertain-probability rule).
Control hierarchy (HAL-2300) | inherent → protective → information → process. The store-and-forward / local-commit control (RCR-028/030) is inherent-by-design and is the reason a link failure is not a data-loss failure; warnings-only control of a safety-alert hazard is challenged (HAL-2200 §6.4).
• P2 is often the governing, uncertain factor (HAL-2100 §2): whether a delayed/absent alert actually reaches harm depends on downstream, out-of-boundary human checks (receipt QA, independent shipping indicators) captured as contributing conditions, not controlled barriers.
• Residual risk is re-estimated only after a control is defined (DO-…), implemented, and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (connectivity prompt list)
1. Alert timeliness: Can a real excursion alert be delayed beyond the window in which a responder could segregate/re-ice/quarantine the shipment? Is latency bounded, or purely connectivity-bound?
2. Alert delivery: Can an alert be lost outright (never delivered) rather than merely deferred? Is there acknowledgement and escalation, or does the device assume "sent = received"?
3. Link availability in the operating environment: Can the RF link fail to close inside the metallised, insulated shipper (the intended operating enclosure)? Across markets/bands?
4. Message integrity/ordering: Can messages arrive out-of-order, duplicated, or corrupted, producing a wrong or stale device state at the cloud?
5. Session enablers: Can certificate expiry, loss of trusted time, or DNS/endpoint failure block the upload/alert session even when radio coverage exists?
6. Coverage continuity: Can a roaming/band-plan gap at a border silently drop coverage mid-trip?
7. Congestion/backpressure: Can network congestion, backoff, or PSM/eDRX timing defer a safety alert?
8. Association integrity: Can telemetry be attributed to the wrong device/trip at the network/registry boundary, so a real excursion is recorded against the wrong shipment?
6. Hazard Analysis Register (HA-COMM records)
HA-COMM-001 — DELAYED excursion alert (offline store-and-forward latency) → no time to segregate/re-ice → compromised dose administered ★ TOP-adjacent (D2)
Hazard ID (library) | HAZ-COMM (deferred safety event) → information-integrity harm (HA-BND-003 pathway)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (in-transit, no bearer) · System State: Monitoring/logging, buffering · Operational Mode: Armed / store-and-forward
Perspective | Beneficiary/patient (ultimate harmed party via a wrong "accept"); responder/consignee (denied the window to act); mission/programme
Source | D2 (alert latency is connectivity-bound, communication-architecture §4); SYS-112; HA-BND-003; HAL-2000 §7 ("delayed/missing alert = S4")
Initiating Cause/Event | An excursion occurs while the device has no available bearer (in-transit, deep in a shipper, out of coverage); the immediate alert cannot egress and is buffered.
Contributing Conditions | Offline-first architecture means the alert is delivered only when a bearer next appears (§4); trips can run long between coverage windows; PSM/eDRX sleep adds latency; the actionable window to re-ice/segregate a shipment is finite; receipt QA is an imperfect, out-of-boundary backstop.
Foreseeable Sequence of Events | Excursion begins → alert committed locally instantly (RCR-028) → no bearer → alert buffered → hours/days pass before coverage → alert finally delivered after the shipment has been received and possibly already dispositioned → responder learns of the excursion too late to segregate → compromised dose administered.
Hazard | Safety-classed alert deferred beyond the actionable window.
Hazardous Situation | A responder holds/administers a shipment believing it uneventful because the excursion alert has not yet arrived.
Sequence to Harm | Late alert → no segregation → potency-lost dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot/shipment of compromised doses administered to a group of beneficiaries (S4; S5 at population scale).
Severity | S4 (Critical) — governing dimension public-health (HAL-2000 §6 anchor; §7 "delayed/missing alert → no time to act = S4"). S5 credible worst-case at lot/population scale.
Probability (P1×P2) | P1=3 (Occasional — in-transit no-coverage windows are a routine, expected condition on global lanes) × P2=3 (Occasional — a late alert reaches harm only if it lands after disposition and receipt QA misses; downstream checks partially mitigate) ⇒ product 9 ⇒ Medium band (PM) — Basis: judgment + D2 offline-first design reality; Medium uncertainty. Alert-latency distribution across representative lanes is Evidence-Pending — closed by the alert-latency + store-and-forward test (PV-COMM-001) and PRD-VER-005 (in-container margin).
Detectability/Warning | The device commits the excursion locally (RCR-028) and the OLED shows the verdict on receipt read even offline (edge read, AD-5), so a receipt-QA reader can catch it at the point of receipt — this is the compensating detection. The gap is the remote real-time alert, which is deferred.
Initial Risk | AP High (region R — S4 × PM under conservative posture)
Risk Acceptability | Unacceptable without control (R region cannot be accepted; HAL-2200 §6).
Risk-Control Strategy | 1 (inherent — local commit + store-and-forward) + 2 (protective — alert-first, ack-required, escalation)
Risk-Control Requirement | RCR-028 — every excursion event committed locally independent of connectivity (no excursion depends on a live link to exist). (+ RCR-029 alerts sent ahead of bulk, ack-required, escalate if unacknowledged; + RCR-030 on connectivity loss degrade to store-and-forward, never record loss, deliver alert-first on reconnection.)
Implementation Reference | DO-SAFE-009 (→ DO-FW-011 comms + DO-CLD-007 rules + DO-CLD-008 notify) · CMP-041/042 (log + comms FW), CMP-048 (cloud alert engine)
Verification Reference | PV-COMM-001 (Evidence-Pending) — alert-latency + store-and-forward test across representative coverage/no-coverage lanes (proves alert-first on reconnection, no record loss); PRD-VER-005 (Evidence-Pending) — RF-in-container margin bounds the no-bearer duration
Validation Reference | VAL-COMM-001 (Evidence-Pending) — field-representative validation that on-receipt edge read + escalated alert together close the actionable window (ties RCR-040 decision-support usability)
Residual Probability | projected Low (PL) after alert-first/escalation + on-receipt edge read verified (residual driven by the finite in-transit dark window, not by data loss) · Residual Severity: S4 (unchanged — outcome unchanged if it occurred) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off; Evidence-Pending until PV-COMM-001/PRD-VER-005
New Risks Introduced | Alert-first prioritisation + escalation can raise notification volume → alarm-fatigue path (controlled by RCR-029 consolidation + RCR-047), tracked in HAL-2700.
Overall Residual Contribution | Load-bearing D2 residual; the on-device excursion commit (RCR-028) is what makes deferral survivable. Carried in HAL-2900; FTA top-event candidate (Stage 11); ties HA-COMM-003 (RF-in-container), HA-CLOUD-001 (cloud notify), HA-ENRG-006 (silent gap).
Status | Evidence-Pending (latency distribution + in-container margin) — Controlled-pending-verification
Owner | Connectivity Engineer (with Cloud, Firmware)
Rationale/Evidence | D2; communication-architecture §4/§5; SYS-112; RCR-028/029/030; DO-SAFE-009. Severity per HAL-2000 §7.
Assumptions | On-receipt edge read (OLED verdict, offline) is available and used by receipt QA as a compensating check; the actionable window (segregate/re-ice) is long enough that a same-lane reconnection alert is still useful (Evidence-Pending, lane-dependent).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-002 — MISSING excursion alert (never delivered) → excursion unknown → compromised dose administered
Hazard ID (library) | HAZ-COMM (undelivered safety event) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / receipt · System State: Monitoring, attempting delivery · Operational Mode: Armed
Perspective | Beneficiary/patient (via wrong "accept"); responder (never notified); mission
Source | SYS-112; SRS-CLOUD-023 (no dropped safety event); HA-BND-003; HAL-2000 §7 ("missing alert = S4")
Initiating Cause/Event | An excursion alert is lost outright — not merely deferred — because a delivery attempt fails silently (dropped session, buffer eviction, "fire-and-forget" assumption that sent = received) and is never retried or reconciled.
Contributing Conditions | A naive comms design would treat a transmit as success; without acknowledgement/escalation and end-to-end reconciliation, a lost datagram vanishes; buffer wear/overflow could evict an un-acked alert; the difference from HA-COMM-001 is that here the alert never arrives at all.
Foreseeable Sequence of Events | Excursion → alert committed locally (RCR-028) → transmit attempted → session drops mid-send / cloud never acks → no retry or reconciliation → alert never re-sent → device sleeps → shipment received with no excursion notification → compromised dose administered.
Hazard | Permanent non-delivery of a safety-classed alert.
Hazardous Situation | A responder disposes of a shipment as clean because the excursion was never communicated.
Sequence to Harm | No alert ever received → no segregation → potency-lost dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered (S4; S5 at population scale).
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7). S5 credible worst-case.
Probability (P1×P2) | P1=2 (Remote — the architecture mandates ack + reconciliation, so true loss requires the control to fail) × P2=4 (Probable — a truly lost alert with no on-receipt read reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + SRS-CLOUD-023 design intent; Medium uncertainty. Closed by fault-injection on the delivery path (PV-COMM-002).
Detectability/Warning | End-to-end delivery reconciliation (SYS-176) and the ack cursor make an undelivered alert detectable by the fleet (device shows un-acked backlog); the on-receipt edge read is the last-line compensating check. Without reconciliation the loss is silent — the core danger.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; a safety event may never be silently discarded (SRS-CLOUD-023).
Risk-Control Strategy | 1 (inherent — local commit) + 2 (protective — ack/retry/reconciliation, non-silent-discard)
Risk-Control Requirement | RCR-028 (local commit) + RCR-029 (ack-required, escalate if unacknowledged) + RCR-035 (the alert path shall not be silently suppressible — suppression/loss itself raises a detectable event and escalates).
Implementation Reference | DO-SAFE-009 (→ DO-FW-011 + DO-CLD-007/008) · SYS-176 reconciliation · CMP-042 (comms FW), CMP-046/048 (ingest, alert engine)
Verification Reference | PV-COMM-002 (Evidence-Pending) — delivery-path fault-injection (drop sessions/acks; confirm retry, reconciliation, and escalation recover every alert; confirm no un-acked eviction)
Validation Reference | VAL-COMM-002 (Evidence-Pending) — fleet-operator validation that an un-acked/undelivered alert is visibly surfaced and escalated
Residual Probability | projected Very Low (PVL) after ack + end-to-end reconciliation verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None new (shares alert-path controls with HA-COMM-001; escalation volume tracked under RCR-047).
Overall Residual Contribution | After ack/reconciliation, small residual; the reconciliation guarantee (SYS-176) is what converts silent loss into detectable backlog. Carried in HAL-2900; ties HA-CLOUD-001 and HA-CYBER-002 (suppressed alarm).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Connectivity Engineer (with Cloud, Product-Security)
Rationale/Evidence | SYS-112/176; SRS-CLOUD-023; RCR-028/029/035; DO-SAFE-009. Severity per HAL-2000 §7.
Assumptions | Buffer is a view over the committed log (SYS-168), so an alert is never evicted before delivery is confirmed; reconciliation runs on every session.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-003 — In-container RF attenuation → no live link → no live alert/telemetry ★ TOP HAZARD (RF-in-container, antenna Evidence-Pending)
Hazard ID (library) | HAZ-COMM (link unavailability) → information-integrity harm (HA-BND-003/007)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport, device inside metallised/insulated shipper · System State: Monitoring, attempting egress · Operational Mode: Armed / in-container
Perspective | Beneficiary/patient (via delayed/undelivered alert → wrong "accept"); responder; mission
Source | D2/D4; communication-architecture §10 (principal open architectural risk); SYS-104/109/117; HA-BND-007; HAL-000 §7 (carried TOP hazard)
Initiating Cause/Event | The device operates inside a metallised, insulated shipper (the intended operating enclosure), whose RF attenuation can exceed the fade margin of the low-power LTE-M WAN radio, so no bearer link closes from inside the container.
Contributing Conditions | Metallised/insulated shipper is RF-hostile (block diagram I-08); GNSS is the most attenuated band and each fix is energy-costly; the integral-vs-external antenna decision (integral CMP-028 baseline vs IP67-sealed external CMP-029 fallback) is unresolved; numeric attenuation/link-margin/delivery-success are not yet quantified (SYS-104/109/117) — the headline Evidence-Pending item.
Foreseeable Sequence of Events | Device sealed in shipper → LTE-M cannot close the in-container budget → excursion occurs → alert cannot egress live → alert buffered (RCR-030) → shipment travels with no live telemetry/alert → if it reaches a depot Wi-Fi/BLE gateway the backlog drains, else the alert stays buffered until receipt → responder learns late or via on-receipt read → potential wrong disposition if the window closed.
Hazard | RF link unavailability inside the operating enclosure (loss of the egress opportunity).
Hazardous Situation | An excursion with no live external notification because the container blocks the radio.
Sequence to Harm | No live link → delayed/undelivered alert (degrades to HA-COMM-001/002) → late/absent segregation → compromised dose administered.
Harm/Loss | A lot of compromised doses administered (S4; S5 population scale) via the delayed/missing-alert chain.
Severity | S4 (Critical) — public-health via the delayed/missing-alert pathway (HAL-2000 §6/§7). S5 credible worst-case.
Probability (P1×P2) | P1=3→4 (Occasional–Probable — attenuation exceeding fade margin inside a metallised shipper is a credible, possibly frequent in-transit condition; numeric attenuation Evidence-Pending) × P2=3 (Occasional — store-and-forward + depot drain + on-receipt read partly mitigate reach-to-harm) ⇒ product 9–12 ⇒ Medium–High band, escalated to High (PH) under the uncertain-probability rule (HAL-2200 §6.2) — Basis: RF link-budget analysis + judgment; High uncertainty (antenna Evidence-Pending) — closed by the RF-in-container qualification test PRD-VER-005.
Detectability/Warning | Connectivity diagnostics (SYS-113) make a silently-attenuated device visible to the fleet (capturing but not draining); the on-device OLED edge read gives the verdict at receipt even with zero connectivity. Live external alerting is the capability lost.
Initial Risk | AP High (region R — S4 × PH under conservative escalation)
Risk Acceptability | Unacceptable without control; the uncertain-probability rule drives design by severity while the antenna number is pending (HAL-2200 §6.2).
Risk-Control Strategy | 1 (inherent — close the in-container link budget by design, or provide a qualified external-antenna fallback) + 2 (protective — store-and-forward + multi-bearer opportunism bound the consequence)
Risk-Control Requirement | RCR-031 — the RF link (integral antenna) shall close the in-container budget across markets, or an IP-sealed external-antenna fallback shall be provided; RF-in-container performance qualified. (+ RCR-030 store-and-forward degrade, alert-first on reconnection — bounds the consequence independent of the pending number.)
Implementation Reference | DO-IF (antenna) + DO-EE · DO-SAFE-009 (store-and-forward mitigation) · CMP-028 (integral antenna), CMP-029 (external-antenna option), CMP-024 (LTE-M modem), CMP-041/042 (buffer + comms FW)
Verification Reference | PRD-VER-005 (Evidence-Pending) — RF-in-container qualification: in-container attenuation, achievable LTE-M/Wi-Fi/BLE link margin, per-trip delivery-success, GNSS in-container behaviour, integral-vs-external antenna decision; PV-COMM-003 (Evidence-Pending) — store-and-forward + depot-drain test under simulated attenuation
Validation Reference | VAL-COMM-003 (Evidence-Pending) — representative-lane validation that a shipper that never reached LTE-M in transit is safely drained/read at receipt within the actionable window
Residual Probability | projected Medium→Low (PM→PL) only after PRD-VER-005 quantifies margin and the antenna strategy is fixed; if the integral antenna cannot close the budget, the external-antenna variant (CMP-029) is the fallback that lowers P1 · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; stays Open / Evidence-Pending until PRD-VER-005 closes
New Risks Introduced | External-antenna option adds an IP-sealed connector (ingress/seal path → HA-ENV-003, mechanical neighbour) and a per-market hardware variant (configuration/provisioning → HA-MFG-006/RCR-056); tracked in HAL-2700.
Overall Residual Contribution | Carried headline unknown (HAL-000 §7). The consequence is bounded by store-and-forward regardless of the number, but the link margin itself remains the single largest open connectivity item; carried explicitly in HAL-2900 benefit-risk. FTA top event (Stage 11). Neighbour flags: RF/antenna engineer (link budget), Mechanical (sealed connector, keep-out), Regulatory EM-ENG-006 (per-market radiated-power, carrier coverage).
Status | Evidence-Pending (antenna / in-container margin) — Open until PRD-VER-005 reports
Owner | Connectivity Engineer (with RF/Antenna, Mechanical, Regulatory)
Rationale/Evidence | communication-architecture §10; SYS-104/109/117; RCR-031/030; DO-IF/DO-SAFE-009. Severity per HAL-2000 §7. Store-and-forward mitigation is fixed by architecture independent of the pending number.
Assumptions | Store-and-forward (RCR-030) guarantees zero data loss regardless of attenuation (the record always survives); a depot Wi-Fi/BLE gateway or on-receipt edge read is the realistic egress/detection point when LTE-M never closes in-container.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-004 — Out-of-order / duplicate / corrupted messages → wrong or duplicated cloud state
Hazard ID (library) | HAZ-DATA (transport integrity) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Upload/alert delivery · System State: Transmitting over lossy link · Operational Mode: Store-and-forward drain
Perspective | Beneficiary/patient (via a wrong cloud-side verdict/state); responder; auditor (record defensibility)
Source | SRS-CLOUD-003 (ordered idempotent ingest); SYS-111/174/176; HA-BND-003; judgment
Initiating Cause/Event | Over a lossy/interrupted link, records/alerts arrive out-of-order, duplicated, or corrupted (retries after timeout, partial sessions, bit errors), so the cloud could build a wrong or duplicated view of the trip.
Contributing Conditions | Bounded retry/backoff re-sends after ambiguous acks (SYS-174); multiple interruptions across a long trip; without idempotent (exactly-once) ingest and integrity checks a duplicate alert could double-count or a reordered sequence could misstate excursion timing.
Foreseeable Sequence of Events | Lossy link → session interrupted mid-batch → device retries from cursor → some records duplicate / arrive out of order / one is corrupted in transit → cloud (if not idempotent/integrity-checked) records a wrong state → dashboard/verdict misstates the trip → wrong disposition or loss of record defensibility.
Hazard | Transport-induced disordering/duplication/corruption of records and alerts.
Hazardous Situation | A cloud trip record that misstates excursion presence, timing, or count.
Sequence to Harm | Wrong cloud state → wrong verdict shown → wrong disposition (or a defensibility deviation on audit).
Harm/Loss | Wrong disposition of a limited quantity of doses; reportable data-integrity deviation (S3).
Severity | S3 (Serious) — governing dimension mixed public-health/data-integrity; a limited-quantity wrong disposition or a reportable deviation (HAL-2000 §5, S3). Does not reach S4 because capture-time integrity anchoring (SYS-132) and cloud recompute (DO-CLD-006) cross-check the device-committed truth.
Probability (P1×P2) | P1=3 (Occasional — lossy links and retries are routine) × P2=2 (Remote — idempotent ingest + capture-time integrity make a wrong final state unlikely) ⇒ product 6 ⇒ Medium band (PM) — Basis: SRS-CLOUD-003 design + judgment; Medium uncertainty. Closed by ingest fault-injection / reorder-duplicate-corrupt test (PV-COMM-004).
Detectability/Warning | Idempotency key on (device identity, record sequence), integrity/tamper-evidence carried unbroken (SYS-132), and end-to-end reconciliation (SYS-176) detect and re-request divergent records; corrupted records fail the integrity check and are re-delivered.
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce as far as practicable via protective design.
Risk-Control Strategy | 2 (protective — idempotent ordered integrity-preserving ingest + reconciliation)
Risk-Control Requirement | RCR-030 (store-and-forward, ordered resume, no loss) + RCR-032 (cloud shall not drop a safety event; dead-letter/retry; end-to-end delivery reconciliation) — together enforce exactly-once, in-order, integrity-checked ingest.
Implementation Reference | DO-SAFE-009 (store-and-forward) · DO-CLD-007 (ingest/reconcile, SRS-CLOUD-003) · CMP-046 (ingest), CMP-041/042 (device buffer/comms), CMP-016 (integrity anchor)
Verification Reference | PV-COMM-004 (Evidence-Pending) — reorder/duplicate/corrupt injection on ingest (confirm exactly-once, ordered, integrity-checked, reconciled state)
Validation Reference | N/A (integrity verification carries the assurance)
Residual Probability | projected Very Low (PVL) after idempotent ingest + reconciliation verified · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small; the capture-time integrity anchor (SYS-132) makes transport the copy, not the truth. Ties HA-CLOUD-004 (record store integrity) and HA-CYBER-004 (replay). Carried in HAL-2900.
Status | Controlled-pending-verification
Owner | Connectivity Engineer (with Cloud)
Rationale/Evidence | SRS-CLOUD-003; SYS-132/174/176; RCR-030/032. Severity per HAL-2000 §5.
Assumptions | Capture-time tamper-evidence is the authoritative integrity anchor; the cloud recompute (DO-CLD-006) cross-checks the device verdict, so transport disordering cannot silently invert a verdict.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-005 — Certificate expiry / time-sync failure / DNS-endpoint failure → upload/alert session blocked
Hazard ID (library) | HAZ-COMM (session-enabler failure) → information-integrity harm (HA-BND-003/007)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (session establishment) · System State: Attempting mutual-TLS session · Operational Mode: Armed / upload window
Perspective | Beneficiary/patient (via blocked alert → delayed disposition info); responder; fleet operator
Source | SYS-137 (mutual TLS); SYS-118 (time sync); HA-BND-007; communication-architecture §8/§9; judgment
Initiating Cause/Event | The device has radio coverage but cannot establish a session — an expired/invalid device or server certificate, loss of trusted time (a TLS validity check needs a plausible clock), or a DNS/endpoint-resolution failure blocks the mutual-TLS connection.
Contributing Conditions | Long device life across cert-rotation cycles; RTC drift/loss over sleep/brown-out at −30 °C (CMP-017); endpoint/region routing depends on name resolution; a fleet-wide cert or endpoint misconfiguration can block many devices at once (common-cause).
Foreseeable Sequence of Events | Coverage present → device attempts mutual TLS → cert expired / time untrusted / endpoint unresolved → session rejected → alert/upload cannot egress despite coverage → alert deferred (degrades to HA-COMM-001) → if fleet-wide, many devices go dark → wrong/late disposition.
Hazard | Session-enabler failure blocking the authenticated link even under coverage.
Hazardous Situation | A device unable to deliver a safety alert because it cannot authenticate/resolve/timestamp the session.
Sequence to Harm | Blocked session → deferred/undelivered alert → late/absent segregation → compromised dose (or, if isolated, mere delay).
Harm/Loss | Delayed disposition info on a shipment (S3) up to a lot of compromised doses if fleet-wide and unnoticed (S4).
Severity | S3/S4 — S3 for an isolated device (delay, HAL-2000 §5); S4 for a fleet-wide common-cause cert/endpoint failure that silently blinds many devices (public-health, HAL-2000 §7). Record anchors at S4 (credible common-cause worst-case).
Probability (P1×P2) | P1=2 (Remote — cert lifecycle + time discipline + endpoint config are managed) × P2=3 (Occasional — a common-cause misconfig reaches many devices before it is noticed) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + PKI/operations reality; Medium uncertainty. Closed by cert-expiry/time-loss/endpoint-failure fault-injection (PV-COMM-005) + cloud monitoring test.
Detectability/Warning | Cloud-side observability monitors auth/handshake failures and backlog growth (SRS-CLOUD-024) → a spike is detectable; the device flags an un-acked backlog; disclosed-drift handling (SYS-118) keeps timestamps defensible while offline. RTC re-disciplines on reconnection.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control for the fleet-wide case.
Risk-Control Strategy | 2 (protective — robust time/cert/endpoint handling + fleet monitoring) + 4 (process — cert-rotation & endpoint-config operations)
Risk-Control Requirement | RCR-025 — maintain trusted monotonic UTC time across sleep/brown-out/service; drift bounded and disclosed when time is untrusted. (+ RCR-032 cloud non-silent-discard + monitoring of auth/handshake/backlog events so a session-enabler failure is detected fleet-wide; + RCR-039 cloud resilience/monitoring.)
Implementation Reference | DO-EE (RTC) + DO-FW (time discipline) · DO-CLD-007 (ingest/monitoring, SRS-CLOUD-024) · CMP-017 (RTC), CMP-016 (device identity/cert), CMP-046 (ingest endpoint)
Verification Reference | PV-COMM-005 (Evidence-Pending) — cert-expiry, time-loss, and DNS/endpoint-failure fault-injection (confirm graceful degrade, disclosed drift, backlog surfaced, monitoring alerts)
Validation Reference | VAL-COMM-005 (Evidence-Pending) — operations validation that a cert/endpoint incident is detected and remediated before fleet-wide blinding
Residual Probability | projected Low (PL) after time discipline + fleet monitoring verified · Residual Severity: S4 (unchanged for common-cause) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Automated cert rotation adds an OTA/config dependency (ties HA-CLOUD-007 / HA-CYBER-005 OTA); tracked in HAL-2700.
Overall Residual Contribution | Moderate; the material risk is the fleet-wide common-cause, mitigated by cloud monitoring (RCR-032/039). Carried in HAL-2900; ties HA-COMM-001, HA-CLOUD-008, HA-SW-005 (device time).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Connectivity Engineer (with Cloud, Product-Security)
Rationale/Evidence | SYS-118/137; communication-architecture §8/§9; RCR-025/032; SRS-CLOUD-024. Severity per HAL-2000 §5/§7.
Assumptions | Trusted time can be re-established on any reconnection; cert/endpoint config is under change control with staged rollout to avoid a single fleet-wide flip.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-006 — Roaming / regional band-plan gap → coverage loss cross-border
Hazard ID (library) | HAZ-COMM (coverage discontinuity) → information-integrity harm (HA-BND-007)
Product/Config | SCM all SKUs (multi-market) · Lifecycle Phase: Armed transport crossing carrier/national boundaries · System State: Monitoring, seeking network · Operational Mode: Armed / roaming
Perspective | Beneficiary/patient (via delayed alert); responder; mission (cross-border lanes)
Source | D4 (multi-market); SYS-115 (regional radio conformance), SYS-116 (roaming, Evidence-Pending); communication-architecture §7; HA-BND-007
Initiating Cause/Event | Crossing a carrier or national boundary mid-trip, the device loses WAN coverage because the roaming SIM/eSIM has no permitted network there, or the device's configured band plan/regulatory profile does not match the region.
Contributing Conditions | Global vaccine lanes cross boundaries mid-trip (D4); per-market carrier coverage and the SIM/eSIM strategy are Evidence-Pending (SYS-116); some regions may need per-market hardware/config variants (PRD-REG-004); a device configured for the wrong region shall not transmit there (SYS-115), so a config error also causes silence.
Foreseeable Sequence of Events | Trip crosses border → no permitted roaming partner / band mismatch → WAN coverage lost for a stretch → excursion during the gap cannot egress live → alert buffered until coverage/depot → late alert (degrades to HA-COMM-001/003).
Hazard | Loss of WAN coverage due to a roaming or band-plan gap.
Hazardous Situation | An excursion during a cross-border coverage gap with no live notification.
Sequence to Harm | Coverage gap → deferred alert → late segregation → compromised dose administered.
Harm/Loss | A lot of compromised doses administered on a cross-border lane (S4).
Severity | S4 (Critical) — public-health via delayed alert (HAL-2000 §7). Same anchor as HA-COMM-001.
Probability (P1×P2) | P1=3 (Occasional — cross-border lanes are common; roaming/band-plan gaps are a real multi-market exposure, Evidence-Pending) × P2=3 (Occasional — store-and-forward + depot drain + on-receipt read mitigate reach-to-harm) ⇒ product 9 ⇒ Medium band (PM), escalated to High under roaming uncertainty — Basis: judgment + D4 multi-market reality; High uncertainty (per-market coverage/SIM Evidence-Pending) — closed by roaming/coverage characterisation + config-verification test (PV-COMM-006) and PRD-VER-005.
Detectability/Warning | Connectivity diagnostics log the network/carrier used (SYS-113), so a device that cannot find a network is visible to the fleet; store-and-forward preserves the record; on-receipt edge read is the last-line check.
Initial Risk | AP High (region R — S4 × PM/PH under uncertainty)
Risk Acceptability | Unacceptable without control on cross-border lanes.
Risk-Control Strategy | 1 (inherent — roaming-capable SIM/eSIM selecting any permitted network) + 4 (process — correct regional config provisioned/verified) + 2 (protective — store-and-forward buffering)
Risk-Control Requirement | RCR-031 (RF link closes across markets, external-antenna fallback) + RCR-056 (correct regional/firmware config — band plan, market profile — provisioned and verified at EOL). (+ RCR-030 store-and-forward.)
Implementation Reference | DO-IF (antenna/band) · DO-MFG + DO-FW (regional config, RCR-056) · CMP-027 (SIM/eSIM roaming), CMP-024 (LTE-M modem), CMP-035 (regional firmware/config)
Verification Reference | PV-COMM-006 (Evidence-Pending) — roaming/coverage characterisation across representative lanes + regional band-plan config-verification test; PRD-VER-005 (in-container margin interacts with coverage)
Validation Reference | VAL-COMM-006 (Evidence-Pending) — representative cross-border-lane validation of coverage continuity and alert delivery
Residual Probability | projected Low (PL) after roaming strategy + config verification proven (residual driven by genuinely unserved geographies, not config error) · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; Evidence-Pending until per-market coverage known
New Risks Introduced | Per-market hardware/config variants add a provisioning-error path (HA-MFG-006 / RCR-056) and a wrong-region-transmit compliance risk (SYS-115); tracked in HAL-2700 and flagged to Regulatory (EM-ENG-006).
Overall Residual Contribution | Moderate on cross-border lanes; ties HA-COMM-001/003 and HA-MFG-006. Carried in HAL-2900. Neighbour flag: Regulatory (per-market carrier coverage, radiated-power conformance).
Status | Evidence-Pending (per-market coverage/SIM) — Controlled-pending-verification
Owner | Connectivity Engineer (with Regulatory, Manufacturing)
Rationale/Evidence | D4; communication-architecture §7; SYS-115/116; RCR-031/056/030. Severity per HAL-2000 §7.
Assumptions | A roaming SIM/eSIM can reach at least one permitted network on the intended lanes (Evidence-Pending); store-and-forward covers unavoidable gaps; correct regional config is verified at EOL.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-007 — Network congestion / backoff → alert deferred behind retry timing
Hazard ID (library) | HAZ-COMM (transient deferral) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (constrained/congested network) · System State: Attempting delivery under backoff · Operational Mode: Armed / retry
Perspective | Beneficiary/patient (via delayed alert); responder
Source | SYS-174 (bounded retry/backoff); SYS-112 (alert-first); communication-architecture §5; judgment
Initiating Cause/Event | A congested or intermittently available network causes repeated session failures; bounded exponential backoff then spaces retries further apart, so an excursion alert is deferred by the backoff schedule even though brief coverage windows exist.
Contributing Conditions | Backoff protects the energy budget by avoiding tight retry loops (SYS-174) — an intentional trade that can lengthen alert latency; congestion in disaster/outbreak scenarios (exactly when the device matters most); PSM/eDRX sleep interacts with retry timing.
Foreseeable Sequence of Events | Excursion → alert promoted to head of queue (SYS-112) → congested network fails the session → backoff lengthens the interval → next brief window missed while backed off → alert delivered later than a shorter cadence would have → delayed disposition info.
Hazard | Alert deferral induced by retry/backoff timing under congestion.
Hazardous Situation | A real excursion whose alert is delayed by backoff during a period of intermittent coverage.
Sequence to Harm | Backoff-lengthened latency → late segregation window → limited-quantity wrong disposition.
Harm/Loss | Delayed disposition info; wrong disposition of a limited quantity (S3).
Severity | S3 (Serious) — a deferral on top of available (if intermittent) coverage; a bounded delay of limited quantity (HAL-2000 §5). Distinct from HA-COMM-001 (no coverage at all) and lower than S4 because the alert-first priority and short-first backoff for the safety plane limit the added latency.
Probability (P1×P2) | P1=3 (Occasional — congestion is common, especially in the scenarios that matter) × P2=2 (Remote — alert-first priority + a bounded, alert-tuned backoff limit the added latency to harm) ⇒ product 6 ⇒ Medium band (PM) — Basis: SYS-174 design + judgment; Medium uncertainty. Closed by congestion/backoff alert-latency test (PV-COMM-007).
Detectability/Warning | Alert-first queue discipline (SYS-112) keeps the alert at the head; the un-acked cursor is visible to the fleet; backoff for the safety plane can be tuned shorter than for bulk.
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce as far as practicable.
Risk-Control Strategy | 2 (protective — alert-first priority + safety-plane-tuned backoff) + 1 (inherent — separate alert/bulk planes)
Risk-Control Requirement | RCR-029 (alerts sent ahead of bulk, ack-required, escalate if unacknowledged) + RCR-030 (store-and-forward, alert-first on reconnection) — the safety plane never queues behind bulk and its retry cadence is tuned for latency.
Implementation Reference | DO-SAFE-009 (→ DO-FW-011 comms priority + DO-CLD-008 notify) · CMP-042 (comms FW), CMP-024 (modem), CMP-048 (alert engine)
Verification Reference | PV-COMM-007 (Evidence-Pending) — congestion/backoff alert-latency test (confirm alert-plane latency under congestion meets target and beats bulk)
Validation Reference | N/A
Residual Probability | projected Low (PL) after alert-first + tuned backoff verified · Residual Severity: S3 · Residual Risk: AP Low→Medium (region A/B, S3×PL → A Med) — ALARP
New Risks Introduced | Shorter safety-plane backoff spends more energy under congestion (energy↔latency trade vs D3 budget); tracked in HAL-2700.
Overall Residual Contribution | Small; ties HA-COMM-001 and HA-CLOUD-005 (queue backlog). Carried in HAL-2900.
Status | Controlled-pending-verification
Owner | Connectivity Engineer (with Firmware)
Rationale/Evidence | SYS-112/174; communication-architecture §5; RCR-029/030. Severity per HAL-2000 §5.
Assumptions | The safety plane's backoff can be tuned independently of the bulk plane without breaching the D3 energy budget (Evidence-Pending, power budget MR-VAL-006).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-COMM-008 — Wrong device association at network/registry → telemetry mis-attributed
Hazard ID (library) | HAZ-DATA (mis-attribution at the device↔registry boundary) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning / arming / first-connect · System State: Associating identity with trip/owner · Operational Mode: Provisioning / armed
Perspective | Beneficiary/patient (via a real excursion recorded against the wrong shipment, or missed on the right one); responder (wrong or no alert); mission
Source | SYS-134 (device identity); HA-BND-005 (trust boundary); ties HA-CLOUD-003; RCR-033; judgment
Initiating Cause/Event | At the network/registry boundary the device is associated with the wrong trip, owner, or tenant — a provisioning error, duplicate identity, or stale binding — so its telemetry and alerts are attributed to a different shipment/responder than the physical one.
Contributing Conditions | Reusable device rebound to a new trip each deployment (D1); binding depends on the registry (CMP-047); a duplicate/cloned identity (ties HA-CYBER-006) or a not-cleared prior binding causes mis-attribution; a wrong binding routes the alert to the wrong responder and leaves the right shipment unmonitored in the record.
Foreseeable Sequence of Events | Device provisioned/armed with wrong trip/owner binding → real excursion on shipment A → telemetry/alert attributed to shipment/responder B → responder for A never alerted and A's record shows no excursion → A's compromised dose administered; B may also be wrongly flagged.
Hazard | Authoritative-binding error mis-attributing telemetry and alerts.
Hazardous Situation | A shipment whose real excursion is invisible on its own record while the alert (if any) went elsewhere.
Sequence to Harm | Mis-attribution → right responder not alerted / right record clean → wrong disposition of the physically-affected lot.
Harm/Loss | A lot of compromised doses administered (S4); a valid lot possibly wrongly flagged on the other side.
Severity | S4 (Critical) — public-health; a mis-attributed excursion both hides a real one and misroutes the alert (HAL-2000 §6/§7). Ties HA-CLOUD-003.
Probability (P1×P2) | P1=2 (Remote — authoritative binding + identity uniqueness are required by design) × P2=4 (Probable — a mis-attributed excursion reaches harm readily; the affected shipment looks clean) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + registry design intent; Medium uncertainty. Closed by association/routing test (PV-COMM-008, shared with HA-CLOUD-003) + identity-uniqueness check.
Detectability/Warning | Authoritative device↔trip↔owner↔responder binding (RCR-033) and unique cryptographic identity (RCR-055) prevent duplicate/stale bindings; provisioning validates the association; a device reporting against no active trip is flagged.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 1 (inherent — authoritative, unique binding) + 4 (process — provisioning validation)
Risk-Control Requirement | RCR-033 — device↔trip↔owner↔responder binding shall be authoritative and correct; alerts routed to the right responder; wrong-tenant binding prevented. (+ RCR-055 unique cryptographic identity/serial, duplicate prevented.)
Implementation Reference | DO-CLD (registry) · CMP-047 (device registry), CMP-016 (device identity/secure element)
Verification Reference | PV-COMM-008 (Evidence-Pending) — association/routing test (wrong-binding, stale-binding, duplicate-identity injection; confirm correct attribution and routing) — shared with HA-CLOUD-003
Validation Reference | VAL-COMM-008 (Evidence-Pending) — provisioning-workflow validation with representative depot operators
Residual Probability | projected Very Low (PVL) after authoritative binding + identity uniqueness verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None new (shares controls with HA-CLOUD-003, HA-CYBER-006).
Overall Residual Contribution | After authoritative binding, small residual; the mis-attribution failure is doubly harmful (hides one excursion, misroutes another), so it is a release-relevant record. Carried in HAL-2900; ties HA-CLOUD-003 (wrong tenant/device binding) and HA-CYBER-006 (spoof/clone).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Connectivity Engineer (with Cloud, Product-Security)
Rationale/Evidence | SYS-134; HA-BND-005; RCR-033/055. Severity per HAL-2000 §7. Ties HA-CLOUD-003 (same control, cloud-side view).
Assumptions | Each device carries a unique, non-cloneable cryptographic identity (RCR-055); provisioning clears any prior binding before re-arming a reused device.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, severity anchor §6/§7), HAL-2100 (P1×P2, basis taxonomy, Evidence-Pending posture §7), and HAL-2200 (matrix, regions R/A/B, AP, severity-first + uncertain-probability rules).
Controls | RCR-025, RCR-028, RCR-029, RCR-030, RCR-031, RCR-032, RCR-033, RCR-035, RCR-055, RCR-056 (HAL-2400) are the connectivity-module controls; RCR-028 (local excursion commit) is the load-bearing, release-critical inherent control that makes every delivery failure a latency — not a data-loss — hazard.
Implementation | DO-SAFE-009 (alert delivery, DOL-1200), DO-IF (antenna), DO-CLD-007/008 (ingest/notify), DO-FW-011 (comms), DO-EE (RTC) are the Implementation References; CMP-016/017/024/025/026/027/028/029/041/ 042/046/047/048 are the allocated components.
Cross-module ties | HA-COMM-001/003 → HA-CLOUD-001 (cloud notify) and HA-ENRG-006 (silent gap) via the HA-BND-003 pathway; HA-COMM-002 → HA-CYBER-002 (suppressed alarm); HA-COMM-004 → HA-CLOUD-004 (record integrity) + HA-CYBER-004 (replay); HA-COMM-005 → HA-CLOUD-008 (outage) + HA-SW-005 (device time); HA-COMM-006 → HA-MFG-006 (regional config); HA-COMM-008 ↔ HA-CLOUD-003 (same authoritative-binding control, RCR-033) and HA-CYBER-006 (spoof/clone).
Downstream | S4 records (HA-COMM-001/002/003/005/006/008) become FTA top events in Stage 11 (ERFKL) and drive deeper connectivity/cloud analyses; all PV-/PRD-VER- references are Evidence-Pending and feed Stages 13–14. HA-COMM-003 (RF-in-container) is the carried headline unknown per HAL-000 §7, closed by PRD-VER-005.
8. Residual-risk / closure notes
Carried TOP hazard | HA-COMM-003 (RF-in-container attenuation) is severity-driven while the antenna / in-container link margin is Evidence-Pending; it cannot be closed until PRD-VER-005 reports and the integral-vs-external antenna decision (CMP-028 vs CMP-029) is fixed. Its consequence is bounded independent of the pending number by store-and-forward (RCR-030) — the load-bearing degradation-to-buffered path.
Load-bearing control | RCR-028 (local commit of every excursion) is the reason a missed live alert is survivable; it is release-critical and its verification (offline-excursion + store-and-forward test) is a release gate.
Release gate (HAL-2200 §8) | no R region connectivity risk may remain open at release; every remaining region-A S4 connectivity risk carries a benefit-risk statement and residual-risk sign-off; every Evidence-Pending probability has a named closing test (PV-COMM-001…008, PRD-VER-005).
Overall contribution | the connectivity module's residual is dominated by HA-COMM-003 (RF-in-container) and the delayed/missing-alert pair (HA-COMM-001/002); these are the principal connectivity inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
New-risk watch (HAL-2700) | the external-antenna variant (sealed connector, per-market variants), the safety-plane backoff energy↔latency trade, and automated cert rotation's OTA dependency are tracked as control-introduced risks.
Neighbour flags (CLAUDE.md rule 2) | RF/Antenna engineer (in-container link budget, GNSS), Mechanical (sealed external-antenna connector, keep-out), Regulatory EM-ENG-006 (per-market radiated-power, carrier coverage), Product-Security (jamming/DoS → HA-CYBER-007, alarm suppression → HA-CYBER-002).
HAL-1200 (HA-CLOUD) — Cloud-Platform Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1200, record code HA-CLOUD · Owner: Principal Connectivity/Cloud/Product-Security Engineer, with System-Safety (lead) and Quality/Regulatory (EM-ENG-006). This module analyses cloud-platform hazards — the ways the ingest→rules→record→dashboard→OTA back end can fail to raise, deliver, attribute, retain, or process a safety event, following the ISO 14971 harm-and-exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM's cloud platform — the second half of the information pathway HA-BND-003 (device → cloud record/alert/dashboard → human decision → beneficiary) — can fail such that a real excursion is not flagged, not delivered, mis-attributed, lost from the evidentiary record, processed too late, or fleet-wide invisible, and a human consequently makes a wrong disposition of vaccine. The governing architectural invariant is SRS-CLOUD-023 / RCR-032: the cloud shall not drop a safety event — dead-letter/retry, never silent discard, with end-to-end delivery reconciliation. The cloud is a distinct exposure surface from connectivity (HAL-1100): even with a perfect link, a rules-engine misconfig, a mis-binding, a lost record, or a bungled OTA rollout can defeat the whole chain.
2. Scope
All SCM SKUs (served by the shared platform). In-scope cloud services: ingest (CMP-046), device registry (CMP-047), rules/alert engine (CMP-048), record store (CMP-049), dashboard (CMP-050), OTA service (CMP-051), API gateway (CMP-052), RBAC/SSO (CMP-053), audit store (CMP-054). In-scope behaviours: event processing, alerting/escalation, tenant/device binding, record durability & backup/restore, queue/backpressure handling, OTA campaign coordination, and platform/regional availability & observability. Security-induced cloud compromise (ransomware, auth bypass, malicious OTA) is analysed in HAL-1300 (HA-CYBER) and cross-referenced; the device↔cloud transport contract is in HAL-1100 (HA-COMM). Based on the system architecture (S8), SRS-CLOUD requirements, DOL-1200 (DO-SAFE-009 alert delivery, DO-SAFE-014 OTA), DO-CLD outputs, and HA-BND-003/007.
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• IEC 62443 / AAMI TIR57 — security→safety (a compromised/mis-operated platform is a safety hazard; cross-ref HAL-1300) — cited by designation.
• ISO/IEC 27001 concepts — availability, backup/restore, change management, monitoring — designation only.
• 21 CFR Part 11 / ALCOA+ — electronic-record integrity, durability, attributability, audit trail (record store / audit store) — designation only.
Upstream | SRS-CLOUD-003 (idempotent ordered ingest), SRS-CLOUD-008 (event processing, no dropped safety event), SRS-CLOUD-009 (notification/escalation), SRS-CLOUD-023 (graceful logged handling / dead-letter), SRS-CLOUD-024 (observability), SYS-175 (platform reliability/RPO/RTO, Evidence-Pending), DOL-1200 (DO-SAFE-009/014), HAL-2400 (RCR register), fault-management-diagnostics.md §8, HA-BND-003/005/007.
Evidence honesty | platform SLA, RPO/RTO, queue-depth thresholds, and regional-failover behaviour are Evidence-Pending (SYS-175), closed by the cloud reliability program, DR test, and fault-injection on ingest/notify. No availability figure or latency number is asserted as fact here — each is Evidence-Pending with its named closing test.
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000 §6/§7): a cloud failure that hides or delays an excursion, mis-attributes it, or loses the evidentiary record anchors at S4 (public-health); fleet-wide compromise/outage can reach S5.
• Probability is decomposed P1×P2 (HAL-2100). Cloud common-cause failures (a bad deploy, a rules misconfig, an outage) can reach many devices at once — P1 is judged at the fleet level and P2 reflects the imperfect out-of-boundary human checks (receipt QA, on-receipt edge read).
Control hierarchy (HAL-2300) | inherent → protective → information → process. The non-silent-discard / dead-letter guarantee (RCR-032) and cross-checked device/cloud verdicts (RCR-021) are protective; warnings-only control of a safety-event-loss hazard is challenged (HAL-2200 §6.4).
• Residual risk is re-estimated only after a control is defined (DO-…), implemented, and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (cloud prompt list)
1. Notification delivery: Once an excursion reaches the cloud, can the notification still fail to reach a responder (notify-service failure, no escalation, wrong channel)?
2. Rules correctness: Can the rules/alert engine be mis-configured so an excursion is not flagged, or the wrong product-stability policy is applied?
3. Attribution: Can a device/trip be bound to the wrong tenant/responder so alerts and records are mis-attributed?
4. Record durability: Can the evidentiary record be corrupted, lost, or non-recoverable (DB corruption, failed backup/restore) so a disposition or audit is made on missing data?
5. Timeliness under load: Can queue backlog / delayed processing make an alert late even though it was ingested?
6. Disaster recovery: Can a backup fail to restore after an incident, losing records permanently?
7. Update coordination: Can an uncoordinated OTA rollout brick or blind a fleet, or leave version skew?
8. Availability & observability: Can a cloud/regional outage or a monitoring failure leave the fleet blind and alerts undelivered without anyone noticing?
6. Hazard Analysis Register (HA-CLOUD records)
HA-CLOUD-001 — Alert notification not delivered (cloud rules/notify failure) → responder unaware → compromised dose
Hazard ID (library) | HAZ-COMM/HAZ-CLOUD (notification non-delivery) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / receipt (excursion ingested) · System State: Cloud processing an excursion event · Operational Mode: Rules/notify pipeline
Perspective | Beneficiary/patient (via wrong "accept"); responder (never notified); mission
Source | SRS-CLOUD-008/009/023; DO-CLD-007/008; HA-BND-003; HAL-2000 §7 ("delayed/missing alert = S4")
Initiating Cause/Event | An excursion event reaches the cloud but the notification is never delivered — the notify service fails, the channel (SMS/email/push) bounces, or there is no escalation on non-delivery — so the responder is never told.
Contributing Conditions | Notification depends on external channels; a "fire-and-forget" notify with no ack/escalation loses the alert at the last hop; the device may have delivered its part perfectly (this is distinct from HA-COMM-002, which loses the alert on the wire); consolidation logic could suppress a real alert if mis-tuned.
Foreseeable Sequence of Events | Excursion ingested → rules engine raises alert → notify service fails / channel bounces / no escalation → responder never receives it → shipment received and dispositioned as clean → compromised dose administered.
Hazard | Non-delivery of a raised notification at the cloud→responder hop.
Hazardous Situation | A responder disposes of a shipment as clean because the raised alert never reached them.
Sequence to Harm | Notification lost → no segregation → potency-lost dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered (S4; S5 population scale).
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7). S5 credible worst-case.
Probability (P1×P2) | P1=2 (Remote — ack-required delivery + escalation are mandated) × P2=4 (Probable — an undelivered alert reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: SRS-CLOUD-009 design + judgment; Medium uncertainty. Closed by notify fault-injection + escalation test (PV-CLOUD-001).
Detectability/Warning | Ack-required notification with escalation on non-acknowledgement (RCR-029); cloud observability monitors notify failures and backlog (SRS-CLOUD-024) so a delivery failure is detectable; the on-receipt device edge read is the last-line compensating check.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (a raised safety alert must not be lost).
Risk-Control Strategy | 2 (protective — ack-required delivery, escalation, non-silent-discard, monitoring)
Risk-Control Requirement | RCR-032 — the cloud shall not drop a safety event: dead-letter/retry on failure, never silent discard; end-to-end delivery reconciliation. (+ RCR-029 alerts ack-required, escalate if unacknowledged; cloud-side consolidation avoids fatigue without suppressing a real alert.)
Implementation Reference | DO-CLD-007 (rules, SRS-CLOUD-023) + DO-CLD-008 (notify/escalation) · DO-SAFE-009 (alert delivery) · CMP-048 (rules/alert engine), CMP-046 (ingest)
Verification Reference | PV-CLOUD-001 (Evidence-Pending) — notify fault-injection + escalation test (fail channels/services; confirm dead-letter, retry, escalation, and monitoring surface every raised alert)
Validation Reference | VAL-CLOUD-001 (Evidence-Pending) — responder-workflow validation that an escalated alert reaches and is acknowledged by a responder within the actionable window
Residual Probability | projected Very Low (PVL) after ack + escalation + dead-letter verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | Escalation across channels raises volume → alarm-fatigue path (controlled by RCR-047 consolidation); tracked in HAL-2700.
Overall Residual Contribution | After ack/escalation/dead-letter, small residual; the non-silent-discard guarantee is load-bearing. Carried in HAL-2900; ties HA-COMM-001/002 (device-side alert) and HA-CYBER-002 (suppressed alarm). FTA top-event candidate (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Cloud Engineer (with Connectivity, Product-Security)
Rationale/Evidence | SRS-CLOUD-008/009/023; DO-CLD-007/008; RCR-032/029. Severity per HAL-2000 §7.
Assumptions | At least one notification channel per responder is reachable and acknowledged; consolidation never suppresses a genuinely new excursion.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-002 — Rules-engine misconfiguration → excursion not flagged / wrong product-stability policy applied
Hazard ID (library) | HAZ-CLOUD (logic/config error) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (rules evaluation) · System State: Cloud rules evaluating telemetry · Operational Mode: Rules/alert engine
Perspective | Beneficiary/patient (via wrong "accept" or wrong "reject"); responder; mission/programme (fleet-wide if a shared rule)
Source | SRS-CLOUD-008; DO-SAFE-006 (verdict consistency, DO-CLD-006 recompute); RCR-032/036; judgment
Initiating Cause/Event | The rules/alert engine is mis-configured — wrong excursion thresholds, wrong product-stability profile, a disabled or shadowed rule, or a logic error — so a real excursion is not flagged (or a valid trip is wrongly flagged).
Contributing Conditions | Multiple product-stability profiles across vaccines; shared rules across a tenant/fleet (common-cause reach); config change without staged rollout/validation; a cloud verdict that disagrees with the device verdict must be reconciled, not silently overridden.
Foreseeable Sequence of Events | Rule mis-configured (wrong threshold/profile) → real excursion evaluated against the wrong rule → not flagged → dashboard/cloud verdict shows clean → responder accepts → compromised dose administered; or a valid trip wrongly flagged → false reject → stock-out.
Hazard | Incorrect excursion-detection policy applied at the cloud.
Hazardous Situation | A trip evaluated against a wrong or absent rule, shown as clean (or wrongly failed).
Sequence to Harm | Wrong policy → missed flag → wrong accept → compromised dose (or false reject → stock-out).
Harm/Loss | A lot of compromised doses administered, or a valid lot wrongly destroyed (S4).
Severity | S4 (Critical) — public-health, both disposition-error directions (HAL-2000 §6/§7). Fleet-wide shared-rule error is S5-capable but anchors at S4 per-lot.
Probability (P1×P2) | P1=2 (Remote — authenticated/authorised/bounded/audited config changes required, RCR-036) × P2=4 (Probable — a wrong rule silently mis-verdicts every trip it touches) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + config-governance design; Medium uncertainty. Closed by rules-config test + verdict-consistency test (PV-CLOUD-002).
Detectability/Warning | Device↔cloud verdict cross-check (RCR-021 / DO-CLD-006) detects a cloud verdict that diverges from the device's independent computation; config changes are bounded, authenticated, and audited (RCR-036); profile provisioning is validated against the product-stability library (RCR-048).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 2 (protective — verdict cross-check, bounded/audited config) + 4 (process — staged config rollout + profile validation)
Risk-Control Requirement | RCR-032 (cloud non-silent-discard / integrity of the event pipeline) + RCR-036 (threshold/config changes authenticated, authorised, bounded, and audited; unauthorised/erroneous change prevented). (+ RCR-021 device/cloud verdict consistency; + RCR-048 profile-validation.)
Implementation Reference | DO-CLD-007 (rules) + DO-CLD-006 (recompute/cross-check) · DO-SEC + DO-CLD (audit) · CMP-048 (rules engine), CMP-053 (RBAC/SSO), CMP-054 (audit store)
Verification Reference | PV-CLOUD-002 (Evidence-Pending) — rules-config test (inject wrong threshold/profile/disabled rule; confirm verdict cross-check catches divergence and config change is bounded/blocked/audited)
Validation Reference | VAL-CLOUD-002 (Evidence-Pending) — config-management validation with representative operators (profile assignment, change control)
Residual Probability | projected Very Low (PVL) after verdict cross-check + bounded/audited config verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | Strict config bounds could block a legitimate urgent policy change (availability trade); handled by process, tracked in HAL-2700.
Overall Residual Contribution | After verdict cross-check, small residual; the independent device/cloud verdict (RCR-021) is the key defeater of a single-sided rules error. Carried in HAL-2900; ties HA-CYBER-003 (altered thresholds — malicious version of the same failure) and HA-SW-002 (device verdict). FTA top-event candidate.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Cloud Engineer (with Product-Security, Firmware for verdict cross-check)
Rationale/Evidence | SRS-CLOUD-008; DO-SAFE-006/DO-CLD-006; RCR-032/036/021/048. Severity per HAL-2000 §7.
Assumptions | Device and cloud compute the verdict independently so a single-sided config error is caught by divergence; profiles are validated against the stability library.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-003 — Wrong tenant/device binding → alert to wrong responder / record mis-attributed
Hazard ID (library) | HAZ-DATA (attribution error) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs (multi-tenant platform) · Lifecycle Phase: Provisioning / arming / operation · System State: Registry binding device↔trip↔owner↔responder · Operational Mode: Registry / routing
Perspective | Beneficiary/patient (real excursion mis-attributed → wrong "accept" on the affected lot); responder (wrong or no alert); other tenant (privacy/wrong-flag)
Source | SRS-CLOUD (registry); HA-BND-005; RCR-033; ties HA-COMM-008; judgment
Initiating Cause/Event | The device registry binds a device/trip to the wrong tenant or responder — stale binding on a reused device, a provisioning error, or a tenant-isolation flaw — so alerts and records are attributed to the wrong owner/responder.
Contributing Conditions | Multi-tenant platform (CMP-047/053); reusable devices rebound each trip (D1); data-residency/tenant isolation (SYS-125); a wrong binding both misroutes the alert and mis-attributes the evidentiary record; ties the device-side view HA-COMM-008.
Foreseeable Sequence of Events | Device bound to wrong tenant/responder → real excursion on shipment A → alert routed to tenant B (or nobody), A's record attributed to B → responder for A never alerted, A's record shows no excursion → A's compromised dose administered; B sees a spurious/foreign record.
Hazard | Registry mis-binding mis-attributing alerts and records across tenants/devices.
Hazardous Situation | A shipment whose real excursion is invisible to its owner while routed to the wrong party.
Sequence to Harm | Mis-binding → right responder not alerted / right record clean → wrong disposition of the affected lot.
Harm/Loss | A lot of compromised doses administered (S4); cross-tenant data exposure (data-integrity/privacy).
Severity | S4 (Critical) — public-health (mis-attributed excursion, HAL-2000 §6/§7) plus a data-integrity/privacy dimension; take the highest: S4. Ties HA-COMM-008.
Probability (P1×P2) | P1=2 (Remote — authoritative binding + tenant isolation required) × P2=4 (Probable — a mis-attributed excursion reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + registry/isolation design; Medium uncertainty. Closed by association/routing + tenant-isolation test (PV-CLOUD-003, shared with HA-COMM-008).
Detectability/Warning | Authoritative device↔trip↔owner↔responder binding (RCR-033); tenant isolation (SYS-125) + RBAC (CMP-053); provisioning validates and clears prior bindings; a device reporting against no active trip is flagged.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 1 (inherent — authoritative binding + tenant isolation) + 4 (process — provisioning validation)
Risk-Control Requirement | RCR-033 — device↔trip↔owner↔responder binding shall be authoritative and correct; alerts routed to the right responder; wrong-tenant binding prevented. (+ RCR-055 unique identity; + RCR-039 RBAC/tenant-isolation resilience.)
Implementation Reference | DO-CLD (registry) · CMP-047 (device registry), CMP-053 (RBAC/SSO), CMP-016 (device identity)
Verification Reference | PV-CLOUD-003 (Evidence-Pending) — association/routing + tenant-isolation test (wrong/stale binding, cross-tenant leakage injection; confirm correct attribution and isolation) — shared with HA-COMM-008
Validation Reference | VAL-CLOUD-003 (Evidence-Pending) — multi-tenant provisioning-workflow validation
Residual Probability | projected Very Low (PVL) after authoritative binding + isolation verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None new (shares controls with HA-COMM-008, HA-CYBER-006).
Overall Residual Contribution | After authoritative binding, small residual; the doubly-harmful nature (hides one excursion, exposes another tenant) makes it release-relevant. Carried in HAL-2900; ties HA-COMM-008 (device-side, same RCR-033) and HA-CYBER-006 (spoof/clone). FTA top-event candidate.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Cloud Engineer (with Connectivity, Product-Security)
Rationale/Evidence | HA-BND-005; SYS-125; RCR-033/055/039. Severity per HAL-2000 §7. Ties HA-COMM-008.
Assumptions | Tenant isolation is enforced at data and routing layers; provisioning clears prior bindings before re-arming a reused device.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-004 — DB corruption / lost record → evidentiary record lost → wrong disposition or indefensible audit
Hazard ID (library) | HAZ-DATA (record durability) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Record retention (post-ingest, pre/post-disposition) · System State: Record store persisting/serving trip records · Operational Mode: Record store
Perspective | Beneficiary/patient (disposition on missing data); auditor/programme (indefensible record); mission
Source | SRS-CLOUD (record store); 21 CFR Part 11 / ALCOA+; RCR-039/024; HA-BND-003; judgment
Initiating Cause/Event | The record store corrupts or loses a trip record — storage-layer corruption, a bad migration, silent bit-rot, or a delete without durability guarantees — so the evidentiary record is wrong or gone.
Contributing Conditions | Long-retention regulatory records (ALCOA+); cloud is the record-of-truth copy (device is capture-time truth, SYS-132); a corruption that also destroys the device copy's window (post-drain) loses the evidence; a disposition or audit made on a missing/corrupt record is wrong or indefensible.
Foreseeable Sequence of Events | Record store corrupts/loses a trip record → dashboard shows missing/wrong data → responder or auditor cannot see the excursion (or sees wrong data) → wrong disposition or a reportable data-integrity deviation on audit.
Hazard | Loss/corruption of the evidentiary trip record in the cloud.
Hazardous Situation | A disposition or audit performed against a missing or corrupted record.
Sequence to Harm | Lost/corrupt record → wrong or undefendable decision → compromised dose administered, or regulatory finding.
Harm/Loss | Wrong disposition of a lot (S4); major data-integrity violation / regulatory finding (S4).
Severity | S4 (Critical) — public-health and data-integrity (HAL-2000 §5 S3 "incomplete record" escalates to §7 S4 when a lot's disposition or a recall-forcing finding turns on it). Take S4.
Probability (P1×P2) | P1=2 (Remote — atomic/durable storage + backups + integrity checks required) × P2=3 (Occasional — a lost record reaches harm if it is the sole copy and the gap is not flagged) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + storage-durability design; Medium uncertainty. Closed by storage-corruption fault-injection + integrity-check test (PV-CLOUD-004).
Detectability/Warning | Record integrity checks / tamper-evidence carried from capture (SYS-132) detect corruption; end-to-end reconciliation (SYS-176) and the device's retained copy allow re-delivery; observability alerts on integrity failures (SRS-CLOUD-024); a missing record is flagged, never shown as a clean pass.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 2 (protective — durable/atomic storage, integrity checks, tested backup/restore) + 4 (process — retention & migration change control)
Risk-Control Requirement | RCR-039 — the cloud shall resist compromise/loss (tested backup/restore; monitoring of integrity events) so a platform incident cannot corrupt records silently. (+ RCR-024 power-fail-safe / no torn/lost/fabricated record — capture-side integrity anchor; + RCR-042 export preserves integrity/gaps.)
Implementation Reference | DO-CLD (record store, backup/restore) · DO-SEC (integrity) · CMP-049 (record store), CMP-054 (audit store)
Verification Reference | PV-CLOUD-004 (Evidence-Pending) — storage-corruption fault-injection + integrity-check test + backup/restore (DR) test (confirm corruption detected, record recoverable, gap flagged not hidden)
Validation Reference | VAL-CLOUD-004 (Evidence-Pending) — record-defensibility validation against ALCOA+ criteria
Residual Probability | projected Very Low (PVL) after integrity checks + tested restore verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None (integrity checks are inert; backup adds a key/remanence concern → ties HA-CYBER-008, tracked in HAL-2700).
Overall Residual Contribution | After integrity + tested restore, small residual; the device-retained copy + capture-time integrity (SYS-132) make the cloud a recoverable copy, not the sole truth. Carried in HAL-2900; ties HA-CLOUD-006 (backup/restore), HA-COMM-004 (transport integrity), HA-CYBER-009 (ransomware). FTA top-event candidate.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Cloud Engineer (with Product-Security)
Rationale/Evidence | ALCOA+/21 CFR Part 11; SYS-132/176; RCR-039/024/042. Severity per HAL-2000 §5/§7.
Assumptions | Capture-time tamper-evidence is the authoritative integrity anchor; the device retains its copy until cloud durability is confirmed (SYS-176 reconciliation).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-005 — Queue backlog / delayed processing → alert late despite ingest
Hazard ID (library) | HAZ-CLOUD (processing latency) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / receipt (load event) · System State: Ingest/processing queue under load · Operational Mode: Ingest → rules → notify pipeline
Perspective | Beneficiary/patient (via late alert); responder; fleet operator
Source | SRS-CLOUD-008 (event processing under load); SYS-175 (platform reliability, Evidence-Pending); communication-architecture §5; judgment
Initiating Cause/Event | Under a load spike (mass upload after a coverage window, a fleet-wide event, an outbreak surge), the ingest/processing queue backs up, so an already-ingested excursion is processed and notified late.
Contributing Conditions | Bursty backlog drains (store-and-forward means many devices deliver at once when coverage returns); shared platform capacity; without safety-event prioritisation in the pipeline, a safety alert queues behind bulk telemetry inside the cloud.
Foreseeable Sequence of Events | Load spike → ingest/rules/notify queue backs up → excursion event sits in queue → notification delayed → responder informed late → limited window to act missed on some shipments.
Hazard | Processing-latency deferral of a safety event inside the cloud.
Hazardous Situation | An ingested excursion whose notification is delayed by cloud backlog.
Sequence to Harm | Processing delay → late alert → late/absent segregation → limited-quantity wrong disposition (up to a lot if severe and prolonged).
Harm/Loss | Delayed disposition info; wrong disposition of a limited quantity (S3), up to a lot if backlog is severe/fleet-wide (S4).
Severity | S3/S4 — S3 for bounded backlog; S4 if a prolonged fleet-wide backlog delays many alerts past the window (public-health). Record anchors at S4 (credible fleet-wide worst-case), noting S3 typical.
Probability (P1×P2) | P1=3 (Occasional — bursty store-and-forward drains + surge scenarios are expected) × P2=2 (Remote — safety-event prioritisation + autoscale + monitoring limit reach-to-harm) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + SRS-CLOUD-008 design; Medium uncertainty. Closed by load/backpressure test (PV-CLOUD-005).
Detectability/Warning | Backlog-growth monitoring (SRS-CLOUD-024) alerts operators; safety events are prioritised in the pipeline (SRS-CLOUD-008); dead-letter/retry (RCR-032) prevents drop even when delayed.
Initial Risk | AP High (region R — S4 × PM under fleet-wide posture)
Risk Acceptability | Unacceptable without control for the fleet-wide case.
Risk-Control Strategy | 2 (protective — safety-event pipeline prioritisation, autoscale, backlog monitoring)
Risk-Control Requirement | RCR-032 — the cloud shall not drop a safety event (dead-letter/retry, never silent discard) and shall prioritise/monitor safety events so backlog does not defer them past the actionable window (SRS-CLOUD-008/024).
Implementation Reference | DO-CLD-007 (ingest/rules, SRS-CLOUD-008/023/024) · CMP-046 (ingest), CMP-048 (rules/alert engine)
Verification Reference | PV-CLOUD-005 (Evidence-Pending) — load/backpressure test (surge backlog; confirm safety-event latency stays within target, monitoring alerts, no drop)
Validation Reference | N/A
Residual Probability | projected Low (PL) after prioritisation + autoscale + monitoring verified · Residual Severity: S4 (unchanged for fleet-wide) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; Evidence-Pending until platform SLA/RPO known (SYS-175)
New Risks Introduced | Autoscale cost/capacity trade (availability vs cost); handled operationally, tracked in HAL-2700.
Overall Residual Contribution | Moderate; ties HA-COMM-007 (device backoff) and HA-CLOUD-001 (notify). Carried in HAL-2900.
Status | Evidence-Pending (platform SLA/RPO/RTO, SYS-175) — Controlled-pending-verification
Owner | Cloud Engineer
Rationale/Evidence | SRS-CLOUD-008/024; SYS-175; RCR-032. Severity per HAL-2000 §5/§7.
Assumptions | Safety events are prioritised end-to-end within the cloud pipeline; platform capacity/autoscale bounds backlog latency (Evidence-Pending, SYS-175).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-006 — Failed backup/restore → record loss on incident
Hazard ID (library) | HAZ-DATA (disaster recovery) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Post-incident recovery · System State: Restoring records after a data-loss incident · Operational Mode: DR / restore
Perspective | Auditor/programme (permanent record loss); beneficiary/patient (disposition on missing data); mission
Source | SRS-CLOUD (backup/restore); SYS-175 (RPO/RTO, Evidence-Pending); RCR-039; 21 CFR Part 11; judgment
Initiating Cause/Event | After a data-loss incident (corruption, deletion, region failure, ransomware — ties HA-CYBER-009), the backup fails to restore — untested backups, corrupt/incomplete backups, or an RPO gap — so records are permanently lost.
Contributing Conditions | Backups only help if tested and complete; an RPO gap loses recent records; a restore that itself is not integrity-checked can restore corrupt data; long-retention regulatory obligation makes permanent loss a compliance event.
Foreseeable Sequence of Events | Data-loss incident → attempt restore → backup incomplete/corrupt/untested → records permanently lost or restored wrong → dispositions/audits on missing data → wrong decisions + regulatory finding.
Hazard | Non-recoverable record loss due to backup/restore failure.
Hazardous Situation | A permanent gap in the evidentiary record base after an incident.
Sequence to Harm | Failed restore → lost records → wrong disposition + indefensible audit.
Harm/Loss | Loss of a lot's (or many lots') evidentiary records → wrong disposition / recall / major data-integrity finding (S4; S5 if systemic).
Severity | S4 (Critical) — data-integrity + public-health (HAL-2000 §5/§7); systemic loss reaches S5 (§5 "systemic falsification/loss → programme certification"). Anchor S4, noting S5 systemic worst-case.
Probability (P1×P2) | P1=2 (Remote — tested backup/restore + DR program required, but untested backups are a classic failure) × P2=3 (Occasional — a failed restore reaches harm once an incident occurs) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + DR practice; Medium uncertainty. Closed by DR / backup-restore test (PV-CLOUD-006).
Detectability/Warning | Regular tested restores + integrity verification of backups (RCR-039); the device-retained copy + reconciliation (SYS-176) provide an independent recovery source for recent records; monitoring of backup success (SRS-CLOUD-024).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 2 (protective — tested, integrity-verified backups + independent device copy) + 4 (process — DR program, RPO/RTO targets)
Risk-Control Requirement | RCR-039 — the cloud shall have tested backup/restore so a platform incident cannot lose records; recovery integrity-verified; RPO/RTO defined. (+ RCR-024 device-side record integrity as independent recovery source; + RCR-042 export-integrity.)
Implementation Reference | DO-CLD (backup/restore, DR) · DO-SEC · CMP-049 (record store), CMP-054 (audit store)
Verification Reference | PV-CLOUD-006 (Evidence-Pending) — DR test: periodic restore from backup with integrity verification, RPO/RTO measurement (confirm complete, integrity-checked recovery within target)
Validation Reference | VAL-CLOUD-006 (Evidence-Pending) — DR runbook validation / incident tabletop
Residual Probability | projected Very Low (PVL) after tested restore + device-copy recovery verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off; Evidence-Pending until RPO/RTO closed (SYS-175)
New Risks Introduced | Backup copies extend the data footprint / key-remanence surface (ties HA-CYBER-008); tracked in HAL-2700.
Overall Residual Contribution | Moderate; ties HA-CLOUD-004 (corruption) and HA-CYBER-009 (ransomware). Carried in HAL-2900.
Status | Evidence-Pending (RPO/RTO, SYS-175) — Controlled-pending-verification
Owner | Cloud Engineer (with Product-Security)
Rationale/Evidence | RCR-039/024; SYS-175/176; ALCOA+. Severity per HAL-2000 §5/§7.
Assumptions | Backups are tested on a schedule, not assumed; the device-retained copy covers the RPO gap for recent records.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-007 — Uncoordinated OTA rollout → fleet bricked / unmonitored
Hazard ID (library) | HAZ-CLOUD/HAZ-OTA (campaign coordination) → information-integrity harm (HA-BND-004)
Product/Config | SCM all SKUs · Lifecycle Phase: Fleet OTA campaign · System State: Devices applying an update · Operational Mode: OTA rollout
Perspective | Beneficiary/patient (via unmonitored/bricked devices → missed excursions fleet-wide); mission/programme; fleet operator
Source | DO-SAFE-014 (OTA never-brick); AD-7 (signed, fail-safe OTA); RCR-037; ties HA-SW-010; HA-CYBER-005; judgment
Initiating Cause/Event | A cloud-orchestrated OTA campaign is uncoordinated — a bad build pushed fleet-wide at once, no staged rollout/canary, or a version-skew between device firmware and cloud contract — so devices brick or become unmonitored during/after the update.
Contributing Conditions | Fleet-wide reach (common-cause); armed devices mid-trip must not lose monitoring during an update; a cloud/device contract mismatch (version skew) can break ingest for updated devices; unmonitored windows must be disclosed, not silent. Distinct from HA-CYBER-005 (malicious firmware) — this is an operational coordination failure.
Foreseeable Sequence of Events | Bad/uncoordinated OTA pushed to many devices → devices brick or reboot into an unmonitored/incompatible state → armed shipments go unmonitored → excursions during the window missed → compromised doses administered across the fleet.
Hazard | Fleet-scale loss of monitoring due to an uncoordinated update campaign.
Hazardous Situation | Many armed shipments unmonitored simultaneously during a botched rollout.
Sequence to Harm | Uncoordinated OTA → fleet unmonitored/bricked → missed excursions → wrong "accept" across many lots.
Harm/Loss | Fleet-wide missed excursions → many compromised doses administered (S4; S5 at campaign scale).
Severity | S4 (Critical) — public-health, fleet-wide (HAL-2000 §6/§7); campaign-scale reaches S5. Anchor S4, noting S5 worst-case. Ties HA-SW-010.
Probability (P1×P2) | P1=2 (Remote — staged rollout, A/B fail-safe, auto-rollback required, RCR-037) × P2=4 (Probable — a fleet-wide unmonitored window reaches harm across many shipments) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + AD-7/DO-SAFE-014 design; Medium uncertainty. Closed by OTA rollback + version-skew + staged-rollout test (PV-CLOUD-007).
Detectability/Warning | Fail-safe A/B apply with automatic rollback (RCR-037) so a bad image self-recovers; staged/canary rollout limits blast radius; unmonitored windows disclosed (RCR-037); version-skew handled by contract compatibility; fleet observability tracks update health (SRS-CLOUD-024).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (fleet-wide common-cause).
Risk-Control Strategy | 2 (protective — A/B fail-safe + auto-rollback + version-skew handling) + 4 (process — staged/canary rollout, campaign coordination)
Risk-Control Requirement | RCR-037 — OTA shall be signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; never bricks a device; unmonitored windows disclosed; staged rollout limits fleet exposure. (Cloud-orchestration side of DO-SAFE-014.)
Implementation Reference | DO-SAFE-014 (→ DO-FW-014/015 + DO-CLD-010 OTA orchestration) · CMP-051 (OTA service), CMP-043 (device OTA client), CMP-035 (firmware)
Verification Reference | PV-CLOUD-007 (Evidence-Pending) — OTA rollback + version-skew + staged-rollout test (inject bad build; confirm auto-rollback, bounded blast radius, disclosed unmonitored window, contract compatibility)
Validation Reference | VAL-CLOUD-007 (Evidence-Pending) — OTA-campaign runbook validation on a representative fleet segment
Residual Probability | projected Very Low (PVL) after A/B rollback + staged rollout verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off (EM-ENG-006 + EM-ENG-001 + EM-ENG-008)
New Risks Introduced | Staged rollout leaves a mixed-version fleet transiently (version-skew handled by contract); auto-rollback adds a rollback-attack surface (ties HA-CYBER-005, RCR-037 rollback-protection); tracked in HAL-2700.
Overall Residual Contribution | After A/B rollback + staging, small residual; the never-brick + disclosed-window guarantee is load-bearing. Carried in HAL-2900; ties HA-SW-010 (device OTA) and HA-CYBER-005 (malicious OTA). FTA top-event candidate.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Cloud Engineer (with Firmware, Product-Security)
Rationale/Evidence | AD-7; DO-SAFE-014; RCR-037. Severity per HAL-2000 §7. Ties HA-SW-010.
Assumptions | A/B fail-safe with auto-rollback is genuinely independent of the pushed image; armed devices preserve trip data and monitoring across an update or disclose the gap.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CLOUD-008 — Cloud / regional outage or monitoring failure → fleet blind, alerts undelivered
Hazard ID (library) | HAZ-CLOUD (availability/observability) → information-integrity harm (HA-BND-003/007)
Product/Config | SCM all SKUs · Lifecycle Phase: Any operational · System State: Cloud degraded/unavailable or unmonitored · Operational Mode: Platform outage
Perspective | Beneficiary/patient (undelivered alerts fleet-wide); mission/programme; fleet operator (blind)
Source | SRS-CLOUD-024 (observability); SYS-175 (platform reliability, Evidence-Pending); HA-BND-007; RCR-032/039; judgment
Initiating Cause/Event | A cloud or regional outage makes the platform unavailable to receive/process/notify, or a monitoring/observability failure lets a degradation go unnoticed — so the fleet is blind and alerts are undelivered, possibly without anyone knowing.
Contributing Conditions | Single-region dependency; observability itself can fail (monitoring the monitor); an outage during a load spike compounds (ties HA-CLOUD-005); the device store-and-forward buffers the record but cannot deliver the live alert while the cloud is down.
Foreseeable Sequence of Events | Regional outage / monitoring blind spot → ingest/rules/notify unavailable → devices buffer (record safe) but live alerts undelivered → excursions during the outage not notified → if unnoticed, prolonged fleet-wide blindness → compromised doses administered across many shipments.
Hazard | Platform unavailability / undetected degradation blinding the fleet.
Hazardous Situation | A fleet-wide outage during which real excursions go un-notified and possibly unnoticed.
Sequence to Harm | Outage/blind spot → undelivered alerts → late/absent segregation → wrong "accept" across many lots.
Harm/Loss | Fleet-wide undelivered alerts → many compromised doses administered (S4; S5 at scale/prolonged).
Severity | S4 (Critical) — public-health, fleet-wide (HAL-2000 §6/§7); prolonged/systemic reaches S5. Anchor S4, noting S5 worst-case.
Probability (P1×P2) | P1=2 (Remote — resilient multi-AZ/region + monitoring required) × P2=3 (Occasional — an outage reaches harm across shipments in its window; store-and-forward + on-receipt read mitigate) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + cloud-resilience design; Medium uncertainty. Closed by outage/failover + observability fault-injection (PV-CLOUD-008) + DR test.
Detectability/Warning | Observability with alerting on service health, auth/integrity failures, and backlog growth (SRS-CLOUD-024); resilient architecture (multi-AZ/region failover); on recovery, buffered alerts drain alert-first (RCR-030); device store-and-forward guarantees no record loss during the outage.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (fleet-wide).
Risk-Control Strategy | 2 (protective — resilient platform + independent observability + failover) + 4 (process — incident response, DR)
Risk-Control Requirement | RCR-039 — the cloud shall resist compromise/outage (resilient architecture, monitoring of health/auth/integrity/backlog, tested backup/restore) so a platform incident cannot silently blind the fleet. (+ RCR-032 non-silent-discard so buffered/undelivered safety events are recovered and reconciled on recovery.)
Implementation Reference | DO-CLD (resilience/observability, SRS-CLOUD-024) · DO-SEC · CMP-046/048/049/050 (ingest/rules/record/dashboard), CMP-053/054 (RBAC/audit)
Verification Reference | PV-CLOUD-008 (Evidence-Pending) — outage/failover + observability fault-injection (simulate regional outage and monitoring blind spot; confirm failover, alerting on degradation, buffered-alert recovery on restore) + DR test
Validation Reference | VAL-CLOUD-008 (Evidence-Pending) — incident-response / fleet-blind tabletop validation
Residual Probability | projected Low (PL) after failover + independent observability verified · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; Evidence-Pending until platform SLA/RPO/RTO closed (SYS-175)
New Risks Introduced | Multi-region failover adds data-residency complexity (SYS-125/GDPR) and a wider attack surface (ties HA-CYBER-009); tracked in HAL-2700.
Overall Residual Contribution | Moderate–high; the store-and-forward record guarantee bounds data loss, but live-alert delivery genuinely depends on the platform. Carried in HAL-2900; ties HA-CLOUD-005 (backlog), HA-COMM-005 (session enablers), HA-CYBER-009 (ransomware/cloud compromise). FTA top-event candidate.
Status | Evidence-Pending (platform SLA/RPO/RTO, SYS-175) — Controlled-pending-verification
Owner | Cloud Engineer (with Product-Security)
Rationale/Evidence | SRS-CLOUD-024; SYS-175; RCR-032/039. Severity per HAL-2000 §7.
Assumptions | Observability is independent enough to detect its own degradation; device store-and-forward covers the record for the outage duration; buffered alerts drain alert-first on recovery.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, severity anchor §6/§7), HAL-2100 (P1×P2, basis taxonomy, Evidence-Pending posture §7), and HAL-2200 (matrix, regions R/A/B, AP, severity-first rules; fleet-common-cause posture on P1).
Controls | RCR-024, RCR-029, RCR-032, RCR-033, RCR-036, RCR-037, RCR-039, RCR-042, RCR-048, RCR-055 (HAL-2400) are the cloud-module controls; RCR-032 (no dropped safety event / SRS-CLOUD-023) and RCR-039 (cloud resilience/backup) are the load-bearing platform controls.
Implementation | DO-CLD-006/007/008/010 (recompute/rules/notify/OTA), DO-SAFE-009 (alert delivery), DO-SAFE-014 (OTA) are the Implementation References; CMP-046/047/048/049/050/051/052/053/054 are the allocated cloud components.
Cross-module ties | HA-CLOUD-001 ↔ HA-COMM-001/002 (alert path) + HA-CYBER-002 (suppressed alarm); HA-CLOUD-002 ↔ HA-CYBER-003 (altered thresholds) + HA-SW-002 (device verdict); HA-CLOUD-003 ↔ HA-COMM-008 (same RCR-033 binding control) + HA-CYBER-006 (spoof); HA-CLOUD-004/006 ↔ HA-CYBER-009 (ransomware) + HA-COMM-004 (transport integrity); HA-CLOUD-007 ↔ HA-SW-010 (device OTA) + HA-CYBER-005 (malicious OTA); HA-CLOUD-005/008 ↔ HA-COMM-005/007 (latency/session).
Downstream | S4 records (all eight) become FTA top events in Stage 11 (ERFKL) and drive deeper cloud/SW analyses; all PV-CLOUD- references are Evidence-Pending and feed Stages 13–14. Platform SLA/RPO/RTO (SYS-175) is the carried Evidence-Pending item across HA-CLOUD-005/006/008.
8. Residual-risk / closure notes
Load-bearing controls | RCR-032 (cloud never drops a safety event; SRS-CLOUD-023) and RCR-039 (cloud resilience + tested backup/restore) are release-critical; their verification (ingest/notify fault-injection, DR test) is a release gate. Device store-and-forward (RCR-028/030) is the independent backstop that keeps the record safe through any cloud failure — the cloud can lose availability without losing evidence.
Evidence-Pending platform parameters | platform SLA, RPO, and RTO (SYS-175) are unresolved and gate the residual of HA-CLOUD-005/006/008; they close on the cloud reliability program and DR test.
Release gate (HAL-2200 §8) | no R region cloud risk may remain open at release; every remaining region-A S4 cloud risk carries a benefit-risk statement and residual-risk sign-off (S4 → EM-ENG-006 + EM-ENG-001, with EM-ENG-008 for OTA/security); every Evidence-Pending probability has a named closing test (PV-CLOUD-001…008, DR test).
Overall contribution | the cloud module's residual is dominated by the notification-delivery / availability set (HA-CLOUD-001/005/008) and the record-durability set (HA-CLOUD-004/006); these are the principal cloud inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
New-risk watch (HAL-2700) | OTA rollback-attack surface, multi-region residency complexity, backup key-remanence, and autoscale cost/capacity trades are tracked as control-introduced risks.
Neighbour flags (CLAUDE.md rule 2) | Product-Security EM-ENG-008 (RBAC/SSO, ransomware, malicious OTA → HAL-1300), Regulatory EM-ENG-006 (data residency, record defensibility), Firmware EM-ENG-007 (device-side OTA and verdict cross-check).
HAL-1300 (HA-CYBER) — Cybersecurity-Induced Safety Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1300, record code HA-CYBER · Owner: Principal Connectivity/Cloud/Product-Security Engineer (EM-ENG-008), with System-Safety (lead) and Quality/Regulatory (EM-ENG-006). This module analyses security-induced SAFETY hazards — the ways an adversary crossing a trust boundary (HA-BND-005) can falsify data, suppress an alarm, alter thresholds, replay data, compromise firmware/OTA, spoof a device, exfiltrate keys, deny service, or compromise the cloud, such that the outcome is a wrong disposition of vaccine and a harmed beneficiary. Per IEC 62443 / AAMI TIR57, a security compromise of a public-health data-integrity device enters the hazard chain and is a safety hazard, not merely a confidentiality/availability concern. It is not an FMEA and not a standalone threat model (that is SRS-SEC / EM-ENG-008): each record is a complete security→safety risk chain scored severity-first per CLAUDE.md rule 3. Standards cited by designation.
---
1. Purpose
Identify, chain, and control every credible way a security compromise of the SCM device, link, or cloud reaches the beneficiary through a wrong disposition — the information pathway HA-BND-003 attacked at a trust boundary HA-BND-005. This is the security→safety bridge required by TIR57: each record shows how a threat becomes a hazard (falsified "clean" record, suppressed alarm, defeated detection, malicious firmware, compromised platform) and how the risk control is a safety control. Four records are headline security hazards that enter the hazard chain and become Stage-11 FTA top events:
• HA-CYBER-001 — falsified sensor/telemetry data → false "clean" record → compromised dose (S4/S5).
• HA-CYBER-002 — suppressed/blocked alarm → excursion hidden (S4).
• HA-CYBER-005 — compromised OTA / supply-chain → malicious firmware → fleet unsafe/blind (S5).
• HA-CYBER-009 — cloud compromise / ransomware → fleet blind + records altered/withheld (S5).
2. Scope
All SCM SKUs and the shared platform. In-scope trust boundaries (HA-BND-005): device↔cloud (mutual TLS), device↔mobile/BLE (authenticated pairing), cloud↔responder notification, USB-C data/OTA, debug port (locked in production), RF (jamming/DoS). In-scope assets: telemetry & record integrity (CMP-016 secure element, CMP-033/041 log, CMP-049 record store), the alarm/alert path (CMP-042/048), thresholds/config (CMP-044/053/054), firmware & OTA (CMP-035/043/051), device identity (CMP-016), keys & remanence (CMP-016/044), and the cloud (CMP-053 RBAC/SSO, CMP-054 audit). Out of scope here: the formal cryptographic threat model, cipher-suite/key-lifecycle detail (SYS-140/143, SRS-SEC, EM-ENG-008, Evidence-Pending) — referenced, not duplicated. Based on HA-BND-005, DOL-1100 (DO-SEC), communication-architecture §8, and the fault taxonomy F-TMP/F-OTA (fault-management §2).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• IEC 62443 — industrial/IoT security (zones/conduits, security levels) — cited by designation.
• AAMI TIR57 — security risk management for medical devices; the security→safety bridge (a security compromise that can cause harm is a safety hazard) — the governing framing of this module.
• ISO/IEC 27001 concepts — cloud information-security management (access control, monitoring, incident, backup) — designation only.
• 3GPP LTE-M security / TLS 1.2+/1.3 — link authenticity/confidentiality (SYS-137) — designation only.
Upstream | HA-BND-005 (trust boundaries/attack surface), DOL-1100 DO-SEC (integrity, secure boot, debug-lock, tamper-evidence), DO-SAFE-008 (record integrity), DO-SAFE-009 (alert), DO-SAFE-014 (OTA), SYS-132/134/137/139/140/142/143/144, SRS-CLOUD-019 (mutual auth), HAL-2400 (RCR register), HA-BND-003.
Evidence honesty | attack probabilities are Judgment anchored to controls + TIR57 exposure, all Evidence-Pending, closed by penetration testing and security fault-injection; the formal threat model / key lifecycle (SYS-140/143) is owned by SRS-SEC and Evidence-Pending. No exploit probability or attack-rate figure is asserted as fact.
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
Security→safety framing (TIR57) | each record's Harm is a safety/public-health outcome (wrong disposition → beneficiary harm), not a data-loss abstraction. Severity is fixed by that outcome (HAL-2000); falsified-data / suppressed-alarm / malicious-OTA / cloud-compromise anchor S4/S5.
• Probability is decomposed P1×P2 (HAL-2100). P1 (attacker reaches the situation) is Judgment against the deployed controls and is Evidence-Pending (closed by pen-test); P2 (situation → harm) is high because a successful falsification/suppression is designed to be undetectable and reaches harm readily.
Control hierarchy (HAL-2300) | inherent (root of trust, secure boot, capture-time integrity) → protective (signed OTA, non-suppressible alarm, RBAC/audit) → process (key management, disposal). Warnings are never a sufficient control for a security→safety hazard (HAL-2200 §6.4).
• Residual risk is re-estimated only after a control is defined (DO-SEC/…), implemented, and verified by penetration test / security fault-injection (PV-…); residuals below are projected, Evidence-Pending.
• S5 records can never be broadly acceptable (HAL-2200 §6.1) and carry EM-ENG-001 + EM-ENG-006 + EM-ENG-008 benefit-risk sign-off + post-market monitoring.
5. Categories & guiding questions (security→safety prompt list)
1. Data authenticity: Can an adversary inject/falsify telemetry or the record so a spoiled trip reads "clean"? Is integrity anchored at capture, or only in transit?
2. Alarm integrity: Can an adversary suppress/mute/drop a real alarm? Does suppression itself raise a detectable event?
3. Config/threshold integrity: Can excursion thresholds/config be altered unauthorised to defeat detection? Is every change authenticated, bounded, audited?
4. Replay/freshness: Can old "good" data be replayed over a bad trip to force a false accept?
5. Firmware/supply-chain trust: Can malicious firmware be introduced via OTA or the supply chain? Is boot rooted in hardware trust; is OTA signed, rollback-protected, fail-safe?
6. Identity/impersonation: Can a device be spoofed/cloned, or a debug port exploited, to forge records?
7. Denial of service: Can an adversary jam RF or flood the cloud to suppress alerts fleet-wide?
8. Remanence/disposal: Can data/keys persist at end-of-life to enable repudiation or key compromise?
9. Platform compromise: Can the cloud be compromised (ransomware, auth bypass) to blind the fleet and alter/withhold records?
6. Hazard Analysis Register (HA-CYBER records)
HA-CYBER-001 — Falsified sensor/telemetry data injected → false "clean" record → compromised dose administered ★ HEADLINE (security→safety)
Hazard ID (library) | HAZ-CYBER/HAZ-DATA (data-authenticity attack) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / upload (data in motion or at rest) · System State: Capturing/transmitting/storing telemetry · Operational Mode: Armed / ingest
Perspective | Beneficiary/patient (via a falsified "accept"); responder; mission/programme (public-health data base)
Source | TIR57 (security→safety); HA-BND-005 (device↔cloud trust boundary); SYS-132 (capture-time integrity); DO-SAFE-008; DO-SEC; RCR-034
Initiating Cause/Event | An adversary injects or falsifies telemetry/records — forging in-transit data, tampering with the device log, or writing fabricated "in-range" readings — so a trip that actually excursed is recorded as clean.
Contributing Conditions | The device is a public-health data-integrity device — falsified data is the highest-value attack; if integrity were anchored only in transit (TLS) a compromised endpoint or a tampered log would defeat it; capture-time authenticity (SYS-132) is what makes the record self-verifying; secure element (CMP-016) roots the integrity.
Foreseeable Sequence of Events | Attacker gains a foothold (link, endpoint, or device) → injects/forges "in-range" telemetry or edits the log → cloud/dashboard shows a clean trip → responder accepts → compromised dose administered → beneficiary not protected → preventable disease.
Hazard | Loss of data authenticity — a falsified record that misrepresents a real excursion as clean.
Hazardous Situation | A responder disposes of a spoiled shipment as clean because the record was falsified.
Sequence to Harm | Falsified "clean" record → wrong "accept" → potency-lost dose administered → beneficiary harmed (at lot/population scale → outbreak).
Harm/Loss | A lot of compromised doses administered (S4); systemic falsification at population scale → outbreak / programme-certification withdrawal (S5).
Severity | S4/S5 — S4 per-lot (public-health, HAL-2000 §6/§7 — HA-CYBER-001 listed among S5-worst/S4-typical); S5 for systemic falsification of the record base (§5). Anchor S4, headline S5 worst-case.
Probability (P1×P2) | P1=2 (Remote — capture-time integrity + secure element + mutual TLS + tamper-evidence must all be defeated; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — a successful falsification is designed to be undetectable and reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: expert judgment against DO-SEC controls; Medium uncertainty. Closed by integrity / penetration test (PV-CYBER-001).
Detectability/Warning | Capture-time cryptographic tamper-evidence (SYS-132) makes falsified/edited records detectable and rejectable at the cloud (RCR-034); device↔cloud verdict cross-check (RCR-021) flags divergence; security monitoring raises integrity-failure events (SYS-142). Without capture-time integrity the falsification is silent — the core danger.
Initial Risk | AP High (region R — S4/S5 × PM; S5 catastrophic-severity rule forces High)
Risk Acceptability | Unacceptable without control; S5 can never be broadly acceptable (HAL-2200 §6.1).
Risk-Control Strategy | 1 (inherent — capture-time authenticity rooted in hardware trust) + 2 (protective — integrity verification/rejection, verdict cross-check, monitoring)
Risk-Control Requirement | RCR-034 — telemetry and the record shall carry capture-time authenticity/integrity (tamper-evidence); falsified/injected data shall be detectable and rejected. (+ RCR-021 device/cloud verdict cross-check; + RCR-038 secure boot / root of trust anchors the capture-time signing key.)
Implementation Reference | DO-SEC (integrity) + DO-SAFE-008 (record integrity) · CMP-016 (secure element / root of trust), CMP-044 (security config), CMP-049 (record store)
Verification Reference | PV-CYBER-001 (Evidence-Pending) — integrity + penetration test (attempt data injection/forgery in transit, at endpoint, and in the log; confirm detection/rejection and verdict-divergence flag)
Validation Reference | VAL-CYBER-001 (Evidence-Pending) — record-defensibility validation that a tampered record is never presented as a clean pass
Residual Probability | projected Very Low (PVL) after capture-time integrity + verdict cross-check verified · Residual Severity: S4/S5 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) / High→A for S5 with EM-ENG-001 + EM-ENG-006 + EM-ENG-008 benefit-risk + post-market monitoring
New Risks Introduced | Cryptographic verification adds a key-management / remanence surface (ties HA-CYBER-008) and a compute/energy cost (D3 budget); tracked in HAL-2700.
Overall Residual Contribution | Headline security→safety residual; capture-time integrity (SYS-132/RCR-034) is the load-bearing control that makes the record self-verifying. Carried in HAL-2900; FTA top event (Stage 11). Ties HA-CYBER-004 (replay), HA-COMM-004 (transport integrity), HA-CLOUD-004 (record store), HA-SW-002 (verdict).
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008)
Rationale/Evidence | TIR57; SYS-132; DO-SEC/DO-SAFE-008; RCR-034/021/038. Severity per HAL-2000 §5/§6/§7.
Assumptions | Integrity is anchored at capture in the secure element (not only in transit); the capture-time signing key is protected by the hardware root of trust.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-002 — Suppressed/blocked alarm (attacker mutes/drops alert) → excursion hidden ★ HEADLINE
Hazard ID (library) | HAZ-CYBER (alarm-integrity attack) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (excursion + alarm) · System State: Alarm/alert path active · Operational Mode: Armed / alerting
Perspective | Beneficiary/patient (via hidden excursion → wrong "accept"); responder (never alerted); mission
Source | TIR57; HA-BND-005; SYS-112 (alert path); RCR-035; ties HA-COMM-002 / HA-CLOUD-001
Initiating Cause/Event | An adversary suppresses, mutes, blocks, or drops a real excursion alarm/alert — at the device (disable alert task), on the wire (block the alert plane), or in the cloud (mute the notification) — so the excursion is hidden.
Contributing Conditions | The alert path is the safety-critical channel; a naive design lets suppression be silent; the malicious counterpart of HA-COMM-002/HA-CLOUD-001 — here the non-delivery is caused, not accidental; the control principle is that suppression must itself raise a detectable, escalating event.
Foreseeable Sequence of Events | Excursion → attacker suppresses/blocks/mutes the alarm → no notification and no visible suppression event → responder unaware → shipment accepted → compromised dose administered.
Hazard | Loss of alarm integrity — a real alarm silently suppressed.
Hazardous Situation | A responder disposes of a spoiled shipment as clean because the alarm was suppressed.
Sequence to Harm | Suppressed alarm → no segregation → potency-lost dose administered → beneficiary harmed.
Harm/Loss | A lot of compromised doses administered (S4; S5 at scale).
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7). S5 credible worst-case if fleet-wide.
Probability (P1×P2) | P1=2 (Remote — alarm path integrity + non-silent-suppression + monitoring required; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — a suppressed alarm reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: expert judgment against controls; Medium uncertainty. Closed by alarm-suppression penetration test (PV-CYBER-002).
Detectability/Warning | The alert path shall not be silently suppressible (RCR-035) — suppression/blocking itself raises a detectable event and escalates; ack-required delivery with escalation (RCR-029) and end-to-end reconciliation (SYS-176) surface a missing expected alert; security monitoring flags anomalous alert-path behaviour (SYS-142).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; a safety alarm may never be silently suppressible.
Risk-Control Strategy | 2 (protective — non-suppressible alarm path, suppression raises event, ack/escalation, monitoring)
Risk-Control Requirement | RCR-035 — the alarm/alert path shall not be silently suppressible; suppression/blocking shall itself raise a detectable event and escalate. (+ RCR-029 ack-required/escalation; + RCR-032 cloud non-silent-discard.)
Implementation Reference | DO-SEC + DO-SAFE-009 (alert) · CMP-048 (rules/alert engine), CMP-042 (device comms/alert FW)
Verification Reference | PV-CYBER-002 (Evidence-Pending) — alarm-suppression penetration test (attempt to mute/drop/block at device, wire, and cloud; confirm suppression raises a detectable escalating event and reconciliation surfaces the gap)
Validation Reference | VAL-CYBER-002 (Evidence-Pending) — operations validation that a suppression event is detected and escalated within the actionable window
Residual Probability | projected Very Low (PVL) after non-suppressible-alarm + reconciliation verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) with benefit-risk sign-off (incl. EM-ENG-008)
New Risks Introduced | Suppression-detection events add to notification volume (alarm-fatigue, RCR-047); tracked in HAL-2700.
Overall Residual Contribution | Headline residual; the non-silent-suppression guarantee is load-bearing. Carried in HAL-2900; FTA top event (Stage 11). Ties HA-COMM-002 (accidental non-delivery) and HA-CLOUD-001 (notify failure) — RCR-035 is the malicious-case control over the same path.
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Cloud/Connectivity
Rationale/Evidence | TIR57; SYS-112/142/176; RCR-035/029/032; DO-SAFE-009. Severity per HAL-2000 §7.
Assumptions | The absence of an expected alarm is itself observable (reconciliation of expected vs received), so suppression cannot be perfectly silent.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-003 — Altered excursion thresholds/config (unauthorised) → detection defeated
Hazard ID (library) | HAZ-CYBER (config-integrity attack) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning / operation (config change) · System State: Threshold/config evaluation · Operational Mode: Config / rules
Perspective | Beneficiary/patient (via defeated detection → wrong "accept"); responder; mission (fleet-wide if shared config)
Source | TIR57; HA-BND-005; RCR-036; ties HA-CLOUD-002 (accidental) / HA-USE-003; DO-SEC
Initiating Cause/Event | An adversary alters excursion thresholds or configuration without authorisation — widening thresholds, disabling a rule, or loading a permissive profile — so real excursions fall outside the (now defeated) detection.
Contributing Conditions | Thresholds/config are the detection policy; unauthorised change (device, cloud, or provisioning app) defeats detection silently; the malicious counterpart of HA-CLOUD-002; shared config reaches the fleet (common-cause); control principle is authenticated, authorised, bounded, audited change.
Foreseeable Sequence of Events | Attacker changes thresholds/config → real excursion evaluated against defeated policy → not flagged → clean verdict → responder accepts → compromised dose administered.
Hazard | Loss of config integrity — detection policy silently weakened.
Hazardous Situation | Trips evaluated against attacker-widened thresholds, shown clean.
Sequence to Harm | Defeated detection → missed flag → wrong "accept" → potency-lost dose administered.
Harm/Loss | A lot of compromised doses administered (S4; S5 if fleet-wide config).
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7). S5 credible if a shared fleet config is altered.
Probability (P1×P2) | P1=2 (Remote — authenticated/authorised/bounded/audited config required, RCR-036; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — defeated detection silently mis-verdicts every affected trip) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment against config-governance; Medium uncertainty. Closed by config-tamper test (PV-CYBER-003).
Detectability/Warning | Config changes authenticated, authorised, bounded to safe ranges, and audited (RCR-036; CMP-053 RBAC, CMP-054 audit); out-of-bound or unauthorised changes rejected and logged; device↔cloud verdict cross-check (RCR-021) catches a one-sided threshold change; profile validation against the stability library (RCR-048).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 2 (protective — authenticated/bounded/audited config, verdict cross-check) + 4 (process — change control)
Risk-Control Requirement | RCR-036 — excursion thresholds/config changes shall be authenticated, authorised, bounded, and audited; unauthorised change prevented. (+ RCR-021 verdict cross-check; + RCR-048 profile validation.)
Implementation Reference | DO-SEC + DO-CLD (audit) · CMP-044 (security config), CMP-053 (RBAC/SSO), CMP-054 (audit store)
Verification Reference | PV-CYBER-003 (Evidence-Pending) — config-tamper test (attempt unauthorised/out-of-bound threshold/config change at device, cloud, app; confirm rejection, audit, and verdict-divergence flag)
Validation Reference | VAL-CYBER-003 (Evidence-Pending) — change-control / RBAC validation with representative admins
Residual Probability | projected Very Low (PVL) after authenticated/bounded/audited config + cross-check verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) with benefit-risk sign-off (incl. EM-ENG-008)
New Risks Introduced | Strict bounds could block a legitimate urgent config change (availability trade); handled by process, tracked in HAL-2700.
Overall Residual Contribution | After bounded/audited config + cross-check, small residual. Carried in HAL-2900; ties HA-CLOUD-002 (accidental misconfig — same detection defeat) and HA-USE-003 (wrong profile at provisioning).
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Cloud
Rationale/Evidence | TIR57; RCR-036/021/048; CMP-044/053/054. Severity per HAL-2000 §7.
Assumptions | Thresholds are bounded to a safe range so even an authorised change cannot silently disable detection; every change is audited and cross-checked.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-004 — Replay of good data over a bad trip → false accept
Hazard ID (library) | HAZ-CYBER (freshness/replay attack) → information-integrity harm (HA-BND-003/005)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / upload · System State: Transmitting/ingesting telemetry · Operational Mode: Armed / ingest
Perspective | Beneficiary/patient (via replayed "clean" record → wrong "accept"); responder; mission
Source | TIR57; HA-BND-005; SYS-132/133 (attribution/time); RCR-034; ties HA-CYBER-001
Initiating Cause/Event | An adversary replays a previously captured "good" (in-range) data set over a trip that actually excursed — re-sending old authentic-looking records so freshness/binding is broken and the bad trip reads clean.
Contributing Conditions | A replay uses genuinely-signed old data, so capture-time integrity alone (RCR-034) is necessary but not sufficient — freshness/binding to (device, trip, time, sequence) is required; a subtler variant of HA-CYBER-001 (falsification) using authentic-but-stale data.
Foreseeable Sequence of Events | Attacker captures a clean trip's records → during a bad trip replays them (right device identity, wrong trip/time) → if freshness/trip-binding not enforced, cloud accepts the replay as the current trip → clean verdict → compromised dose administered.
Hazard | Loss of data freshness — stale authentic records replayed as current.
Hazardous Situation | A bad trip whose record is a replay of an old clean trip.
Sequence to Harm | Replayed clean record → wrong "accept" → potency-lost dose administered → beneficiary harmed.
Harm/Loss | A lot of compromised doses administered (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7). Same anchor as HA-CYBER-001.
Probability (P1×P2) | P1=2 (Remote — nonce/sequence/trip-binding + monotonic time defeat replay; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — an accepted replay reaches harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment against freshness controls; Medium uncertainty. Closed by replay penetration test (PV-CYBER-004).
Detectability/Warning | Idempotency key on (device identity, record sequence) (SRS-CLOUD-003) + monotonic UTC time (RCR-025) + trip-binding + tamper-evidence attribution (SYS-133) reject a replay as duplicate/stale/wrong-trip; verdict cross-check (RCR-021) flags divergence.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 2 (protective — freshness/nonce/sequence + trip-binding + monotonic time)
Risk-Control Requirement | RCR-034 — capture-time authenticity/integrity with freshness (records bound to device, trip, time, and sequence); replayed/stale data detectable and rejected. (+ RCR-025 monotonic UTC time; + RCR-033 trip-binding; + RCR-032 idempotent ingest.)
Implementation Reference | DO-SEC (freshness/nonce) + DO-SAFE-008 · CMP-044 (security config), CMP-016 (identity), CMP-017 (RTC), CMP-049 (record store)
Verification Reference | PV-CYBER-004 (Evidence-Pending) — replay penetration test (replay captured records across trips/times; confirm rejection via freshness/sequence/trip-binding)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after freshness + trip-binding verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) with benefit-risk sign-off
New Risks Introduced | Nonce/sequence state adds storage/sync complexity; tracked in HAL-2700.
Overall Residual Contribution | After freshness controls, small residual; complements RCR-034 integrity. Carried in HAL-2900; ties HA-CYBER-001 (falsification), HA-COMM-004 (duplicate/reorder), HA-COMM-008/HA-CLOUD-003 (binding).
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008)
Rationale/Evidence | TIR57; SYS-133; SRS-CLOUD-003; RCR-034/025/033/032. Severity per HAL-2000 §7.
Assumptions | Records are bound to (device, trip, monotonic time, sequence) so authentic-but-stale data cannot pass as current.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-005 — Compromised OTA / supply-chain → malicious firmware → fleet unsafe/blind ★ HEADLINE
Hazard ID (library) | HAZ-CYBER/HAZ-OTA (firmware-trust attack) → information-integrity + physical-safety harm (HA-BND-004/005)
Product/Config | SCM all SKUs · Lifecycle Phase: OTA campaign / manufacturing (supply-chain) · System State: Applying/booting firmware · Operational Mode: OTA / boot
Perspective | Beneficiary/patient (via fleet-wide falsified/missed excursions); mission/programme; potentially physical safety (a malicious image could defeat charge-safety FW defense-in-depth — HW interlock still holds, RCR-001)
Source | TIR57; AD-7 (signed fail-safe OTA); HA-BND-004/005; DO-SAFE-014; RCR-037/038; ties HA-CLOUD-007 / HA-SW-010
Initiating Cause/Event | An adversary introduces malicious firmware — a compromised OTA (unsigned/forged image, rollback to a vulnerable version) or a supply-chain implant — so devices run attacker-controlled code fleet-wide.
Contributing Conditions | Firmware controls detection, verdict, alerting, and record integrity — a malicious image can falsify or suppress everything at once, fleet-wide (the broadest-reach attack, hence S5); the HW charge-interlock (RCR-001) is independent of FW so physical charge-safety survives, but the information safety functions do not; control principle is signed, rollback-protected, fail-safe OTA rooted in hardware trust.
Foreseeable Sequence of Events | Attacker forges/rolls-back an OTA image or implants firmware in supply-chain → devices boot malicious code → detection/verdict/alerting/record silently subverted fleet-wide → excursions falsified/hidden across the fleet → many compromised doses administered → outbreak-scale harm.
Hazard | Loss of firmware trust — attacker-controlled code on the fleet.
Hazardous Situation | A fleet running malicious firmware that falsifies/suppresses safety information.
Sequence to Harm | Malicious firmware → fleet-wide falsified/missed excursions → mass wrong "accept" → population-scale beneficiary harm.
Harm/Loss | Fleet-wide falsified/missed excursions → mass administration of compromised doses → outbreak / programme-certification withdrawal (S5).
Severity | S5 (Catastrophic) — public-health at population scale + systemic record falsification (HAL-2000 §5; HA-CYBER-005 anchored S5). Broadest-reach security→safety hazard.
Probability (P1×P2) | P1=1→2 (Improbable–Remote — secure boot + signed rollback-protected OTA + supply-chain controls; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — malicious firmware reaches harm across the fleet) ⇒ product 4–8 ⇒ Low–Medium band, escalated to High by the S5 catastrophic-severity rule (HAL-2200 §6.1) — Basis: judgment against DO-SEC/DO-SAFE-014; Medium uncertainty. Closed by OTA rollback + secure-boot penetration test (PV-CYBER-005).
Detectability/Warning | Secure boot against a hardware root of trust rejects unsigned/forged/rolled-back images (RCR-038); OTA signed + integrity- + rollback-protected + fail-safe A/B (RCR-037); supply-chain provisioning integrity (RCR-054/055); unmonitored windows disclosed; a booted-but-untrusted image is refused, not run.
Initial Risk | AP High (region R — S5 × Low–Med; catastrophic-severity rule forces High regardless)
Risk Acceptability | Unacceptable without control; S5 can never be broadly acceptable (HAL-2200 §6.1); maximum practicable reduction + benefit-risk + disclosure required.
Risk-Control Strategy | 1 (inherent — secure boot rooted in hardware trust) + 2 (protective — signed rollback-protected fail-safe OTA) + 4 (process — supply-chain integrity)
Risk-Control Requirement | RCR-037 — OTA signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; never bricks; unmonitored windows disclosed. + RCR-038 — secure boot against a hardware root of trust; production debug ports locked; tamper-evident enclosure. (+ RCR-054/055 supply-chain part/identity integrity.)
Implementation Reference | DO-SAFE-014 (→ DO-FW-014/015 + DO-CLD-010) + DO-SEC-006 (secure boot) · CMP-035 (firmware), CMP-043 (OTA client), CMP-051 (OTA service), CMP-016 (root of trust)
Verification Reference | PV-CYBER-005 (Evidence-Pending) — OTA rollback + secure-boot penetration test (attempt unsigned/forged/rolled-back image + supply-chain implant; confirm boot refuses untrusted code, OTA rejects/rolls back) + version-skew test
Validation Reference | VAL-CYBER-005 (Evidence-Pending) — secure supply-chain / provisioning audit
Residual Probability | projected Very Low (PVL) after secure boot + signed rollback-protected OTA verified · Residual Severity: S5 (unchanged) · Residual Risk: AP High→A (region A, S5×PVL) — per §6.1 stays region A with EM-ENG-001 + EM-ENG-006 + EM-ENG-008 benefit-risk sign-off + post-market monitoring
New Risks Introduced | Auto-rollback adds a rollback-abuse surface (rollback-protection required, RCR-037); key-management for signing adds a remanence/compromise surface (ties HA-CYBER-008); tracked in HAL-2700.
Overall Residual Contribution | Headline S5 residual — the broadest-reach security→safety hazard; secure boot + signed OTA are load-bearing. Mandatory benefit-risk line in HAL-2900; FTA top event (Stage 11). Ties HA-CLOUD-007 (uncoordinated OTA — operational counterpart) and HA-SW-010 (device OTA).
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Firmware, Manufacturing
Rationale/Evidence | TIR57; AD-7; DO-SAFE-014/DO-SEC-006; RCR-037/038/054/055. Severity per HAL-2000 §5. HW charge-interlock (RCR-001) survives malicious FW — physical charge-safety not defeated.
Assumptions | Secure boot's root of trust is not itself compromisable by the same OTA path; signing keys are protected (HSM/secure provisioning); supply-chain provisioning is integrity-verified.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-006 — Device impersonation/spoof or debug-port exploit → forged records / cloned device
Hazard ID (library) | HAZ-CYBER (identity/interface attack) → information-integrity harm (HA-BND-005)
Product/Config | SCM all SKUs · Lifecycle Phase: Operation / service / manufacturing · System State: Authenticating / debug interface exposed · Operational Mode: Any
Perspective | Beneficiary/patient (via forged records from a spoofed/cloned device); responder; mission
Source | TIR57; HA-BND-005; SYS-134/144; RCR-038/055; ties HA-COMM-008 / HA-CLOUD-003 / HA-SVC-006
Initiating Cause/Event | An adversary impersonates or clones a device (forged/duplicated identity) or exploits an unlocked debug port to extract keys, forge records, or inject data as a trusted device.
Contributing Conditions | Device identity roots trust (CMP-016); a cloned identity submits forged "clean" records for a real trip; an open debug port defeats secure boot, at-rest encryption, and provisioning (fault-management §7); control principle is unique non-cloneable identity + locked debug ports + tamper-evidence.
Foreseeable Sequence of Events | Attacker clones identity or opens a debug port → extracts keys / forges records / submits as a trusted device → cloud accepts forged "clean" records → responder accepts → compromised dose administered; or the cloned identity collides with a real device (ties HA-COMM-008).
Hazard | Loss of device-identity integrity / exposed debug interface.
Hazardous Situation | A spoofed/cloned device (or an exploited one) submitting forged records accepted as authentic.
Sequence to Harm | Forged authentic-looking records → wrong "accept" → potency-lost dose administered → beneficiary harmed.
Harm/Loss | A lot of compromised doses administered (S4); cloned-device record repudiation.
Severity | S4 (Critical) — public-health (HAL-2000 §6/§7); ties HA-COMM-008/HA-CLOUD-003 attribution.
Probability (P1×P2) | P1=2 (Remote — unique cryptographic identity + locked debug ports + secure element; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — forged records from a trusted-looking device reach harm readily) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment against identity/debug controls; Medium uncertainty. Closed by secure-boot + debug-lock penetration test (PV-CYBER-006) + identity-uniqueness check.
Detectability/Warning | Secure boot + locked/disabled production debug ports + tamper-evident enclosure (RCR-038); unique cryptographic identity/serial, duplicate prevented (RCR-055); mutual TLS admits only registered authenticated devices (SRS-CLOUD-001); a duplicate identity is detected at the registry.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control.
Risk-Control Strategy | 1 (inherent — non-cloneable identity + hardware root of trust) + 2 (protective — locked debug, tamper-evidence) + 4 (process — provisioning uniqueness)
Risk-Control Requirement | RCR-038 — secure boot against a hardware root of trust; production debug ports locked/disabled; tamper-evident enclosure. + RCR-055 — each device carries a unique cryptographic identity and serial; duplicate prevented. (+ RCR-051 service-mode lock.)
Implementation Reference | DO-SEC-006 (secure boot/debug-lock) + DO-MFG (provisioning) · CMP-016 (secure element/identity), CMP-044 (security config)
Verification Reference | PV-CYBER-006 (Evidence-Pending) — secure-boot + debug-lock penetration test + identity-uniqueness check (attempt clone/spoof/debug-extract; confirm identity uniqueness, debug lockdown, key protection)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after identity uniqueness + debug lockdown verified · Residual Severity: S4 (unchanged) · Residual Risk: AP Medium (region A, S4×PVL) with benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | After identity + debug controls, small residual. Carried in HAL-2900; ties HA-COMM-008 / HA-CLOUD-003 (binding/attribution), HA-SVC-006 (unlocked service mode), HA-CYBER-001 (forged data).
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Manufacturing
Rationale/Evidence | TIR57; SYS-134/144; RCR-038/055/051. Severity per HAL-2000 §7.
Assumptions | Debug ports are locked before field deployment (RCR-051); each device identity is unique and non-cloneable (secure element).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-007 — Denial-of-service (jam RF / flood cloud) → missing alerts fleet-wide
Hazard ID (library) | HAZ-CYBER (availability attack) → information-integrity harm (HA-BND-003/005/007)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / operation · System State: Attempting delivery under attack · Operational Mode: Armed / ingest
Perspective | Beneficiary/patient (via undelivered alerts); mission/programme; fleet operator
Source | TIR57; HA-BND-005/007; RCR-035/039; ties HA-COMM-001/003 / HA-CLOUD-008
Initiating Cause/Event | An adversary denies service — jams RF locally (blocking a device's egress) or floods the cloud (blocking ingest/notify fleet-wide) — so real excursion alerts are not delivered.
Contributing Conditions | RF jamming is a local physical attack (bounded blast radius) but defeats the live link like the container does (ties HA-COMM-003); cloud flooding is fleet-wide (ties HA-CLOUD-008); store-and-forward keeps the record safe, but the live alert is denied; the malicious counterpart of coverage loss / outage.
Foreseeable Sequence of Events | Attacker jams RF near a shipment (or floods the cloud) → excursion alert cannot egress/be processed → responder not notified in the window → shipment accepted → compromised dose administered; fleet-wide if the cloud is flooded.
Hazard | Denial of the alert-delivery capability (device egress or cloud availability).
Hazardous Situation | A real excursion whose alert is blocked by a DoS attack.
Sequence to Harm | DoS → undelivered/late alert → late/absent segregation → wrong "accept".
Harm/Loss | Compromised doses administered — a lot (local jam) up to fleet-wide (cloud flood) (S4).
Severity | S4 (Critical) — public-health via undelivered alert (HAL-2000 §6/§7); fleet-wide cloud flood approaches S5.
Probability (P1×P2) | P1=2 (Remote — RF jamming is local/opportunistic; cloud flooding must beat DDoS protection + autoscale; Judgment, Evidence-Pending pen-test) × P2=3 (Occasional — store-and-forward + on-receipt read + monitoring mitigate reach-to-harm) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment against availability controls; Medium uncertainty. Closed by DoS/flood penetration test (PV-CYBER-007).
Detectability/Warning | Suppression/blocking raises a detectable event (RCR-035); the un-acked backlog and connectivity diagnostics (SYS-113) make a jammed device visible; cloud DDoS protection, autoscale, and backlog monitoring (RCR-039/SRS-CLOUD-024); store-and-forward + on-receipt edge read are the compensating checks.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control for the fleet-wide case.
Risk-Control Strategy | 2 (protective — DoS resistance, non-silent-suppression, monitoring) + 1 (inherent — store-and-forward keeps the record)
Risk-Control Requirement | RCR-035 (suppression/blocking raises a detectable escalating event) + RCR-039 (cloud resists compromise/flooding: DDoS protection, autoscale, monitoring). (+ RCR-030 store-and-forward; + RCR-031 external-antenna/margin reduces jam susceptibility.)
Implementation Reference | DO-SEC + DO-CLD (resilience) + DO-SAFE-009 (store-and-forward) · CMP-024 (modem), CMP-046/048 (ingest/rules), CMP-052 (API gateway)
Verification Reference | PV-CYBER-007 (Evidence-Pending) — DoS/flood penetration test (RF jam + cloud flood; confirm record preserved, suppression event raised, cloud resists/scales, monitoring alerts)
Validation Reference | VAL-CYBER-007 (Evidence-Pending) — incident-response validation for a sustained DoS
Residual Probability | projected Low (PL) after DoS resistance + non-silent-suppression + monitoring verified · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) with benefit-risk sign-off
New Risks Introduced | DDoS protection adds a legitimate-traffic false-positive risk (availability trade); tracked in HAL-2700.
Overall Residual Contribution | Moderate; the record is always safe (store-and-forward), the live alert is the denied capability. Carried in HAL-2900; ties HA-COMM-001/003 (link loss) and HA-CLOUD-008 (outage) — RCR-035/039 are the malicious-case controls over the same paths.
Status | Evidence-Pending (pen-test) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Cloud, Connectivity
Rationale/Evidence | TIR57; SYS-113/142; RCR-035/039/030. Severity per HAL-2000 §7.
Assumptions | Store-and-forward guarantees zero record loss under any DoS; a jammed/flooded state is detectable (backlog/monitoring) so it is not silent.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-008 — Data/key remanence at disposal → record repudiation / key compromise
Hazard ID (library) | HAZ-CYBER (remanence/end-of-life) → information-integrity harm (HA-BND-005)
Product/Config | SCM all SKUs · Lifecycle Phase: End-of-life / disposal / service (device or backup media) · System State: Powered-down / decommissioned · Operational Mode: Disposal / retire
Perspective | Mission/programme (record repudiation, fleet key compromise); auditor; downstream beneficiary (if a compromised key enables future forgery)
Source | TIR57; HA-BND-005; RCR-052/038; ties HA-SVC-007 / HA-CYBER-001/006; ISO 27001 (media sanitisation)
Initiating Cause/Event | At disposal/decommission, data or keys persist in device storage or backup media (no secure erase), so an adversary recovering a discarded device/media can extract keys or records.
Contributing Conditions | Reusable fleet with EOL retirement (D1); a recovered signing/identity key enables future record forgery or impersonation (feeds HA-CYBER-001/006); recovered records are a privacy/repudiation issue; control principle is designed-in data/key erasure at EOL.
Foreseeable Sequence of Events | Device/media disposed without secure erase → adversary recovers it → extracts device key/identity or records → uses the key to forge future records or impersonate → later wrong "accept", or repudiation of past records.
Hazard | Data/key remanence enabling later forgery/impersonation/repudiation.
Hazardous Situation | A discarded device/media whose keys or records are recoverable.
Sequence to Harm | Recovered key → future forged records (→ eventual wrong "accept"), or record repudiation undermining the evidence base.
Harm/Loss | Record repudiation / reportable data-integrity deviation; a compromised key can seed future S4 forgery (this record scores the direct remanence event) (S3).
Severity | S3 (Serious) — data-integrity/repudiation (HAL-2000 §5); the escalation to forgery/impersonation harm is carried by HA-CYBER-001/006, not double-counted here.
Probability (P1×P2) | P1=2 (Remote — designed EOL erasure + disposal procedure required, RCR-052) × P2=2 (Remote — recovered key/record must then be exploited) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + ISO 27001 sanitisation practice; Medium uncertainty. Closed by disposal-procedure + erasure-verification review (PV-CYBER-008).
Detectability/Warning | Designed data/key erasure at EOL (RCR-052); key protection in the secure element with revocation (RCR-038); disposal procedure supported by design; a revoked key cannot be used even if recovered.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP — reduce as far as practicable via inherent erasure + process.
Risk-Control Strategy | 3 (information/process — disposal instructions) + 1 (inherent — designed erasure + key revocation)
Risk-Control Requirement | RCR-052 — safe end-of-life instructions (battery removal/recycling, data/key erasure) provided and supported by design. (+ RCR-038 key protection/revocation so a recovered key is unusable.)
Implementation Reference | DO-SVC + DO-LBL + DO-SEC (erasure/revocation) · CMP-016 (secure element/keys), CMP-044 (security config)
Verification Reference | PV-CYBER-008 (Evidence-Pending) — disposal-procedure + erasure-verification review (confirm data/key erasure at EOL; confirm key revocation renders recovered keys unusable)
Validation Reference | VAL-CYBER-008 (Evidence-Pending) — disposal-workflow validation with representative depot/retirement operators
Residual Probability | projected Very Low (PVL) after erasure + revocation verified · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small directly; the material escalation (forgery/impersonation) is carried by HA-CYBER-001/006. Carried in HAL-2900; ties HA-SVC-007 (disposal) and HA-CYBER-001/006.
Status | Evidence-Pending — Controlled-pending-verification; enters the hazard chain (TIR57) via key compromise
Owner | Product-Security Engineer (EM-ENG-008) with Service
Rationale/Evidence | TIR57; ISO 27001 sanitisation; RCR-052/038. Severity per HAL-2000 §5.
Assumptions | Keys are revocable so a recovered key can be invalidated; EOL erasure is designed-in, not procedure-only.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-CYBER-009 — Cloud compromise / ransomware → fleet blind + records altered/withheld ★ HEADLINE
Hazard ID (library) | HAZ-CYBER (platform compromise) → information-integrity + availability harm (HA-BND-003/005)
Product/Config | SCM all SKUs (shared platform) · Lifecycle Phase: Any operational · System State: Cloud compromised / encrypted / auth-bypassed · Operational Mode: Platform under attack
Perspective | Beneficiary/patient (fleet-wide blinded/altered records → mass wrong "accept"); mission/programme; auditor
Source | TIR57; HA-BND-005; RCR-039; ties HA-CLOUD-004/006/008 / HA-CYBER-001; ISO 27001
Initiating Cause/Event | An adversary compromises the cloud platform — ransomware encrypting/withholding records, an auth/RBAC bypass, or an insider — so the fleet is blinded and records are altered, withheld, or destroyed at scale.
Contributing Conditions | The cloud holds the fleet's record base, rules, and notification (CMP-046…054); a platform compromise is the broadest-reach cloud attack (hence S5); the malicious counterpart of HA-CLOUD-004/006/008; capture-time integrity (RCR-034) and device-retained copies limit falsification but not availability/withholding; control principle is RBAC/SSO hardening, monitoring, and tested backup/restore.
Foreseeable Sequence of Events | Attacker compromises platform → ransomware encrypts/withholds records + rules/notify disabled → fleet blind, records altered/withheld → excursions across the fleet not flagged/notified → mass compromised doses administered; audit base corrupted → programme-certification risk.
Hazard | Platform compromise blinding the fleet and altering/withholding records.
Hazardous Situation | A fleet whose safety back end is under adversary control.
Sequence to Harm | Cloud compromise → fleet-wide missed/altered excursions → mass wrong "accept" → population-scale beneficiary harm.
Harm/Loss | Fleet-wide missed/altered excursions → mass administration of compromised doses; systemic record loss/falsification → programme-certification withdrawal (S5).
Severity | S5 (Catastrophic) — public-health at population scale + systemic record base compromise (HAL-2000 §5; HA-CYBER-009 anchored S5).
Probability (P1×P2) | P1=1→2 (Improbable–Remote — hardened RBAC/SSO + monitoring + tested backup/restore; Judgment, Evidence-Pending pen-test) × P2=4 (Probable — a compromised platform reaches harm fleet-wide) ⇒ product 4–8 ⇒ Low–Medium band, escalated to High by the S5 catastrophic-severity rule (HAL-2200 §6.1) — Basis: judgment against DO-CLD/DO-SEC; Medium uncertainty. Closed by cloud-security review + penetration test + DR test (PV-CYBER-009).
Detectability/Warning | RBAC/SSO + least-privilege (CMP-053); security monitoring of auth/integrity/backlog/anomaly events (SRS-CLOUD-024, SYS-142); tested backup/restore recovers from ransomware (RCR-039); capture-time integrity + device-retained copies allow re-establishing an authentic record base; audit store (CMP-054) is tamper-evident.
Initial Risk | AP High (region R — S5 × Low–Med; catastrophic-severity rule forces High)
Risk Acceptability | Unacceptable without control; S5 can never be broadly acceptable (HAL-2200 §6.1); maximum practicable reduction + benefit-risk + disclosure + post-market monitoring.
Risk-Control Strategy | 2 (protective — hardened access, monitoring, tested backup/restore) + 4 (process — incident response, DR)
Risk-Control Requirement | RCR-039 — the cloud shall resist compromise (RBAC/SSO, monitoring of auth/integrity/backlog events, tested backup/restore) so a platform incident cannot blind the fleet or corrupt records silently. (+ RCR-034 capture-time integrity so a compromised platform cannot silently falsify; + RCR-024 device-retained recovery source.)
Implementation Reference | DO-CLD (resilience/RBAC/backup) + DO-SEC · CMP-053 (RBAC/SSO), CMP-054 (audit store), CMP-049 (record store), CMP-046/048 (ingest/rules)
Verification Reference | PV-CYBER-009 (Evidence-Pending) — cloud-security review + penetration test (auth bypass, privilege escalation, ransomware simulation) + DR test (confirm monitoring detects, backup restores, capture-time integrity + device copies re-establish an authentic base)
Validation Reference | VAL-CYBER-009 (Evidence-Pending) — incident-response / ransomware tabletop validation
Residual Probability | projected Very Low (PVL) after hardening + monitoring + tested restore verified · Residual Severity: S5 (unchanged) · Residual Risk: AP High→A (region A, S5×PVL) — per §6.1 stays region A with EM-ENG-001 + EM-ENG-006 + EM-ENG-008 benefit-risk sign-off + post-market monitoring
New Risks Introduced | Backup copies extend the data/key footprint (ties HA-CYBER-008); wider monitoring adds an insider-visibility trade; tracked in HAL-2700.
Overall Residual Contribution | Headline S5 residual — the broadest-reach cloud security→safety hazard; RBAC/SSO + monitoring + tested backup/restore are load-bearing, backstopped by capture-time integrity + device copies. Mandatory benefit-risk line in HAL-2900; FTA top event (Stage 11). Ties HA-CLOUD-004/006/008 (accidental counterparts) and HA-CYBER-001 (falsification).
Status | Evidence-Pending (pen-test / DR) — Controlled-pending-verification; enters the hazard chain (TIR57)
Owner | Product-Security Engineer (EM-ENG-008) with Cloud
Rationale/Evidence | TIR57; ISO 27001; SRS-CLOUD-024; RCR-039/034/024. Severity per HAL-2000 §5.
Assumptions | Backups are offline/immutable enough to survive ransomware; capture-time integrity + device-retained copies provide an independent authentic record source; RBAC/SSO enforces least privilege.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Security→safety framing | every record is a TIR57 security→safety chain — the Harm is a public-health outcome (wrong disposition → beneficiary harm), so severity maps to HAL-2000 §6/§7 (S4/S5), probability to HAL-2100 (P1 Judgment/Evidence-Pending against controls; P2 high), acceptability to HAL-2200 (severity-first; S5 catastrophic-severity rule §6.1 forces High for HA-CYBER-005/009).
Controls | RCR-034, RCR-035, RCR-036, RCR-037, RCR-038, RCR-039, RCR-052, RCR-055 (HAL-2400) are the cybersecurity-induced-safety controls; RCR-034 (capture-time integrity), RCR-035 (non-suppressible alarm), RCR-037/038 (signed OTA + secure boot), and RCR-039 (cloud resilience) are the load-bearing security→safety controls.
Implementation | DO-SEC / DO-SEC-006 (integrity, secure boot, debug-lock, erasure), DO-SAFE-008 (record integrity), DO-SAFE-009 (alert), DO-SAFE-014 (OTA), DO-CLD (RBAC/backup) are the Implementation References; CMP-016/035/043/044/046/048/049/051/053/054 are the allocated components.
Cross-module ties | HA-CYBER-001 ↔ HA-COMM-004 / HA-CLOUD-004 / HA-SW-002 (integrity/verdict); HA-CYBER-002 ↔ HA-COMM-002 / HA-CLOUD-001 (alert path — malicious vs accidental); HA-CYBER-003 ↔ HA-CLOUD-002 / HA-USE-003 (thresholds/config); HA-CYBER-004 ↔ HA-COMM-004/008 / HA-CLOUD-003 (freshness/ binding); HA-CYBER-005 ↔ HA-CLOUD-007 / HA-SW-010 (OTA — malicious vs operational); HA-CYBER-006 ↔ HA-COMM-008 / HA-CLOUD-003 / HA-SVC-006 (identity/attribution/service mode); HA-CYBER-007 ↔ HA-COMM-001/003 / HA-CLOUD-008 (DoS vs link/outage); HA-CYBER-008 ↔ HA-SVC-007 (disposal); HA-CYBER-009 ↔ HA-CLOUD-004/006/ 008 (platform).
Downstream | all S4/S5 records enter the hazard chain (TIR57) and become Stage-11 FTA top events (ERFKL), feeding the security threat model (SRS-SEC/EM-ENG-008) and the SW-FMEA; all PV-CYBER- references are Evidence-Pending penetration/fault-injection tests feeding Stages 13–14.
8. Residual-risk / closure notes
Four headline security→safety hazards | HA-CYBER-001 (falsified data), HA-CYBER-002 (suppressed alarm), HA-CYBER-005 (malicious OTA, S5), HA-CYBER-009 (cloud compromise/ransomware, S5). All enter the hazard chain per TIR57 and are Stage-11 FTA top events.
• S5 records (HA-CYBER-005/009) can never be classified broadly acceptable (HAL-2200 §6.1); after secure-boot/signed-OTA (RCR-037/038) and cloud-resilience/backup (RCR-039) evidence they sit in region A with EM-ENG-001 + EM-ENG-006 + EM-ENG-008 benefit-risk sign-off and post-market monitoring.
Load-bearing controls | RCR-034 (capture-time integrity makes the record self-verifying against falsification/replay), RCR-035 (alarm cannot be silently suppressed), RCR-037/038 (firmware trust), and RCR-039 (cloud resilience) are release-critical; their verification is by penetration test / security fault-injection / DR test (PV-CYBER-001…009), all Evidence-Pending.
Release gate (HAL-2200 §8) | no R region cyber risk may remain open at release; every remaining region-A S4/S5 cyber risk carries a benefit-risk statement and residual-risk sign-off; every Evidence-Pending probability has a named closing pen-test.
Overall contribution | the cybersecurity module's residual is dominated by the two S5 records (HA-CYBER-005/009) and the two S4 headliners (HA-CYBER-001/002); these are the principal security→safety inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary. The recurring backstop is that capture-time integrity + device-retained copies + non-suppressible alarms keep the device-side truth recoverable even when the link or cloud is compromised.
New-risk watch (HAL-2700) | signing-key management/remanence (HA-CYBER-008), OTA rollback-abuse surface, DDoS false-positive/availability trade, and monitoring insider-visibility are tracked as control-introduced risks.
Neighbour flags (CLAUDE.md rule 2) | the formal cryptographic threat model, cipher-suite/key-lifecycle detail (SYS-140/143, SRS-SEC) is owned by Product-Security EM-ENG-008 and is Evidence-Pending; Firmware EM-ENG-007 (secure boot, device OTA), Regulatory EM-ENG-006 (data-integrity/GDP, breach reporting), Manufacturing (supply-chain integrity, provisioning uniqueness).
HAL-1400 (HA-DATA) — Data & Decision-Integrity Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1400, record code HA-DATA · Owner: Principal Human-Factors / Data-Decision Safety Engineer, with System-Safety (lead), Cloud/Data Engineering, and Quality/Regulatory (EM-ENG-006). This module analyses the hazards of the information the SCM presents for a human disposition decision — the correctness, completeness, honesty, attribution, timing, and framing of the data on which a nurse/QA officer decides accept / reject / quarantine. It follows the ISO 14971 harm-and-exposure-first method (HAL-000 §4) and the IEC 62366-1 use-related framing. It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM's data can mislead the human disposition decision — by being wrong (biased/incorrect), mis-attributed (another shipment's record), mis-scaled/ mis-timed (units/timestamp/location), incomplete but shown as clean (the automation-bias headline), mis-visualised (a dashboard the eye reads wrong), silently lost (missing evidence), over-trusted when estimated (interpolation/last-known passed off as measured), or wrapped in an unsupported auto-recommendation that manufactures confidence the underlying data does not justify.
The SCM is an information device (HA-CTX-001 / HA-BND-003): it never physically touches a dose, yet its data is the proximate cause of the dominant harm — a wrong disposition → a compromised dose administered (false accept) or a potent dose destroyed → stock-out (false reject) → beneficiary harm (HAL-2000 §6 severity anchor). This module owns the data-integrity and decision-framing layer of that pathway; the firmware that computes detection/verdict is owned by HAL-1000 (HA-SW) and the human act of reading and deciding is owned by HAL-1500 (HA-USE) — this module sits between them, on the artefact the human reads. The governing invariant it defends, jointly with HAL-1000, is "honest data, no false reassurance": the record must never present itself as more complete, more measured, more attributable, or more confident than it truly is.
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope data/decision elements:
• The on-device record and verdict surface (CMP-039 detection output, CMP-040 verdict, CMP-030 OLED presentation) as data an operator reads.
• The cloud record, dashboard, analytics, and export (CMP-049 cloud recompute/record, CMP-050 dashboard, CMP-052 export/API), the device↔trip↔owner association registry (CMP-047), the GNSS/location and gap-fill estimation path (CMP-023), and the time base (CMP-017) as it appears in the record.
• The mobile/BLE app (CMP-055) where it presents record/verdict data to a field user.
In-scope lifecycle phases: armed monitoring (data capture), receipt/disposition read (the decision), audit/recall investigation (record defensibility), and export/hand-off to external QA systems.
Boundaries with adjacent modules (stay in lane, name the neighbours — CLAUDE.md rule 2):
• The firmware defect that produces a wrong detection/verdict/timestamp is owned by HAL-1000 (HA-SW-001/005/007/008/013/014); this module owns the data-as-presented consequence and the cross-check/qualifier controls that keep the artefact honest (HA-DATA-001/003/004).
• The human misreading a correctly-presented verdict/visualisation is owned by HAL-1500 (HA-USE-001); this module owns the case where the data or its framing is itself misleading (HA-DATA-005).
• Malicious falsification/injection/suppression of data is owned by HAL-1300 (HA-CYBER) (RCR-034/035); this module owns the non-malicious integrity-loss twins and cross-references the cyber record.
• Delivery of data across the network/cloud is owned by HAL-1100 (HA-COMM) / HAL-1200 (HA-CLOUD); this module owns the content-integrity of what is delivered (attribution, gaps, estimation flags).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000, wrong-disposition anchor §6/§7); P1×P2 decomposition (HAL-2100); acceptability, severity-first (HAL-2200).
• IEC 62366-1 — usability engineering; use-related risk where the information artefact is what the human acts on — data-integrity and decision-framing errors are use-related risk sources (the record, dashboard, and export are the "user interface" for the disposition decision); summative usability applies to the decision-support surfaces (RCR-040).
• WHO PQS E006 — cold-chain monitor performance and the data/record expectations for a disposition decision (attribution, completeness, defensible record); WHO/UNICEF EVM (record integrity underpins effective vaccine management).
• ALCOA+ data-integrity principles (attributable, legible, contemporaneous, original, accurate, + complete/consistent/enduring/available) — the reference frame for a defensible cold-chain record; cited by designation.
Upstream | UN-SAFE-001/003 (no misleading, no false reassurance); PRD-SAFE-001/003 (do-not-mislead, honest gaps); HA-IUE-001/003/006/012/013 (safety depends on correct/complete information; automation bias; unmonitored-believed-monitored); HA-CTX-001/005 (information device; vulnerable last-mile user); DO-SAFE-004/005/006/008; DO-HF-002/007; HAL-2400 (RCR register, RCR-019/020/021/022/024/025/026/033/ 034/040/041/042); HAL-1000 (HA-SW twins).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000). Data records inherit the SCM severity anchor: a wrong disposition is S4 typical (a lot/group of beneficiaries) and S5 worst (population scale); a record that only degrades defensibility or a limited-quantity call is S3.
• Probability is decomposed P1 × P2 with an explicit basis and uncertainty flag (HAL-2100). The data signature mirrors software: P1 (data defect → hazardous situation) is reduced to Remote/Occasional by integrity controls and cross-checks, but P2 (situation → harm) is Probable because the corrupted/ incomplete/mis-framed data is invisible to the reader and is acted on (automation bias, HA-IUE-012). Many records therefore stay Medium/High and firmly severity-first. Probabilities that turn on human reading of a surface are Evidence-Pending, closed by the summative usability study (VAL-USE-###); never invented as fact (HAL-000 §7).
• No-false-reassurance / honest-data invariant is the governing acceptance test: a data control is only accepted if the record can never (a) present an incomplete/gapped record as a clean pass, (b) present estimated/interpolated/last-known data as measured/authoritative, (c) attach the wrong shipment's data to a decision, or (d) silently lose evidence a decision or audit depends on (RCR-021/022/040/041/042/024/033).
Control hierarchy (HAL-2300) + warnings-challenge rule (HAL-2200 §6.4) | inherent-by-design (make the misleading state unrepresentable — suppress a clean pass over a gap; bind data to its subject by construction) → protective (integrity checks, device↔cloud cross-check, dead-letter/reconciliation) → design-enforced information (qualifiers/flags the surface cannot omit) → process (export format, verification). Passive "read the small print" labelling is not accepted where the data surface can be made honest by construction — especially for the low-training last-mile user (HA-CTX-005).
• Residual risk is re-estimated only after a control is defined, implemented (DO-…), and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (data & decision prompt list)
1. Accuracy: Can the recorded value be biased/incorrect (sensor/probe/calibration) yet look valid, so the decision rests on wrong numbers? (HA-DATA-001)
2. Attribution: Can a decision be made on another shipment's data — wrong device↔trip↔owner binding, duplicated/swapped record? (HA-DATA-002)
3. Scale & context: Can units, timestamp, or location be wrong/ambiguous so an excursion is mis-judged or the record is indefensible at audit? (HA-DATA-003)
4. Completeness: Can a monitoring gap or fault interval be absent from the picture so an incomplete record reads as a clean pass — the automation-bias false accept? (HA-DATA-004)
5. Presentation: Can a chart/dashboard/summary be technically correct but perceptually misleading so the user reads the wrong status? (HA-DATA-005)
6. Persistence: Can record or export data be silently lost so the evidence a decision/audit needs is missing? (HA-DATA-006)
7. Provenance of estimates: Can estimated/interpolated/last-known data (GNSS last-fix, gap-fill) be treated as measured/authoritative? (HA-DATA-007)
8. Decision framing: Can an auto-recommendation/analytic manufacture confidence the underlying data does not support, nudging an unjustified accept? (HA-DATA-008)
6. Hazard Analysis Register (HA-DATA records)
HA-DATA-001 — Wrong data: biased/incorrect reading (sensor / probe / calibration) presented as valid → wrong disposition
Hazard ID (library) | HAZ-DATA (measurement-accuracy failure) → HAZ-FUNC — information-integrity harm via HA-BND-003
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Armed monitoring → receipt/disposition · System State: Sampling/Logging then Verdict · Operational Mode: Armed / logging → disposition read
Perspective | Beneficiary/patient (ultimate harmed party via a wrong "accept"/"reject"); disposition decision-maker; mission/programme
Source | UN-SAFE-001; PRD-SAFE-001; HA-IUE-001/003; SYS-007/024; DO-SAFE-004/005; HAL-2000 §7 (HA-DATA-001 = S5-worst/S4-typical); judgment
Initiating Cause/Event | The value the operator ultimately reads is biased or incorrect — a mis-calibrated or drifted sensor, an out-of-tolerance/wrong probe, a stale/corrupt calibration coefficient, or a plausible-but-wrong sensor read — but it is presented in the record as a valid product-zone measurement.
Contributing Conditions | A reusable device carries calibration across many trips and cold/thermal cycles; an out-of-tolerance value can still fall inside an acceptance window (a small bias never trips a flag); the harmful artefact is invisible — a biased reading produces a record that looks clean and defensible; the last-mile user cannot independently verify the number (HA-CTX-005).
Foreseeable Sequence of Events | Sensor/probe/calibration is biased → conditioning/detection runs on wrong numbers → excursion mis-judged (missed, or falsely declared) → record shows plausible values → human reads "accept" (or "reject") → compromised dose administered (or potent lot destroyed → stock-out).
Hazard | Incorrect measurement data presented as valid.
Hazardous Situation | A disposition decision resting on numbers that do not represent the true product-zone condition.
Sequence to Harm | Biased data → wrong verdict acted on → compromised dose administered → beneficiary not protected → preventable disease (false accept); or potent dose destroyed → stock-out → unvaccinated beneficiary (false reject).
Harm/Loss | A lot/shipment of compromised doses administered, or a valid lot wrongly destroyed (S4); population-scale on a systemic calibration bias across the fleet (S5 worst).
Severity | S4 (Critical) typical / S5 (Catastrophic) worst — governing dimension public-health (HAL-2000 §6 anchor; §7 lists HA-DATA-001 as S5-worst/S4-typical). Record anchor S4, S5 credible worst-case retained (systemic fleet-wide calibration bias).
Probability (P1×P2) | P1=2 (Remote — calibration is integrity-protected, expired/invalid calibration blocks arming, and probe/plausibility faults are detected) × P2=4 (Probable — a plausible-but-wrong value reads as a clean trip and is acted on; the reader cannot detect the bias) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment consistent with HAL-2100 §7 information-harm posture; Medium uncertainty, Evidence-Pending — closed by EOL/interval calibration-coverage + probe-fault-injection tests (PV-DATA-001, ties PV-SW-008/PV-MFG-001).
Detectability/Warning | Not detectable by the user (a biased-but-valid value has no signature). Defences are upstream and design-enforced: two-sided detection integrity (RCR-019), probe/plausibility (absent/open/short/wrong probe never read as valid, RCR-020), calibration integrity + expiry pre-arm gate (RCR-026), and capture-time authenticity/integrity so tampered/injected data is rejected (RCR-034, cyber twin).
Initial Risk | AP High (region R — S4 × PM; the S5 worst-case invokes the catastrophic-severity rule, HAL-2200 §6.1)
Risk Acceptability | Unacceptable without control (R region cannot be accepted; HAL-2200 §6).
Risk-Control Strategy | 2 (protective — calibration integrity/expiry gate, probe plausibility, capture-time integrity) + 1 (inherent — a detected fault/invalid value is never read as valid by construction) + 4 (process — EOL calibration coverage)
Risk-Control Requirement | RCR-019 (two-sided detection, no undetected-excursion path) + RCR-020 (absent/open/short/wrong probe and implausible reads never read as a valid product-zone temperature) + RCR-026 (calibration integrity-protected; expired/invalid calibration blocks arming) + RCR-034 (capture-time authenticity/integrity; falsified/injected data detectable and rejected).
Implementation Reference | DO-SAFE-004 (detection → DO-FW-006 + DO-EE-013) · DO-SAFE-005 (probe/plausibility → DO-FW-007 + DO-EE-013) · DO-SEC (integrity, RCR-034) · CMP-038/039/040 (drivers/detection/verdict), CMP-020 (probe), CMP-049 (cloud recompute)
Verification Reference | PV-DATA-001 (Evidence-Pending) — calibration-coverage/interval verification + probe-fault-injection + integrity-check test that a biased/invalid/injected value is flagged, not presented as valid; ties PV-SW-008 (probe), PV-MFG-001 (EOL calibration)
Validation Reference | VAL-USE-001 (Evidence-Pending) — summative usability confirming a flagged data-quality fault is understood and not acted on as a clean value (ties HA-USE-001); otherwise integrity is design-enforced
Residual Probability | projected Low (PL) after calibration/probe/integrity controls verified (P1 → Improbable; residual is P2-dominated, harm remains human-mediated) · Residual Severity: S4 (unchanged outcome) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off; Evidence-Pending until PV-DATA-001 closes
New Risks Introduced | Conservative flagging of marginal-tolerance values can raise false-reject/quarantine → wastage/stock-out (ties HA-USE-009 / HA-MPE-003) — the safer direction, noted for HAL-2700.
Overall Residual Contribution | Headline data-accuracy residual; FTA top event (Stage 11); ties HA-SW-001 (detection), HA-SW-008 (probe plausibility), HA-SW-014 (units/scale), HA-USE-005/006 (probe placement/channel), HA-CYBER-001 (falsified data). Carried in HAL-2900.
Status | Evidence-Pending (calibration/probe/integrity verification) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer (with Metrology and Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1; WHO PQS E006; ALCOA+ (accurate); RCR-019/020/026/034; DO-SAFE-004/005; HAL-2000 §6/§7; HAL-2100 §7.
Assumptions | A traceable calibration birth-record and validity interval exist (RCR-053); probe presence/open/short HW detect is paired with FW plausibility (SYS-007); the cloud recompute uses the same calibration/profile as the device.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-002 — Wrong subject/trip attribution → decision made on another shipment's data
Hazard ID (library) | HAZ-DATA (attribution/association failure) — HA-BND-003; ALCOA+ (attributable)
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning/association → receipt/disposition · System State: Registry binding / Verdict display · Operational Mode: Disposition read (device, app, or dashboard)
Perspective | Beneficiary/patient (via a decision on the wrong record); disposition decision-maker; mission/programme
Source | PRD-SAFE-001/003; HA-IUE-003; SYS (registry/association); DO-CLD (registry); HAL-2000 §6; judgment
Initiating Cause/Event | The record used for a disposition is bound to the wrong subject — a device↔trip↔owner↔responder mis-association, a duplicated/reused trip ID, a swapped shipment, a wrong-tenant binding, or a dashboard row that shows shipment B's data under shipment A's label — so the decision is made on another shipment's data.
Contributing Conditions | A reusable fleet re-associates devices to new trips each cycle; many near-identical shipments move in parallel; a plausible record for the wrong shipment looks entirely valid; the last-mile user has no way to cross-check identity in the field (HA-CTX-005); cloud multi-tenancy raises wrong-tenant exposure.
Foreseeable Sequence of Events | Device re-armed for trip B but record still bound to trip A (or vice-versa) → operator scans/opens what they believe is this shipment's record → sees a valid record that is not this shipment's → accepts a shipment whose true (unshown) record is a reject → compromised dose administered (or rejects a good shipment on a bad other-record → wastage).
Hazard | Data-to-subject mis-attribution.
Hazardous Situation | A disposition decision made on a record that belongs to a different shipment.
Sequence to Harm | Wrong record → wrong verdict acted on → compromised dose administered (or potent lot destroyed) → beneficiary harm.
Harm/Loss | A lot of compromised doses released, or a valid lot destroyed, on another shipment's record (S4).
Severity | S4 (Critical) — public-health (a decision on the wrong subject is outcome-equivalent to a false accept/reject).
Probability (P1×P2) | P1=2 (Remote — device↔trip↔owner binding is authoritative and validated; duplicate identity/serial prevented) × P2=4 (Probable — a wrong-but-valid record is acted on with high likelihood; the user cannot detect the swap) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by association/routing + identity-uniqueness test (PV-DATA-002, ties PV-CLOUD-003).
Detectability/Warning | Authoritative, correct device↔trip↔owner↔responder binding with wrong-tenant prevention (RCR-033); unique cryptographic identity/serial per device (RCR-055) so a duplicate/swapped ID is caught; the disposition surface shows the shipment/trip identity for confirmation (DO-HF-002/007). Without these, undetectable to the field user.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 1 (inherent — bind data to its subject authoritatively so a mis-attributed decision is unrepresentable) + 2 (protective — identity uniqueness, routing validation)
Risk-Control Requirement | RCR-033 — device↔trip↔owner↔responder binding shall be authoritative and correct; alerts routed to the right responder; wrong-tenant binding prevented. (+ RCR-055 unique cryptographic identity/serial; duplicate prevented.)
Implementation Reference | DO-CLD (registry/association, RCR-033) · DO-MFG (provisioning) + DO-SEC (identity, RCR-055) · CMP-047 (association registry), CMP-050/055 (dashboard/app surfaces), CMP-016 (device identity)
Verification Reference | PV-DATA-002 (Evidence-Pending) — association/routing test (correct binding across re-arm cycles; wrong-tenant rejected) + identity-uniqueness check; ties PV-CLOUD-003
Validation Reference | VAL-USE-008 (Evidence-Pending) — usability that the operator confirms shipment/trip identity before acting (subset of the disposition critical task)
Residual Probability | projected Very Low (PVL) after authoritative binding + identity uniqueness verified (mis-attribution made unrepresentable by construction) · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | Strict binding could block a legitimate re-association in the field (availability/logistics cost, not safety) — service-supported re-binding required.
Overall Residual Contribution | FTA top event (Stage 11); the non-malicious twin of HA-CYBER-003 (wrong-tenant) and HA-CLOUD-003 (routing). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cloud/Platform)
Rationale/Evidence | ISO 14971; ALCOA+ (attributable); WHO PQS E006; RCR-033/055; HAL-2000 §6.
Assumptions | The registry is the single source of truth for device↔trip binding; re-association is a controlled provisioning step (ties HA-USE-003); dashboard/app read identity from the registry, not a cache.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-003 — Wrong units / timestamp / location → excursion mis-judged or record indefensible
Hazard ID (library) | HAZ-DATA (scale/time/geo-context integrity) — HA-BND-003; ALCOA+ (contemporaneous/accurate)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring → disposition / audit · System State: Sampling/Logging then Verdict/Record · Operational Mode: Armed / logging → disposition & audit
Perspective | Disposition decision-maker; mission/programme (record defensibility); indirectly beneficiary
Source | PRD-SAFE-003; HA-IUE-006; SYS-118 (trusted time); DO-CLD; RCR-025/041; judgment
Initiating Cause/Event | The record carries a wrong or ambiguous unit (°C vs °F, mislabelled axis), a wrong/untrusted timestamp (RTC drift/reset, timezone ambiguity), or a wrong/last-known location so an excursion is mis-judged (fell inside vs outside the armed window; wrong severity) or the record is indefensible at audit/recall.
Contributing Conditions | Global deployment mixes unit conventions and timezones; long 90-day trips accumulate clock drift in deep cold; location is often a last-known GNSS fix (ties HA-DATA-007); a single-glance reader assumes the displayed unit/time is correct; a mis-timed excursion can appear to fall in an allowed pre-arm interval.
Foreseeable Sequence of Events | Unit/time/location wrong or ambiguous → excursion window mis-attributed or magnitude mis-read → operator judges a real excursion as in-spec (or an in-spec trip as excursed) → wrong disposition of a limited quantity, or a monitored trip released on a time-indefensible record that later fails audit.
Hazard | Loss of trusted scale/time/geo context in the record.
Hazardous Situation | A disposition or audit decision resting on mis-scaled, mis-timed, or mis-located data.
Sequence to Harm | Mis-judged excursion → wrong disposition of a limited quantity (wastage or a bad release), and/or an indefensible record forcing conservative rejection → wastage; reportable data-integrity deviation at audit.
Harm/Loss | Wrong disposition of a limited quantity of doses; significant avoidable wastage; a monitored trip released on an incomplete/indefensible record; reportable data-integrity deviation (S3).
Severity | S3 (Serious) — governing dimension data-integrity / limited-quantity disposition (HAL-2000 §5 S3; RCR-025 allocated S3). Escalates toward S4 only if a unit/time error directly causes a lot-scale missed excursion — bounded by the two-sided detection integrity (RCR-019) and disclosure-when-untrusted (RCR-025).
Probability (P1×P2) | P1=3 (Occasional — unit/timezone ambiguity and cold-drift are real, expected tensions) × P2=2 (Remote — magnitude usually dominates over exact timing/unit for a pass/fail call; a gross error is often self-evident) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + oscillator-drift physics; Medium uncertainty, Evidence-Pending — closed by clock-drift/discipline + unit/locale-rendering + geo-provenance test (PV-DATA-003, ties PV-SW-005).
Detectability/Warning | Trusted monotonic UTC maintained, drift bounded and disclosed when untrusted (RCR-025); estimated/last-known location flagged as estimated, never as measured/authoritative (RCR-041); unit and locale rendered explicitly and consistently across device/app/dashboard/export (DO-HF-003/007, RCR-042 export preserves time/context).
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce as far as practicable (bounded drift + disclosure + explicit units).
Risk-Control Strategy | 2 (protective — bounded/monotonic time, re-discipline) + 3 (design-enforced information — disclose untrusted time, flag estimated location, explicit units)
Risk-Control Requirement | RCR-025 — trusted monotonic UTC across sleep/brown-out/service; drift bounded and disclosed when untrusted. (+ RCR-041 estimated/last-known data flagged as such, never presented as measured; + RCR-042 export preserves time/context.)
Implementation Reference | DO-EE (RTC + backup) / DO-FW (time discipline) · DO-CLD/DO-HF (units, geo-provenance, RCR-041) · CMP-017 (RTC), CMP-023 (GNSS/estimation), CMP-050 (dashboard), CMP-052 (export)
Verification Reference | PV-DATA-003 (Evidence-Pending) — clock-drift/discipline over cold-endurance profile + untrusted-time disclosure + unit/locale rendering review + geo-provenance (measured vs last-known) test
Validation Reference | VAL-USE-008 (Evidence-Pending) — usability that units/time/location are read correctly (subset of the disposition/record-comprehension critical task)
Residual Probability | projected Low (PL) after bounded-drift/disclosure/explicit-units verified · Residual Severity: S3 · Residual Risk: AP Medium (region A, S3×PL) — ALARP; residual accepted with disclosure
New Risks Introduced | Frequent time re-discipline uses radio energy (minor budget cost, ties HA-ENRG-006) — not a safety hazard.
Overall Residual Contribution | Moderate; the software time-integrity twin is HA-SW-005; ties HA-DATA-007 (estimation) and HA-USE-008 (labelling/language). Not an FTA top event by itself (S3), but a contributor to record-defensibility (HA-MPE record-integrity).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Embedded/Software Safety and Cloud)
Rationale/Evidence | ISO 14971; ALCOA+ (contemporaneous/accurate); WHO PQS E006; RCR-025/041/042; HAL-2000 §5.
Assumptions | A backup time source survives brown-out (CMP-017); connectivity windows allow re-discipline within the drift bound; export consumers honour the disclosed units/provenance.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-004 — Incomplete context: monitoring gap not shown → false reassurance → false accept ★ TOP HAZARD (headline · automation bias)
Hazard ID (library) | HAZ-DATA (completeness/false-reassurance failure) — HA-BND-003; IEC 62366-1 (automation bias)
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition · System State: Verdict / record presentation (device, app, dashboard) · Operational Mode: Disposition read
Perspective | Beneficiary/patient (via a false accept on an incomplete record); disposition decision-maker; mission/programme
Source | UN-SAFE-003; PRD-SAFE-003; HA-IUE-012/013 (automation bias; unmonitored-believed-monitored); DO-SAFE-006; HAL-2000 §7 (HA-DATA-004 = S4); judgment
Initiating Cause/Event | A trip contains a monitoring gap, outage, sensor-fault interval, or lapse (from any cause — dead-in-cold, brown-out, probe fault, buffered outage), but the record/verdict surface does not make the gap visible — it presents an unqualified clean pass over an incomplete record.
Contributing Conditions | Automation bias is real (HA-IUE-012): a busy, low-training user over-trusts a "clean" record and does not seek independent shipping indicators; a gap is an absence of data, which is perceptually easy to omit from a summary; an "accept" is the default a rushed operator acts on; device and cloud must both refuse to show a clean pass over a gap (SYS-151); the reader has no way to know what was not recorded.
Foreseeable Sequence of Events | Gap/outage/fault interval occurs → surface omits the gap (shows only recorded samples as if the whole trip was covered) → operator sees a clean "accept" → releases a lot whose unmonitored interval may hide a real excursion → compromised dose administered.
Hazard | Presentation of an incomplete record as complete (false reassurance).
Hazardous Situation | A lot presented as unconditionally acceptable when its monitoring record is incomplete.
Sequence to Harm | Hidden gap → false accept → compromised/unverifiable dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot of compromised/unverifiable doses released and administered on an incomplete record (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §7, "false accept on incomplete/faulted record = S4").
Probability (P1×P2) | P1=2 (Remote — the verdict suppresses a false accept on any gap and the incomplete-monitoring qualifier is design-enforced and always shown) × P2=4 (Probable — a clean-looking accept is acted on with high likelihood; automation bias amplifies this for the low-training user) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment consistent with HAL-2100 §7 (information-harm) and HA-IUE-012; Medium uncertainty, Evidence-Pending — closed by gap-exposure + verdict-suppression test (PV-DATA-004, ties PV-SW-002) and the summative usability decision-support study (VAL-USE-004).
Detectability/Warning | Design-enforced: any gap/outage/lapse made visible in the record and on-device status — an incomplete record never shows an unqualified clean pass (RCR-022); the verdict suppresses a false accept on any gap/fault, device and cloud consistent (RCR-021); the disposition surface shows the incomplete-monitoring qualifier and never presents a gap as a clean pass (RCR-040). Availability self-check flags a device not monitoring (RCR-027).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; the honest-gap / no-false-reassurance invariant is non-defeatable (HAL-2400 §4).
Risk-Control Strategy | 1 (inherent — a clean pass over a gap is made unrepresentable; verdict suppresses false accept by construction) + 3 (design-enforced information — the incomplete-monitoring qualifier is always shown, survives the warnings-challenge rule)
Risk-Control Requirement | RCR-040 — the disposition surface shall show the incomplete-monitoring qualifier and never present a gap/fault as a clean pass. (+ RCR-021 verdict suppresses a false accept on any gap/fault, device=cloud; + RCR-022 any gap made visible in record and on-device status.)
Implementation Reference | DO-SAFE-006 (verdict integrity → DO-FW-008 + DO-CLD-006) + DO-HF-002 (verdict/qualifier presentation) · CMP-040 (verdict), CMP-050 (dashboard), CMP-030 (OLED), CMP-055 (app)
Verification Reference | PV-DATA-004 (Evidence-Pending) — gap-exposure test (inject gapped/faulted/incomplete records; confirm the surface never shows an unqualified clean pass and verdict never returns an unqualified accept; device=cloud); ties PV-SW-002
Validation Reference | VAL-USE-004 (Evidence-Pending) — summative usability critical task (IEC 62366-1): an incomplete record is never read as a clean pass by the low-training user; the qualifier changes the decision (ties HA-USE-004/007)
Residual Probability | projected Very Low (PVL) after suppression + always-shown qualifier + device/cloud cross-check verified (false-accept-over-gap made unrepresentable) · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off
New Risks Introduced | Conservative gap-marking converts marginal gaps into reject/quarantine → wastage/stock-out pressure (ties HA-USE-009) — accepted as the safer direction, noted for HAL-2700.
Overall Residual Contribution | The headline data residual (automation bias). Carried in HAL-2900 with benefit-risk; FTA top event (Stage 11). The data-surface twin of HA-SW-002 (verdict integrity) and HA-SW-013 (silent gap); ties HA-USE-004 (void trip believed monitored) and HA-USE-007 (override).
Status | Evidence-Pending (gap-exposure + summative usability) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer (with Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1 (automation bias, summative usability); UN-SAFE-003; PRD-SAFE-003; HA-IUE-012/013; RCR-040/021/022; DO-SAFE-006; HAL-2000 §7.
Assumptions | Gap/fault flags are correctly written into the record (storage integrity HA-DATA-006 / RCR-024); device and cloud recompute from the same quality-annotated record; the qualifier is legible/language-independent for the last-mile user (ties RCR-043).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-005 — Misleading visualization / dashboard → user misreads status → wrong disposition
Hazard ID (library) | HAZ-DATA (presentation/perception integrity) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition; fleet monitoring · System State: Dashboard/record visualisation · Operational Mode: Disposition read / responder triage
Perspective | Disposition decision-maker / fleet responder (acts on a mis-perceived chart); beneficiary; mission
Source | PRD-SAFE-001/003; HA-IUE-010/012; DO-HF-007; RCR-040/043; judgment
Initiating Cause/Event | A visualisation is technically correct but perceptually misleading — a truncated/auto-scaled y-axis that hides an excursion, a colour scheme that reads "green/OK" over a marginal trip, an aggregated "all-clear" summary that averages away a spike, a legend/threshold line omitted, or an alert count that buries the one real excursion among informational events.
Contributing Conditions | The dashboard is read under time pressure by a responder and, in the app, by a low-training field user in low light (HA-IUE-010); charts default to auto-scaling that can hide magnitude; colour-only coding fails for colour-vision and low-light reading; a compelling summary invites automation bias (HA-IUE-012); the underlying data may be correct while the rendering misleads.
Foreseeable Sequence of Events | Correct data rendered with a misleading scale/colour/summary → user perceives "OK" (or misjudges severity) → releases/accepts a lot that a faithful rendering would flag → compromised dose administered (or a good lot needlessly rejected).
Hazard | Perceptually misleading presentation of correct data.
Hazardous Situation | A disposition/triage decision made on a chart or summary the user reads as the opposite of its true status.
Sequence to Harm | Misread status → wrong verdict acted on → compromised dose administered (S4), or avoidable wastage (S3).
Harm/Loss | A lot released on a misread dashboard (S4 worst) / limited-quantity misjudgement or wastage (S3 typical).
Severity | S4 (Critical) worst (a mis-perceived "all-clear" over a real excursion = a false accept) / S3 (Serious) typical (limited-quantity misjudgement/wastage). Record anchor S3/S4, governing dimension public-health.
Probability (P1×P2) | P1=2 (Remote — decision-support visualisation is designed to HF standards, no misleading "all-clear," no colour-only coding, faithful scaling) × P2=3 (Occasional — a misleading rendering is sometimes but not always acted on; a faithful design and device↔cloud verdict provide catches) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by decision-support visualisation review + summative usability (PV-DATA-005 / VAL-USE-005).
Detectability/Warning | The disposition surface never presents a gap/fault as a clean pass and shows the incomplete-monitoring qualifier (RCR-040); the verdict is legible and language-independent, not colour-only (RCR-043); dashboard HF requires overdue/stray flagged and no misleading "all clear" (DO-HF-007); device↔cloud verdict consistency (RCR-021) backstops the authoritative call regardless of chart.
Initial Risk | AP Medium/High (region A at S3×PM; region R at the S4 worst-case) — carried at the S4 worst-case: AP High (region R)
Risk Acceptability | Unacceptable at the S4 worst-case without control; ALARP at S3.
Risk-Control Strategy | 3 (design-enforced information — faithful, unambiguous visualisation; qualifier always shown) + 1 (inherent — the authoritative verdict is a legible non-colour-only artefact independent of the chart)
Risk-Control Requirement | RCR-040 — disposition surface shows the incomplete-monitoring qualifier and never presents a gap/fault as a clean pass. (+ RCR-043 verdict legible and language-independent, not colour-only.)
Implementation Reference | DO-HF-007 (dashboard reporting HF — no misleading all-clear) + DO-HF-002/003 (verdict/iconography) + DO-SAFE-006 (verdict integrity) · CMP-050 (dashboard), CMP-055 (app), CMP-030 (OLED)
Verification Reference | PV-DATA-005 (Evidence-Pending) — visualisation review against HF criteria (faithful scaling, non-colour-only, no misleading all-clear, legend/threshold present) + record-comprehension test
Validation Reference | VAL-USE-005 (Evidence-Pending) — summative usability that responders/field users correctly read status from the dashboard/app under time pressure and low light (IEC 62366-1)
Residual Probability | projected Low (PL) after HF-faithful visualisation + legible verdict verified · Residual Severity: S3/S4 · Residual Risk: AP Medium (region A, S3/S4×PL) — ALARP with benefit-risk sign-off for the S4 contribution
New Risks Introduced | Over-conservative "flag everything" visualisation can add to alert load → fatigue (ties HA-USE-002) — balanced by consolidation (RCR-029/047).
Overall Residual Contribution | Moderate–high; ties HA-USE-001 (misread verdict) and HA-DATA-004 (incomplete context); an S4-worst contributor to the information-harm pathway. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cloud/Platform)
Rationale/Evidence | ISO 14971; IEC 62366-1; DO-HF-007; RCR-040/043; HAL-2000 §6/§7.
Assumptions | The dashboard/app renders from the authoritative verdict/record, not a stale cache; HF visualisation criteria are defined in DO-HF-007 and testable.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-006 — Silent data loss (record / export) → missing evidence → wrong disposition
Hazard ID (library) | HAZ-DATA (persistence/availability integrity) — HA-BND-003; ALCOA+ (enduring/available/complete)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring → receipt/disposition → export/audit · System State: Storage/commit; ingest; export · Operational Mode: Logging / export / audit
Perspective | Disposition decision-maker; mission/programme (evidence base); indirectly beneficiary
Source | PRD-SAFE-003/010; HA-IUE-006; SYS-168 (power-fail-safe); DO-SAFE-008; RCR-024/042; judgment
Initiating Cause/Event | Record data is silently lost — a torn/lost write on power loss or brown-out, an overwrite of unsent buffered data, a cloud ingest drop, or an export/API that drops rows or attribution/gap/tamper metadata — so the evidence a disposition or audit needs is missing without notice.
Contributing Conditions | The device sees power loss/brown-out in deep cold; store-and-forward buffers must never be overwritten before delivery; cloud ingest can fail under load; export to external QA systems is a lossy hand-off if the format is not integrity-preserving; a missing record can be mistaken for "no excursion" (an absence read as an all-clear).
Foreseeable Sequence of Events | Data lost at write/ingest/export with no error surfaced → the record (or export) is incomplete but not flagged → operator/auditor treats the available fragment as the whole → releases a lot whose lost interval hid an excursion (or an audit cannot defend a good lot → forced wastage).
Hazard | Silent loss of record/export data.
Hazardous Situation | A disposition or audit decision made on a record/export that is missing evidence without disclosure.
Sequence to Harm | Missing evidence read as all-clear → false accept → compromised dose administered (S4); or indefensible record → conservative wastage / regulatory finding.
Harm/Loss | A lot released on a record with silently lost data (S4); major data-integrity violation forcing recall/hold (S4 information/regulatory).
Severity | S4 (Critical) — public-health / data-integrity (a silent loss can produce a false accept and is a major GDP/data-integrity violation; HAL-2000 §5/§7).
Probability (P1×P2) | P1=2 (Remote — power-fail-safe atomic writes/brown-out barrier; no overwrite of unsent data; cloud dead-letter/retry; integrity-preserving export) × P2=3 (Occasional — a silent loss can produce a false accept but the committed-local record, honest-gap qualifier, and reconciliation provide catches) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by power-fail-during-write + export-integrity + ingest-reconciliation test (PV-DATA-006, ties PV-SW-013/PV-CLOUD).
Detectability/Warning | Power-fail-safe storage (atomic writes, brown-out barrier, no torn/lost/fabricated record, no overwrite of unsent data, tamper-evident) — RCR-024; a loss becomes an honest gap, never a silent absence (RCR-022); cloud never drops a safety event (dead-letter/retry, end-to-end reconciliation) — RCR-032; export/API preserves attribution/time/gaps/tamper-evidence — RCR-042.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — power-fail-safe storage, dead-letter/retry, reconciliation) + 3 (design-enforced information — a loss surfaces as an honest gap) + 4 (process — integrity-preserving export format)
Risk-Control Requirement | RCR-024 — power-fail-safe storage; no torn/lost/fabricated record; no overwrite of unsent data; tamper-evident. (+ RCR-042 export/API preserves attribution, time, gaps, tamper-evidence; no silent data loss; + RCR-032 cloud dead-letter/reconciliation; + RCR-022 loss surfaces as honest gap.)
Implementation Reference | DO-SAFE-008 (record-integrity/power-fail-safe → DO-FW-010 + DO-EE-002 + DO-SEC-006) + DO-CLD (API/audit, RCR-042/032) · CMP-041/033 (storage/NV), CMP-049 (cloud record), CMP-052 (export)
Verification Reference | PV-DATA-006 (Evidence-Pending) — power-fail-during-write test + export-integrity test (attribution/time/gaps/tamper preserved) + cloud ingest fault-injection/reconciliation; ties PV-SW-013, PV-CLOUD
Validation Reference | N/A (integrity is design-enforced; a surfaced honest gap is validated under HA-DATA-004/VAL-USE-004)
Residual Probability | projected Low (PL) after power-fail-safe + reconciliation + integrity-export verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Conservative "treat any loss as a gap/reject" increases wastage (ties HA-USE-009) — the safer direction, noted for HAL-2700.
Overall Residual Contribution | FTA top event (Stage 11); the data-persistence twin of HA-SW-013 (silent gap) and ties HA-CLOUD-001/002 (drop) and HA-DATA-004 (loss-as-gap). Contributes to record-integrity aggregate (RCR-057). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Embedded/Software Safety and Cloud)
Rationale/Evidence | ISO 14971; ALCOA+ (enduring/available/complete); WHO PQS E006; RCR-024/042/032/022; DO-SAFE-008; HAL-2000 §5/§7.
Assumptions | Buffered unsent data is retained until acknowledged delivery; export consumers preserve the delivered integrity metadata; brown-out barrier is validated over the cold-endurance profile.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-007 — Overconfidence in estimated/interpolated data (GNSS last-known, gap-fill) treated as measured
Hazard ID (library) | HAZ-DATA (provenance/estimation integrity) — HA-BND-003; ALCOA+ (original/accurate)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring → disposition / investigation · System State: Record/visualisation with estimated fields · Operational Mode: Disposition read / investigation
Perspective | Disposition/investigation decision-maker; mission/programme; indirectly beneficiary
Source | PRD-SAFE-003; HA-IUE-012; RCR-041; DO-CLD/DO-HF; judgment
Initiating Cause/Event | Estimated/interpolated/last-known data — a GNSS last-fix location shown as current, a gap-fill/interpolated temperature drawn as a continuous measured line, a modelled value — is presented as measured/authoritative, so the user places measurement-grade confidence in an estimate.
Contributing Conditions | In-container RF/GNSS attenuation means location is often a stale last-fix; gap-fill smooths a chart to look continuous; a continuous line invites the reader to assume continuous measurement; automation bias raises trust in a polished visualisation (HA-IUE-012); an investigator may site an excursion at a last-known (wrong) location.
Foreseeable Sequence of Events | Estimate shown as measured → user/investigator treats it as ground truth → a disposition or a root-cause/recall conclusion is drawn from an estimate that is actually uncertain → wrong disposition of a limited quantity, or a mis-directed investigation/wastage.
Hazard | Estimated data misrepresented as measured.
Hazardous Situation | A decision or investigation resting on an estimate believed to be a measurement.
Sequence to Harm | Over-trusted estimate → wrong/limited-quantity disposition or mis-sited excursion → avoidable wastage / indefensible conclusion; escalates only where an estimate masks a real excursion window.
Harm/Loss | Wrong disposition of a limited quantity; significant avoidable wastage; a conclusion drawn on an estimate (S3).
Severity | S3 (Serious) — data-integrity / limited-quantity (HAL-2000 §5; RCR-041 allocated S3). Bounded from S4 because temperature excursion detection itself runs on measured samples (RCR-019/020), not on gap-fill; estimation affects location/context and chart perception.
Probability (P1×P2) | P1=3 (Occasional — last-known GNSS and gap-fill are common given in-container attenuation) × P2=2 (Remote — an estimate rarely flips a pass/fail temperature verdict; it mainly affects location/context and investigation) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by visualisation/provenance review (PV-DATA-007) and summative usability comprehension.
Detectability/Warning | Estimated/interpolated/last-known data flagged as such, never presented as measured/authoritative (RCR-041) — distinct styling for estimated fields, explicit "last-known" labelling, gap-fill shown as interpolated not measured; export preserves the estimated/measured distinction (RCR-042).
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce as far as practicable via provenance flagging.
Risk-Control Strategy | 3 (design-enforced information — estimated data flagged and visually distinct by construction) + 4 (process — provenance preserved in export)
Risk-Control Requirement | RCR-041 — estimated/interpolated/last-known data (GNSS, gap-fill) shall be flagged as such, never presented as measured/authoritative. (+ RCR-042 export preserves the provenance distinction.)
Implementation Reference | DO-CLD / DO-HF (provenance flagging, visualisation) · CMP-023 (GNSS/estimation), CMP-050 (dashboard), CMP-055 (app)
Verification Reference | PV-DATA-007 (Evidence-Pending) — visualisation/provenance review (estimated vs measured distinct; last-known labelled; gap-fill non-measured styling) + export-provenance check
Validation Reference | VAL-USE-005 (Evidence-Pending) — usability that users distinguish estimated from measured data and do not over-trust an estimate (subset of decision-support study)
Residual Probability | projected Low (PL) after provenance flagging verified · Residual Severity: S3 · Residual Risk: AP Medium (region A, S3×PL) — ALARP; residual accepted with disclosure
New Risks Introduced | Excessive "estimated" flagging can erode trust in genuinely measured data (a boy-who-cried-wolf perception cost) — balanced by clear, proportionate styling.
Overall Residual Contribution | Moderate; ties HA-DATA-003 (location/context), HA-COMM-003 (in-container RF/GNSS attenuation), HA-DATA-004 (gap-fill vs honest gap). Not an FTA top event by itself (S3). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cloud/Platform)
Rationale/Evidence | ISO 14971; ALCOA+ (original/accurate); RCR-041/042; HAL-2000 §5; HA-IUE-012.
Assumptions | Gap-fill is never used to replace a missing sample in the excursion verdict (a gap stays a gap, HA-DATA-004); location estimation does not feed the pass/fail temperature verdict.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-DATA-008 — Unsupported auto-recommendation / analytics → unjustified accept
Hazard ID (library) | HAZ-DATA (decision-framing / over-automation) — HA-BND-003; IEC 62366-1 (automation bias)
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition; fleet analytics · System State: Verdict/recommendation presentation · Operational Mode: Disposition read / triage
Perspective | Disposition decision-maker (nudged by a recommendation); beneficiary; mission/programme
Source | PRD-SAFE-001/003; HA-IUE-007/012 (override; automation bias); RCR-021/041; DO-HF-002; judgment
Initiating Cause/Event | An auto-recommendation or analytic ("likely OK," a predicted disposition, a confidence score, a suggested accept) is presented that manufactures confidence the underlying data does not justify — e.g., a recommendation computed over an incomplete/estimated record, or an ML/heuristic output presented without its uncertainty or basis — nudging an unjustified accept.
Contributing Conditions | A recommendation is a strong anchor for a busy, low-training user (automation bias, HA-IUE-012); an "accept" recommendation is the path of least resistance under production/stock-out pressure (ties HA-USE-007); a confidence figure implies rigour even when the input is a gapped/estimated record; the device's core call is a rules-based verdict (RCR-021) — an additional analytic must not undermine it.
Foreseeable Sequence of Events | Analytic suggests "accept / likely OK" over a record that the honest verdict would qualify → user defers to the recommendation → releases a lot the underlying data does not support → compromised dose administered.
Hazard | An unsupported recommendation overriding an honest verdict/record.
Hazardous Situation | A disposition decision anchored on a recommendation not justified by the data quality.
Sequence to Harm | Unjustified accept → compromised dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot released on an unsupported recommendation (S4).
Severity | S4 (Critical) — public-health (an unjustified accept is outcome-equivalent to a false accept).
Probability (P1×P2) | P1=2 (Remote — the authoritative rules-based verdict suppresses a false accept and no recommendation may present higher confidence than the data supports; recommendations are advisory and their basis/uncertainty is shown) × P2=3 (Occasional — a recommendation is influential but the honest verdict/qualifier and recorded-override control provide catches) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by decision-support usability + recommendation-provenance review (PV-DATA-008 / VAL-USE-007).
Detectability/Warning | The authoritative verdict suppresses a false accept on any gap/fault (RCR-021) and is the source of truth over any analytic; estimated/modelled outputs flagged as such with basis/uncertainty, never as measured/authoritative (RCR-041); accepting against an honest qualifier requires a deliberate, recorded override (RCR-044, ties HA-USE-007); the incomplete-monitoring qualifier is always shown (RCR-040).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 1 (inherent — the honest rules-based verdict is authoritative; a recommendation cannot present unjustified confidence by construction) + 3 (design-enforced information — recommendation basis/uncertainty disclosed) + 2 (protective — recorded override to act against the honest verdict)
Risk-Control Requirement | RCR-021 — verdict suppresses a false accept; device=cloud (authoritative over analytics). (+ RCR-041 estimated/modelled outputs flagged, never authoritative; + RCR-044 deliberate recorded override to accept against an honest reject/qualifier.)
Implementation Reference | DO-SAFE-006 (verdict integrity) + DO-HF-002 (verdict presentation) + DO-CLD (analytics provenance) · CMP-040 (verdict), CMP-049/050 (cloud analytics/dashboard), CMP-055 (app)
Verification Reference | PV-DATA-008 (Evidence-Pending) — recommendation-provenance review (no confidence beyond data support; basis/uncertainty shown; verdict authoritative) + override-recording test
Validation Reference | VAL-USE-007 (Evidence-Pending) — summative usability that users do not defer to an unsupported recommendation over an honest reject/qualifier (ties HA-USE-007, IEC 62366-1)
Residual Probability | projected Low (PL) after authoritative-verdict + provenance + recorded-override verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Suppressing helpful analytics entirely could reduce decision support (a usability cost) — balanced by advisory-with-provenance framing rather than removal.
Overall Residual Contribution | FTA top event (Stage 11); the analytics twin of HA-DATA-004 (false reassurance) and ties HA-USE-007 (override/automation bias). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cloud/Platform and Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1 (automation bias); HA-IUE-007/012; RCR-021/041/044; DO-SAFE-006/DO-HF-002; HAL-2000 §6/§7.
Assumptions | Any analytic is advisory only; the rules-based verdict remains the authoritative disposition input; a recommendation cannot compute over a record without its quality flags.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale; wrong-disposition anchor §6/§7 — S4 typical / S5 worst; HA-DATA-001 named as the S5-worst/S4-typical archetype), HAL-2100 (P1×P2, basis taxonomy; the data signature is P1 Remote/Occasional (integrity controls + cross-checks) × P2 Probable/Occasional (invisible, human-mediated, automation-biased harm) ⇒ many records Medium), and HAL-2200 (matrix, regions R/A/B, AP; severity-first + uncertain-probability + warnings-challenge rules). All eight records begin in region R (S4) or A (S3) and none is closable as broadly acceptable while its probability is Evidence-Pending.
Controls (HAL-2400) | RCR-019/020 (detection / probe-plausibility accuracy), RCR-021/022/040 (verdict integrity / honest gap / decision-surface qualifier — the release-critical, non-defeatable set, HAL-2400 §4), RCR-024 (power-fail-safe storage), RCR-025 (time), RCR-026 (calibration integrity/expiry), RCR-033/055 (attribution/identity), RCR-041 (estimation provenance), RCR-042 (export integrity), RCR-032 (cloud no-drop/reconciliation), RCR-034 (capture-time authenticity — cyber twin), RCR-044 (recorded override), RCR-057 (record-integrity aggregate).
Implementation | DO-SAFE-004/005/006/008 (DOL-1200) and DO-HF-002/003/007 (DOL-1300) are the primary Implementation References; each points to concrete DO-FW/DO-CLD/DO-EE/DO-SEC outputs and to CMP-017/020/023/ 030/033/038/039/040/041/047/049/050/052/055.
Cross-module ties | HA-DATA-001 ↔ HA-SW-001/008/014, HA-USE-005/006, HA-CYBER-001; HA-DATA-002 ↔ HA-CLOUD-003, HA-CYBER-003; HA-DATA-003 ↔ HA-SW-005, HA-USE-008; HA-DATA-004 ↔ HA-SW-002/013, HA-USE-004/007 (the automation-bias headline cluster); HA-DATA-005 ↔ HA-USE-001; HA-DATA-006 ↔ HA-SW-013, HA-CLOUD-001/002; HA-DATA-007 ↔ HA-COMM-003, HA-DATA-003/004; HA-DATA-008 ↔ HA-USE-007, HA-DATA-004. The malicious twins (falsified/injected data, wrong-tenant, suppressed evidence) live in HAL-1300 (RCR-034/035/036/039).
Downstream | the S4/S5 records (HA-DATA-001/002/004/005/006/008) become FTA top events in Stage 11 (ERFKL) and feed the deeper SW-FMEA / data-flow FMEA; all PV-DATA-###/VAL-USE-### references are Evidence-Pending and feed Stages 13–14 (verification: calibration/probe/integrity, association/routing, clock-drift, gap-exposure, visualisation review, power-fail/export-integrity, provenance; validation: the summative usability decision-support study, VAL-USE-001/004/005/007/008).
8. Residual-risk / closure notes
Headline data hazard | HA-DATA-004 (incomplete context shown as a clean pass → false accept) is the module headline and the data-layer face of the automation-bias risk that dominates the SCM's information-harm pathway (HA-IUE-012). It is severity-driven while probability is Evidence-Pending and cannot close until the gap-exposure test and the summative usability critical task (VAL-USE-004) report.
• The governing acceptance invariant — honest data, no false reassurance — is realised across RCR-021/022/040 (verdict/gap/qualifier), RCR-024/042/032 (persistence/export/reconciliation), RCR-033/055 (attribution/identity), and RCR-041 (estimation provenance). A data control is not accepted if the record can present itself as more complete, more measured, more attributable, or more confident than it truly is; passive labelling is rejected for the low-training user (warnings-challenge, HAL-2200 §6.4).
S5-capable record | HA-DATA-001 carries a credible S5 worst case (a systemic, fleet-wide calibration bias releasing potency-lost lots at population scale). Per HAL-2200 §6.1 an S5 residual can never be broadly acceptable; after verification it sits in region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off and post-market monitoring.
Release gate (HAL-2200 §8) | no R region data risk may remain open at release; every remaining region-A S4/S5 data risk carries a benefit-risk statement and residual-risk sign-off, and every Evidence-Pending probability has a named closing PV-DATA-### / VAL-USE-### test listed above.
New-risk watch (HAL-2700) | the safety↔availability trade recurs — conservative flagging/gap-marking/ suppression (the safer direction for the beneficiary) increases false-reject/quarantine/wastage and stock-out pressure (ties HA-USE-009, HA-MPE-003); tracked as a control-induced risk balanced by detection/ alarm/consolidation tuning (RCR-019/029/047).
Overall contribution | with HAL-1000 (software) and HAL-1500 (use-related), this module is one of the three domains that carry the SCM's dominant harm. The data layer is where a correct computation can still mislead a human, and where the honest-data invariant is enforced on the artefact the human actually reads. Its residuals are principal inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
HAL-1500 (HA-USE) — Use-Related Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1500, record code HA-USE · Owner: Principal Human-Factors / Data-Decision Safety Engineer, with System-Safety (lead), Human-Factors/HMI, and Quality/Regulatory (EM-ENG-006). This module analyses the use-related hazards of the SCM — the ways a foreseeable human interaction with the device, app, or record produces, masks, or fails to prevent a wrong disposition. It applies the IEC 62366-1 use-related risk process (use specification, use scenarios, use errors, summative usability) on top of the ISO 14971 harm-and- exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible use error — normal-use slip/lapse, mistake, or foreseeable misuse — by which a human interacting with the SCM causes a wrong disposition (a compromised dose administered / a potent dose wasted → beneficiary harm, HAL-2000 §6 anchor). The module is organised around the critical user of this device: the last-mile clinic nurse/health worker — per HA-CTX-005 and HA-IUE-009 the least-trained, most safety-critical operator, often gloved, offline, in deep cold, in low light, with low literacy or reading a non-primary UI language, who reads the device once and acts. Two human phenomena dominate this module and are named explicitly in every relevant record: automation bias (over-trust in a "clean" record, HA-IUE-012) and alarm fatigue (HA-IUE-011 / UN-SAFE-004).
The overriding design rule for this module is the warnings-challenge rule (HAL-2200 §6.4 / HAL-000 §4): for the low-training last-mile user, training, labels, and IFU CANNOT be the primary control. Where a design (inherent or protective) measure is feasible — a legible language-independent verdict, a pre-arm gate that blocks an unready device, a deliberate recorded override — it is required, and a warnings-only mitigation is escalated, not accepted. Every S4 record here is a summative-usability critical task (VAL-USE-###) and a Stage-11 use-related risk / FTA contributor.
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope use interactions:
On-device HMI | OLED verdict/status (CMP-030), button (CMP-031), status LED (CMP-032) — read/actuated gloved, cold, low-light.
Critical tasks (IEC 62366-1 use specification) | provisioning/trip-profile configuration, arming, probe placement, disposition read (accept/reject/quarantine), responding to / acknowledging an alert, verdict override, charging/redeploy, cleaning, recalibration/reuse.
App & cloud touchpoints for field use | mobile/BLE provisioning & disposition app (CMP-055), dashboard read (CMP-050), alert notification/ack (CMP-048) where a human acts.
User set (HA-CTX-005) | dispatch/QA, logistics/3PL, depot/service technician, fleet/IT admin, and the last-mile nurse (vulnerable user) who governs this module.
In-scope lifecycle phases: provisioning/arming, transport handling, receipt/disposition read, alert response, depot recharge/triage, cleaning, recalibration/reuse over service life (HAL-300).
Boundaries with adjacent modules (stay in lane, name the neighbours — CLAUDE.md rule 2):
• The device presenting a wrong/misleading verdict, visualisation, or record is owned by HAL-1000 (HA-SW-007) / HAL-1400 (HA-DATA-004/005); this module owns the human misreading a correctly- presented artefact (HA-USE-001) and the use errors around it.
• The environmental cause of illegibility (cold OLED response, low light) is shared with HAL-900 (HA-ENV-002); this module owns the use consequence (misread → wrong disposition).
• Battery/charge physical hazards are owned by HAL-500 (HA-ENRG); this module owns the use error of mistaking a cold charge-inhibited device for a dead one (HA-USE-009) and the discard/non-deploy consequence.
• Malicious disabling/falsification is owned by HAL-1300 (HA-CYBER) (RCR-036/044); this module owns the non-malicious deliberate bypass / production-pressure twin (HA-USE-012).
• Contamination/cleaning and calibration/service lifecycle are shared with HAL-800 (HA-TFC-005) and HAL-1600/1700 (HA-SVC/HA-MFG); this module owns the user-action face (missed cleaning, reuse-without-recalibration).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000, wrong-disposition anchor §6/§7); P1×P2 (HAL-2100); acceptability, severity-first + warnings-challenge (HAL-2200 §6.4).
• IEC 62366-1 — usability engineering for medical/adjacent devices: use specification, user profiles, use scenarios, identification of hazard-related use scenarios, use errors, and summative usability evaluation of critical tasks. This module is the SCM's use-related risk analysis; every S4 record names a summative-usability critical task (VAL-USE-###) whose use-error probability is Evidence-Pending until the study reports.
• WHO PQS E006 — cold-chain monitor use in the immunisation programme; the last-mile disposition context and the low-training/offline operating environment.
Upstream | UN-SAFE-005/006 (use-error protection; prevent unmonitored trip start); PRD-SAFE-005/006 (use-error protection; mis-arm prevention); HA-IUE-007/008/009/010/011/012/013 (override; misuse; training gaps; accessibility; alarms; automation bias; unmonitored-believed-monitored); HA-CTX-003/005 (foreseeable misuse; vulnerable user); DO-SAFE-013 (state legibility); DO-HF-001…010 (HMI, verdict, states, task-flow); DO-LBL (labelling/IFU); DO-SVC (service); HAL-2400 (RCR register, esp. RCR-043/044/ 045/046/047/048, and RCR-019/020/021/022/026/029/036/049/050).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000). A use error by the vulnerable last-mile user leading to a compromised dose is S4 (HAL-2000 §6 anchor; task severity quick-ref); a false reject → stock-out is S4 (S5 campaign-scale); a use error affecting only defensibility/limited quantity or a contamination path is S2/S3.
• Probability is decomposed P1 × P2 with an explicit basis and uncertainty flag (HAL-2100). The use-error signature is distinctive: P1 (use error → hazardous situation) is often Occasional/Probable because the operator is minimally trained, gloved, offline, and reads once (HA-IUE-009) — it is not reduced to Remote the way verified software is — so records stay severity-first and frequently Medium/ High. All use-error probabilities are Evidence-Pending and closed by the summative usability study (VAL-USE-###); a use-error probability is never invented as fact (HAL-000 §7).
Warnings-challenge rule (the governing method of this module, HAL-2200 §6.4) | for the low-training user, the primary control must be inherent or protective (legible language-independent verdict, pre-arm gate, deliberate recorded override, alarm consolidation), not training/labels/IFU. Any record whose only proposed mitigation is a warning where design prevention is feasible is escalated, not accepted; information-for-safety here is design-enforced (the wrong state is unrepresentable or made unmistakable), which is what lets it survive the rule.
Control hierarchy (HAL-2300) | inherent (make the use error impossible or its consequence unrepresentable) → protective (gates, confirmations, consolidation) → design-enforced information (legible states/qualifiers) → process (training/IFU as supplement only).
• Residual risk is re-estimated only after a control is defined, implemented (DO-…), and verified (PV-…) and the summative usability study validates the critical task; residual probabilities below are projected, Evidence-Pending until VAL-USE-### reports.
5. Categories & guiding questions (use-related prompt list)
1. Read the verdict: Can the last-mile user, gloved/cold/low-light/low-literacy/non-primary-language, misread the accept/reject/quarantine verdict? (HA-USE-001)
2. Respond to the alert: Can a real excursion alert be ignored/muted through alarm fatigue? (HA-USE-002)
3. Configure & arm: Can the wrong trip profile be set (HA-USE-003) or the device not armed yet believed monitored (HA-USE-004)?
4. Place the sensor: Can the probe be left out of the product zone (HA-USE-005) or the wrong probe/ channel used (HA-USE-006)?
5. Override the verdict: Can a reject be overridden under automation bias / production pressure / stock-out fear? (HA-USE-007)
6. Understand the labelling: Can confusing labelling/language cause misconfiguration or misread? (HA-USE-008)
7. Interpret device state: Can a cold charge-inhibited device be mistaken for dead and discarded/ not deployed? (HA-USE-009)
8. Reuse safely: Can cleaning be missed (HA-USE-010) or the device reused without recalibration? (HA-USE-011)
9. Deliberate bypass: Can a user deliberately disable alerts / falsify to pass? (HA-USE-012)
6. Hazard Analysis Register (HA-USE records)
HA-USE-001 — Misread OLED verdict (cold / gloved / low-literacy / low-light / non-primary language) → wrong disposition ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-USE (verdict-comprehension use error) — information-harm via HA-BND-003; IEC 62366-1 critical task
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition · System State: Verdict display (OLED) · Operational Mode: Disposition read
Perspective | Last-mile nurse (vulnerable user) who reads and acts; beneficiary/patient; mission/programme
Source | UN-SAFE-005; PRD-SAFE-005; HA-IUE-009/010/011; HA-CTX-005; DO-SAFE-013; DO-HF-002; HAL-2000 §6 anchor; judgment
Initiating Cause/Event | The last-mile user misreads the accept/reject/quarantine verdict on the OLED — a dim/slow cold display, glare/low light, a gloved hand obscuring, low literacy, or a non-primary UI language makes an "accept" read for a "reject" (or the reverse), or the verdict is not comprehended at all.
Contributing Conditions | The user is minimally trained, reads the device once and acts (HA-IUE-009); deep cold slows OLED response and the reader's cognition; low light and glare degrade contrast (HA-IUE-010); text-only or colour-only coding fails for low literacy / non-primary language / colour-vision; a rushed clinic makes a glance-decision; there is no second reader to catch a misread.
Foreseeable Sequence of Events | Verdict computed correctly and shown → user misreads it (cold/gloved/low-light/language) → acts on the wrong verdict → releases a rejected lot (compromised dose administered) or discards an accepted lot (potent dose wasted → stock-out).
Hazard | Verdict-comprehension use error.
Hazardous Situation | A disposition decision made on a misread verdict.
Sequence to Harm | Misread verdict → wrong action → compromised dose administered (beneficiary not protected) / potent dose wasted (unvaccinated beneficiary).
Harm/Loss | A lot of compromised doses administered, or a valid lot wrongly destroyed, on a misread verdict (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §6 anchor; a use error by the vulnerable last-mile user leading to a compromised dose = S4).
Probability (P1×P2) | P1=3 (Occasional — foreseeable misread by a gloved/cold/low-literacy/non-primary-language user; usability Evidence-Pending) × P2=3 (Occasional — a misread verdict is acted on, though independent shipping indicators/receipt QA are partial catches) ⇒ product 9 ⇒ Medium band (PM) — Basis: HAL-2100 §7 ("use-error: misread OLED verdict — Occasional × Occasional ⇒ Medium"); High uncertainty, Evidence-Pending — closed by the summative usability study critical task (VAL-USE-001) + cold-display test (PV-USE-001).
Detectability/Warning | The user cannot self-detect a misread in the moment. The primary defence is design: the verdict is legible and language-independent (icon + colour + word), readable gloved, in low light, and at −30 °C within the responsiveness budget (RCR-043) — not colour-only, not text-only; cold-legibility display characteristics specified (DO-HF-010). Training/IFU is supplement only (warnings-challenge rule).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (R region; and warnings-only would be escalated).
Risk-Control Strategy | 1 (inherent — legible, language-independent, multi-redundant verdict encoding readable in the actual use environment; not a warnings-only control)
Risk-Control Requirement | RCR-043 — the accept/reject/quarantine verdict shall be legible and language-independent (icon + colour + word), readable gloved, in low light, and at −30 °C within the responsiveness budget.
Implementation Reference | DO-SAFE-013 (state/verdict legibility) → DO-HF-002 (verdict presentation) + DO-HF-010 (cold-display characteristics) + DO-HF-003 (iconography) · CMP-030 (OLED), CMP-032 (LED)
Verification Reference | PV-USE-001 (Evidence-Pending) — cold-display legibility test (−30 °C, response time, contrast, viewing angle, character size) + iconography comprehension test
Validation Reference | VAL-USE-001 (Evidence-Pending) — summative usability critical task (IEC 62366-1): representative low-training users read the verdict correctly gloved, cold, low-light, low-literacy/non-primary-language and cannot be led to the wrong disposition
Residual Probability | projected Low (PL) after legible language-independent verdict verified + summative usability passes (P1 → Remote for the misread; residual P2-dominated) · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off; Evidence-Pending until VAL-USE-001 closes
New Risks Introduced | A very salient reject encoding could contribute to alarm/attention load (ties HA-USE-002) — balanced by consistent, proportionate iconography (DO-HF-003).
Overall Residual Contribution | The headline use residual. Carried in HAL-2900 with benefit-risk; FTA top event (Stage 11) and the primary summative-usability critical task. Ties HA-ENV-002 (cold/low-light legibility), HA-SW-007 (wrong verdict displayed), HA-DATA-005 (misleading visualisation), HA-USE-008 (language/labelling).
Status | Evidence-Pending (cold-display + summative usability) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer
Rationale/Evidence | ISO 14971; IEC 62366-1; WHO PQS E006; UN-SAFE-005; PRD-SAFE-005; HA-IUE-009/010/011; RCR-043; DO-SAFE-013/DO-HF-002/010; HAL-2000 §6.
Assumptions | The computed verdict is correct and consistently rendered (HA-SW-007); the icon/colour/word set is validated for the deployment locales; OLED cold response meets the responsiveness budget (Evidence-Pending, DO-HF-010).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-002 — Ignored / muted alert (alarm fatigue) → real excursion not acted on ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-USE (alarm-response use error / alarm fatigue) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Armed monitoring → excursion response · System State: Alerting / notification · Operational Mode: Alert response
Perspective | Responder / clinic staff (loses time-to-act); beneficiary; mission/programme
Source | UN-SAFE-004; PRD-SAFE-004; HA-IUE-011; DO-SAFE-004; DO-HF-008; HAL-2000 §7; judgment
Initiating Cause/Event | A real excursion alert is ignored, dismissed, or muted because the user is habituated by alarm fatigue — too many prior false or low-value alerts have trained the responder to disregard notifications — so no cold-chain intervention is taken while there is still time.
Contributing Conditions | Over-sensitive alerting or un-consolidated per-device notifications flood the responder (HA-IUE-011 / UN-SAFE-004 conflict C8); a busy clinic/logistics operation defers "yet another" alert; muting/snoozing is an easy action; the harmful case is indistinguishable from the many benign ones at a glance; the window to re-ice/segregate/reject is finite.
Foreseeable Sequence of Events | Genuine excursion → alert fires among many prior false/low-value alerts → responder mutes/ignores it → no intervention within the actionable window → compromised product proceeds to point-of-care → compromised dose administered.
Hazard | Failure to act on a genuine excursion alert (alarm fatigue).
Hazardous Situation | A real, actionable excursion with an alert that is muted or disregarded.
Sequence to Harm | Ignored alert → no timely intervention → compromised dose released/administered → beneficiary not protected.
Harm/Loss | Delayed/missing intervention on a lot → compromised doses administered (S4).
Severity | S4 (Critical) — public-health (delayed/missing action on a real excursion, HAL-2000 §7).
Probability (P1×P2) | P1=4 (Probable — alarm fatigue is a well-documented, expected human response when alerting is not tuned/consolidated; HA-IUE-011) × P2=3 (Occasional — an ignored alert often but not always reaches harm; escalation/ack and receipt QA are partial catches) ⇒ product 12 ⇒ High band (PH) — Basis: judgment + IEC 62366-1 alarm-fatigue literature; High uncertainty, Evidence-Pending — closed by alarm-management validation to false-alarm/missed-excursion targets (PV-USE-002) + summative usability (VAL-USE-002).
Detectability/Warning | Design defences, not exhortation: alerting tuned to validated false-alarm/missed-excursion targets that avoid fatigue, with cloud-side consolidation (RCR-047); excursion alerts sent ahead of bulk, ack-required, and escalate if unacknowledged so an unactioned alert self-escalates to another responder (RCR-029); the alarm path cannot be silently suppressed (RCR-035, cyber twin); salient, prioritized alert presentation (DO-HF-008).
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control; a training-only "pay attention" mitigation is escalated (warnings-challenge).
Risk-Control Strategy | 2 (protective — alarm tuning/consolidation, ack-required, escalate-if-unacknowledged) + 1 (inherent — an unactioned excursion escalates automatically rather than depending on one habituated responder)
Risk-Control Requirement | RCR-047 — alerting tuned to validated false-alarm/missed-excursion targets that avoid alarm fatigue muting a real excursion; consolidation applied. (+ RCR-029 alerts ahead of bulk, ack-required, escalate if unacknowledged.)
Implementation Reference | DO-SAFE-004 (detection/alarm tuning) + DO-CLD-003 (alarm management) + DO-HF-008 (alert presentation) + DO-SAFE-009 (escalation) · CMP-039 (detection), CMP-048 (cloud notify), CMP-042 (comms)
Verification Reference | PV-USE-002 (Evidence-Pending) — alarm-management validation against false-alarm/missed-excursion targets + escalation-on-no-ack test
Validation Reference | VAL-USE-002 (Evidence-Pending) — summative usability that responders act on a genuine excursion alert within the actionable window under representative alert load (IEC 62366-1)
Residual Probability | projected Low (PL) after tuning/consolidation + ack/escalation verified (P1 → Occasional/Remote; escalation removes single-responder dependence) · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Under-sensitive tuning to reduce fatigue risks a missed excursion (the opposite C8 error, ties HA-SW-001/HA-SW-006) — balanced by the validated two-sided targets (RCR-019/047).
Overall Residual Contribution | Headline use residual; FTA top event (Stage 11); the human twin of HA-COMM-001/002 (delivery) and HA-SW-006 (generation). Carried in HAL-2900.
Status | Evidence-Pending (alarm-management + summative usability) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer (with Cloud and Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1 (alarm fatigue); UN-SAFE-004; PRD-SAFE-004; HA-IUE-011; RCR-047/029; DO-SAFE-004/009. Severity per HAL-2000 §7.
Assumptions | The cloud escalation policy is configured and monitored (HA-CLOUD); consolidation does not delay a genuine excursion alert beyond the actionable window (ties RCR-029).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-003 — Wrong trip profile configured at provisioning → wrong thresholds → missed excursion
Hazard ID (library) | HAZ-USE (configuration use error) → HAZ-FUNC — HA-BND-003/004; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning / arming · System State: Provisioning/Arming · Operational Mode: Trip setup
Perspective | Beneficiary/patient; disposition decision-maker; dispatch/QA (provisioner); mission
Source | UN-SAFE-006; PRD-SAFE-006; HA-IUE-008; HA-CTX-003; DO-SAFE-007; DO-HF-006; judgment
Initiating Cause/Event | The provisioner selects/enters the wrong product-stability trip profile (or a mismatched threshold set) in the app — e.g., a heat-tolerant profile for a freeze-sensitive vaccine — so the device monitors against thresholds that do not protect the actual product.
Contributing Conditions | Many vaccines with different stability profiles share the fleet; profile selection is a human step; a plausible-but-wrong profile passes silently unless validated against the product-stability library; provisioning may be done quickly by non-specialist staff; the device faithfully executes an incorrect rule set (this is the use-error face; the software config-integrity twin is HA-SW-003).
Foreseeable Sequence of Events | Wrong profile selected → device arms against wrong thresholds → a real excursion for the actual product does not breach the wrong thresholds → excursion not declared → clean record → false accept → compromised dose administered.
Hazard | Wrong-profile configuration use error.
Hazardous Situation | A device monitoring a product against thresholds that do not protect that product.
Sequence to Harm | Wrong thresholds → missed excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered on a mis-configured trip (S4).
Severity | S4 (Critical) — public-health (same anchor as the missed-excursion archetype).
Probability (P1×P2) | P1=3 (Occasional — a foreseeable provisioning slip/mistake; usability Evidence-Pending) × P2=3 (Occasional — a wrong-profile miss is invisible downstream, though profile validation catches implausible pairings) ⇒ product 9 ⇒ Medium band (PM) — Basis: judgment; High uncertainty, Evidence-Pending — closed by profile-validation + pre-arm gate test (PV-USE-003) + provisioning summative usability (VAL-USE-003).
Detectability/Warning | Provisioning validates the selected profile against the product-stability library and flags an implausible/mismatched profile (RCR-048); the pre-arm readiness gate blocks arming on an invalid/missing profile and the armed device shows the active profile for confirmation (RCR-045); provisioning task-flow includes confirmations (DO-HF-006).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — profile validation) + 1 (inherent — pre-arm gate makes an unvalidated/void arming state unrepresentable/unmistakable)
Risk-Control Requirement | RCR-048 — provisioning shall validate the selected trip profile against the product-stability library and flag an implausible/mismatched profile. (+ RCR-045 pre-arm gate blocks arming unless a valid firmware/profile is present.)
Implementation Reference | DO-APP / DO-CLD (profile validation, RCR-048) + DO-SAFE-007 / DO-FW-032 (pre-arm gate) + DO-HF-006 (provisioning task-flow) · CMP-055 (app/profile), CMP-045 (readiness gate), CMP-040 (verdict)
Verification Reference | PV-USE-003 (Evidence-Pending) — profile-validation (mismatched/implausible rejected) + pre-arm-gate blocks-on-invalid-profile test; ties PV-SW-003
Validation Reference | VAL-USE-003 (Evidence-Pending) — provisioning-workflow summative usability that the correct profile is selectable and confirmable and a wrong one is caught (IEC 62366-1)
Residual Probability | projected Low (PL) after validation + gate + task-flow verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Over-strict validation could block legitimate arming (availability/stock-out pressure) — balanced against safety, noted for HAL-2700.
Overall Residual Contribution | FTA top event (Stage 11); the use-error twin of HA-SW-003; ties HA-USE-004 (arming) and HA-USE-008 (labelling/language). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cloud/App and Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-006; PRD-SAFE-006; HA-IUE-008; RCR-048/045; DO-SAFE-007/DO-HF-006.
Assumptions | A canonical product-stability library exists as the validation reference; cloud and device share the same profile version (ties HA-SW-004).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-004 — Forgot-to-arm / void trip believed monitored → unmonitored trip trusted ★ TOP HAZARD (headline · false reassurance)
Hazard ID (library) | HAZ-USE (arming use error / false reassurance) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning/dispatch → receipt/disposition · System State: Not-armed / void vs believed-armed · Operational Mode: Dispatch and disposition
Perspective | Dispatch/QA and last-mile user (both may trust an unmonitored trip); beneficiary; mission
Source | UN-SAFE-006; PRD-SAFE-006; HA-IUE-013 (unmonitored-believed-monitored); HA-CTX-003; DO-SAFE-007; DO-HF-005/006; judgment
Initiating Cause/Event | The device is dispatched without being armed (forgot-to-arm), or arming failed/voided silently, but the trip is believed monitored — so an unmonitored shipment is trusted and later "accepted" as if it had a complete monitoring record.
Contributing Conditions | Arming is a human step under time pressure; a not-armed device can look similar to an armed one without an unmistakable indication; automation bias means a device present in the shipper is assumed to be working (HA-IUE-012/013); the harmful case is an absence of monitoring, which is perceptually easy to overlook; this is called out as a top hazard (HA-IUE-013).
Foreseeable Sequence of Events | Device placed but not armed (or armed then voided) → no unmistakable "not monitoring" signal acted on → shipment transported unmonitored → at receipt the trip is treated as monitored → an excursion that occurred is invisible → false accept → compromised dose administered.
Hazard | Unmonitored trip believed monitored (false reassurance).
Hazardous Situation | A shipment trusted as monitored when it was never (or only partly) monitored.
Sequence to Harm | Unmonitored trip trusted → hidden excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses released on an unmonitored trip (S4).
Severity | S4 (Critical) — public-health (false reassurance / false accept on an incomplete record, HAL-2000 §7).
Probability (P1×P2) | P1=3 (Occasional — forgot-to-arm is a foreseeable dispatch slip; usability Evidence-Pending) × P2=4 (Probable — an unmonitored trip believed monitored proceeds to a false accept with high likelihood; the absence is invisible) ⇒ product 12 ⇒ High band (PH) — Basis: judgment consistent with HA-IUE-013 top-hazard framing; High uncertainty, Evidence-Pending — closed by pre-arm gate + arming/void test (PV-USE-004) + summative usability (VAL-USE-004).
Detectability/Warning | Design-enforced: the pre-arm readiness gate blocks a credited "monitored" trip unless armed with a valid profile, and an un-armed/void state is unmistakable on device (LED/OLED) and in the record (RCR-045); any monitoring gap/lapse (including a void trip) is made visible and never shows an unqualified clean pass (RCR-022); the verdict suppresses a false accept over an incomplete/absent record (RCR-021); availability self-check flags not-monitoring (RCR-027).
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control; a "remember to arm" instruction alone is escalated (warnings-challenge).
Risk-Control Strategy | 1 (inherent — an un-armed/void state is unmistakable and a void trip can never present as a clean monitored pass) + 2 (protective — pre-arm gate; availability self-check)
Risk-Control Requirement | RCR-045 — the pre-arm readiness gate shall block arming unless SoC/calibration/firmware/profile/probe all pass, and an un-armed/void state shall be unmistakable. (+ RCR-022 any gap/void made visible; an incomplete record never shows an unqualified clean pass.)
Implementation Reference | DO-SAFE-007 / DO-FW-032 (pre-arm gate) + DO-HF-005 (state legibility) + DO-HF-006 (arming task-flow) + DO-SAFE-006 (verdict suppresses false accept) · CMP-045 (readiness gate), CMP-031 (button/arm), CMP-030/032 (OLED/LED state), CMP-040 (verdict)
Verification Reference | PV-USE-004 (Evidence-Pending) — pre-arm gate + arming/void-state test (un-armed/void unmistakable; void trip never shows a clean pass); ties PV-SW-003
Validation Reference | VAL-USE-004 (Evidence-Pending) — summative usability that users cannot dispatch believing a device is armed when it is not, and never read a void/unmonitored trip as a clean pass (IEC 62366-1; ties HA-DATA-004)
Residual Probability | projected Low (PL) after unmistakable-state + gate + honest-void verified (an unmonitored-believed-monitored trip made unrepresentable) · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Strict arming gates could delay dispatch (availability/logistics cost) — balanced against safety, noted for HAL-2700.
Overall Residual Contribution | Headline use residual (false reassurance); FTA top event (Stage 11); the use-error face of HA-DATA-004 (incomplete context) and HA-SW-013 (silent gap); ties HA-SW-003 (profile). Carried in HAL-2900.
Status | Evidence-Pending (pre-arm gate + summative usability) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer (with Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-006; PRD-SAFE-006; HA-IUE-013; RCR-045/022/021; DO-SAFE-006/007/DO-HF-005/006. Severity per HAL-2000 §7.
Assumptions | The pre-arm gate and armed/void indication are unmistakable to the low-training user (validated in VAL-USE-004); the record marks a void trip as unmonitored (ties HA-DATA-004).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-005 — Probe not placed in product zone (left in air / against wall) → unrepresentative temperature → wrong disposition
Hazard ID (library) | HAZ-USE (sensor-placement use error) → HAZ-DATA/FUNC — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs (probe-equipped) · Lifecycle Phase: Provisioning/loading → armed monitoring · System State: Probe placement / Sampling · Operational Mode: Trip setup / armed logging
Perspective | Beneficiary/patient; disposition decision-maker; provisioner/loader; mission
Source | UN-SAFE-005; PRD-SAFE-005; HA-IUE-008; HA-CTX-003; SYS-007; DO-SAFE-005; judgment
Initiating Cause/Event | The external probe (CMP-020) is not placed in the product zone — left in the air gap, against the shipper wall, or coiled outside the payload — so it measures an unrepresentative temperature that does not reflect the doses' condition.
Contributing Conditions | Correct product-zone placement is a manual step; a probe in air/against-a-wall reads plausible in-range values (defeating any obvious error signal); loading may be done by non-specialist staff; once sealed, placement is not observed; a mis-placed-but-plausible reading defeats both detection and verdict (the physical twin of the plausibility hazard HA-SW-008).
Foreseeable Sequence of Events | Probe left in air/against wall → measures air/wall temperature, not product → a real product excursion is not seen by the probe → no excursion declared → clean record → false accept → compromised dose administered.
Hazard | Unrepresentative sensor placement.
Hazardous Situation | A device reporting valid-looking temperatures it is not measuring at the product.
Sequence to Harm | Unrepresentative reading → missed excursion → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered on unrepresentative data (S4).
Severity | S4 (Critical) — public-health (false accept archetype; RCR-020 allocated S4).
Probability (P1×P2) | P1=3 (Occasional — a foreseeable placement slip by loading staff; usability Evidence-Pending) × P2=4 (Probable — an in-range unrepresentative reading reads as a clean trip and is acted on) ⇒ product 12 ⇒ High band (PH) — Basis: judgment; High uncertainty, Evidence-Pending — closed by probe-fault/plausibility + placement-usability test (PV-USE-005) + summative usability (VAL-USE-005).
Detectability/Warning | Firmware plausibility (internal-vs-external divergence, rate-of-change, stuck-value) flags an implausible/unrepresentative read so it is never read as a valid product-zone temperature (RCR-020); the pre-arm gate requires probe-present and a plausible baseline (RCR-045); placement is supported by physical design and instruction (DO-HF-006/009) — but design plausibility, not instruction alone, is the primary control (warnings-challenge).
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — plausibility detection of unrepresentative placement) + 1 (inherent — an implausible read is never accepted as valid; pre-arm probe-present gate) + 4 (process — placement instruction as supplement)
Risk-Control Requirement | RCR-020 — absent/open/short/wrong probe and implausible sensor reads shall be detected and flagged, never read as a valid product-zone temperature. (+ RCR-045 pre-arm gate requires probe-present.)
Implementation Reference | DO-SAFE-005 (probe/plausibility → DO-FW-007 + DO-EE-013) + DO-SAFE-007 (pre-arm gate) + DO-HF-006/009 (placement task/instruction) · CMP-020 (probe), CMP-038 (drivers/plausibility), CMP-045 (readiness gate)
Verification Reference | PV-USE-005 (Evidence-Pending) — plausibility test for unrepresentative placement (air/wall) + pre-arm probe-present test; ties PV-SW-008
Validation Reference | VAL-USE-005 (Evidence-Pending) — summative usability that users place the probe in the product zone and an unrepresentative placement is caught (IEC 62366-1)
Residual Probability | projected Low (PL) after plausibility + probe-present gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Aggressive plausibility voiding could block legitimate arming or flag valid but unusual profiles (availability) — balanced by tuned plausibility bands.
Overall Residual Contribution | FTA top event (Stage 11); the physical/placement twin of HA-SW-008 (plausibility) and ties HA-DATA-001 (wrong data); shares controls with HA-USE-006. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Embedded/Software Safety and Sensing)
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-005; PRD-SAFE-005; HA-IUE-008; SYS-007; RCR-020/045; DO-SAFE-005.
Assumptions | Internal-vs-external divergence provides a plausibility signal for air/wall placement; a plausible baseline is checkable at pre-arm; instruction/design guide correct placement (supplement only).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-006 — Wrong probe / wrong channel used → wrong reading → wrong disposition
Hazard ID (library) | HAZ-USE (sensor-selection use error) → HAZ-DATA — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs (multi-probe/channel configs) · Lifecycle Phase: Provisioning/loading → armed monitoring · System State: Probe/channel selection / Sampling · Operational Mode: Trip setup / armed logging
Perspective | Beneficiary/patient; disposition decision-maker; provisioner; mission
Source | UN-SAFE-005; PRD-SAFE-005; HA-IUE-008; HA-CTX-009; DO-SAFE-005; judgment
Initiating Cause/Event | The wrong probe type is fitted or the wrong channel is assigned/read — e.g., an out-of-range or uncalibrated probe used, an internal-sensor value read as the product-zone value, or a channel mapped to the wrong physical sensor — so the value acted on is wrong for the intended measurement.
Contributing Conditions | Accessories include an external probe + lead (HA-CTX-009); a fleet may have variant probes; channel/probe assignment is a config step; a wrong-but-plausible probe/channel reads in-range; the reusable device may retain a prior probe/channel mapping; a mis-selected channel is invisible to the field user.
Foreseeable Sequence of Events | Wrong probe/channel selected or fitted → reading does not represent the product zone (or is out of calibrated range) → detection/verdict computed on the wrong signal → excursion mis-judged → wrong disposition.
Hazard | Wrong-probe / wrong-channel use error.
Hazardous Situation | A disposition resting on the wrong sensor's reading.
Sequence to Harm | Wrong reading → missed excursion (false accept → compromised dose) or false excursion (false reject → wastage) → beneficiary harm.
Harm/Loss | A lot of compromised doses administered, or a valid lot destroyed (S4 worst); limited-quantity misjudgement (S3 typical).
Severity | S3/S4 — public-health; S4 (Critical) where a wrong channel drives a lot-scale false accept, S3 where it degrades a limited-quantity call. Record anchor S4 conservatively.
Probability (P1×P2) | P1=2 (Remote — wrong/absent/mis-mapped probe and out-of-range reads are detected by presence/plausibility and pre-arm checks; usability Evidence-Pending) × P2=3 (Occasional — a wrong-but-plausible channel reads clean and is acted on) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium–High uncertainty, Evidence-Pending — closed by probe/channel-identity + plausibility test (PV-USE-006) + summative usability.
Detectability/Warning | Absent/wrong/open/short probe and implausible reads detected and never read as valid (RCR-020); probe/channel identity and calibrated range checked; pre-arm gate requires a valid, present, calibrated probe on the expected channel (RCR-045/RCR-026); the active channel/probe shown for confirmation (DO-HF-005/006).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control at the S4 anchor
Risk-Control Strategy | 2 (protective — probe/channel identity + plausibility) + 1 (inherent — an unidentified/out-of-range probe/channel cannot arm/be read as valid)
Risk-Control Requirement | RCR-020 — absent/open/short/wrong probe and implausible reads detected and never read as valid. (+ RCR-045 pre-arm probe-present/valid; + RCR-026 calibration validity.)
Implementation Reference | DO-SAFE-005 (probe/plausibility → DO-FW-007 + DO-EE-013 presence/identity) + DO-SAFE-007 (pre-arm) + DO-HF-005/006 · CMP-005 (probe accessory), CMP-020 (probe), CMP-038 (drivers), CMP-045 (readiness gate)
Verification Reference | PV-USE-006 (Evidence-Pending) — probe/channel-identity + wrong-probe-injection + calibrated-range test
Validation Reference | VAL-USE-006 (Evidence-Pending) — usability that the correct probe/channel is fitted/selected and a wrong one is caught (subset of provisioning study)
Residual Probability | projected Low (PL) after identity + plausibility + pre-arm verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Strict probe-identity gating could block legitimate spare probes (availability) — managed via a controlled approved-probe list.
Overall Residual Contribution | FTA top event (Stage 11) at the S4 anchor; shares controls with HA-USE-005 and ties HA-DATA-001 (wrong data), HA-SW-008 (plausibility). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Sensing and Embedded/Software Safety)
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-005; HA-IUE-008; HA-CTX-009; RCR-020/045/026; DO-SAFE-005.
Assumptions | Probes carry identifiable type/calibration; the expected probe/channel for a profile is defined; an approved-probe list is maintained (ties HA-SVC).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-007 — Override of a "reject" verdict (automation bias / production pressure / stock-out fear) → compromised dose administered ★ TOP HAZARD (headline)
Hazard ID (library) | HAZ-USE (override use error / decision pressure) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition · System State: Verdict override · Operational Mode: Disposition read + override
Perspective | Disposition decision-maker (under pressure); beneficiary; mission/programme
Source | UN-SAFE-001; PRD-SAFE-001; HA-IUE-007/012 (override; automation bias); HA-CTX-003; DO-HF (override); judgment
Initiating Cause/Event | A user overrides a "reject"/"quarantine" verdict and accepts a compromised lot — driven by production pressure, stock-out fear ("we have no other doses"), or automation-bias in reverse (distrust of the device, or over-trust that "it's probably fine") — administering doses the device rejected.
Contributing Conditions | Overriding is a foreseeable, high-consequence use error explicitly identified (HA-IUE-007); stock-out fear and campaign pressure are strong real-world drivers; a busy user may treat the verdict as advisory; accepting is the path of least resistance under scarcity; without a deliberate, recorded barrier an accidental or casual override is easy.
Foreseeable Sequence of Events | Device computes reject/quarantine → user, under pressure, overrides to accept → compromised lot released → compromised dose administered → beneficiary not protected.
Hazard | Override of a protective reject verdict.
Hazardous Situation | A rejected/quarantined lot accepted by human override.
Sequence to Harm | Override reject → compromised dose administered → beneficiary not protected → preventable disease.
Harm/Loss | A lot of compromised doses administered by override (S4).
Severity | S4 (Critical) — public-health (HAL-2000 §7 lists HA-USE-007 as the S4 false-accept-by-override archetype).
Probability (P1×P2) | P1=3 (Occasional — production/stock-out pressure makes override a foreseeable, recurring temptation; usability Evidence-Pending) × P2=4 (Probable — an override is the act of administering the compromised dose; harm follows directly) ⇒ product 12 ⇒ High band (PH) — Basis: judgment consistent with HA-IUE-007; High uncertainty, Evidence-Pending — closed by override-workflow usability + audit test (VAL-USE-007 / PV-USE-007).
Detectability/Warning | An override cannot be accidental: it requires a deliberate, confirmed, recorded action, and accidental acceptance of a compromised dose is prevented (RCR-044); the honest reject/qualifier is legible and language-independent so the user knows they are overriding a safety verdict (RCR-043); the override is audited (accountability). Design barrier + record, not a warning label, is the control (warnings-challenge).
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control; note a legitimate clinical override may exist by policy, but it must be deliberate, recorded, and accountable — never accidental or frictionless.
Risk-Control Strategy | 2 (protective — deliberate, confirmed, recorded override barrier) + 1 (inherent — accidental acceptance of a reject made unrepresentable) + 4 (process — audit/accountability, programme disposition policy)
Risk-Control Requirement | RCR-044 — a verdict-override shall require a deliberate, confirmed, recorded action; accidental acceptance of a compromised dose shall be prevented.
Implementation Reference | DO-HF (override workflow) + DO-CLD (audit) + DO-SAFE-006 (verdict) · CMP-040 (verdict), CMP-055 (app), CMP-031 (button/confirm), CMP-050 (dashboard/audit)
Verification Reference | PV-USE-007 (Evidence-Pending) — override-workflow test (no accidental override; confirmation required; override recorded and audited)
Validation Reference | VAL-USE-007 (Evidence-Pending) — summative usability that a reject cannot be accidentally overridden and a deliberate override is understood as accepting a compromised dose (IEC 62366-1; ties HA-DATA-008)
Residual Probability | projected Low (PL) after deliberate-recorded-override barrier verified (accidental override removed; a deliberate override is a conscious, accountable, policy-bounded act) · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off; residual carries a programme disposition-policy dependency (outside device boundary)
New Risks Introduced | Too much override friction could block a legitimate time-critical clinical override (an availability/clinical-judgment tension) — balanced by a deliberate-but-achievable recorded workflow, not a lockout.
Overall Residual Contribution | Headline use residual; FTA top event (Stage 11); the human twin of HA-DATA-004/008 (false reassurance / unsupported recommendation) and HA-SW-002 (verdict integrity). A residual dependent on out-of-device disposition policy — flagged for the Risk-Management Summary. Carried in HAL-2900.
Status | Evidence-Pending (override usability) — Controlled-pending-verification
Owner | Human-Factors / Data Engineer
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-001; PRD-SAFE-001; HA-IUE-007/012; RCR-044/043; DO-HF. Severity per HAL-2000 §7.
Assumptions | A programme disposition policy governs when a clinical override is permitted; the override record is preserved and auditable (ties HA-DATA-006/RCR-042); the reject verdict is legible so the override is informed.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-008 — Confusing labelling / language → misconfiguration or misread
Hazard ID (library) | HAZ-USE (labelling/language comprehension) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs (all locales) · Lifecycle Phase: Provisioning/arming; receipt/disposition · System State: Any user-reading state · Operational Mode: Configure / read
Perspective | Last-mile user (vulnerable) and provisioner; beneficiary; mission
Source | UN-SAFE-005; PRD-SAFE-005; HA-IUE-010 (accessibility/language); HA-CTX-005; DO-LBL/DO-HF; judgment
Initiating Cause/Event | Confusing labelling, iconography, or a non-primary/absent-translation UI language causes the user to misconfigure (wrong profile/setting) or misread (verdict/state) — a mistranslated verdict word, an ambiguous icon, an untranslated menu, or a label inconsistent across device/app/packaging.
Contributing Conditions | Global deployment across many languages and literacy levels (HA-IUE-010 / HA-CTX-005); text-dependent UI fails for low literacy or non-primary language; inconsistent semantics across device/app/label breed error; the last-mile user reads once and acts; localisation may lag hardware.
Foreseeable Sequence of Events | Ambiguous/untranslated label or icon → user misconfigures or misreads → wrong profile set (→ missed excursion) or verdict misread (→ wrong disposition).
Hazard | Labelling/language comprehension failure.
Hazardous Situation | A configuration or disposition decision made on a misunderstood label/language.
Sequence to Harm | Misconfigure/misread → missed excursion or wrong disposition → compromised dose administered (S4) or limited-quantity error (S3).
Harm/Loss | A lot of compromised doses administered (S4 worst, where a misread verdict/mis-set profile drives a lot-scale error) / limited-quantity misjudgement or wastage (S3 typical).
Severity | S3/S4 — public-health; S4 where it drives a verdict misread or profile mis-set at lot scale (feeds HA-USE-001/003), S3 for a limited-quantity/defensibility effect. Record anchor S4 conservatively for the verdict/profile path.
Probability (P1×P2) | P1=3 (Occasional — foreseeable in low-literacy/non-primary-language deployments; usability Evidence-Pending) × P2=3 (Occasional — a misread/mis-set is acted on, with validation/legibility catches) ⇒ product 9 ⇒ Medium band (PM) — Basis: judgment; High uncertainty, Evidence-Pending — closed by symbol/label comprehension + localisation review (PV-USE-008) + summative usability across locales (VAL-USE-008).
Detectability/Warning | The verdict/state is language-independent (icon + colour + word) so comprehension does not depend on reading text (RCR-043); iconography/symbol semantics are consistent across device/app/label and comprehensible without training (DO-HF-003); profile validation catches an implausible pairing from a mis-set label (RCR-048); labelling/IFU is localised (DO-LBL) — but as a supplement, since the primary control is language-independent design (warnings-challenge).
Initial Risk | AP High (region R — S4 × PM at the conservative anchor)
Risk Acceptability | Unacceptable without control at the S4 path; ALARP at S3.
Risk-Control Strategy | 1 (inherent — language-independent verdict/state; consistent iconography) + 2 (protective — profile validation) + 4 (process — localised labelling/IFU as supplement)
Risk-Control Requirement | RCR-043 — verdict legible and language-independent (icon + colour + word). (+ RCR-048 profile validation catches a mis-set profile from a misread label.)
Implementation Reference | DO-HF-003 (iconography/symbols) + DO-HF-002 (verdict) + DO-LBL (localised labelling/IFU) + DO-APP (app localisation) · CMP-030 (OLED), CMP-032 (LED), CMP-055 (app)
Verification Reference | PV-USE-008 (Evidence-Pending) — symbol/label comprehension test + cross-surface consistency + localisation completeness review
Validation Reference | VAL-USE-008 (Evidence-Pending) — summative usability across representative locales/literacy levels (IEC 62366-1)
Residual Probability | projected Low (PL) after language-independent design + localisation verified · Residual Severity: S3/S4 · Residual Risk: AP Medium (region A, S3/S4×PL) — ALARP with benefit-risk sign-off for the S4 contribution
New Risks Introduced | Icon-only design without any word could itself be ambiguous in some cultures — mitigated by the icon+colour+word triple and comprehension testing.
Overall Residual Contribution | Feeds HA-USE-001 (verdict misread) and HA-USE-003 (profile mis-set) as a contributing cause; ties HA-DATA-003 (units/locale). An S4-path contributor to the information-harm pathway. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Labelling/Regulatory)
Rationale/Evidence | ISO 14971; IEC 62366-1; UN-SAFE-005; HA-IUE-010; HA-CTX-005; RCR-043/048; DO-HF-003/DO-LBL.
Assumptions | Deployment locales/languages are enumerated; icon+colour+word set is validated per locale; localisation is available before deployment in each market.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-009 — Charge-inhibit (cold) mistaken for a dead device → device discarded or not deployed → stock-out / lost monitoring
Hazard ID (library) | HAZ-USE (state-interpretation use error) — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs · Lifecycle Phase: Depot recharge/triage; pre-deployment · System State: Cold charge-inhibited / unready · Operational Mode: Charging / triage
Perspective | Depot/service technician; dispatch; programme (availability/stock-out); indirectly beneficiary
Source | PRD-SAFE-007; HA-IUE-011 (state misinterpretation); HA-CTX-003; DO-SAFE-013; RCR-046; judgment
Initiating Cause/Event | A cold device that has correctly inhibited charging (charge-inhibit-below-0 °C, the D3 battery-safety control, HA-ENRG-001) shows "not charging"; a technician mistakes it for a dead/faulty device and discards it or fails to deploy it — when it simply needs to warm before charging.
Contributing Conditions | Cold-returned devices routinely arrive charge-inhibited; a "won't charge" symptom is easily read as "dead"; without an explicit cause indication the correct action (warm, then charge) is not obvious; discarding a Li-device is also an environmental/disposal issue (ties HA-SVC-007); a not-deployed device removes fleet capacity → stock-out pressure or an unmonitored shipment substituted.
Foreseeable Sequence of Events | Cold device inhibits charge → shows "not charging" without cause → technician judges it dead → discards / sets aside / does not deploy → fleet capacity lost → shipment sent unmonitored (ties HA-USE-004) or stock-out pressure → downstream wrong-disposition or unvaccinated-beneficiary risk.
Hazard | Misinterpretation of a safe charge-inhibit as device death.
Hazardous Situation | A serviceable device removed from service (discarded/not deployed) because its safe state was misread.
Sequence to Harm | Device wrongly retired → lost monitoring capacity → unmonitored trip or stock-out → wrong disposition / unvaccinated beneficiary (S4), plus Li-disposal environmental harm (S3).
Harm/Loss | Reduced fleet leading to an unmonitored trip or stock-out (S4 worst, via the disposition/availability path); needless device loss + improper Li-disposal (S3). Record anchor S3/S4.
Severity | S3/S4 — S4 (Critical) where the lost device leads to an unmonitored trip / stock-out affecting a lot/population (HAL-2000 false-reject/stock-out anchor); S3 for the device-loss/disposal effect alone.
Probability (P1×P2) | P1=3 (Occasional — cold-returned charge-inhibited devices are common and "won't charge = dead" is an intuitive misread; usability Evidence-Pending) × P2=2 (Remote — a discarded device usually causes availability/stock-out pressure rather than a direct wrong dose; downstream logistics/QA mitigate) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium–High uncertainty, Evidence-Pending — closed by state-legibility test (PV-USE-009) + summative usability (VAL-USE-009).
Detectability/Warning | Charge-inhibited/fault/unready states are made legible with cause (LED/OLED) so a cold, charge-inhibited device is not mistaken for a dead/discardable device (RCR-046); state-and-cause presentation (DO-HF-005) tells the technician to warm-then-charge; honest low-battery/dead-in-cold vs charge-inhibited are distinguished (DO-SAFE-013). Legible cause, not a manual note, is the control (warnings-challenge).
Initial Risk | AP Medium/High (region A at S3×PM; region R at the S4 stock-out path) — carried at the S4 path: AP High (region R)
Risk Acceptability | Unacceptable at the S4 path without control; ALARP at S3.
Risk-Control Strategy | 3 (design-enforced information — charge-inhibit state shown with cause, distinct from dead) + 1 (inherent — the state is legible/unmistakable by construction)
Risk-Control Requirement | RCR-046 — charge-inhibited/fault/unready states shall be made legible with cause (LED/OLED); a cold, charge-inhibited device shall not be mistaken for a dead/discardable device.
Implementation Reference | DO-SAFE-013 (state legibility) → DO-HF-005 (device-state presentation) + DO-EE-008 (fuel gauge) · CMP-032 (LED), CMP-030 (OLED)
Verification Reference | PV-USE-009 (Evidence-Pending) — state-legibility test (charge-inhibited-with-cause vs dead distinguishable, cold, low-light)
Validation Reference | VAL-USE-009 (Evidence-Pending) — summative usability that a technician correctly identifies a charge-inhibited device and warms-then-charges rather than discarding (IEC 62366-1)
Residual Probability | projected Low (PL) after legible-cause state verified · Residual Severity: S3/S4 · Residual Risk: AP Medium (region A, S3/S4×PL) — ALARP with benefit-risk sign-off for the S4 contribution
New Risks Introduced | A "warm before charging" instruction handled poorly could tempt bypass of the charge-inhibit (self-heating misuse) — governed by RCR-005/001 (self-heat cannot bypass inhibit).
Overall Residual Contribution | Ties HA-ENRG-001 (charge-inhibit control), HA-USE-004 (unmonitored substitute), HA-SVC-007 (Li-disposal); an S4-path availability/stock-out contributor. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Battery/Power)
Rationale/Evidence | ISO 14971; IEC 62366-1; PRD-SAFE-007; HA-IUE-011; RCR-046; DO-SAFE-013/DO-HF-005; HAL-2000 (false-reject/stock-out anchor).
Assumptions | The device can distinguish and display charge-inhibit-with-cause vs true dead/fault; depot procedure supports warm-then-charge (supplement); LED/OLED legible in depot conditions.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-010 — Cleaning between trips missed → contamination / residue on reused device
Hazard ID (library) | HAZ-USE (reuse-hygiene use error) — HA-BND-005 (contamination); IEC 62366-1
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Return / depot reconditioning → redeploy · System State: Reconditioning · Operational Mode: Between-trip turnaround
Perspective | Depot/service technician; next-trip handlers; environment/hygiene; indirectly beneficiary
Source | HA-CTX-002 (reuse); HA-IUE-015; DO-SVC; RCR-050; judgment; ties HA-TFC-005
Initiating Cause/Event | Cleaning/decontamination between trips is missed or inadequate, leaving contamination or chemical residue on the reused device/probe carried into the next shipper.
Contributing Conditions | High-throughput reuse turnaround; cleaning is a manual reconditioning step; a device may return from varied environments; probe/lead contacts the payload zone; residue from harsh cleaning agents can also degrade materials/legibility over reuse life (ties HA-ENV-003/HA-TFC-005).
Foreseeable Sequence of Events | Device returns → cleaning step skipped/inadequate → contaminant/residue remains → device placed in next shipper near doses → contamination of the payload zone / residue transfer.
Hazard | Contamination/residue on a reused device.
Hazardous Situation | A contaminated device introduced into a vaccine shipper.
Sequence to Harm | Contamination/residue near payload → potential contamination of packaging/handling → hygiene concern; the SCM has no direct patient contact (HA-IUE-005), so the harm is a hygiene/quality path bounded by packaging, not a direct-dose path.
Harm/Loss | Small avoidable wastage / re-inspection / hygiene non-conformance (S2); a contamination event requiring lot quarantine (S3 worst).
Severity | S2/S3 — governing dimension hygiene/quality; S2 (Minor) typical (nuisance re-inspection / correctable), S3 (Serious) worst (a contamination event forcing quarantine of a limited quantity). Bounded below the disposition-harm anchor because there is no direct dose contact (HA-IUE-005).
Probability (P1×P2) | P1=3 (Occasional — a foreseeable turnaround lapse under throughput pressure) × P2=2 (Remote — residue rarely reaches a dose given packaging separation and receipt inspection) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by post-service BIST/reconditioning-procedure verification (PV-USE-010).
Detectability/Warning | Service shall restore seal/config/firmware and a post-service BIST + pre-arm gate verify integrity before redeploy (RCR-050), with cleaning as a defined reconditioning step; materials/seal tolerate approved cleaning agents (ties RCR-011/015); contamination is caught at reconditioning inspection, not at point-of-care.
Initial Risk | AP Medium (region A — S3 × PM at the worst-case; S2×PM = A Med)
Risk Acceptability | ALARP — reduce as far as practicable via a defined reconditioning step and post-service check.
Risk-Control Strategy | 4 (process — defined cleaning/reconditioning step) + 2 (protective — post-service BIST/inspection before redeploy)
Risk-Control Requirement | RCR-050 — service shall restore seal, configuration, and firmware; a post-service BIST + pre-arm gate shall verify integrity before redeploy (cleaning included in the defined reconditioning procedure).
Implementation Reference | DO-SVC (reconditioning/cleaning procedure) + DO-TST (post-service BIST) · CMP-045 (readiness/BIST) · materials/seal DO-MECH-005 (agent compatibility)
Verification Reference | PV-USE-010 (Evidence-Pending) — reconditioning-procedure + post-service BIST verification; cleaning-agent material-compatibility review (ties PV for RCR-011/015)
Validation Reference | N/A (process control; no summative-usability critical task — not a disposition decision)
Residual Probability | projected Low (PL) after defined reconditioning + post-service check verified · Residual Severity: S2/S3 · Residual Risk: AP Low/Medium (region B/A, S2/S3×PL) — ALARP; low residual
New Risks Introduced | Harsh/incorrect cleaning agents could degrade seals/OLED legibility over reuse life (ties HA-ENV-003) — governed by approved-agent list and material-compatibility (RCR-011/015).
Overall Residual Contribution | Low; ties HA-TFC-005 (contamination/materials) and HA-SVC (reconditioning). Not an FTA top event (S2/S3, no direct-dose path). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Service and Materials)
Rationale/Evidence | ISO 14971; IEC 62366-1; HA-IUE-005/015; RCR-050; DO-SVC; HAL-2000 §5 (S2/S3).
Assumptions | No direct patient/dose contact (HA-IUE-005); packaging separates device from doses; an approved cleaning-agent list and reconditioning procedure exist.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-011 — Reused without recalibration → biased reading escapes → wrong disposition
Hazard ID (library) | HAZ-USE (recalibration-lifecycle use error) → HAZ-DATA — HA-BND-003; IEC 62366-1
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Return/reconditioning → redeploy/arming · System State: Reconditioning / pre-arm · Operational Mode: Reuse turnaround / arming
Perspective | Beneficiary/patient; depot technician; disposition decision-maker; mission
Source | HA-CTX-008 (reuse life); HA-IUE-008; DO-FW-007/DO-SVC; RCR-049/026/045; judgment; ties HA-SVC-001/HA-MFG-001
Initiating Cause/Event | The reusable device is redeployed without recalibration past its calibration validity interval, so an accumulated sensor bias escapes and biased readings feed detection/verdict on the next trip.
Contributing Conditions | Multi-trip/multi-year reuse (HA-CTX-008) accumulates drift across cold/thermal cycles; recalibration is a scheduled service step that can be skipped under throughput pressure; a modestly biased device still reads plausibly; the metrology/calibration lifecycle twin is HA-SVC-001/HA-MFG-001.
Foreseeable Sequence of Events | Calibration interval lapses → device redeployed without recalibration → readings biased → excursion mis-judged (missed or false) → wrong disposition.
Hazard | Reuse with expired calibration (bias escape).
Hazardous Situation | A device monitoring with an out-of-tolerance bias it presents as valid.
Sequence to Harm | Biased readings → missed excursion → false accept → compromised dose administered (or false reject → wastage) → beneficiary harm.
Harm/Loss | A lot of compromised doses administered, or a valid lot destroyed, on a mis-calibrated reused device (S4).
Severity | S4 (Critical) — public-health (miscalibration-escape → wrong disposition; HAL-2000 §7).
Probability (P1×P2) | P1=2 (Remote — the pre-arm expiry gate blocks arming an out-of-calibration device; recalibration interval enforced) × P2=4 (Probable — a biased-but-plausible read is invisible and acted on) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by interval-enforcement + calibration-expiry-gate test (PV-USE-011, ties PV-SW-004).
Detectability/Warning | Recalibration interval enforced by the pre-arm expiry gate — an out-of-calibration device cannot be armed (RCR-049); expired/invalid calibration blocks arming (RCR-026); the pre-arm readiness gate checks valid/unexpired calibration among its conditions (RCR-045). A hard gate, not a service reminder, is the control (warnings-challenge).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 1 (inherent — an out-of-calibration device cannot be armed, so bias-escape-on-reuse is unrepresentable) + 4 (process — enforced recalibration lifecycle)
Risk-Control Requirement | RCR-049 — recalibration interval shall be enforced by the pre-arm expiry gate (RCR-026); an out-of-calibration device cannot be armed. (+ RCR-045 pre-arm gate checks valid/unexpired calibration.)
Implementation Reference | DO-FW / DO-SVC (interval enforcement, RCR-049) + DO-FW-007 (calibration) + DO-SAFE-007 (pre-arm gate) · CMP-045 (readiness/expiry gate), CMP-038 (calibration path)
Verification Reference | PV-USE-011 (Evidence-Pending) — calibration-interval enforcement + expiry-gate blocks-arming test; ties PV-SW-004, PV-MFG-001
Validation Reference | N/A (expiry gate is design-enforced; a blocked arm is not a disposition critical task) — service-procedure review only
Residual Probability | projected Low (PL) after expiry gate + interval enforcement verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Aggressive expiry voids serviceable devices (availability/service cost, not safety) — a logistics trade, noted for HAL-2700.
Overall Residual Contribution | FTA top event (Stage 11); the use/lifecycle face of HA-DATA-001 (wrong data) and HA-SW-004/014 (calibration integrity/units); ties HA-SVC-001, HA-MFG-001 (calibration lifecycle). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Metrology and Service)
Rationale/Evidence | ISO 14971; IEC 62366-1; HA-CTX-008; HA-IUE-008; RCR-049/026/045; DO-FW-007/DO-SVC. Severity per HAL-2000 §7.
Assumptions | A calibration validity interval and birth-record baseline exist (RCR-053); the pre-arm gate can read calibration status; service schedules recalibration within interval.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-USE-012 — Deliberate bypass (disable alerts / falsify to pass) → excursion hidden → wrong disposition
Hazard ID (library) | HAZ-USE (deliberate misuse / non-malicious circumvention) — HA-BND-003; IEC 62366-1; ties HAL-1300
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning/monitoring/disposition · System State: Config/alert/verdict · Operational Mode: Any operational mode
Perspective | Programme/beneficiary (harmed by a hidden excursion); QA/regulator (record integrity); mission
Source | HA-CTX-003 (foreseeable misuse); HA-IUE-007 (override misuse); DO-SEC/DO-HF; RCR-044/036; judgment
Initiating Cause/Event | A user deliberately bypasses a control to make a trip "pass" — disabling/muting alerts, altering thresholds to hide an excursion, or overriding/falsifying the verdict — under production/stock-out pressure or to avoid a rejection, hiding a real excursion. (Non-malicious-insider face; the malicious/cyber twin is HAL-1300.)
Contributing Conditions | Production and stock-out pressure create incentives to make a trip pass (ties HA-USE-007); if alerts are silently mutable or thresholds freely editable, circumvention is easy; a falsified "pass" is indistinguishable downstream without tamper-evidence; accountability is absent without auditing.
Foreseeable Sequence of Events | User disables alerts / edits thresholds / overrides verdict to force a pass → a real excursion is hidden → lot released → compromised dose administered; and the record is falsified (data-integrity violation).
Hazard | Deliberate circumvention of a safety control.
Hazardous Situation | A trip with a hidden/suppressed excursion presented as a pass.
Sequence to Harm | Hidden excursion → false accept → compromised dose administered → beneficiary not protected; plus a data-integrity/regulatory violation.
Harm/Loss | A lot of compromised doses administered on a falsified pass (S4); major data-integrity violation (S4 information/regulatory).
Severity | S4 (Critical) — public-health / data-integrity.
Probability (P1×P2) | P1=2 (Remote — controls resist casual circumvention: alarms not silently suppressible, config changes authenticated/authorised/audited, override deliberate and recorded) × P2=4 (Probable — a forced/falsified pass leads to a dose administered) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by suppression/config-tamper/override-audit test (PV-USE-012, ties PV-CYBER RCR-035/036).
Detectability/Warning | The alarm/alert path is not silently suppressible — suppression/blocking itself raises a detectable, escalating event (RCR-035); threshold/config changes are authenticated, authorised, bounded, and audited (RCR-036); a verdict override is deliberate, confirmed, and recorded (RCR-044); capture-time tamper-evidence detects falsification (RCR-034). Design + audit, not policy alone, is the control (warnings-challenge).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 2 (protective — non-suppressible alarms, authenticated/audited config, recorded override, tamper-evidence) + 4 (process — accountability/audit, programme policy)
Risk-Control Requirement | RCR-044 — verdict-override deliberate, confirmed, recorded. (+ RCR-036 config/threshold changes authenticated, authorised, bounded, audited; + RCR-035 alarm path not silently suppressible; + RCR-034 tamper-evidence.)
Implementation Reference | DO-HF (override) + DO-SEC (config auth/audit, non-suppressible alarm, tamper-evidence) + DO-CLD (audit) · CMP-040 (verdict), CMP-048 (alert), CMP-044 (security/crypto), CMP-053/054 (cloud config/audit)
Verification Reference | PV-USE-012 (Evidence-Pending) — alarm-suppression/config-tamper/override-audit test; ties PV-CYBER (RCR-034/035/036)
Validation Reference | VAL-USE-007 (Evidence-Pending) — override-workflow usability confirming a deliberate, recorded override is required (shared with HA-USE-007)
Residual Probability | projected Low (PL) after non-suppressible alarms + audited config + recorded override + tamper-evidence verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; residual carries a programme-policy/accountability dependency
New Risks Introduced | Strict config authorisation could impede legitimate field re-configuration (availability) — balanced by role-based authorised change (RCR-036).
Overall Residual Contribution | FTA top event (Stage 11); the non-malicious twin of HA-CYBER-002/003 (suppressed alarm / unauthorised config) and ties HA-USE-007 (override), HA-DATA-006 (falsified/lost evidence). A residual dependent on out-of-device accountability policy — flagged for the Risk-Management Summary. Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Human-Factors / Data Engineer (with Cybersecurity and Cloud)
Rationale/Evidence | ISO 14971; IEC 62366-1; HA-CTX-003; HA-IUE-007; RCR-044/036/035/034; DO-SEC/DO-HF/DO-CLD. Severity per HAL-2000 §7.
Assumptions | An audit trail is preserved and reviewed; a programme accountability policy governs configuration/override authority; tamper-evidence is cryptographically sound (ties HAL-1300).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale; wrong-disposition anchor §6/§7 — a use error by the vulnerable last-mile user leading to a compromised dose = S4), HAL-2100 (P1×P2; the use-error signature is P1 Occasional/Probable (low-training, gloved, offline, reads-once) × P2 Occasional/Probable (invisible, human-mediated, automation-biased harm) ⇒ many records Medium/High; HAL-2100 §7 anchors the misread-verdict record at Occasional×Occasional⇒Medium), and HAL-2200 (matrix, regions R/A/B, AP; severity-first + warnings-challenge). All twelve records begin in region R (S4) or A (S2/S3) and none is closable while its use-error probability is Evidence-Pending — every S4 record is a summative-usability critical task (VAL-USE-###).
Controls (HAL-2400) | RCR-043 (legible language-independent verdict), RCR-044 (deliberate recorded override), RCR-045 (pre-arm gate / unmistakable void), RCR-046 (charge-inhibit-with-cause legibility), RCR-047 (alarm tuning/consolidation), RCR-048 (profile validation), plus RCR-019/020 (detection/probe plausibility), RCR-021/022 (verdict integrity / honest gap), RCR-026/049 (calibration expiry/interval), RCR-029 (alert escalation), RCR-034/035/036 (tamper/suppression/config — cyber-adjacent), RCR-050 (reconditioning). RCR-043/044/045 are the module's release-critical, non-defeatable / not-warnings-only set (HAL-2400 §4; warnings-challenge, HAL-2200 §6.4).
Implementation | DO-SAFE-013 (state/verdict legibility) and DO-HF-002/003/005/006/010 (verdict, iconography, states, arming task-flow, cold display), with DO-LBL (labelling/IFU) and DO-SVC (reconditioning/calibration), are the primary Implementation References; each points to concrete DO-FW/ DO-CLD/DO-EE outputs and CMP-030/031/032/020/038/040/045/048/050/055.
Cross-module ties | HA-USE-001 ↔ HA-ENV-002 (cold/low-light legibility), HA-SW-007 (wrong verdict displayed), HA-DATA-005 (misleading visualisation); HA-USE-002 ↔ HA-COMM-001/002, HA-SW-006, HA-CLOUD-001; HA-USE-003 ↔ HA-SW-003; HA-USE-004 ↔ HA-DATA-004, HA-SW-013; HA-USE-005/006 ↔ HA-SW-008, HA-DATA-001, HA-ELEC-008; HA-USE-007 ↔ HA-DATA-004/008, HA-SW-002, HA-IUE-007; HA-USE-008 ↔ HA-DATA-003, HA-USE-001/003; HA-USE-009 ↔ HA-ENRG-001, HA-USE-004, HA-SVC-007; HA-USE-010 ↔ HA-TFC-005, HA-SVC; HA-USE-011 ↔ HA-SVC-001, HA-MFG-001, HA-DATA-001; HA-USE-012 ↔ HA-CYBER-002/003, HA-USE-007, HA-DATA-006.
Downstream | the S4 records (HA-USE-001/002/003/004/005/006/007/011/012, and the S4-path HA-USE-008/009) become FTA top events / use-related risk contributors in Stage 11 (ERFKL) and drive the summative usability engineering file; all PV-USE-###/VAL-USE-### references are Evidence-Pending and feed Stages 13–14 (verification: cold-display, alarm-management, profile-validation, pre-arm gate, probe plausibility, override-audit, state-legibility, calibration-expiry, reconditioning; validation: the summative usability study critical tasks VAL-USE-001…012).
8. Residual-risk / closure notes
• Five headline use hazards anchor this module and the SCM's use-related harm pathway: HA-USE-001 (misread verdict), HA-USE-002 (ignored/muted alert — alarm fatigue), HA-USE-004 (void trip believed monitored — false reassurance), HA-USE-007 (override of a reject — automation bias / production pressure), and, as the physical-safety-adjacent availability path, HA-USE-009 (charge-inhibit mistaken for dead). All are severity-driven while use-error probability is Evidence-Pending and none can close until its named summative usability critical task (VAL-USE-###) reports.
The warnings-challenge rule is the governing method here (HAL-2200 §6.4) | for the low-training last-mile user, training/labels/IFU are never the primary control. Every headline is controlled by an inherent or protective measure — legible language-independent verdict (RCR-043), pre-arm gate / unmistakable void (RCR-045), deliberate recorded override (RCR-044), alarm tuning + escalation (RCR-047/029), charge-inhibit- with-cause legibility (RCR-046) — with labelling/IFU only as a supplement. A record proposing a warnings-only mitigation where design prevention is feasible is escalated, not accepted.
• Automation bias and alarm fatigue are treated as design problems, not user failings: the honest-gap qualifier and false-accept suppression (RCR-021/022/040, shared with HAL-1400) defend against over-trust; validated alarm tuning + consolidation + escalation (RCR-047/029) defend against habituation.
Residuals dependent on out-of-device policy | HA-USE-007 (clinical override) and HA-USE-012 (deliberate bypass) retain a residual dependency on programme disposition/accountability policy outside the device boundary; these are flagged for the benefit-risk determination in the Risk-Management Summary (HAL-2900) and cannot be fully closed by device design alone.
Release gate (HAL-2200 §8) | no R region use risk may remain open at release; every remaining region-A S4 use risk carries a benefit-risk statement and residual-risk sign-off, and every Evidence-Pending use-error probability has a named closing PV-USE-### / VAL-USE-### critical task.
New-risk watch (HAL-2700) | the safety↔availability trade recurs — strict gates (pre-arm, profile validation, calibration expiry, override friction) reduce wrong-disposition risk but can delay dispatch or void serviceable devices (availability/stock-out pressure, ties HA-USE-009 / HA-MPE-003); tuning for fatigue vs missed-excursion (C8) is the recurring alarm trade; both are tracked as control-induced risks.
Overall contribution | with HAL-1000 (software) and HAL-1400 (data), this module completes the SCM's dominant harm pathway — the human who reads the artefact and acts. The SCM is an information device operated at its most safety-critical point by its least-trained user; this module's residuals are principal inputs to the overall-residual-risk evaluation (HAL-2900), the summative usability engineering file, and the benefit-risk determination in the Risk-Management Summary.
HAL-1600 (HA-SVC) — Service & Maintenance Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1600, record code HA-SVC · Owner: Principal Service/Maintenance Engineer, with System-Safety (lead), Quality/Regulatory (EM-ENG-006), and the Manufacturing Engineer. This module analyses the hazards introduced or amplified in the reuse loop — return → clean → recharge → recalibrate → reprovision → redeploy — and in depot service, maintenance, and decommission. It follows the ISO 14971 harm-and-exposure-first method (HAL-000 §4); it is not an FMEA. Each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way that servicing, maintaining, recalibrating, reprovisioning, or decommissioning a reusable SCM can leave it in a state that later causes harm — dominantly a wrong vaccine-disposition decision, secondarily a technician injury or an environmental release. The reuse loop is unique to the reusable (D1) SCM and is a hazard amplifier (HAL-300 §5): a service escape does not harm one trip, it is baked into the device and re-exposed on every subsequent trip until caught by the pre-arm readiness gate. The module headline is HA-SVC-001 — a miscalibration that escapes depot service and biases every reading until the next recalibration.
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope service/maintenance elements and lifecycle phases (HAL-300): cleaning between trips (HA-LIFE-012), recharge at depot (HA-LIFE-013), maintenance/recalibration (HA-LIFE-014), software update at the depot (HA-LIFE-015 wired-OTA path), and decommission / disposal (HA-LIFE-017/018). In-scope components: diagnostics/health & pre-arm readiness firmware (CMP-045), the internal sensor + signal chain and external probe as the recalibration targets (CMP-019/020), bootloader & OTA agent for depot firmware (CMP-035/043), seals/gaskets and USB-C port that service must restore (CMP-002/004), the BMS/protection and cell as the stored-energy source during service (CMP-006/008), the security/crypto module and debug interface (CMP-044), and the diagnostic/developer mode. Out of scope: field service beyond probe/charger (prohibited per HA-CTX-004); the EOL manufacturing calibration and birth record (owned by HAL-1700, HA-MFG-001, though HA-SVC-001 shares its recalibration control RCR-049/026/053). Physical battery-service injury shares controls with HAL-500 (HA-ENRG-004); disposal/environmental consequence is aggregated in HAL-1800 (HA-MPE-005); cyber exposure via left-open debug/data-remanence is shared with HAL-1300 (HA-CYBER-006/008).
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• ISO 13485 / WHO GDP (Good Distribution Practice) — controlled servicing, calibration, and record of maintenance for the reusable fleet; calibrated-status control.
• WHO PQS E006 — cold-chain monitoring device qualification (recalibration and accuracy retention over reuse life).
• IEC 62368-1 — electrical/battery safety during service handling (stored-energy source classification).
• IEC 62304 — software lifecycle for depot software update / configuration (OTA at service).
• IEC 62443 / AAMI TIR57 — security↔safety: locked debug/diagnostic modes, secure decommission, data/key erasure (ties HA-CYBER-006/008).
• UN 38.3 / IATA DGR — lithium-battery handling and transport at decommission/return.
Upstream | D1 (reusable), HAL-300 (lifecycle sweep, reuse-loop amplification §5), UN-1200 (service), PRD-1500/1600 (mfg/service), HA-CTX-004 (field service prohibited), HAL-2400 (RCR register), DOL-1200 (DO-SAFE controls), component-list.md.
Evidence honesty | no service-escape rate is asserted as fact; each P1/P2 records its basis and uncertainty, with process-escape probabilities reduced by the inspection/coverage control named (post-service BIST, pre-arm gate, calibration-expiry gate) and closed by a named PV-/VAL- test (Evidence-Pending).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000). For service records the governing dimension is usually public-health (a wrong disposition) at S4, not the technician-injury dimension.
• Probability is decomposed P1×P2 with an explicit basis and uncertainty flag (HAL-2100). A service-escape is a process probability: P1 (escape past service) is reduced by the post-service BIST and the pre-arm readiness gate; P2 (escape → wrong disposition) is governed by the reuse-loop amplification (the error rides every subsequent trip) and by external receipt-QA checks (contributing, not controlled). Where the closing test is not yet run the record is Evidence-Pending and driven by severity (HAL-2200 §6.2).
Reuse-loop amplification (D1) | per HAL-300 §5, HA-SVC-001 (miscalibration) and HA-SVC-005 (wrong config) are explicitly amplified — a single service escape is re-exposed on every trip until the pre-arm readiness gate (RCR-045) catches it. That gate is the last-defence control before redeploy.
Control hierarchy (HAL-2300) | inherent-by-design → protective → information → process. Because service is intrinsically procedural, service controls lean on process (4) but are always backstopped by a design-enforced gate (RCR-045/026/049/050) so no safety outcome rests on procedure alone (HAL-2200 §6.4).
• Residual risk is re-estimated only after the control is defined, implemented (DO-…), and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (service prompt list)
1. Calibration integrity: Can a recalibration be missed, mis-applied, or drift-uncorrected such that the device reads with a systematic bias — and does anything block arming an out-of-calibration device?
2. Firmware/config restoration: Can service load the wrong firmware, wrong region/band profile, or fail to restore the trip configuration, so the device applies the wrong detection logic or is left unarmed?
3. Seal / ingress restoration: Can service leave the enclosure, port, or probe seal un-restored, admitting moisture that later corrupts the record-integrity electronics?
4. Service-handling energy: Can stored charge expose the technician to arc/short/burn during service?
5. Mode / access hygiene: Can a diagnostic, developer, or debug mode be left enabled after service, creating a security/integrity exposure or a bypass of safety logic?
6. Decommission / disposal: Can a retired device's lithium cell be improperly disposed (fire/environment) or its residual data/keys leak?
6. Hazard Analysis Register (HA-SVC records)
HA-SVC-001 — Miscalibration at depot recalibration escapes → systematic reading bias → wrong disposition ★ HEADLINE (reuse amplifier, D1)
Hazard ID (library) | HAZ-INST (service/recalibration escape) → information-integrity harm (wrong disposition)
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Maintenance / recalibration (HA-LIFE-014) · System State: Off-trip, at service bench · Operational Mode: Recalibration / reprovision
Perspective | Depot/service technician (initiator); beneficiary/patient (ultimate harmed party via a compromised or wrongly-rejected dose); mission/programme
Source | D1 (reuse loop); HAL-300 §5 (recalibration amplifies miscalibration→false-accept); UN-1200; PRD-1600; WHO PQS E006 (accuracy retention); ties HA-MFG-001 (shared calibration control)
Initiating Cause/Event | A scheduled recalibration is skipped, performed against an un-traceable/expired reference, applied to the wrong channel, or its coefficients are not committed — so the sensor signal chain (CMP-019 internal / CMP-020 probe) carries a systematic offset.
Contributing Conditions | High-throughput depot turnaround; recalibration is a manual/procedural step; sensor accuracy class and offset/time-constant are Evidence-Pending (CMP-019/020); a small systematic bias is invisible to the operator (unlike a random fault). Reuse-loop amplifier: the bias rides every subsequent trip until the next recalibration.
Foreseeable Sequence of Events | Recalibration missed/mis-applied → biased coefficients stored → device armed and dispatched → reads e.g. +2 °C low → a real excursion sits just under the threshold and is not detected (or a non-excursion reads as one) → clean-looking verdict → repeated across many trips of the same device.
Hazard | A systematically biased measurement chain (miscalibrated device).
Hazardous Situation | A compromised dose at point-of-care believed acceptable because the biased device under-reported the temperature (or a valid lot wrongly rejected).
Sequence to Harm | Systematic bias → missed excursion → false "accept" → potency-lost dose administered → beneficiary not protected → preventable disease (or false reject → stock-out).
Harm/Loss | A lot/shipment of compromised doses administered to a group of beneficiaries (S4; S5 at population scale as the same device biases many trips).
Severity | S4 (Critical) — governing dimension public-health (a lot of compromised doses administered → beneficiaries unprotected, HAL-2000 §6 severity anchor; §7 "miscalibration escape at EOL = S4"). S5 credible worst-case because one biased reusable device silently affects many trips.
Probability (P1×P2) | P1=2 (Remote — recalibration is a controlled, scheduled step; Test — EOL/service calibration coverage) × P2=4 (Probable — a systematic bias reliably yields a wrong disposition, reuse-amplified across trips; downstream receipt QA is an imperfect external check) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + calibration-coverage; Medium uncertainty, Evidence-Pending — closed by service/EOL calibration coverage audit (PV-SVC-001) and the calibration-expiry pre-arm gate test (PV-SVC-002).
Detectability/Warning | A systematic bias is not directly detectable by the last-mile user; the control is a design-enforced pre-arm gate that blocks arming an out-of-calibration device (RCR-026/049), plus integrity-protected coefficients (RCR-026). Post-service BIST (RCR-050) checks the calibration record before redeploy.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (R region cannot be accepted; HAL-2200 §6).
Risk-Control Strategy | 4 (process — controlled recalibration + coverage) backstopped by 1/2 (design-enforced pre-arm expiry gate that cannot be bypassed)
Risk-Control Requirement | RCR-049 — recalibration interval enforced by the pre-arm expiry gate; an out-of-calibration device cannot be armed. (+ RCR-026 calibration coefficients integrity-protected; expired/invalid calibration blocks arming; + RCR-053 service recalibration meets defined coverage/accuracy and records an updated calibration/birth record.)
Implementation Reference | DO-SVC (recalibration procedure + coverage), DO-TST (calibration coverage), DO-FW-007 / DO-SAFE-007 (pre-arm expiry gate) · CMP-045 (diagnostics/pre-arm), CMP-019/020 (calibration targets), CMP-038 (coefficient application)
Verification Reference | PV-SVC-001 (Evidence-Pending) — service/EOL calibration coverage & accuracy audit; PV-SVC-002 (Evidence-Pending) — calibration-expiry pre-arm gate test (forced expired/invalid coefficients must block arming)
Validation Reference | VAL-SVC-001 (Evidence-Pending) — depot-representative validation that recalibration is completed and traceable before redeploy
Residual Probability | projected Low (PL) after the pre-arm expiry gate is verified (an out-of-cal device cannot be armed; residual is a within-tolerance drift, not an escape) · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off
New Risks Introduced | A strict expiry gate can void a device mid-turnaround (availability/nuisance, not safety) — managed as fleet logistics.
Overall Residual Contribution | Module headline residual and a reuse-loop amplifier; feeds HAL-2900 with HA-MFG-001 as the paired calibration-integrity residual. FTA top event (Stage 11); traces to PFMEA calibration-station controls (Stage 11).
Status | Evidence-Pending (calibration coverage; probability closed by PV-SVC-001/002) — Controlled-pending-verification
Owner | Service Engineer (with Manufacturing and Firmware)
Rationale/Evidence | WHO PQS E006; ISO 13485/GDP calibrated-status control; HAL-300 §5 reuse amplification; RCR-026/049/053; DO-SAFE-007. Severity per HAL-2000 §6/§7.
Assumptions | A calibration reference chain traceable to a standard exists at the depot; the pre-arm gate reads a trustworthy calibration timestamp (RCR-025 trusted time). Final sensor accuracy class Evidence-Pending (CMP-019/020).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-002 — Wrong firmware loaded at depot → wrong detection logic / wrong region → missed excursion or coverage loss
Hazard ID (library) | HAZ-INST (service software escape) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Software update at depot (HA-LIFE-015, wired-OTA/USB-C path) · System State: Service / update · Operational Mode: Firmware load / reprovision
Perspective | Depot technician (initiator); beneficiary/patient; mission
Source | HAL-300 (HA-LIFE-015); IEC 62304 (software lifecycle); ties HA-SW-010 (OTA) and HA-MFG-006 (regional config); PRD-1600
Initiating Cause/Event | A technician loads an incorrect, outdated, or wrong-market firmware image (or a wrong-region band/market profile) onto a device during depot service.
Contributing Conditions | Manual selection of an image/profile at the bench; multiple firmware/market variants in circulation; wired update path bypasses some cloud-side campaign controls; version skew across a partially-updated fleet.
Foreseeable Sequence of Events | Wrong image/profile selected → loaded and applied → device runs mismatched detection logic or a wrong band plan → in the field it either misjudges an excursion (wrong stability logic) or cannot connect on that market's bands (coverage loss) → missed/incorrect/delayed excursion decision.
Hazard | A device running the wrong detection logic or the wrong regional configuration.
Hazardous Situation | A dose monitored by a device applying the wrong stability rules, or a device that cannot deliver an alert in its market.
Sequence to Harm | Wrong logic/coverage → missed excursion or undelivered alert → false "accept" or no time to act → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered / undelivered alerts (S4).
Severity | S4 (Critical) — public-health (wrong disposition), HAL-2000 §6.
Probability (P1×P2) | P1=2 (Remote — signed/authenticated OTA and config verification limit wrong loads; Test) × P2=4 (Probable — a wrong-logic device reliably yields wrong dispositions until caught) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + IEC 62304 update controls; Medium uncertainty, Evidence-Pending — closed by config/version verification test (PV-SVC-003) and the pre-arm valid-firmware/profile gate test (PV-SVC-002).
Detectability/Warning | OTA is signed, integrity- and rollback-protected (RCR-037); the pre-arm gate blocks arming without valid firmware/profile (RCR-045); post-service BIST checks the loaded version (RCR-050). Wrong load is caught before arming, not by the field user.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 2 (protective — signed/verified update) + 4 (process — controlled image/profile selection) backstopped by the pre-arm gate (design-enforced)
Risk-Control Requirement | RCR-037 — OTA signed, integrity- and rollback-protected, fail-safe A/B; never bricks; unmonitored windows disclosed. (+ RCR-056 correct regional/firmware configuration provisioned and verified; + RCR-045 pre-arm valid-firmware/profile gate.)
Implementation Reference | DO-SAFE-014 (→ DO-FW-014/015 + DO-CLD-010), DO-MFG/DO-FW (config verification), DO-SVC (service update procedure) · CMP-035 (bootloader), CMP-043 (OTA agent), CMP-027 (SIM/band), CMP-045 (pre-arm)
Verification Reference | PV-SVC-003 (Evidence-Pending) — firmware/region config verification test at service; PV-SVC-002 (Evidence-Pending) — pre-arm valid-firmware/profile gate test; OTA rollback + version-skew test (with HA-SW-010)
Validation Reference | VAL-SVC-003 (Evidence-Pending) — depot update-procedure validation (correct variant selected for market)
Residual Probability | projected Low (PL) after signed update + pre-arm gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | A blocked/void device on a version mismatch (availability, not safety).
Overall Residual Contribution | Ties HA-SW-010 (OTA) and HA-MFG-006 (regional config); carried in HAL-2900. FTA top-event contributor (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Service Engineer (with Firmware and Manufacturing)
Rationale/Evidence | IEC 62304; RCR-037/045/056; DO-SAFE-014. Severity per HAL-2000 §6.
Assumptions | Signed OTA also governs the wired/USB-C depot update path; band/market profile is part of the verified configuration set.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-003 — Incomplete service: seal/gasket not restored → ingress in field → record-integrity electronics fault
Hazard ID (library) | HAZ-ENV (ingress) → HAZ-ELEC / information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Maintenance / reassembly after service (HA-LIFE-014) · System State: Reassembled, returning to service · Operational Mode: Redeploy after service
Perspective | Technician (initiator); beneficiary/patient (via a later record-integrity fault); mission
Source | HAL-300 (HA-LIFE-014); IEC 60529 (IP); ties HA-ELEC-006/007, HA-ENV-003, HA-MFG-004 (assembly-time seal damage); PRD-MECH
Initiating Cause/Event | After opening the device for service, a seal/gasket (enclosure joint, USB-C port, or probe channel) is pinched, mis-seated, damaged, or omitted on reassembly.
Contributing Conditions | Reuse life subjects seals to repeated open/close and compression set (CMP-002); cold-to-warm cycling drives condensation into any breach; the fault is latent — the device works at the bench and fails later in the field.
Foreseeable Sequence of Events | Seal not restored → device redeployed → cold-to-warm/humid cycling drives moisture ingress → condensation/leakage/corrosion on the PCB and signal chain → record-integrity electronics fault → wrong/lost readings or a monitoring gap → wrong disposition.
Hazard | Moisture ingress reaching record-integrity electronics.
Hazardous Situation | A dose monitored by a device whose measurement/record electronics are being corrupted by ingress.
Sequence to Harm | Ingress → corrupted/lost record → false "accept" on a gap/faulted read → compromised dose administered, or a monitored trip released on an incomplete record.
Harm/Loss | Wrong disposition of a limited quantity of doses / a trip on an incomplete record (S3; escalates toward S4 if it produces a clean-looking false accept).
Severity | S3 (Serious) — public-health, "a monitored trip released on an incomplete record" / limited-quantity wrong disposition (HAL-2000 §5 S3). (If it manifests as a silent false-accept it escalates via HA-SW-002/HA-DATA-004; scored here at the direct-outcome S3.)
Probability (P1×P2) | P1=2 (Remote — post-service BIST + IP practice; but manual reassembly over reuse life) × P2=3 (Occasional — ingress → record fault → wrong disposition, partly caught by honest-gap logic) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + IEC 60529; Medium uncertainty, Evidence-Pending — closed by post-service IP/BIST test (PV-SVC-004).
Detectability/Warning | Post-service BIST + pre-arm gate (RCR-050/045) check integrity before redeploy but cannot fully prove a seal; honest-gap and verdict-integrity logic (RCR-022/021) prevent a resulting fault from reading as a clean pass. IP integrity is best assured by restoring the seal correctly (design/process).
Initial Risk | AP Medium (region A/R — S3 × PM → region A per HAL-2200 matrix S3×PM)
Risk Acceptability | ALARP — reduce as far as practicable via process (seal restoration) + protective (IP design + BIST).
Risk-Control Strategy | 4 (process — service restores seal + post-service BIST) + 1/2 (inherent IP + verdict/gap integrity as backstop)
Risk-Control Requirement | RCR-050 — service restores seal, configuration, and firmware; a post-service BIST + pre-arm gate verify integrity before redeploy. (+ RCR-011 enclosure + seals maintain ingress rating through reuse life; + RCR-022/021 honest gap / verdict integrity as information backstop.)
Implementation Reference | DO-SVC (reassembly + BIST), DO-SAFE-010/011 (→ DO-MECH-005 seal + DO-PCB-010 coat) · CMP-002 (seals), CMP-004 (USB-C port), CMP-005 (probe channel), CMP-034 (PCB/coat), CMP-045 (BIST/pre-arm)
Verification Reference | PV-SVC-004 (Evidence-Pending) — post-service IP re-test + BIST integrity check; seal re-seat/compression-set over reuse-cycle test
Validation Reference | VAL-SVC-004 (Evidence-Pending) — service-reassembly procedure usability (seal correctly restored by representative technicians)
Residual Probability | projected Low (PL) after BIST + IP design verified · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B, S3×PL) — ALARP
New Risks Introduced | None (shares seal control with HA-MFG-004 / HA-ENV-003).
Overall Residual Contribution | Small; ties HA-ELEC-006/007 (ingress→electronics) and HA-MFG-004 (assembly-time seal). Carried in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Service Engineer (with Mechanical)
Rationale/Evidence | IEC 60529; RCR-011/050; DO-SAFE-011. Severity per HAL-2000 §5.
Assumptions | Post-service BIST covers a functional check but not a full IP proof; a periodic seal-replacement policy exists over reuse life.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-004 — Service while energized → arc / short / burn to technician
Hazard ID (library) | HAZ-THERM / HAZ-ELEC (stored electrical energy, low-voltage high-current)
Product/Config | SCM all SKUs · Lifecycle Phase: Service / maintenance (board rework, battery/pack service) · System State: Powered-down but stored energy present · Operational Mode: Service
Perspective | Depot / repair technician
Source | IEC 62368-1 (energy-source classification); ties HA-ENRG-004 (shared control); PRD-1600 (service procedure)
Initiating Cause/Event | During service a tool or conductor bridges the charged cell/pack terminals or a power net, releasing high short-circuit current (arc/spark).
Contributing Conditions | The cell is a SELV but high-energy source (large short-circuit current); service may proceed without a battery-safe (de-energize/isolate) procedure; sealed enclosure means only trained depot service opens the device (HA-CTX-004).
Foreseeable Sequence of Events | Technician opens an energized device → tool bridges terminals/nets → arc / high short-circuit current → localized heating/spark → burn or startle → possible cell damage feeding HA-ENRG-002.
Hazard | Stored electrical energy (short-circuit current / arc) during service.
Hazardous Situation | A technician handling an energized pack/board with exposed conductors.
Sequence to Harm | Arc/short → burn to hand / eye flash / startle → moderate reversible injury.
Harm/Loss | Moderate reversible burn requiring first aid (S3).
Severity | S3 (Serious) — human injury, moderate reversible burn (HAL-2000 §5 S3 anchor).
Probability (P1×P2) | P1=2 (Remote — controlled depot service with a battery-safe procedure) × P2=2 (Remote — arc → treatable burn) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62368-1; Medium uncertainty — closed by service-short fault-injection (PV-SVC-005).
Detectability/Warning | Energized state indicated (LED, CMP-032); the secondary non-resettable fuse (RCR-002) limits sustained fault current; battery-safe procedure de-energizes before rework.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP — reduce via protective (fuse limits energy) + process (battery-safe service procedure).
Risk-Control Strategy | 2 (protective — fuse limits fault current; SELV) + 4 (process — battery-safe service procedure)
Risk-Control Requirement | RCR-002 — autonomous protection with a non-resettable secondary backstop (fuse limits fault current). (+ RCR-013 SELV: no accessible hazardous voltage; + RCR-050 service procedure / post-service BIST.)
Implementation Reference | DO-SAFE-002 (→ DO-EE-005 BMS + DO-EE-010 fuse), DO-SVC (battery-safe procedure) · CMP-008 (BMS), CMP-009 (fuse), CMP-006/007 (cell/pack)
Verification Reference | PV-SVC-005 (Evidence-Pending) — service-short fault-injection (confirm fuse limits current); service-procedure review
Validation Reference | VAL-SVC-005 (Evidence-Pending) — service-procedure usability with representative technicians
Residual Probability | projected Very Low (PVL) with fuse + procedure · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small; occupational-safety residual noted in HAL-2900; shares control with HA-ENRG-004.
Status | Controlled-pending-verification
Owner | Service Engineer (with Battery/Power)
Rationale/Evidence | IEC 62368-1 energy-source classification; RCR-002/013; DO-SAFE-002. Severity per HAL-2000 §5.
Assumptions | Field service prohibited (HA-CTX-004); only trained depot service opens the device with a de-energize step.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-005 — Configuration not restored after service → wrong profile / device left un-armed → unmonitored or mis-monitored trip
Hazard ID (library) | HAZ-INST (service config escape) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Reprovision / redeploy after service (HA-LIFE-014→007) · System State: Reprovisioned, dispatched · Operational Mode: Redeploy / arm
Perspective | Depot technician / dispatch (initiator); beneficiary/patient; mission
Source | HAL-300 §5 (reprovisioning amplifies wrong-config); ties HA-USE-003/004, HA-SW-003; PRD-1600
Initiating Cause/Event | After service the trip/product-stability profile is not restored (left at a default/wrong profile), or the device is returned to stock un-armed while believed armed/ready.
Contributing Conditions | Reuse loop requires reprovisioning each cycle (D1); manual configuration step; a wrong profile or an un-armed device is not obvious without a status check. Reuse-loop amplifier: wrong config persists across the device's next deployment.
Foreseeable Sequence of Events | Config not restored → device dispatched with wrong stability thresholds or un-armed → in the field it either applies the wrong excursion logic or logs nothing → excursion undetected → clean-looking or absent record → wrong disposition.
Hazard | A device with the wrong trip configuration or left un-armed.
Hazardous Situation | A dose in transit monitored against the wrong rules or not monitored at all, believed monitored.
Sequence to Harm | Wrong/absent monitoring → missed excursion → false "accept" → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered / an unmonitored trip (S4).
Severity | S4 (Critical) — public-health (wrong disposition / unmonitored trip), HAL-2000 §6.
Probability (P1×P2) | P1=2 (Remote — reprovision is a controlled step; pre-arm gate + profile validation) × P2=4 (Probable — a wrong/absent config reliably yields a wrong or missing decision) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment; Medium uncertainty, Evidence-Pending — closed by pre-arm readiness gate test (PV-SVC-002) and profile-validation test (PV-SVC-006).
Detectability/Warning | The pre-arm readiness gate blocks arming unless a valid/unexpired profile, firmware, SoC-for-trip, and probe-present all pass, and makes an un-armed/void state unmistakable (RCR-045); profile validation flags an implausible/mismatched profile (RCR-048). This gate is the last-defence reuse-loop control (HAL-300 §5).
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — restore config in service) backstopped by 1/2 (design-enforced pre-arm gate + profile validation)
Risk-Control Requirement | RCR-050 — service restores seal, configuration, and firmware; post-service BIST + pre-arm gate verify before redeploy. (+ RCR-045 pre-arm readiness gate blocks arming and makes un-armed/void unmistakable; + RCR-048 profile validation.)
Implementation Reference | DO-SVC (reprovision procedure + BIST), DO-SAFE-007 (pre-arm gate → DO-FW-032), DO-APP/DO-CLD (profile validation) · CMP-045 (pre-arm/diagnostics), CMP-055 (mobile provisioning), CMP-031/032 (arming button / status)
Verification Reference | PV-SVC-002 (Evidence-Pending) — pre-arm readiness gate + arming/void test; PV-SVC-006 (Evidence-Pending) — profile-validation test (implausible profile flagged)
Validation Reference | VAL-SVC-005 (Evidence-Pending) — reprovision/arming procedure usability (correct profile restored; un-armed state noticed)
Residual Probability | projected Low (PL) after pre-arm gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Frequent void/blocked states could drive workarounds (use-error) — reconciled with HA-USE-003/004 and RCR-046 legibility.
Overall Residual Contribution | Reuse-loop residual paired with HA-SVC-001; ties HA-USE-003/004. Carried in HAL-2900. FTA top-event contributor (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Service Engineer (with Firmware and Human-Factors)
Rationale/Evidence | HAL-300 §5 reuse amplification; RCR-045/048/050; DO-SAFE-007. Severity per HAL-2000 §6.
Assumptions | The pre-arm gate is not user-overridable for a missing profile; profile library is current for the market.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-006 — Debug / diagnostic / developer mode left enabled after service → security & integrity exposure → falsified record / suppressed alarm
Hazard ID (library) | HAZ-CYBER (open access mode) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Post-service redeploy (HA-LIFE-014→016) · System State: Fielded with an unexited service mode · Operational Mode: Armed / logging with debug active
Perspective | Beneficiary/patient; mission; organization (fleet security); ties HA-CYBER-006
Source | IEC 62443 / AAMI TIR57 (secure development/operation); ties HA-CYBER-006 (debug ports / secure boot); PRD-SEC; RCR-051/038
Initiating Cause/Event | A diagnostic, developer, or debug mode (or an unlocked debug port) used during service is not disabled/re-locked before the device returns to the field.
Contributing Conditions | Service legitimately needs elevated access; re-locking is a manual step; an open debug path can bypass integrity/authentication controls (record signing, alarm path, config authorization).
Foreseeable Sequence of Events | Debug/diagnostic mode left on → fielded device exposes an unauthenticated access/override path → an actor (or a fault) alters thresholds, injects/falsifies telemetry, or suppresses the alarm path → a real excursion is hidden → clean-looking verdict → wrong disposition.
Hazard | An unlocked service/debug access path in a fielded device.
Hazardous Situation | A fielded, monitoring device whose record/alarm integrity can be bypassed via a left-open mode.
Sequence to Harm | Integrity bypass → falsified record / suppressed alarm → missed excursion → false "accept" → compromised dose administered → beneficiary not protected.
Harm/Loss | Falsified/suppressed monitoring → a lot of compromised doses administered (S4).
Severity | S4 (Critical) — public-health (falsified telemetry / suppressed alarm → wrong disposition), HAL-2000 §7 (cyber archetype = S4).
Probability (P1×P2) | P1=2 (Remote — mode-lock verification + secure-boot/debug-lock; Pen-test) × P2=3 (Occasional — exploitation of an open mode → undetected → wrong disposition) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + pen-test posture; Medium uncertainty, Evidence-Pending — closed by mode-lock verification (PV-SVC-007) and debug-lock pen-test (with HA-CYBER-006).
Detectability/Warning | An unexited service mode is detected (RCR-051) and post-service BIST/pre-arm checks the security posture (RCR-050); secure boot + debug-port lock (RCR-038) are the design backstops. An open mode is otherwise invisible to the field user.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 2 (protective — mode-lock + secure boot + debug-port lock) + 4 (process — service re-lock step and its verification)
Risk-Control Requirement | RCR-051 — diagnostic/developer/debug modes locked/disabled before field deployment; an unexited service mode is detected. (+ RCR-038 secure boot against a hardware root of trust; production debug ports locked/disabled; tamper-evident enclosure.)
Implementation Reference | DO-SEC-006, DO-SVC (re-lock + verification), DO-FW (mode control) · CMP-044 (security/crypto module), CMP-016 (secure element/root of trust), CMP-045 (diagnostics/BIST)
Verification Reference | PV-SVC-007 (Evidence-Pending) — mode-lock verification (no service mode escapes to field); secure-boot + debug-lock pen-test (shared with HA-CYBER-006)
Validation Reference | N/A (assurance carried by pen-test/verification)
Residual Probability | projected Very Low (PVL) after mode-lock detection + debug-lock verified · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off (with EM-ENG-008)
New Risks Introduced | A hard mode-lock could impede legitimate field diagnostics (serviceability, not safety) — managed by an authenticated re-entry path.
Overall Residual Contribution | Ties HA-CYBER-006 (debug/secure-boot); jointly carried in HAL-2900 with Product Security (EM-ENG-008). FTA top-event contributor (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Service Engineer (with Product Security EM-ENG-008)
Rationale/Evidence | IEC 62443 / AAMI TIR57; RCR-038/051; DO-SEC-006. Severity per HAL-2000 §7.
Assumptions | Elevated service access is always time-boxed and auto-relocks; secure boot detects an unauthorized/left-open image or mode at next boot.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-SVC-007 — Improper lithium-battery disposal at decommission → fire / environmental release (and data/key remanence)
Hazard ID (library) | HAZ-FIRE / HAZ-ECOL (Li-battery / e-waste at end-of-life) + HAZ-CYBER (data/key remanence)
Product/Config | SCM all SKUs (end-of-life) · Lifecycle Phase: Decommission / disposal / recycling (HA-LIFE-017/018) · System State: Retired / powered-off · Operational Mode: Decommission
Perspective | Waste-handler / bystander (incl. child near a discarded device); environment; organization (data remanence); ties HA-MPE-005, HA-CYBER-008
Source | UN 38.3 / IATA DGR (Li-battery handling); local e-waste/battery regulation; IEC 62443 (secure decommission/erasure); ties HA-MPE-005 (fleet-scale environmental) and HA-CYBER-008 (data/key remanence); RCR-052
Initiating Cause/Event | A retired device (or its lithium cell) is discarded to general waste rather than through battery recycling, and/or is decommissioned without data/key erasure.
Contributing Conditions | Reusable fleet reaches end-of-life at scale (D1); depots/partners may lack a battery-recycling path; a damaged/deep-discharged cell is more prone to a thermal event; residual keys/records remain in flash/secure element.
Foreseeable Sequence of Events | Device retired → discarded improperly → cell crushed/shorted in waste stream → thermal event / fire or heavy-metal/electrolyte environmental release; separately, residual data/keys recoverable from the discarded device.
Hazard | Stored electrochemical energy + hazardous materials in an uncontrolled waste stream; residual cryptographic identity/records.
Hazardous Situation | A discarded lithium cell in a waste/recycling stream near people; a discarded device retaining recoverable keys/records.
Sequence to Harm | Cell crush/short in waste → fire/venting → burn/environmental contamination; or key/record recovery → identity misuse / record repudiation.
Harm/Loss | Fire/environmental release (S3 direct; S3 aggregate/fleet-scale environmental via HA-MPE-005); data/key compromise (via HA-CYBER-008).
Severity | S3 (Serious) — governing dimension environmental (improper Li-battery disposal), HAL-2000 §7 ("improper Li-battery disposal = S3"). Fleet-scale aggregation carried by HA-MPE-005.
Probability (P1×P2) | P1=3 (Occasional — improper disposal is a real fleet-scale behaviour without a mandated path) × P2=2 (Remote — a single discarded cell → fire/release, mitigated by discharge state and instructions) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + industry Li-ion disposal experience; Medium uncertainty — closed by disposal-procedure review (PV-SVC-008).
Detectability/Warning | Not detectable post-disposal; the control is preventive — end-of-life instructions (battery removal/recycling, data/key erasure) supported by design (RCR-052) and secure decommission (RCR-038/052).
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce via information + process (mandated recycling + erasure) and design-for-disassembly.
Risk-Control Strategy | 3 (information — safe EOL instructions) + 4 (process — mandated battery recycling + secure decommission) + 1 (design-for-disassembly/erasure support)
Risk-Control Requirement | RCR-052 — safe end-of-life instructions (battery removal/recycling, data/key erasure) provided and supported by design. (+ RCR-038 secure decommission; + RCR-007 UN 38.3 / IATA transport of returned cells.)
Implementation Reference | DO-SVC (decommission procedure), DO-LBL (EOL/recycling label), DO-SEC (data/key erasure) · CMP-006 (cell), CMP-044/016 (keys/identity)
Verification Reference | PV-SVC-008 (Evidence-Pending) — disposal/decommission procedure review; data/key-erasure verification (with HA-CYBER-008)
Validation Reference | VAL-SVC-007 (Evidence-Pending) — take-back/recycling program uptake validation (fleet-scale, ties HA-MPE-005)
Residual Probability | projected Low (PL) with mandated take-back + erasure verified · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B, S3×PL) — ALARP
New Risks Introduced | None (disposal handling is inert).
Overall Residual Contribution | Ties HA-MPE-005 (fleet-scale environmental) and HA-CYBER-008 (data/key remanence); the aggregate environmental view is carried in HAL-1800 and HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Service Engineer (with Product Security EM-ENG-008 and Environmental)
Rationale/Evidence | UN 38.3 / IATA DGR; e-waste/battery regulation; RCR-052; HAL-2000 §7. Severity per HAL-2000 §7.
Assumptions | A regional battery take-back/recycling channel is contractually available; secure erasure removes recoverable keys/records at decommission.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, severity anchor §6/§7), HAL-2100 (P1×P2, basis taxonomy, process-escape probabilities reduced by inspection/coverage), and HAL-2200 (matrix, regions R/A/B, AP, severity-first + uncertain-probability rules).
Controls | RCR-049/053/026 (calibration), RCR-037/056/045 (firmware/config/pre-arm), RCR-050/011 (seal + BIST), RCR-002/013 (service energy), RCR-051/038 (mode-lock/secure boot), RCR-052 (decommission/disposal) — all from HAL-2400. The pre-arm readiness gate (RCR-045) is the last-defence reuse-loop control across HA-SVC-001/002/005.
Implementation | DO-SVC / DO-TST are the primary service Implementation References, backstopped by DO-SAFE-007 (pre-arm gate), DO-SAFE-002/010/011 (energy/seal), DO-SAFE-014 (OTA), DO-SEC-006 (mode-lock); each points to CMP-002/004/005/006/008/009/016/019/020/034/035/043/044/045/055.
Cross-module ties | HA-SVC-001 ↔ HA-MFG-001 (shared calibration control, EOL vs service) and ↔ HA-USE-011 (reuse-without-recalibration use-error); HA-SVC-002 ↔ HA-SW-010 / HA-MFG-006 (firmware/region); HA-SVC-003 ↔ HA-ELEC-006/007 / HA-ENV-003 / HA-MFG-004 (ingress/seal); HA-SVC-004 ↔ HA-ENRG-004 (service energy); HA-SVC-005 ↔ HA-USE-003/004 / HA-SW-003 (config/arming); HA-SVC-006 ↔ HA-CYBER-006 (debug/secure-boot); HA-SVC-007 ↔ HA-MPE-005 (environmental) and HA-CYBER-008 (data/key remanence).
Downstream | the S4 records (HA-SVC-001/002/005/006) become FTA top events in Stage 11 (ERFKL) and drive deeper service/SW analysis; the calibration and provisioning escapes trace to the PFMEA and production/service controls (Stage 11) alongside HAL-1700. All PV-/VAL- references are Evidence-Pending and feed Stages 13–14.
8. Residual-risk / closure notes
Module headline | HA-SVC-001 (miscalibration escape) is the carried service headline — a reuse-loop amplifier whose residual is driven by severity while the calibration-coverage probability is Evidence-Pending (PV-SVC-001/002). It is paired in HAL-2900 with HA-MFG-001 (EOL calibration escape) as the calibration-integrity residual.
Reuse-loop set | HA-SVC-001/005 (and HA-USE-011) are the D1 amplifiers; each is backstopped by the design-enforced pre-arm readiness gate (RCR-045/026/049), the last defence before redeploy (HAL-300 §5).
No warnings-only acceptance | no service record is accepted on procedure alone (HAL-2200 §6.4); each procedural control is backstopped by a design-enforced gate (pre-arm, calibration-expiry, mode-lock, post-service BIST).
Release gate (HAL-2200 §8) | no R region service risk may remain open at release; every remaining region-A S4 service risk carries a benefit-risk statement and residual-risk sign-off (EM-ENG-006 + EM-ENG-001; EM-ENG-008 for HA-SVC-006); every Evidence-Pending probability has a named closing test above.
Overall contribution | the service module's residual is dominated by the calibration and config/firmware reuse-loop escapes (HA-SVC-001/002/005); these are principal service inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
HAL-1700 (HA-MFG) — Manufacturing & Supply-Chain Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1700, record code HA-MFG · Owner: Principal Manufacturing/Supply-Chain Engineer, with System-Safety (lead), Quality/Regulatory (EM-ENG-006), and the Battery/Power and Firmware engineers. This module analyses the hazards introduced by procurement, assembly, EOL calibration/provisioning, and crypto/config birth — the ways a manufacturing or supply-chain escape becomes a latent defect that later causes harm in the field. It follows the ISO 14971 harm-and-exposure-first method (HAL-000 §4); it is not an FMEA — but every record here traces forward into the PFMEA and production controls (Stage 11). Each record is a complete risk chain, scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way that a manufacturing or supply-chain escape — a wrong part, a wrong build, a calibration/provisioning error, or an inadequate screen — can leave a device that later causes harm, dominantly a wrong vaccine-disposition decision, secondarily a battery/thermal event. Manufacturing escapes are latent and fleet-scaled: one process defect can ship across a lot of devices before it is seen in the field. The module headline is HA-MFG-001 — a miscalibration that passes EOL test and biases every device it escapes on. This register is the harm-and-exposure precursor to the PFMEA (Stage 11): it names the hazards; the PFMEA quantifies the process failure modes and detection at each station.
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope manufacturing/supply-chain elements and lifecycle phases (HAL-300): procurement/incoming (HA-LIFE-001), assembly (HA-LIFE-002), EOL calibration & provisioning (HA-LIFE-003), packaging/labeling (HA-LIFE-004). In-scope components: the low-temp cell and pack (CMP-006/007), BMS/protection IC, secondary fuse, charge-input front-end (CMP-008/009/014), internal sensor & external probe as EOL calibration targets (CMP-019/020), secure element / crypto identity (CMP-016), SIM/band & bootloader for regional/firmware config (CMP-027/035), seals/gaskets (CMP-002), and the PCB/solder stack-up (CMP-034). In-scope EOL steps: calibration + birth record, crypto identity provisioning, and regional/firmware configuration. Out of scope: depot recalibration/reprovision in the reuse loop (owned by HAL-1600, though HA-MFG-001 shares its calibration control RCR-053/026/049 with HA-SVC-001); the mission-level aggregation of these escapes (owned by HAL-1800). The chemistry (CMP-006) and antenna (CMP-028/029) remain the two Evidence-Pending architectural unknowns; no chemistry, capacity, datasheet, or RF figure is invented here.
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• ISO 13485 / WHO GDP — controlled manufacture, calibrated-status control, device history/birth record, supplier control and incoming inspection.
• IEC 62133-2 / UN 38.3 — safety and transport qualification of the rechargeable lithium cell/pack (correct chemistry, protection, abuse tolerance).
• IEC 62368-1 — battery/electrical safety of the built device (protection component orientation/rating).
• IEC 62443 / AAMI TIR57 — unique cryptographic identity provisioning, secure supply chain (counterfeit and duplicate-identity prevention).
• IPC-A-610 / IPC-6012 (workmanship/PCB) and AOI/ICT coverage practice — solder/PCB defect screening.
• WHO PQS E006 — device qualification / EOL accuracy verification for cold-chain monitors.
Upstream | HAL-300 (HA-LIFE-001/002/003), PRD-1500 (mfg), component-list.md (CMP status/Evidence-Pending), HAL-2400 (RCR register), DOL-1200 (DO-SAFE), and the PFMEA (Stage 11, ERFKL) as the downstream owner of process failure-mode detail.
Evidence honesty | no process-escape rate is asserted as fact; each P1/P2 records its basis and uncertainty, with process-escape probabilities reduced by the named inspection/coverage control (incoming inspection, AOI/ICT, poka-yoke, EOL calibration coverage) and closed by a named PV-/VAL- test (Evidence-Pending).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000). Calibration/config escapes govern at S4 public-health; battery/protection build escapes govern at S4 human-injury (S5 credible worst-case via the fire chain, cross-referenced to HAL-500/800, not double-counted).
• Probability is a process-escape probability decomposed P1×P2 (HAL-2100): P1 (defect escapes the line) is reduced by the station control (incoming inspection, AOI/ICT, poka-yoke, EOL calibration coverage); P2 (escape → harm) is governed by fleet-scale exposure and by the field backstop (pre-arm gate, protection backstop, plausibility/verdict integrity). Where the closing test is not yet run the record is Evidence-Pending and driven by severity (HAL-2200 §6.2).
Single-fault rule (HAL-2200 §6.5) | for the S4/S5 battery/protection build escapes (HA-MFG-002/003) the device must remain single-fault-safe — the independent protective backstop (RCR-002) must survive a wrong or reversed primary protection part, and correct parts are verified by poka-yoke + AOI (RCR-054).
Control hierarchy (HAL-2300) | inherent-by-design → protective → information → process. Manufacturing controls are predominantly process (4) but are always paired with a design/field backstop so no safety outcome rests on a single inspection (HAL-2200 §6.4).
PFMEA hand-off (Stage 11) | each record names the process step and the station control it relies on; the PFMEA quantifies the process failure mode, occurrence, and detection at that station. This module is the hazard/harm precursor; the PFMEA is the process-quantification.
• Residual risk is re-estimated only after the control is defined, implemented (DO-…), and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (manufacturing/supply-chain prompt list)
1. Calibration & birth record: Can a miscalibrated device pass EOL, and does anything block arming it later? Is the birth record complete and traceable?
2. Cell & protection integrity: Can a wrong/counterfeit cell chemistry or a reversed/wrong-rated protection component be built in, defeating the battery-safety design?
3. Enclosure & seal build: Can a seal be damaged at assembly, admitting field ingress?
4. Identity & configuration: Can a device be given a wrong/duplicate crypto identity or a wrong regional/firmware configuration at provisioning?
5. Test coverage: Is EOL test coverage adequate to catch safety-critical latent defects before shipment?
6. Latent workmanship: Can a solder/PCB defect from thermal/mechanical assembly stress escape and fail intermittently in the field, creating a silent monitoring gap?
6. Hazard Analysis Register (HA-MFG records)
HA-MFG-001 — EOL calibration escape (miscalibrated unit passes EOL) → systematic reading bias → wrong disposition ★ HEADLINE
Hazard ID (library) | HAZ-MFG (calibration/test escape) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: EOL calibration & provisioning (HA-LIFE-003) · System State: End-of-line test · Operational Mode: Calibration / birth record
Perspective | Manufacturing (initiator); beneficiary/patient (ultimate harmed party); mission/programme
Source | HAL-300 (HA-LIFE-003); WHO PQS E006 (EOL accuracy verification); HAL-2100 §7 ("miscalibration escape at EOL"); ties HA-SVC-001 (shared calibration control); PRD-1500
Initiating Cause/Event | A device is calibrated against an inadequate reference, with insufficient coverage across the range, or its coefficients/birth record are mis-committed — and it passes EOL carrying a systematic offset.
Contributing Conditions | Calibration coverage/accuracy is a defined but Evidence-Pending target; sensor accuracy class is Evidence-Pending (CMP-019/020); a systematic bias is invisible on a pass/fail functional test unless the calibration coverage is designed to catch it. Fleet-scale: the same station escapes onto a lot of devices.
Foreseeable Sequence of Events | Inadequate calibration → biased coefficients pass EOL → device armed in the field → reads systematically low/high → real excursions sit under threshold and go undetected (or false excursions) → clean-looking verdicts across many devices/trips.
Hazard | A fleet of systematically biased measurement chains.
Hazardous Situation | Compromised doses at point-of-care believed acceptable because the biased devices under-reported temperature (or valid lots wrongly rejected).
Sequence to Harm | Systematic bias → missed excursion → false "accept" → potency-lost dose administered → beneficiary not protected → preventable disease (or false reject → stock-out).
Harm/Loss | Lots of compromised doses administered across many beneficiaries (S4; S5 at population/fleet scale).
Severity | S4 (Critical) — public-health (a lot of compromised doses administered, HAL-2000 §6/§7 "miscalibration escape at EOL = S4"). S5 credible worst-case because an EOL escape scales across a manufacturing lot.
Probability (P1×P2) | P1=2 (Remote — EOL calibration coverage + birth record; Test — EOL calibration coverage audit) × P2=4 (Probable — a systematic bias reliably yields a wrong disposition, fleet-amplified) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + EOL coverage; Medium uncertainty, Evidence-Pending — closed by EOL calibration coverage & accuracy audit (PV-MFG-001) and the calibration-expiry/validity pre-arm gate test (PV-MFG-002).
Detectability/Warning | A systematic bias is not field-detectable; the controls are (a) EOL calibration coverage that catches the bias before shipment (RCR-053) and (b) the design-enforced pre-arm gate that blocks arming on invalid/expired calibration (RCR-026/049). The birth record makes the calibration state auditable.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — EOL calibration coverage + birth record) backstopped by 1/2 (integrity-protected coefficients + pre-arm validity gate)
Risk-Control Requirement | RCR-053 — EOL calibration shall meet defined coverage and accuracy and record a birth record; a miscalibrated unit cannot pass EOL. (+ RCR-026 calibration coefficients integrity-protected; expired/invalid blocks arming; + RCR-049 recalibration-interval pre-arm enforcement shared with HA-SVC-001.)
Implementation Reference | DO-TST (EOL calibration station + coverage), DO-MFG (birth record), DO-FW-007 / DO-SAFE-007 (pre-arm validity gate) · CMP-019/020 (calibration targets), CMP-038 (coefficient application), CMP-045 (pre-arm/diagnostics)
Verification Reference | PV-MFG-001 (Evidence-Pending) — EOL calibration coverage & accuracy audit (coverage across the range, reference traceability); PV-MFG-002 (Evidence-Pending) — calibration-validity pre-arm gate test
Validation Reference | VAL-MFG-001 (Evidence-Pending) — production-representative validation that EOL calibration achieves field accuracy over reuse life
Residual Probability | projected Low (PL) after EOL coverage + pre-arm validity gate verified · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off
New Risks Introduced | Tighter EOL calibration adds cycle time/yield loss (cost, not safety).
Overall Residual Contribution | Module headline residual; paired in HAL-2900 with HA-SVC-001 as the calibration-integrity residual. FTA top event (Stage 11); traces to the PFMEA calibration-station controls (Stage 11).
Status | Evidence-Pending (calibration coverage; probability closed by PV-MFG-001/002) — Controlled-pending-verification
Owner | Manufacturing Engineer (with Quality and Firmware)
Rationale/Evidence | WHO PQS E006; ISO 13485/GDP calibrated-status control; RCR-026/049/053; DO-TST. Severity per HAL-2000 §6/§7.
Assumptions | A traceable calibration reference chain exists at EOL; birth record is committed atomically and read by the pre-arm gate. Sensor accuracy class Evidence-Pending (CMP-019/020).
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-002 — Wrong / counterfeit cell chemistry sourced → cold-charge/thermal risk & endurance miss
Hazard ID (library) | HAZ-SUPP (counterfeit/wrong part) → HAZ-THERM / HAZ-FIRE + information-integrity harm (endurance)
Product/Config | SCM all SKUs · Lifecycle Phase: Procurement / incoming (HA-LIFE-001) · System State: Cell/pack build · Operational Mode: Manufacture
Perspective | People near the device (thermal); beneficiary/patient (endurance→silent gap); mission/property
Source | HAL-300 (HA-LIFE-001); IEC 62133-2; UN 38.3; ties HA-ENRG-001/002 (cold-charge/thermal) and HA-ENRG-006 (endurance); component-list (CMP-006 chemistry Evidence-Pending); PRD-1500
Initiating Cause/Event | A wrong-grade or counterfeit lithium cell (not the qualified low-temperature chemistry) is sourced and built into the pack.
Contributing Conditions | The low-temp chemistry itself is Evidence-Pending (CMP-006); counterfeit cells are a known supply-chain risk; a wrong chemistry may not tolerate the cold-charge inhibit assumptions or meet the 90-day/−30 °C budget, and may have inferior abuse tolerance.
Foreseeable Sequence of Events | Wrong/counterfeit cell built in → either (a) it plates/degrades under the cold-charge regime or has poor abuse tolerance → thermal event, or (b) its cold usable capacity misses the budget → dead-in-cold silent gap → wrong disposition.
Hazard | A non-qualified cell chemistry with wrong safety/abuse properties or insufficient cold capacity.
Hazardous Situation | A device with an unqualified cell charging/operating near people, or running out of energy mid-trip.
Sequence to Harm | Thermal path → burn/fire (S4, S5 in occupied hold via HA-TFC-002); endurance path → silent gap → false accept → compromised dose administered (S4).
Harm/Loss | Serious burn / thermal event; or a lot of compromised doses administered via silent gap (S4).
Severity | S4 (Critical) — highest of human-injury (thermal) and public-health (endurance); S5 credible worst-case via the occupied-hold fire chain (HA-TFC-002, cross-referenced, not double-counted).
Probability (P1×P2) | P1=2 (Remote — incoming inspection + supplier control; Test — incoming verification) × P2=3 (Occasional — wrong chemistry → thermal or endurance harm) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + supply-chain counterfeit experience; Medium uncertainty, Evidence-Pending — closed by incoming cell verification (PV-MFG-003) and, on the qualified chemistry, UN 38.3 + cold-capacity/cycle-life (with HA-ENRG-001/006).
Detectability/Warning | Incoming inspection + supplier qualification (RCR-054) verify the cell; the BMS/fuse backstop (RCR-002) and cold-charge interlock (RCR-001) limit the thermal consequence; the energy-reserve/honest-gap control (RCR-006) limits the endurance consequence. A counterfeit is otherwise invisible after build.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6); single-fault rule applies to the thermal path.
Risk-Control Strategy | 4 (process — incoming verification + supplier control) + 2 (protective — BMS/fuse + cold-charge interlock as independent backstops)
Risk-Control Requirement | RCR-054 — correct cell chemistry (and correct-rated/oriented protection) verified (poka-yoke + AOI); counterfeit/wrong parts prevented. (+ RCR-001 cold-charge interlock; + RCR-002 protection backstop; + RCR-006 endurance reserve; + RCR-007 UN 38.3.)
Implementation Reference | DO-MFG (incoming/supplier control), DO-TST (incoming verification), DO-SAFE-001/002 (interlock/protection) · CMP-006 (cell), CMP-007 (pack), CMP-008/009 (protection)
Verification Reference | PV-MFG-003 (Evidence-Pending) — incoming cell verification (chemistry/authenticity, datasheet traceability); UN 38.3 + cold-capacity on the qualified chemistry (ties HA-ENRG)
Validation Reference | N/A (physical/incoming verification carries the assurance)
Residual Probability | projected Very Low (PVL) after incoming verification + supplier control · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | Ties HA-ENRG-001/002/006 (battery thermal/endurance); carried in HAL-2900 with the D3 battery residuals. FTA top-event contributor; traces to PFMEA incoming/cell-build controls (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending; chemistry EP per CMP-006)
Owner | Manufacturing/Supply-Chain Engineer (with Battery/Power)
Rationale/Evidence | IEC 62133-2; UN 38.3; RCR-054/001/002/006; component-list (CMP-006). Severity per HAL-2000 §7.
Assumptions | Supplier qualification + incoming test can distinguish the qualified chemistry from a counterfeit; the interlock/protection design does not depend on the exact chemistry to remain safe.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-003 — Reversed / wrong-rated protection component built in → protection defeated → thermal event
Hazard ID (library) | HAZ-MFG (assembly error) → HAZ-THERM / HAZ-FIRE
Product/Config | SCM all SKUs · Lifecycle Phase: Assembly (HA-LIFE-002) · System State: Board/pack build · Operational Mode: Manufacture
Perspective | People near the device (depot/hold/clinic); mission/property
Source | HAL-300 (HA-LIFE-002); IEC 62368-1; ties HA-ENRG-002 (protection defeated → runaway); PRD-1500
Initiating Cause/Event | A battery-protection component (BMS/AFE, secondary fuse, or charge-input front-end) is fitted reversed, wrong-value, wrong-rated, or omitted.
Contributing Conditions | Fine-pitch/polarity-sensitive parts; the protection nets are safety-critical (CMP-008/009/014, "Preliminary — reconcile with HAL"); a wrong/reversed protection part can silently defeat the single-fault backstop.
Foreseeable Sequence of Events | Wrong/reversed protection part built in → protection does not act (or acts incorrectly) → on a later cell fault (OV/OC/OT/short) the primary + backstop layers do not terminate the hazard → thermal runaway.
Hazard | A device whose battery protection is defeated by a build error.
Hazardous Situation | A cell fault in a device whose protection cannot terminate it, near people.
Sequence to Harm | Cell fault + defeated protection → thermal runaway → fire/heat/offgas → burn (S4), catastrophic in an occupied hold (S5 via HA-TFC-002).
Harm/Loss | Serious burn / thermal event (S4); catastrophic fire (S5 worst-case, cross-referenced).
Severity | S4 (Critical) — human injury (thermal); S5 credible worst-case via the occupied-hold fire chain (HA-TFC-002, not double-counted here).
Probability (P1×P2) | P1=2 (Remote — AOI + poka-yoke on protection nets; Test — AOI/ICT coverage) × P2=2 (Remote — defeated protection reaches harm only on a coincident cell fault) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + AOI/ICT coverage on safety-critical nets; Medium uncertainty, Evidence-Pending — closed by AOI/ICT coverage review (PV-MFG-004) and protection fault-injection (with HA-ENRG-002).
Detectability/Warning | AOI + ICT + poka-yoke verify correct-rated/oriented protection (RCR-054); protection fault-injection at EOL confirms the backstop acts. A wrong part is invisible until a coincident cell fault.
Initial Risk | AP Medium→High (region A/R — S4 × PL → region A per HAL-2200 matrix S4×PL; treated High for design under the single-fault rule)
Risk Acceptability | Unacceptable without control for the safety-critical protection nets (single-fault rule, HAL-2200 §6.5); reduce via AOI/ICT + fault-injection.
Risk-Control Strategy | 4 (process — AOI/ICT + poka-yoke) + 2 (protective — verified independent backstop)
Risk-Control Requirement | RCR-054 — correct-rated/oriented protection components verified (poka-yoke + AOI); wrong/reversed parts prevented. (+ RCR-002 independent non-resettable backstop, single-fault-safe; + RCR-016 latent-defect screening on safety-critical nets.)
Implementation Reference | DO-MFG (poka-yoke), DO-TST (AOI/ICT + protection fault-injection, DO-TST-009), DO-SAFE-002 · CMP-008 (BMS), CMP-009 (fuse), CMP-014 (input front-end)
Verification Reference | PV-MFG-004 (Evidence-Pending) — AOI/ICT coverage review on protection nets; protection fault-injection at EOL (confirm backstop terminates with a forced-failed primary)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after AOI/ICT + fault-injection verified · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | Ties HA-ENRG-002 (protection→runaway); carried in HAL-2900. Traces to PFMEA SMT/assembly + AOI/ICT controls (Stage 11); FTA top-event contributor.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Manufacturing Engineer (with Battery/Power and Electrical)
Rationale/Evidence | IEC 62368-1; RCR-054/002/016; DO-SAFE-002. Severity per HAL-2000 §7.
Assumptions | AOI/ICT coverage includes polarity/value/rating on all protection nets; fault-injection is part of EOL for a sample or 100% per the PFMEA detection plan.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-004 — Damaged seal at assembly → field ingress → record loss
Hazard ID (library) | HAZ-MFG (assembly damage) → HAZ-ENV (ingress) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Assembly (HA-LIFE-002) · System State: Enclosure build · Operational Mode: Manufacture
Perspective | Beneficiary/patient (via later record fault); mission; ties HA-SVC-003, HA-ENV-003, HA-ELEC-006/007
Source | HAL-300 (HA-LIFE-002); IEC 60529 (IP); PRD-MECH; ties HA-SVC-003 (service-time seal)
Initiating Cause/Event | A seal/gasket (enclosure joint, USB-C port, probe channel) is pinched, contaminated, mis-seated, or nicked during assembly.
Contributing Conditions | Manual/automated seal placement; elastomer sensitivity to contamination and compression; the defect is latent (passes a bench test, fails after field cold-to-warm/humid cycling).
Foreseeable Sequence of Events | Seal damaged at assembly → passes EOL functionally → field cold-to-warm/humid cycling drives ingress → condensation/leakage/corrosion on record-integrity electronics → wrong/lost readings or a monitoring gap → wrong disposition.
Hazard | Moisture ingress reaching record-integrity electronics due to an assembly-damaged seal.
Hazardous Situation | A dose monitored by a device whose record electronics are being corrupted by ingress.
Sequence to Harm | Ingress → corrupted/lost record → false "accept" on a gap/faulted read, or a trip released on an incomplete record → compromised dose administered.
Harm/Loss | Wrong disposition of a limited quantity / trip on an incomplete record (S3).
Severity | S3 (Serious) — public-health, "a monitored trip released on an incomplete record" / limited-quantity wrong disposition (HAL-2000 §5). (Escalates via HA-SW-002/HA-DATA-004 if it produces a silent false accept.)
Probability (P1×P2) | P1=2 (Remote — seal inspection + IP sampling at EOL) × P2=3 (Occasional — ingress → record fault → wrong disposition, partly caught by honest-gap logic) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + IEC 60529; Medium uncertainty, Evidence-Pending — closed by EOL IP sampling + seal-inspection audit (PV-MFG-005).
Detectability/Warning | Seal inspection + EOL IP sampling (RCR-054/011); honest-gap and verdict-integrity logic (RCR-022/021) prevent a resulting fault from reading as a clean pass. A pinched seal is best caught by inspection/design.
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce via seal-inspection/IP sampling + robust seal design.
Risk-Control Strategy | 4 (process — seal inspection + IP sampling) + 1/2 (inherent IP design + verdict/gap integrity backstop)
Risk-Control Requirement | RCR-011 — enclosure + seals maintain ingress rating and cell containment through drop/vibration/transport and reuse life. (+ RCR-054 correct assembly verification incl. seal integrity; + RCR-022/021 honest gap / verdict integrity.)
Implementation Reference | DO-MFG (seal placement + inspection), DO-TST (EOL IP sampling), DO-SAFE-010/011 · CMP-002 (seals), CMP-004/005 (port/probe channel), CMP-034 (PCB/coat)
Verification Reference | PV-MFG-005 (Evidence-Pending) — EOL IP sampling + seal-inspection audit; assembly seal-integrity over reuse-cycle test (shared with HA-SVC-003)
Validation Reference | N/A
Residual Probability | projected Low (PL) after IP sampling + seal design verified · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B, S3×PL) — ALARP
New Risks Introduced | None (shares seal control with HA-SVC-003 / HA-ENV-003).
Overall Residual Contribution | Small; ties HA-SVC-003 (service seal), HA-ELEC-006/007 (ingress→electronics). Carried in HAL-2900; traces to PFMEA enclosure/seal station controls (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Manufacturing Engineer (with Mechanical)
Rationale/Evidence | IEC 60529; RCR-011/054; DO-SAFE-011. Severity per HAL-2000 §5.
Assumptions | EOL IP sampling rate is set by the PFMEA detection plan; seal design tolerates the assembly handling.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-005 — Wrong crypto identity / duplicate serial provisioned → mis-routing / record repudiation
Hazard ID (library) | HAZ-MFG (identity provisioning error) → HAZ-CYBER / information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: EOL provisioning (HA-LIFE-003) · System State: Identity/crypto birth · Operational Mode: Provisioning
Perspective | Beneficiary/patient (via mis-routed alert / repudiated record); mission; organization; ties HA-CLOUD-003
Source | HAL-300 (HA-LIFE-003); IEC 62443 (unique identity); ties HA-CLOUD-003 (device↔trip↔owner binding); PRD-SEC; RCR-055
Initiating Cause/Event | A device is provisioned with a wrong, duplicated, or non-unique cryptographic identity/serial at EOL.
Contributing Conditions | Automated crypto-identity provisioning; a duplicate/wrong identity breaks the authoritative device↔trip↔owner↔responder binding; identity is trusted downstream for routing and record attribution.
Foreseeable Sequence of Events | Wrong/duplicate identity provisioned → in the field the cloud binds the device to the wrong trip/owner → an excursion alert is routed to the wrong responder (or not at all), or a record cannot be attributed/authenticated → the true responder does not act, or the record is repudiated → wrong disposition or lost evidence.
Hazard | A device with a wrong/duplicate cryptographic identity.
Hazardous Situation | An excursion alert routed to the wrong party, or a record that cannot be trusted/attributed, for a monitored dose.
Sequence to Harm | Mis-routing/repudiation → no timely action on an excursion → compromised dose administered; or a lot released on a repudiated record.
Harm/Loss | Mis-routed alert → compromised dose administered (S4) / repudiated record → GDP data-integrity deviation (S3).
Severity | S3/S4 — take S4 (Critical) as the credible worst-case (mis-routed alert → wrong disposition, HAL-2000 §6); the record-repudiation path anchors at S3 (audit/data-integrity deviation).
Probability (P1×P2) | P1=2 (Remote — identity-uniqueness check at provisioning; Test) × P2=3 (Occasional — wrong identity → mis-routing/repudiation → wrong disposition) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + provisioning controls; Medium uncertainty, Evidence-Pending — closed by identity-uniqueness check (PV-MFG-006) and association/routing test (with HA-CLOUD-003).
Detectability/Warning | Identity-uniqueness check at provisioning (RCR-055); the cloud registry enforces authoritative binding and rejects duplicate identities (RCR-033). A duplicate would otherwise surface only as a field routing/attribution failure.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — unique-identity provisioning + check) + 2 (protective — cloud registry rejects duplicates / enforces binding)
Risk-Control Requirement | RCR-055 — each device carries a unique cryptographic identity and serial; duplicate identity/serial prevented. (+ RCR-033 authoritative device↔trip↔owner↔responder binding; wrong-tenant binding prevented.)
Implementation Reference | DO-MFG (provisioning), DO-SEC (identity/keys), DO-CLD (registry) · CMP-016 (secure element/identity), CMP-047 (device registry)
Verification Reference | PV-MFG-006 (Evidence-Pending) — identity-uniqueness check (no duplicate/wrong identity ships); association/routing test (with HA-CLOUD-003)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after uniqueness check + registry enforcement verified · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | Ties HA-CLOUD-003 (binding/routing) and HA-DATA-002; carried in HAL-2900; traces to PFMEA provisioning-station controls (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Manufacturing Engineer (with Product Security EM-ENG-008 and Cloud)
Rationale/Evidence | IEC 62443; RCR-055/033. Severity per HAL-2000 §6.
Assumptions | Crypto identities are generated/injected uniquely and the registry is the single authoritative binding source.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-006 — Wrong regional / firmware configuration at provisioning → coverage or logic error in market
Hazard ID (library) | HAZ-MFG (config provisioning error) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: EOL provisioning (HA-LIFE-003) · System State: Config/firmware birth · Operational Mode: Provisioning
Perspective | Beneficiary/patient; mission; ties HA-COMM-006, HA-SVC-002
Source | HAL-300 (HA-LIFE-003); IEC 62304 (config/software); ties HA-COMM-006 (band/coverage) and HA-SVC-002 (depot config); PRD-1500
Initiating Cause/Event | A device is provisioned at EOL with a wrong regional configuration (band plan, market profile, carrier/SIM) or a wrong firmware baseline for its market.
Contributing Conditions | Multiple market variants; SIM/eSIM and band strategy Evidence-Pending (CMP-027); a wrong band plan means the device cannot connect in its market; a wrong market profile means wrong stability logic.
Foreseeable Sequence of Events | Wrong region/firmware provisioned → device deployed in that market → cannot connect on its bands (alert coverage lost) or applies wrong stability logic → missed/delayed/incorrect excursion decision.
Hazard | A device configured for the wrong market/logic.
Hazardous Situation | A dose monitored by a device that cannot deliver its alert, or applies the wrong rules, in-market.
Sequence to Harm | Coverage/logic error → missed excursion or undelivered alert → false "accept" or no time to act → compromised dose administered.
Harm/Loss | A lot of compromised doses / undelivered alerts (S4).
Severity | S4 (Critical) — public-health (wrong disposition / coverage loss), HAL-2000 §6.
Probability (P1×P2) | P1=2 (Remote — config verification at EOL; Test — config-verification) × P2=4 (Probable — wrong config reliably yields wrong/undelivered decisions in market) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + config verification; Medium uncertainty, Evidence-Pending — closed by config-verification test (PV-MFG-007) and pre-arm valid-profile gate (PV-MFG-002).
Detectability/Warning | Config verified at EOL (RCR-056); the pre-arm gate blocks arming without a valid firmware/profile (RCR-045). Wrong config otherwise surfaces only as a field coverage/logic failure.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — config provisioning + verification) backstopped by 1/2 (pre-arm valid-profile gate)
Risk-Control Requirement | RCR-056 — correct regional/firmware configuration provisioned and verified at EOL (band plan, market profile). (+ RCR-045 pre-arm valid-firmware/profile gate; + RCR-037 signed firmware baseline.)
Implementation Reference | DO-MFG (config provisioning), DO-FW (config verification), DO-SAFE-007 (pre-arm gate) · CMP-027 (SIM/band), CMP-035 (bootloader/firmware baseline), CMP-045 (pre-arm)
Verification Reference | PV-MFG-007 (Evidence-Pending) — config-verification test (band plan + market profile correct for SKU/market); pre-arm valid-profile gate test
Validation Reference | VAL-MFG-006 (Evidence-Pending) — market-representative connectivity/logic validation (with HA-COMM-006)
Residual Probability | projected Low (PL) after config verification + pre-arm gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | A blocked/void device on a config mismatch (availability, not safety).
Overall Residual Contribution | Ties HA-COMM-006 (band/coverage) and HA-SVC-002 (depot config); carried in HAL-2900; traces to PFMEA provisioning/config-station controls (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending; SIM/band strategy EP per CMP-027)
Owner | Manufacturing Engineer (with Firmware and Connectivity)
Rationale/Evidence | IEC 62304; RCR-056/045; component-list (CMP-027). Severity per HAL-2000 §6.
Assumptions | Market/band configuration is part of the verified provisioning set; the pre-arm gate reads a trustworthy profile/market tag.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-007 — Inadequate EOL test coverage → latent safety-critical defect escapes to field
Hazard ID (library) | HAZ-MFG (test-coverage gap) → information-integrity harm (or thermal, per defect)
Product/Config | SCM all SKUs · Lifecycle Phase: EOL test (HA-LIFE-003) · System State: End-of-line test · Operational Mode: Final test
Perspective | Beneficiary/patient; people near the device (if thermal defect); mission
Source | HAL-300 (HA-LIFE-003); ISO 13485 (verification of process output); ties HA-MFG-008 (latent solder) and HA-ELEC-008; PRD-1500; RCR-016/053
Initiating Cause/Event | The EOL test suite does not exercise a safety-critical function or net (e.g., probe-fault handling, verdict-gap suppression, protection backstop, calibration coverage), so a latent defect passes.
Contributing Conditions | Test-time/cost pressure; safety-critical functions (detection, verdict, probe-fault, protection) require deliberate fault-injection coverage, not just a functional smoke test; the escaped defect is latent.
Foreseeable Sequence of Events | Coverage gap → a device with a latent safety-critical defect passes EOL → in the field the defect manifests (missed excursion, false accept, or thermal) → wrong disposition or thermal event.
Hazard | A latent safety-critical defect that EOL did not screen.
Hazardous Situation | A fielded device with an unscreened safety-critical defect monitoring a dose.
Sequence to Harm | Latent defect manifests → missed/incorrect excursion → false accept → compromised dose administered (or thermal event).
Harm/Loss | A lot of compromised doses administered / latent failure escapes (S4).
Severity | S4 (Critical) — public-health (wrong disposition), HAL-2000 §6; can be S5 if the unscreened defect is a fleet-scale calibration/protection escape (cross-referenced to HA-MFG-001/003).
Probability (P1×P2) | P1=2 (Remote — defined EOL coverage requirement + fault-injection on safety-critical functions) × P2=4 (Probable — an unscreened safety-critical defect reliably reaches a wrong disposition) ⇒ product 8 ⇒ Medium band (PM) — Basis: judgment + coverage practice; Medium uncertainty, Evidence-Pending — closed by EOL coverage review vs the safety-critical function list (PV-MFG-008).
Detectability/Warning | The control is coverage itself — EOL test coverage must include the safety-critical functions/nets (RCR-016/053); field plausibility/verdict-integrity (RCR-020/021) is a partial backstop for measurement defects.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — defined EOL coverage incl. fault-injection on safety-critical functions) + 2 (protective — field plausibility/verdict backstop)
Risk-Control Requirement | RCR-016 — latent solder/PCB/component defects screened by AOI + EOL test coverage; no latent single-point failure escapes on safety-critical nets. (+ RCR-053 EOL calibration coverage; + RCR-020/021 field probe-fault/verdict backstops.)
Implementation Reference | DO-TST (EOL coverage plan + fault-injection), DO-PCB (AOI) · CMP-019/020/038/039/040 (safety-critical functions), CMP-008/009 (protection), CMP-034 (PCB)
Verification Reference | PV-MFG-008 (Evidence-Pending) — EOL test-coverage review against the safety-critical function/net list (detection, verdict, probe-fault, protection, calibration); fault-injection coverage audit
Validation Reference | N/A
Residual Probability | projected Low (PL) after coverage review closes gaps · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Added EOL coverage increases test time/cost (not safety).
Overall Residual Contribution | Umbrella coverage control for HA-MFG-001/003/008 and HA-ELEC-008; carried in HAL-2900; the EOL coverage plan is a primary PFMEA detection input (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Manufacturing Engineer (with Quality, Firmware, Electrical)
Rationale/Evidence | ISO 13485 verification of output; RCR-016/053; DO-TST. Severity per HAL-2000 §6.
Assumptions | A safety-critical function/net list exists (from this HAL + DOL-1200) that EOL coverage is audited against; fault-injection is feasible in the EOL fixture.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MFG-008 — Latent solder / PCB damage (thermal or mechanical assembly stress) → intermittent field failure → silent monitoring gap
Hazard ID (library) | HAZ-MFG (workmanship/latent defect) → HAZ-ELEC → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Assembly (HA-LIFE-002) · System State: Board build (reflow/press-fit/handling) · Operational Mode: Manufacture
Perspective | Beneficiary/patient; mission; ties HA-ELEC-008
Source | HAL-300 (HA-LIFE-002); IPC-A-610 / IPC-6012 (workmanship/PCB); ties HA-ELEC-008 (intermittent → silent gap); PRD-1500; RCR-016
Initiating Cause/Event | A solder joint, via, or component is damaged or marginal from thermal (reflow) or mechanical (press-fit, depanel, handling) assembly stress, creating a latent/intermittent defect on a safety-critical net.
Contributing Conditions | −30↔+60 °C field thermal cycling + transport vibration exercise a marginal joint into intermittent open/short; the defect passes a static bench test but fails intermittently in the field; the affected net may be a sensor/verdict/record path.
Foreseeable Sequence of Events | Marginal joint/via built in → passes EOL statically → field thermal/vibration cycling → intermittent open/short on a safety-critical net → intermittent wrong reading or a silent monitoring gap → excursion missed → wrong disposition.
Hazard | A latent/intermittent solder/PCB defect on a safety-critical net.
Hazardous Situation | A fielded device intermittently mis-reading or not logging while monitoring a dose.
Sequence to Harm | Intermittent fault → missed/wrong reading or silent gap → false accept → compromised dose administered → beneficiary not protected.
Harm/Loss | A lot of compromised doses administered / silent gap (S4).
Severity | S4 (Critical) — public-health (wrong disposition via silent gap), HAL-2000 §6/§7 ("lost record ... silent = S4").
Probability (P1×P2) | P1=2 (Remote — AOI + ICT + process control; Test — AOI/ICT coverage) × P2=3→4 (Occasional→Probable — an intermittent gap on a safety-critical net reads as uneventful, partly caught by honest-gap/self-check) ⇒ product 6–8 ⇒ Medium band (PM) — Basis: judgment + IPC workmanship + AOI/ICT coverage; Medium uncertainty, Evidence-Pending — closed by AOI/ICT coverage review + assembly-stress screen (PV-MFG-009) and thermal-cycle/vibration reliability (with HA-ELEC-008).
Detectability/Warning | AOI + ICT screen static defects (RCR-016); an availability self-check flags a device "not monitoring" on a sensing-chain outage (RCR-027), and honest-gap/verdict integrity (RCR-022/021) stop a resulting gap from reading as a clean pass. Intermittent defects are the hard-to-catch residual.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 4 (process — AOI/ICT + assembly process control) + 2 (protective — self-check + honest-gap field backstop)
Risk-Control Requirement | RCR-016 — latent solder/PCB/component defects screened by AOI + EOL test coverage; no latent single-point failure escapes on safety-critical nets. (+ RCR-027 availability self-check flags "not monitoring"; + RCR-022/021 honest gap / verdict integrity.)
Implementation Reference | DO-TST (AOI/ICT + assembly-stress screen), DO-PCB (workmanship/stack-up) · CMP-034 (PCB/solder), CMP-045 (self-check), CMP-041 (storage/gap)
Verification Reference | PV-MFG-009 (Evidence-Pending) — AOI/ICT coverage review + post-assembly stress screen (thermal-cycle/vibration on safety-critical nets); intermittent-fault detection via self-check (with HA-ELEC-008)
Validation Reference | N/A
Residual Probability | projected Low (PL) after AOI/ICT + stress screen + self-check verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | Added stress-screen can precipitate infant-mortality removals (yield, not safety).
Overall Residual Contribution | Ties HA-ELEC-008 (intermittent → silent gap); carried in HAL-2900; traces to PFMEA SMT/reflow/assembly + AOI/ICT controls (Stage 11); FTA top-event contributor.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Manufacturing Engineer (with Electrical)
Rationale/Evidence | IPC-A-610 / IPC-6012; RCR-016/027; DO-TST/DO-PCB. Severity per HAL-2000 §6/§7.
Assumptions | AOI/ICT + optional stress screen achieve the coverage set by the PFMEA on safety-critical nets; the self-check can detect a sensing-chain outage but not every intermittent misread.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, severity anchor §6/§7), HAL-2100 (P1×P2, basis taxonomy, process-escape probabilities reduced by inspection/coverage — incoming, AOI/ICT, poka-yoke, EOL calibration coverage), and HAL-2200 (matrix, regions R/A/B, AP, severity-first + uncertain-probability + single-fault rules).
Controls | RCR-053/026/049 (calibration), RCR-054/001/002 (cell/protection), RCR-011 (seal), RCR-055/033 (identity/binding), RCR-056/045 (regional/firmware config + pre-arm), RCR-016 (latent-defect screening) — all from HAL-2400.
Implementation | DO-MFG / DO-TST are the primary manufacturing Implementation References, backstopped by DO-SAFE-001/002/007/010/011/014, DO-PCB, DO-SEC, DO-CLD; each points to CMP-002/004/005/006/007/008/009/014/ 016/019/020/027/034/035/038/039/040/045/047.
Cross-module ties | HA-MFG-001 ↔ HA-SVC-001 (EOL vs service calibration); HA-MFG-002 ↔ HA-ENRG-001/002/006 (battery thermal/endurance); HA-MFG-003 ↔ HA-ENRG-002 (protection defeated); HA-MFG-004 ↔ HA-SVC-003 / HA-ENV-003 / HA-ELEC-006/007 (ingress); HA-MFG-005 ↔ HA-CLOUD-003 / HA-DATA-002 (identity/binding); HA-MFG-006 ↔ HA-COMM-006 / HA-SVC-002 (region/firmware); HA-MFG-008 ↔ HA-ELEC-008 (intermittent → silent gap).
Downstream — PFMEA hand-off (Stage 11) | every record traces forward into the PFMEA and production controls (Stage 11, ERFKL) — the PFMEA quantifies the process failure mode, occurrence, and station detection named here (incoming, SMT/reflow/AOI/ICT, seal station, calibration station, provisioning station, EOL coverage). The S4/S5 records (HA-MFG-001/002/003/005/006/007/008) become FTA top events; all PV-/VAL- references are Evidence-Pending and feed Stages 13–14.
8. Residual-risk / closure notes
Module headline | HA-MFG-001 (EOL calibration escape) is the carried manufacturing headline — driven by severity while the EOL calibration-coverage probability is Evidence-Pending (PV-MFG-001/002); paired in HAL-2900 with HA-SVC-001 as the calibration-integrity residual.
Battery/protection build set | HA-MFG-002/003 are single-fault-governed (HAL-2200 §6.5) — the independent protective backstop (RCR-002) and cold-charge interlock (RCR-001) must survive a wrong/reversed part; chemistry remains Evidence-Pending (CMP-006), tying the D3 battery residuals.
Coverage umbrella | HA-MFG-007 (EOL coverage) and RCR-016 govern HA-MFG-001/003/008 and HA-ELEC-008 — the EOL coverage plan and AOI/ICT coverage are the load-bearing detection controls and the primary PFMEA detection inputs (Stage 11).
No inspection-only acceptance | no manufacturing record is accepted on a single inspection (HAL-2200 §6.4); each process control is paired with a design/field backstop (pre-arm gate, protection backstop, self-check, verdict/gap integrity).
Release gate (HAL-2200 §8) | no R region manufacturing risk may remain open at release; every remaining region-A S4/S5 record carries a benefit-risk statement and residual-risk sign-off (EM-ENG-006 + EM-ENG-001; EM-ENG-008 for HA-MFG-005); every Evidence-Pending probability has a named closing test above.
Overall contribution | the manufacturing module's residual is dominated by the calibration escape (HA-MFG-001), the battery/protection build escapes (HA-MFG-002/003), and the latent workmanship/coverage set (HA-MFG-007/008); these are principal manufacturing inputs to HAL-2900 and the Risk-Management Summary, and the whole module is the harm/exposure precursor to the Stage-11 PFMEA.
HAL-1800 (HA-MPE) — Mission, Property & Environmental Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1800, record code HA-MPE · Owner: Principal Mission-Safety Engineer, with System-Safety (lead), Quality/Regulatory (EM-ENG-006), Chief Systems (EM-ENG-001), and the Service/Environmental engineers. This module takes the aggregate / consequence view: it does not re-derive device-level faults (those live in HAL-500…HAL-1700) but analyses the mission-, property-, and environmental-scale outcomes those faults roll up to — the vaccine lot, the immunisation programme, the fleet, and the environment. It follows the ISO 14971 harm-and-exposure-first method (HAL-000 §4); it is not an FMEA. Each record is a complete risk chain scored severity-first per CLAUDE.md rule 3. The module contains the library's ultimate population-scale top event, HA-MPE-002. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control the aggregate consequences — at the scale of the vaccine lot/shipment, the immunisation programme, the reusable fleet, and the environment — that the SCM's device-level information and energy hazards can produce. The SCM is an information device: its dominant harm is a wrong disposition decision, and at scale that decision is not one dose but a lot, a campaign, or a population. This module is the consequence aggregation of the whole library: it is where the S4 device-level records converge into the S5 population-scale outcome (HA-MPE-002) — the ultimate FTA top event of the entire hazard analysis.
2. Scope
All SCM SKUs, at the mission/property/environmental scale (HA-CTX-006/007 perspectives: organization/programme, mission, environment). In-scope consequence classes: vaccine-lot loss (a valid lot destroyed), cascading public-health harm (mass administration of compromised doses), false-reject stock-out (mass under-vaccination), regulatory/GDP data-integrity violation at scale, environmental release (Li-battery/e-waste at fleet scale), and loss of programme trust / essential immunisation service. In-scope aggregate perspectives: the vaccine lot/shipment (property + public-health), the immunisation programme (mission), the reusable device fleet (property + environment), and the environment. Out of scope / owned elsewhere: the device-level initiating faults (missed excursion → HAL-1000; delayed alert → HAL-1100; cloud drop → HAL-1200; cyber → HAL-1300; use-error → HAL-1500; calibration/config escapes → HAL-1600/1700; battery/thermal → HAL-500/800) — this module aggregates them, it does not replace them. Cross-references name the contributing device-level records.
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); overall residual risk & benefit-risk (HAL-2900); population/scale captured explicitly at the S4/S5 boundary.
• WHO PQS E006 / WHO-UNICEF Effective Vaccine Management (EVM) — public-health consequence of potency loss, freezing, and stock-out; programme-level cold-chain integrity.
• ISO 13485 / WHO GDP — data-integrity/record obligations at scale; ALCOA+; certification consequences.
• IEC 62443 / AAMI TIR57 — where a security event scales to a fleet/programme consequence (via HAL-1300).
• UN 38.3 / IATA DGR + local e-waste/battery regulation — fleet-scale lithium/e-waste environmental release.
Upstream | HA-CTX-006/007 (perspectives incl. programme/mission/environment), HAL-2000 §6/§7 (S4/S5 disposition anchor and population-scale S5), HAL-2400 (RCR register incl. the overarching RCR-057/058), DOL-1200, and every contributing domain register (HAL-500…HAL-1700).
Evidence honesty | outbreak, stock-out, and campaign consequences are outcome-severity statements (population-scale S5 is credible per WHO EVM), not probability claims; each probability records its basis and uncertainty and is Evidence-Pending where the contributing device-level tests are Evidence-Pending. No outbreak/stock-out figure is invented.
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
Aggregate method | each HA-MPE record's probability is inherited from the aggregate of its contributing device-level records (the whole record-integrity chain), not a new bottom-up estimate. It is therefore held Evidence-Pending until the contributing device-level tests close, and driven by severity (HAL-2200 §6.2).
• Severity is fixed by credible worst-case outcome (HAL-2000). Mission records anchor at S4 (a lot/shipment) with S5 credible worst-case (population/campaign scale); HA-MPE-002 anchors at S5 (population-scale) as the top consequence of the library.
Catastrophic-severity rule (HAL-2200 §6.1) | the S5 mission records (HA-MPE-002, and the campaign-scale face of HA-MPE-003) can never be classified broadly acceptable; they stay region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off and post-market monitoring, regardless of how low the aggregate probability.
Control strategy | the aggregate consequence is controlled by the overarching record-integrity control RCR-057 (the whole detection→verdict→record→export chain preserves mission evidence) plus the specific contributing controls; the release-level control is RCR-058 (overall residual & benefit-risk evaluated before release). Warnings-only control of a mission consequence is not accepted (HAL-2200 §6.4).
• Residual risk is re-estimated only after the contributing controls are verified; residuals below are projected, Evidence-Pending until the contributing device-level tests close.
5. Categories & guiding questions (mission/property/environmental prompt list)
1. Lot integrity (property + public-health): Can a valid lot be destroyed by a wrong reject or a lost record, or a compromised lot be released by a wrong accept?
2. Programme/population (mission): Can compromised doses be administered at programme scale → outbreak? Can false rejects cascade into a stock-out → mass under-vaccination?
3. Regulatory/data integrity (mission): Can a data-integrity failure at scale trigger programme certification loss, recall, or hold?
4. Environment (fleet scale): Can end-of-life lithium/e-waste be released at fleet scale?
5. Trust (mission): Can loss of confidence in the record/monitoring erode the essential immunisation service itself?
6. Hazard Analysis Register (HA-MPE records)
HA-MPE-001 — Vaccine-lot loss (mission/property) from a wrong reject or a lost record → destroyed valid product
Hazard ID (library) | HAZ-MISS / HAZ-PROP (aggregate lot-loss consequence) → information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt / disposition (HA-LIFE-011), aggregated over trips · System State: Disposition decision · Operational Mode: Accept/reject/quarantine
Perspective | Mission (the vaccine shipment) and property (high-value lot); organization/programme
Source | HA-CTX-007 (mission/property); HAL-2000 §5 (loss of a full shipment/lot = S4 property; false reject → destroyed valid product); aggregates HA-SW-002 (false verdict), HA-DATA-004/006 (lost/incomplete record), HA-USE-007 (override), HA-ENRG-006 (silent gap)
Initiating Cause/Event | A device or record failure causes a wrong reject of a valid lot, or a lost/incomplete record that forces precautionary destruction of a lot that was in fact potent.
Contributing Conditions | High-value lots; a lost record or an unqualified gap is (correctly) treated conservatively → the valid lot is discarded; the aggregate of every false-reject/lost-record device-level path.
Foreseeable Sequence of Events | Device/record fault (false reject or lost record) → disposition cannot confirm the lot is good → lot quarantined/destroyed → valid, potent vaccine is wasted at lot scale.
Hazard | Loss of trustworthy disposition evidence for a valuable lot.
Hazardous Situation | A potent lot facing destruction because its integrity cannot be confirmed (or was wrongly rejected).
Sequence to Harm | Wrong reject / lost record → valid lot destroyed → property loss + downstream supply reduction (feeds HA-MPE-003 stock-out if it cascades).
Harm/Loss | Loss of a full shipment/lot of vaccine; product wastage (S4 property).
Severity | S4 (Critical) — governing dimension property ("loss of a full shipment/lot of vaccine", HAL-2000 §5); rises toward the public-health face if it cascades to stock-out (HA-MPE-003).
Probability (P1×P2) | aggregate P1=2 (Remote — the record-integrity chain suppresses false-reject/lost-record paths) × P2=3 (Occasional — a lost/unqualified record conservatively drives lot destruction) ⇒ product 6 ⇒ Medium band (PM) — Basis: aggregate of contributing device-level records; High uncertainty, Evidence-Pending — closed when the contributing tests (verdict/gap, record-integrity, endurance) close.
Detectability/Warning | The record-integrity chain (RCR-057) and export-integrity (RCR-042) make gaps visible and preserve evidence so a recoverable record avoids needless destruction; the honest-gap/verdict controls (RCR-021/022) reduce both false-reject and lost-record destruction.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 1/2 (inherent/protective — record-integrity chain preserves recoverable evidence) + 4 (process — disposition policy uses the preserved record)
Risk-Control Requirement | RCR-057 — the record-integrity chain (detection→verdict→record→export) preserves mission evidence such that a wrong disposition/lot-loss cannot arise from data loss or corruption. (+ RCR-021 verdict suppresses false accept/reject on gaps; + RCR-042 export preserves attribution/time/gaps.)
Implementation Reference | DO-SAFE-004…008 (detection/verdict/record chain), DO-CLD (export/API) · CMP-039/040/041/049 (detection/verdict/record), CMP-052 (export)
Verification Reference | PV-MPE-001 (Evidence-Pending) — record-integrity chain traceability + integration test (a recoverable record prevents needless lot destruction); inherits verdict/gap and record-integrity PVs
Validation Reference | VAL-MPE-001 (Evidence-Pending) — disposition-policy validation that preserved records are used before a lot is destroyed
Residual Probability | projected Low (PL) after the record-integrity chain is verified · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-006 + EM-ENG-001 benefit-risk sign-off
New Risks Introduced | Over-conservative disposition (destroy-on-any-gap) trades lot loss against beneficiary safety — a benefit-risk tension carried in HAL-2900.
Overall Residual Contribution | The property/mission face of the disposition-integrity risk; feeds HAL-2900 and pairs with HA-MPE-003 (stock-out) when lot loss cascades. FTA top-event contributor (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending — inherits contributing tests)
Owner | Mission-Safety Engineer (with Quality/Regulatory)
Rationale/Evidence | HAL-2000 §5 (lot-loss property severity); RCR-057/021/042; aggregates HA-SW-002/HA-DATA-004/006/HA-USE-007/HA-ENRG-006.
Assumptions | Disposition policy prefers a recoverable record over destruction; lot value and wastage cost bounded by programme data.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MPE-002 — Cascading public-health harm: outbreak from mass administration of compromised doses across a programme ★ TOP MISSION CONSEQUENCE (population-scale S5 — ultimate library top event)
Hazard ID (library) | HAZ-MISS (population-scale public-health consequence) → the terminal information-integrity harm
Product/Config | SCM all SKUs · Lifecycle Phase: Aggregated across normal/degraded/fault operation → receipt/disposition (HA-LIFE-008…011) · System State: Programme-scale disposition · Operational Mode: Fleet in service
Perspective | Population of beneficiaries (the ultimate harmed party, at scale); mission (the immunisation programme); public infrastructure
Source | HAL-2000 §5/§6/§7 (population-scale S5; wrong-accept anchor); WHO-UNICEF EVM (potency loss → non-protection → outbreak); aggregates every S4 "compromised dose administered" record — HA-SW-001/002, HA-DATA-001, HA-COMM-001/002, HA-CLOUD-001, HA-CYBER-001/002, HA-USE-007, HA-ENRG-006, HA-SVC-001, HA-MFG-001
Initiating Cause/Event | A systematic or fleet-scale failure of the monitoring/disposition chain (a common-mode calibration/config/firmware/cyber/endurance failure, or a systemic false-accept) causes many lots of compromised doses to be released and administered across a programme.
Contributing Conditions | The SCM is used at programme scale; a common-mode escape (miscalibration lot, wrong firmware campaign, cyber campaign, or a systematic false-accept logic defect) affects many devices/lots at once; downstream receipt QA is an imperfect, non-independent check; population immunity depends on doses being potent.
Foreseeable Sequence of Events | Common-mode failure across the fleet → many excursions undetected/false-accepted → many compromised lots released → doses administered across a population → recipients not protected → immunity gap → disease outbreak in the covered population.
Hazard | Systematic failure of the population-scale cold-chain disposition evidence.
Hazardous Situation | A population receiving doses believed protective that are potency-lost.
Sequence to Harm | Mass administration of compromised doses → population not protected → preventable-disease outbreak / vaccine-derived harm at population scale.
Harm/Loss | Population-scale preventable disease — outbreak across many beneficiaries; loss of programme-level public-health protection (S5).
Severity | S5 (Catastrophic) — governing dimension public-health at population scale ("mass administration of potency-lost vaccine across many beneficiaries → outbreak", HAL-2000 §5 S5). This is the ultimate top event of the entire hazard-analysis library.
Probability (P1×P2) | aggregate P1=2 (Remote — requires a common-mode/systematic escape past the whole record-integrity chain and its cross-checks) × P2=3 (Occasional — a systematic false-accept propagates to administration; receipt QA is imperfect) ⇒ product 6 ⇒ Medium band (PM), escalated conservatively — Basis: aggregate of all S4 "compromised dose" records + common-mode analysis; High uncertainty, Evidence-Pending — driven by severity per the catastrophic-severity rule regardless of probability; closes only when the contributing device-level tests and the common-mode/independence analysis close.
Detectability/Warning | No single downstream check reliably catches a systematic, fleet-wide false-accept — which is precisely why the control is prevention of common-mode failure across the record-integrity chain (RCR-057), independent device↔cloud verdict cross-check (RCR-021), calibration/config integrity at scale (RCR-026/053/056), cyber integrity (RCR-034/035), and overall residual/benefit-risk governance (RCR-058).
Initial Risk | AP High (region R — S5 × PM; catastrophic-severity rule forces High regardless of probability, HAL-2200 §6.1)
Risk Acceptability | Unacceptable without control; S5 can never be broadly acceptable (HAL-2200 §6.1). Requires the maximum practicable risk reduction, documented benefit-risk, and post-market monitoring.
Risk-Control Strategy | 1/2 (inherent/protective — record-integrity chain + independent verdict cross-check + calibration/config/cyber integrity, engineered against common-mode) + 4 (process — overall residual & benefit-risk governance, post-market surveillance)
Risk-Control Requirement | RCR-057 — the record-integrity chain preserves mission evidence such that a wrong disposition cannot arise from data loss/corruption (aggregate of RCR-019…024/042). (+ RCR-019 two-sided detection, + RCR-034 telemetry authenticity, + RCR-058 overall residual/benefit-risk evaluated before release; + the calibration/config integrity set RCR-026/053/056 against common-mode escapes.)
Implementation Reference | DO-SAFE-004…008 (detection/verdict/record), DO-CLD-006 (independent cloud recompute), DO-SEC (telemetry integrity), Risk-Management File (RCR-058) · CMP-039/040/041/049 (chain), CMP-016/044 (integrity)
Verification Reference | PV-MPE-002 (Evidence-Pending) — end-to-end record-integrity + device↔cloud verdict-consistency integration test; common-mode / independence analysis across the fleet-scale controls (calibration, config, firmware, cyber)
Validation Reference | VAL-MPE-002 (Evidence-Pending) — programme-scale benefit-risk validation + post-market surveillance plan (ISO 14971) demonstrating a systematic false-accept is detectable and bounded
Residual Probability | projected Low (PL) after the record-integrity chain, independent verdict cross-check, and common-mode controls are verified · Residual Severity: S5 (unchanged) · Residual Risk: AP High→Medium (region A, S5×PL) — per §6.1 an S5 residual stays region A with EM-ENG-001 + EM-ENG-006 sign-off, documented benefit-risk, and mandatory post-market monitoring; never broadly acceptable
New Risks Introduced | Over-conservative fleet-wide reject to avoid false-accept could cascade into stock-out (HA-MPE-003) — the central benefit-risk tension of the programme, carried in HAL-2900.
Overall Residual Contribution | The ultimate population-scale top event of the whole library — the terminal S5 consequence into which every S4 "compromised dose" record converges. Mandatory benefit-risk line in HAL-2900 and the Risk-Management Summary; the top FTA event of Stage 11 (ERFKL), with independence/common-mode analysis as its principal deepening.
Status | Open / Evidence-Pending (S5; driven by severity; closes only with the contributing tests + common-mode analysis + post-market plan)
Owner | Mission-Safety Engineer (with EM-ENG-001 Chief Systems and EM-ENG-006 Quality/Regulatory)
Rationale/Evidence | HAL-2000 §5/§6/§7 (population-scale S5 anchor); WHO-UNICEF EVM; RCR-057/019/034/058; aggregates all S4 "compromised dose administered" records. Population-scale outbreak is an outcome-severity statement, not an asserted probability.
Assumptions | Programme-scale deployment; receipt QA is a contributing (external, imperfect) check, not an independent controlled barrier; population immunity depends on dose potency.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MPE-003 — False-reject stock-out → unvaccinated population → preventable disease
Hazard ID (library) | HAZ-MISS (public-health consequence, under-supply direction)
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt/disposition aggregated → supply (HA-LIFE-011) · System State: Programme-scale disposition · Operational Mode: Fleet in service
Perspective | Population of beneficiaries (unvaccinated); mission (campaign); organization
Source | HAL-2000 §5/§7 (false reject → stock-out → unvaccinated, S4; S5 campaign-scale); WHO EVM (stock-out during outbreak); aggregates HA-USE-009 (charge-inhibit misread → discard), HA-SW-006 (alarm fatigue/false alarm), HA-MPE-001 (lot loss)
Initiating Cause/Event | A systematic false-reject tendency (over-sensitive detection, false alarms, or wrongly discarded devices/lots) destroys or withholds potent vaccine at scale.
Contributing Conditions | Conservative disposition + a false-reject-prone chain; a fleet-scale false-alarm or a mass discard (e.g., cold charge-inhibited devices mistaken for dead) removes potent supply; stock-out during an outbreak leaves a population unprotected.
Foreseeable Sequence of Events | Systematic false reject / mass discard → potent lots destroyed/withheld → local or campaign stock-out → population goes unvaccinated → preventable disease among the unprotected.
Hazard | Systematic loss of potent supply from false rejection.
Hazardous Situation | A population without available vaccine because potent supply was wrongly rejected.
Sequence to Harm | Stock-out → unvaccinated population → preventable disease (individuals; campaign-scale in an outbreak).
Harm/Loss | A population left unvaccinated → preventable disease (S4; S5 campaign-scale during an outbreak).
Severity | S4 (Critical) typical / S5 campaign-scale — governing dimension public-health (HAL-2000 §5/§7 "false reject → stock-out → unvaccinated"). Take S4 anchor, S5 credible worst-case (campaign scale).
Probability (P1×P2) | aggregate P1=2 (Remote — false-alarm/false-reject tuning + legibility controls) × P2=3 (Occasional — stock-out → unvaccinated, mitigated by supply buffers) ⇒ product 6 ⇒ Medium band (PM) — Basis: aggregate of contributing records + judgment; High uncertainty, Evidence-Pending — closed by alarm-management validation and the contributing tests.
Detectability/Warning | False-reject drivers are controlled: alarm tuning to validated false-alarm/missed targets (RCR-047), charge-inhibit legibility so a cold device is not discarded (RCR-046), and the record-integrity chain (RCR-057) so a recoverable record avoids needless rejection.
Initial Risk | AP High (region R — S4 × PM; S5 campaign-scale)
Risk Acceptability | Unacceptable without control; the S5 campaign-scale face can never be broadly acceptable (HAL-2200 §6.1).
Risk-Control Strategy | 2 (protective — alarm tuning + legibility to prevent needless reject/discard) + 1 (inherent — record integrity preserves recoverable evidence) + 4 (process — supply buffering policy)
Risk-Control Requirement | RCR-047 — alerting tuned to validated false-alarm/missed-excursion targets that avoid alarm fatigue and needless reject. (+ RCR-021 verdict avoids false reject on recoverable records; + RCR-046 charge-inhibit legibility; + RCR-057 record integrity.)
Implementation Reference | DO-SAFE-004 (detection tuning), DO-CLD-003 (alarm management), DO-SAFE-013 (legibility) · CMP-039/048 (detection/alert), CMP-032 (status legibility), CMP-049 (record)
Verification Reference | PV-MPE-003 (Evidence-Pending) — alarm-management validation (false-alarm/missed targets); false-reject-rate assessment across the disposition chain
Validation Reference | VAL-MPE-003 (Evidence-Pending) — programme-scale supply-continuity validation (false-reject does not drive campaign stock-out)
Residual Probability | projected Low (PL) after alarm tuning + legibility + record integrity verified · Residual Severity: S4 (S5 campaign-scale unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP; the S5 campaign face carries EM-ENG-001 + EM-ENG-006 benefit-risk sign-off
New Risks Introduced | Loosening detection to cut false rejects would raise false-accept (HA-MPE-002) — the false-accept/false-reject benefit-risk balance, carried in HAL-2900.
Overall Residual Contribution | The under-supply face of the disposition-integrity risk; the counterpart to HA-MPE-002 (over-release). Both are carried in HAL-2900 as the two directions of the central benefit-risk trade. FTA top-event contributor (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Mission-Safety Engineer (with Quality/Regulatory and Human-Factors)
Rationale/Evidence | HAL-2000 §5/§7; WHO EVM (stock-out); RCR-047/021/046/057; aggregates HA-USE-009/HA-SW-006/HA-MPE-001.
Assumptions | Supply buffers exist but are finite; alarm/false-reject targets are validated against field-representative profiles.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MPE-004 — Regulatory / GDP data-integrity violation at scale → programme certification loss / hold
Hazard ID (library) | HAZ-MISS (regulatory/data-integrity consequence) → HAZ-DATA at scale
Product/Config | SCM all SKUs · Lifecycle Phase: Record/export aggregated over the fleet (HA-LIFE-008…011) · System State: Record management / audit · Operational Mode: Fleet in service / audit
Perspective | Mission (programme certification); organization; regulator
Source | HAL-2000 §5 (systemic falsification → withdrawal of certification, S5; major GDP violation → recall/hold, S4); ISO 13485/GDP (ALCOA+); aggregates HA-DATA-006 (export loss), HA-CYBER-009 (platform compromise), HA-CLOUD-004/006 (silent corruption)
Initiating Cause/Event | A systematic data-integrity failure across the fleet — export loss, silent record corruption, missing attribution/time/tamper-evidence, or a platform compromise — undermines the auditable record base at scale.
Contributing Conditions | GDP/ALCOA+ obligations apply to the whole record base; a scaled integrity failure (not a single record) triggers a regulatory finding; the record base underpins programme certification.
Foreseeable Sequence of Events | Systematic integrity failure → audit finds the record base unreliable at scale → regulatory finding → forced hold/recall or programme certification loss.
Hazard | Loss of trustworthy, auditable record integrity at fleet/programme scale.
Hazardous Situation | A programme whose cold-chain record base cannot be relied upon in a regulatory audit.
Sequence to Harm | Systemic integrity failure → certification/regulatory action → programme hold (doses withheld → ties HA-MPE-003) or loss of certification.
Harm/Loss | Major GDP/data-integrity violation forcing recall/hold (S4); systemic falsification → certification withdrawal (S5).
Severity | S4 (Critical) typical / S5 worst — governing dimension information/regulatory ("major GDP/data-integrity violation forcing recall/hold" S4; "systemic falsification → withdrawal of programme certification" S5, HAL-2000 §5). Take S4 anchor, S5 credible worst-case.
Probability (P1×P2) | aggregate P1=2 (Remote — export/audit-integrity + cloud-resilience controls) × P2=3 (Occasional — a scaled integrity failure → regulatory action) ⇒ product 6 ⇒ Medium band (PM) — Basis: aggregate of contributing data/cloud records; High uncertainty, Evidence-Pending — closed by export-integrity + cloud security/DR tests.
Detectability/Warning | Audit-store append-only integrity + export-integrity (RCR-042/RCR-039) and cloud resilience (RCR-039) surface corruption before it scales; tamper-evidence (RCR-034) makes falsification detectable.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6); S5 face never broadly acceptable.
Risk-Control Strategy | 2 (protective — export/audit integrity + cloud resilience + tamper-evidence) + 4 (process — GDP data-governance)
Risk-Control Requirement | RCR-042 — export/API preserves attribution, time, gaps, and tamper-evidence in a documented format; no silent data loss. (+ RCR-034 capture-time authenticity/integrity; + RCR-039 cloud resists compromise, tested backup/restore.)
Implementation Reference | DO-CLD (API/audit), DO-SEC (integrity), DO-CLD (DR) · CMP-052 (export/API), CMP-054 (audit store), CMP-049 (record store)
Verification Reference | PV-MPE-004 (Evidence-Pending) — export-integrity test + cloud security review + DR/backup-restore test; audit-append-only verification
Validation Reference | VAL-MPE-004 (Evidence-Pending) — regulatory/GDP compliance validation of the record base at scale
Residual Probability | projected Low (PL) after export/audit integrity + cloud resilience verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | The regulatory/mission face of data integrity; a programme hold ties HA-MPE-003 (supply); carried in HAL-2900. Aggregates HA-DATA-006/HA-CYBER-009/HA-CLOUD-004/006.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Mission-Safety Engineer (with Quality/Regulatory and Cloud/Product Security)
Rationale/Evidence | HAL-2000 §5; ISO 13485/GDP/ALCOA+; RCR-042/034/039.
Assumptions | GDP/ALCOA+ obligations govern the whole record base; regulators treat a scaled integrity failure as a certification-level finding.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MPE-005 — Environmental release: lithium / e-waste at fleet scale
Hazard ID (library) | HAZ-ECOL / HAZ-FIRE (fleet-scale environmental release)
Product/Config | SCM all SKUs (end-of-life, fleet scale) · Lifecycle Phase: Disposal / recycling aggregated over the fleet (HA-LIFE-018) · System State: Retired · Operational Mode: Fleet decommission
Perspective | Environment; bystander/waste-handler; organization (sustainability); ties HA-SVC-007, HA-ENV-008
Source | HA-CTX-007 (environment); HAL-2000 §7 (improper Li-battery disposal = S3); UN 38.3 / e-waste regulation; aggregates HA-SVC-007 (single-device disposal) and HA-ENV-008
Initiating Cause/Event | At fleet scale, retired reusable devices and their lithium cells enter the waste stream improperly (no take-back/recycling), aggregating single-device disposal (HA-SVC-007) into a fleet-scale release.
Contributing Conditions | Large reusable fleet reaching end-of-life over the programme's life (D1); inconsistent regional recycling infrastructure; lithium + electronics + battery chemicals at volume.
Foreseeable Sequence of Events | Many devices retired without a recycling path → fleet-scale lithium/e-waste enters landfill/uncontrolled waste → cumulative environmental release (heavy metals, electrolyte) and aggregated fire risk in waste streams.
Hazard | Fleet-scale lithium/e-waste and hazardous materials in uncontrolled waste.
Hazardous Situation | A programme discarding a fleet of lithium-bearing devices without controlled recycling.
Sequence to Harm | Fleet-scale improper disposal → cumulative environmental contamination and waste-stream fire risk.
Harm/Loss | Fleet-scale environmental release (S3, aggregated).
Severity | S3 (Serious) — governing dimension environmental (improper Li-battery disposal, HAL-2000 §7); scored at S3 consistent with the single-device HA-SVC-007, aggregated to fleet scale.
Probability (P1×P2) | P1=3 (Occasional — fleet-scale improper disposal is a real behaviour absent a mandated program) × P2=2 (Remote — cumulative environmental harm per unit disposed) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + fleet-scale e-waste experience; Medium uncertainty — closed by take-back/recycling program review (PV-MPE-005).
Detectability/Warning | Not detectable post-disposal; the control is a mandated fleet take-back/recycling program and design-for-disassembly (RCR-052).
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP — reduce via a mandated fleet recycling/take-back program and design-for-disassembly.
Risk-Control Strategy | 4 (process — fleet take-back/recycling program) + 3 (information — EOL instructions) + 1 (design-for-disassembly/recyclability)
Risk-Control Requirement | RCR-052 — safe end-of-life instructions (battery removal/recycling, data/key erasure) provided and supported by design — applied at fleet/program scale (take-back). (+ RCR-007 UN 38.3/IATA for returned cells.)
Implementation Reference | DO-SVC (fleet decommission/take-back), DO-LBL (recycling), DO-MECH (design-for-disassembly) · CMP-006 (cell), CMP-034 (PCB/e-waste)
Verification Reference | PV-MPE-005 (Evidence-Pending) — fleet take-back/recycling program review; recyclability/design-for-disassembly assessment
Validation Reference | VAL-MPE-005 (Evidence-Pending) — fleet-scale take-back uptake validation (shares VAL-SVC-007)
Residual Probability | projected Low (PL) with a mandated take-back program · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B, S3×PL) — ALARP
New Risks Introduced | Reverse-logistics transport of returned cells adds a transport-safety exposure (UN 38.3/IATA, RCR-007) — managed, not a new hazard.
Overall Residual Contribution | The fleet-scale environmental aggregate of HA-SVC-007; carried in HAL-2900 as the programme sustainability residual. Ties HA-ENV-008.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Mission-Safety Engineer (with Service and Environmental)
Rationale/Evidence | HAL-2000 §7; UN 38.3 / e-waste regulation; RCR-052; aggregates HA-SVC-007.
Assumptions | A contractual fleet take-back/recycling channel is established across deployment regions.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
HA-MPE-006 — Loss of programme trust / essential immunisation service
Hazard ID (library) | HAZ-MISS (loss of trust / service continuity consequence)
Product/Config | SCM all SKUs · Lifecycle Phase: Aggregated over the programme's operational life · System State: Programme in service · Operational Mode: Fleet in service
Perspective | Mission (the immunisation programme and trust in the auditable record); organization; population (indirect, via service withdrawal)
Source | HA-CTX-007 (mission = trust in the record); aggregates HA-MPE-001/002/003/004 (repeated visible failures → loss of confidence); RCR-057/058
Initiating Cause/Event | Repeated or high-profile monitoring/disposition failures (visible wrong accepts/rejects, data-integrity findings, or a public outbreak/stock-out) erode stakeholder confidence in the device and the auditable record.
Contributing Conditions | The programme's value rests on trust in an auditable cold-chain record; loss of trust can cause reversion to manual/less-effective monitoring, or abandonment of the device — degrading the essential immunisation service the SCM was meant to protect.
Foreseeable Sequence of Events | Repeated visible failures → stakeholders lose confidence → device abandoned or record distrusted → reversion to weaker monitoring or reduced immunisation coverage → degradation of the essential immunisation service.
Hazard | Loss of confidence in the monitoring/record system.
Hazardous Situation | A programme that can no longer rely on (or has abandoned) its cold-chain monitoring.
Sequence to Harm | Loss of trust → weaker monitoring / reduced coverage → increased downstream disposition error and under-vaccination (feeds HA-MPE-002/003).
Harm/Loss | Loss of an essential immunisation-service capability / programme trust (S4).
Severity | S4 (Critical) — mission (loss of the essential-service capability the device exists to protect); can compound toward S5 by feeding HA-MPE-002/003. Anchored at S4.
Probability (P1×P2) | aggregate P1=2 (Remote — depends on repeated/visible failures escaping the controls) × P2=3 (Occasional — visible failures erode trust and degrade service) ⇒ product 6 ⇒ Medium band (PM) — Basis: aggregate of HA-MPE-001…004 + judgment; High uncertainty, Evidence-Pending — closed as the contributing controls and post-market trust monitoring close.
Detectability/Warning | Trust erosion is measurable via post-market surveillance/field feedback; the control is preventing the visible failures (RCR-057) and governing residual risk transparently (RCR-058), plus honest, legible operation (RCR-040/046) that sustains user confidence.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control (HAL-2200 §6).
Risk-Control Strategy | 1/2 (inherent/protective — reliable, honest record-integrity operation) + 4 (process — transparent residual-risk governance + post-market trust monitoring)
Risk-Control Requirement | RCR-057 — record-integrity chain preserves mission evidence (prevents the visible failures that erode trust). (+ RCR-058 overall residual risk & benefit-risk evaluated before release and monitored post-release; + RCR-040/046 honest, legible decision support and state legibility.)
Implementation Reference | DO-SAFE-004…008 (record integrity), Risk-Management File (RCR-058), DO-HF (honest decision support) · CMP-040/049/050 (verdict/record/dashboard)
Verification Reference | PV-MPE-006 (Evidence-Pending) — inherits the record-integrity PVs; benefit-risk/residual-governance review
Validation Reference | VAL-MPE-006 (Evidence-Pending) — post-market trust/field-confidence monitoring plan (ISO 14971 post-market surveillance)
Residual Probability | projected Low (PL) after record-integrity + governance verified and post-market monitoring in place · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off
New Risks Introduced | None.
Overall Residual Contribution | The programme-trust umbrella consequence; aggregates HA-MPE-001…004 and depends on the whole library's residual. Carried in HAL-2900 and central to the benefit-risk determination.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Mission-Safety Engineer (with EM-ENG-001 and Quality/Regulatory)
Rationale/Evidence | HA-CTX-007 (mission = trust in the record); RCR-057/058/040/046; aggregates HA-MPE-001…004.
Assumptions | Programme value depends on sustained trust in the auditable record; post-market surveillance can detect trust erosion early.
Review History | 2026-07-19 v1.0 initial (ORCH-017)
7. Traceability notes
Aggregation, not duplication | each HA-MPE record names the contributing device-level records it aggregates (HAL-500…HAL-1700) and inherits their probability posture; it does not re-derive device faults.
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, §5/§6/§7 disposition and population-scale anchors), HAL-2100 (aggregate P1×P2, Evidence-Pending inheriting contributing tests), and HAL-2200 (matrix, catastrophic-severity rule §6.1 for the S5 records HA-MPE-002 and the campaign face of HA-MPE-003).
Controls | the overarching RCR-057 (record-integrity chain) and RCR-058 (overall residual & benefit-risk) govern this module, with RCR-047/021 (false-reject balance), RCR-042/034/039 (data/regulatory), and RCR-052 (environmental) as the specific consequence controls.
Implementation | DO-SAFE-004…008 (record-integrity chain), DO-CLD (export/audit/DR), DO-SEC (integrity), DO-SVC/DO-LBL (environmental), and the Risk-Management File (RCR-058).
Cross-module ties | HA-MPE-001 ↔ HA-SW-002/HA-DATA-004/006/HA-USE-007/HA-ENRG-006 (lot loss); HA-MPE-002 ↔ every S4 "compromised dose" record (HA-SW-001/002, HA-DATA-001, HA-COMM-001/002, HA-CLOUD-001, HA-CYBER-001/002, HA-USE-007, HA-ENRG-006, HA-SVC-001, HA-MFG-001); HA-MPE-003 ↔ HA-USE-009/HA-SW-006 (false reject); HA-MPE-004 ↔ HA-DATA-006/HA-CYBER-009/HA-CLOUD-004/006 (regulatory); HA-MPE-005 ↔ HA-SVC-007/HA-ENV-008 (environmental); HA-MPE-006 ↔ HA-MPE-001…004 (trust).
Downstream | the S4/S5 records become FTA top events in Stage 11 (ERFKL); HA-MPE-002 is the top FTA event of the whole library, with common-mode/independence analysis as its principal deepening. All PV-/VAL- references are Evidence-Pending and feed Stages 13–14, plus the ISO 14971 post-market surveillance plan.
8. Residual-risk / closure notes
Ultimate top event | HA-MPE-002 (population-scale outbreak from mass administration of compromised doses, S5) is the terminal consequence of the entire library — the convergence of every S4 "compromised dose" record. Per the catastrophic-severity rule (HAL-2200 §6.1) it can never be broadly acceptable; its residual stays region A (S5×PL) with EM-ENG-001 + EM-ENG-006 sign-off, documented benefit-risk, and mandatory post-market monitoring.
The central benefit-risk trade | HA-MPE-002 (over-release / false-accept) and HA-MPE-003 (under-supply / false-reject) are the two directions of the programme's core benefit-risk balance; both are carried in HAL-2900 and cannot be optimised independently — tightening one loosens the other.
No warnings-only acceptance | no mission consequence is accepted on procedure/labelling alone (HAL-2200 §6.4); each rests on the design-enforced record-integrity chain (RCR-057) plus governance (RCR-058).
Release gate (HAL-2200 §8) | no R region mission risk may remain open at release; every remaining region-A S4/S5 record carries a benefit-risk statement and residual-risk sign-off, and the S5 records carry post-market monitoring; every Evidence-Pending probability inherits the named closing tests of its contributing device-level records.
Overall contribution | this module is the consequence-aggregation input to HAL-2900 and the Risk-Management Summary — HA-MPE-002/003 define the benefit-risk envelope, HA-MPE-004 the regulatory envelope, HA-MPE-005 the environmental residual, and HA-MPE-006 the programme-trust umbrella. HA-MPE-002 is the single most important residual in the whole SCM hazard analysis.
HAL-1900 (HAZ-) — Hazard Taxonomy / Hazard-Class Library
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-1900, class code HAZ- · Owner: Principal System-Safety Engineer, with Quality/Regulatory (EM-ENG-006) and every discipline lead. This module is the library's hazard-class library — a reusable, product-agnostic taxonomy of 23 hazard classes used as a prompt list during hazard identification (HAL-000 §4). It suggests; it never auto-declares a hazard applies: every class carries an explicit SCM applicability line (Applicable / Applicable-limited / Not Applicable (confirm)) with a one-line rationale and the SCM records that instantiate it. Based on ISO 14971, IEC 62366-1, IEC 62368-1, and the domain standards; cited by designation, revisions to confirm.
---
1. Purpose
Provide a single, structured hazard-class library so hazard identification for the SCM (and future products in the org) is systematic and complete: for each of 23 classes it gives a definition, common sources, potential hazardous situations, typical harms, applicable product types, lifecycle phases, related failure mechanisms, typical control strategies, associated standards, and example search questions — plus an SCM applicability line that records whether and how the class applies here, and the SCM records that instantiate it. The taxonomy is the "prompt list" behind HAL-000 §4; it is deliberately broader than the SCM so that a "Not Applicable" is a documented decision, not an omission.
2. Scope
The 23 hazard classes below span the org's target industries (medical devices, consumer electronics, drones, automotive, industrial, aerospace, robotics — CLAUDE.md Coverage). For the SCM specifically, the dominant applicable classes are HAZ-FUNC, HAZ-DATA, HAZ-COMM, HAZ-CYBER, HAZ-USE, HAZ-THERM, HAZ-FIRE, HAZ-ENV, and HAZ-MISS (the information-device + reusable-battery-IoT profile). Applicable-limited and Not-Applicable classes are called out with their confirming rationale. This module is a library, not a register: it declares no risk chains itself (those live in HAL-500…HAL-1800) — sections 5 and 6 hold the class library and the SCM applicability summary.
3. Applicable standards & references
• ISO 14971 — hazard/hazardous-situation/harm definitions; hazard identification uses a hazard-class prompt list (Annex C hazard categories concept).
• IEC 62366-1 — use-related hazards (HAZ-USE), user profiles, use environment.
• IEC 62368-1 — energy-source classification (electrical/thermal/mechanical), fire enclosure (HAZ-ELEC, HAZ-THERM, HAZ-MECH, HAZ-FIRE).
• IEC 62133-2 / UN 38.3 / IATA DGR — lithium-battery energy/thermal/transport (HAZ-THERM, HAZ-FIRE, HAZ-PRES).
• IEC 61000 / CISPR — EMC immunity/emissions (HAZ-EMC, HAZ-RAD).
• IEC 62443 / AAMI TIR57 — cyber-induced safety (HAZ-CYBER).
• IEC 62304 / ISO 13485 / WHO GDP / WHO PQS E006 / WHO-UNICEF EVM — software, quality, data-integrity, and public-health/mission classes (HAZ-FUNC, HAZ-DATA, HAZ-MFG, HAZ-MISS).
• ISO 10218 / ISO 13482 (robotics), DO-178C / ARP4761 (aerospace) — cited for cross-industry class completeness, Not Applicable to the SCM.
Upstream | HAL-000 §4 (prompt-list method), HAL-100 (context), HAL-300 (lifecycle), and every domain register that instantiates a class.
4. Governance & method
• The library suggests; it never auto-declares. Each class is a prompt; applicability is decided per product and recorded on the SCM applicability line with a rationale. A "Not Applicable" is confirmed, not assumed (marked "Not Applicable (confirm)").
Applicability grades | Applicable (instantiated by ≥1 SCM record), Applicable-limited (applies only in a narrowed way, stated), Not Applicable (confirm) (no credible SCM path, with the confirming rationale and the reviewer who confirms).
• No severity/probability here. The taxonomy does not score risk (that is HAL-2000/2100/2200 via the domain registers). It lists typical harms and control strategies to guide identification only.
• Control-strategy hints follow the HAL-2300 hierarchy (inherent → protective → information → process) but are illustrative; the binding control is the RCR assigned in the domain register.
• Sections 5 and 6 hold the class library (the 23 entries) and the SCM applicability summary; sections 7–8 are traceability and closure.
5. Hazard classes (the 23-entry class library)
HAZ-ELEC — Electrical energy
Definition | Hazards from electrical energy — shock, short-circuit current, arc, over/under-voltage on accessible or internal conductors.
Common sources | Mains/SELV rails, battery packs, charge paths, connectors, stored charge during service.
Potential hazardous situations | Person contacts an energised conductor; a tool bridges a high-current node; over-voltage stresses a circuit.
Typical harms | Electric shock, arc-flash burn, secondary injury; component damage → downstream information harm.
Applicable product types | All powered electronics (medical, consumer, automotive, industrial, aerospace, robotics).
Lifecycle phases | Manufacture, operation, charge, service, disposal.
Related failure mechanisms | Insulation breakdown, creepage/clearance loss, connector fault, reverse insertion, latent solder short.
Typical control strategies | Inherent (SELV, clearance/creepage), protective (fusing, OVP, isolation), process (ESD/service procedure).
Associated standards | IEC 62368-1, IEC 61010, IEC 60950 (legacy).
Example search questions | Is any accessible conductor hazardous in normal or single-fault? Can a service tool short a high-current node?
SCM applicability | Applicable — SELV device but a high-energy cell and charge path; instantiated by HA-ELEC-001 (SELV shock), HA-ENRG-004 / HA-SVC-004 (service arc/short), HA-ELEC-006/007 (ingress→electronics).
HAZ-THERM — Thermal energy
Definition | Hazards from heat or extreme temperature — hot surfaces, self-heating, thermal runaway, and (for this product) energy-budget/temperature-driven loss of function.
Common sources | Lithium cells, charge circuits, self-heat elements, resistive joints, hot/cold ambient.
Potential hazardous situations | Cell in runaway near people; hot surface touched; device dead-in-cold; cold-charge plating.
Typical harms | Burn, smoke inhalation, fire; information harm via cold-driven loss of monitoring.
Applicable product types | Battery-powered electronics, EVs, industrial, aerospace.
Lifecycle phases | Charge, transport, operation (cold/hot), storage, service.
Related failure mechanisms | Thermal runaway, lithium plating, over-temp charge, resistive I²R heating, capacity fade.
Typical control strategies | Inherent (charge-inhibit interlock, energy governor), protective (BMS/OT, fuse), information (honest low-battery/inhibit state).
Associated standards | IEC 62133-2, IEC 62368-1, UN 38.3, IEC 60068.
Example search questions | Can the cell be charged below 0 °C? Can it run out of energy silently in deep cold? Any hot accessible surface?
SCM applicability | Applicable (dominant) — the D3 battery/energy tension; instantiated by HA-ENRG-001 (cold-charge plating), HA-ENRG-006/008 (dead-in-cold/fade), HA-TFC-001 (self-heat surface), HA-MFG-002 (wrong chemistry).
HAZ-MECH — Mechanical energy / physical hazards
Definition | Hazards from mechanical energy or geometry — sharp edges, pinch/crush, small parts, drop/impact, stored spring energy.
Common sources | Enclosures, moving parts, fasteners, probes/leads, small components.
Potential hazardous situations | Sharp edge cuts a handler; small part is swallowed by a child; drop crushes the cell.
Typical harms | Laceration, pinch injury, choking/strangulation, cell breach → thermal event.
Applicable product types | All physical products.
Lifecycle phases | Handling, transport, operation, service, disposal.
Related failure mechanisms | Fastener loosening, enclosure fracture, seal failure, small-part detachment.
Typical control strategies | Inherent (no sharp edges, captive/retained parts), protective (retention, cell isolation), information (small-parts warning).
Associated standards | IEC 62368-1 (mechanical), 16 CFR 1501 (small parts), UN 38.3 (shock/crush).
Example search questions | Any accessible sharp/pinch point? Any small part a child could swallow? Can a drop breach the cell?
SCM applicability | Applicable — enclosure, probe, small parts, cell-crush path; instantiated by HA-MECH-001/002 (edges/pinch), HA-MECH-003 (small parts/child), HA-MECH-009 / HA-ENRG-003 (crush→cell breach).
HAZ-PRES — Pressure / stored-pressure energy
Definition | Hazards from pressurised gas/fluid or pressure differentials — vessel rupture, venting, altitude/decompression effects.
Common sources | Pressure vessels, sealed enclosures, cell venting, altitude changes in air transport.
Potential hazardous situations | Cell vents gas in a sealed enclosure; sealed device at altitude; hydraulic burst.
Typical harms | Blast/projectile injury, offgas exposure, enclosure rupture.
Applicable product types | Pneumatic/hydraulic systems, aerospace, battery products (venting), diving/medical gas.
Lifecycle phases | Operation, transport (air), fault.
Related failure mechanisms | Cell overpressure/venting, seal blow-out, altitude pressure differential.
Typical control strategies | Inherent (vent path, pressure relief), protective (burst containment), process (transport limits).
Associated standards | UN 38.3 (T1 altitude), pressure-equipment directives, IEC 62133-2 (venting).
Example search questions | Can the cell vent build pressure in the sealed enclosure? Any altitude pressure-differential concern?
SCM applicability | Applicable-limited — no pressure vessel; only cell venting (in-runaway) and altitude in air transport; instantiated within HA-ENRG-002 / HA-TFC-003 (offgas/venting) and UN 38.3 T1 (via HA-ENRG). Confirm: no other pressure source.
HAZ-CHEM — Chemical hazards
Definition | Hazards from hazardous chemicals — corrosive/toxic substances, electrolyte, cleaning agents, offgassing.
Common sources | Battery electrolyte, cleaning/disinfectant agents, adhesives, offgas from thermal events.
Potential hazardous situations | Electrolyte leak; incompatible cleaning agent attacks the enclosure/seal; toxic offgas.
Typical harms | Chemical burn, respiratory irritation, material degradation → ingress.
Applicable product types | Battery products, medical (reprocessed), industrial chemical handling.
Lifecycle phases | Operation, cleaning, fault, disposal.
Related failure mechanisms | Electrolyte leak, cleaning-agent incompatibility, adhesive outgassing.
Typical control strategies | Inherent (compatible materials), protective (containment), information (approved-agent list), process (cleaning SOP).
Associated standards | IEC 62133-2 (electrolyte), GHS/SDS, cleaning-compatibility test methods.
Example search questions | Can electrolyte leak reach people? Is any cleaning agent incompatible with the enclosure/seal?
SCM applicability | Applicable-limited — no process chemicals; only battery electrolyte (fault) and cleaning-agent compatibility between trips; instantiated by HA-TFC-005 (cleaning-agent ingress/incompatibility) and offgas within HA-TFC-003. Confirm: not patient-contacting.
HAZ-BIO — Biological / biocompatibility / contamination
Definition | Hazards from biological agents or biocompatibility — infection, cross-contamination, cytotoxic/sensitising materials on contacting surfaces.
Common sources | Patient-contacting surfaces, reprocessed devices, contaminated surfaces, residues.
Potential hazardous situations | Cross-contamination via a reused device; sensitising material contacts skin.
Typical harms | Infection, sensitisation, contamination of a sterile/clean field.
Applicable product types | Medical devices (patient-contacting, reprocessed), food-contact.
Lifecycle phases | Cleaning, reuse, operation.
Related failure mechanisms | Inadequate cleaning, incompatible/residual materials, biofilm.
Typical control strategies | Inherent (cleanable surfaces, biocompatible materials), process (validated cleaning), information (cleaning IFU).
Associated standards | ISO 10993 (biocompatibility), AAMI ST/reprocessing standards.
Example search questions | Is the device patient-contacting? Can cleaning residue or a missed cleaning contaminate anything downstream?
SCM applicability | Applicable-limited — not patient-contacting; the only path is cleaning-residue/missed-cleaning contamination of the shipper/vaccine environment between trips; instantiated by HA-USE-010 / HA-TFC-005 (cleaning). Confirm: no biocompatibility (ISO 10993) obligation.
HAZ-RAD — Radiation (ionising / optical / laser / RF-as-radiation)
Definition | Hazards from radiated energy treated as radiation — ionising radiation, laser/optical, and RF/EM exposure to tissue.
Common sources | X-ray/isotope sources, lasers, high-power RF, optical emitters.
Potential hazardous situations | Tissue exposure above limits; eye exposure to a beam.
Typical harms | Radiation injury, burns, eye damage; RF tissue heating (SAR).
Applicable product types | Imaging/therapy medical, laser tools, high-power RF/industrial.
Lifecycle phases | Operation.
Related failure mechanisms | Interlock failure, shielding loss, over-power emission.
Typical control strategies | Inherent (low power, shielding), protective (interlocks), information (labels/limits).
Associated standards | IEC 60601-2-x (radiation), IEC 60825 (laser), RF-exposure/SAR limits, ICNIRP.
Example search questions | Any ionising/laser source? Does RF exposure approach SAR limits?
SCM applicability | Applicable-limited — RF/EM only, at low IoT power well below SAR limits; HAZ-EMC governs the interference concern; instantiated by HA-ENRG-009 (RF/SAR, S2). Confirm: no ionising/laser/optical source.
HAZ-ACOU — Acoustic energy
Definition | Hazards from acoustic energy — noise, ultrasound, infrasound causing hearing or physiological harm.
Common sources | Alarms/sirens, ultrasonic transducers, high-power motors/actuators.
Potential hazardous situations | Sustained high-SPL exposure; ultrasound tissue effects.
Typical harms | Hearing damage, discomfort, physiological effects.
Applicable product types | Industrial machinery, ultrasonic devices, alarms.
Lifecycle phases | Operation.
Related failure mechanisms | Stuck-on alarm, transducer over-drive.
Typical control strategies | Inherent (bounded SPL), protective (limits), information (PPE/labels).
Associated standards | ISO 1999 (noise exposure), IEC 60601-2-37 (ultrasound).
Example search questions | Any acoustic source above harm thresholds?
SCM applicability | Not Applicable (confirm) — the SCM has no acoustic energy source (no siren/buzzer at harmful SPL; alerts are visual/network); confirmed by System-Safety. No SCM record instantiates this class.
HAZ-EMC — Electromagnetic compatibility
Definition | Hazards from electromagnetic interference — susceptibility causing malfunction, or emissions interfering with other equipment.
Common sources | RF radios, switching supplies, ESD, nearby emitters; the device as an emitter.
Potential hazardous situations | EMI induces a false reading/verdict/missed excursion; emissions interfere with aircraft/medical equipment.
Typical harms | Malfunction → wrong disposition (information harm); interference with safety-critical equipment.
Applicable product types | All electronics, especially aerospace/medical co-located equipment.
Lifecycle phases | Operation, transport (aircraft).
Related failure mechanisms | Immunity failure (ESD/RF/surge/EFT), excessive emissions.
Typical control strategies | Protective (immunity design, shielding, filtering), inherent (layout/evidence).
Associated standards | IEC 61000-4-x (immunity), CISPR/FCC (emissions), RTCA DO-160 (airborne).
Example search questions | Can interference cause a false reading or missed excursion? Can emissions affect aircraft equipment?
SCM applicability | Applicable — immunity (no false reading) and emissions (no aircraft/medical interference); instantiated by HA-ELEC-003 (immunity → false reading) and HA-ELEC-004 (emissions → aircraft interference, S5-capable).
HAZ-FIRE — Fire
Definition | Hazards from ignition and fire propagation — enclosure fire, flame spread, in-transit fire.
Common sources | Lithium cells (thermal runaway), electrical faults, resistive hot spots.
Potential hazardous situations | Cell fire in an occupied aircraft hold, depot, or clinic; flame propagation beyond the enclosure.
Typical harms | Burn, smoke inhalation, catastrophic fire (occupied space); property/aircraft loss.
Applicable product types | Battery products, aerospace cargo, consumer electronics.
Lifecycle phases | Charge, transport, operation, disposal.
Related failure mechanisms | Thermal runaway, arc ignition, flame propagation.
Typical control strategies | Inherent (fire enclosure, low-flammability materials), protective (BMS/fuse), process (UN 38.3/IATA transport).
Associated standards | IEC 62368-1 (fire enclosure), UN 38.3, IATA DGR, UL 94 (flammability).
Example search questions | Can a single-cell event propagate flame outside the enclosure? Is transport UN 38.3/IATA-compliant?
SCM applicability | Applicable (dominant) — Li-cell fire is the top physical-safety consequence; instantiated by HA-TFC-002 (transport/occupied-hold fire, S5), HA-TFC-003 (offgas), HA-ENRG-002/003 (runaway), HA-SVC-007 / HA-MPE-005 (disposal fire).
HAZ-FUNC — Functional / systematic (software & logic) failure
Definition | Hazards from functional or systematic failure of logic/software/algorithms — wrong output, missed detection, unsafe state, that produce harm without a component "breaking."
Common sources | Detection logic, verdict engine, firmware, timekeeping, safe-state handling.
Potential hazardous situations | Missed excursion; false "accept"; wrong verdict; unsafe recovery; silent outage.
Typical harms | Wrong disposition → compromised dose administered (information harm); loss of safe state.
Applicable product types | All software-driven/automated products (medical, automotive ADAS, robotics, aerospace).
Lifecycle phases | Operation, degraded/fault, update.
Related failure mechanisms | Logic defect, threshold/hysteresis error, race/brown-out, time drift, requirements error.
Typical control strategies | Inherent (two-sided detection, verdict suppresses false accept), protective (safe state, watchdog, cross-check), process (IEC 62304 lifecycle).
Associated standards | IEC 62304, ISO 14971, ISO 26262 (automotive), DO-178C (aerospace).
Example search questions | Is there any undetected-excursion path? Can a gap/fault ever read as a clean pass? Does a fault reach a safe state?
SCM applicability | Applicable (dominant) — the defining information-device class; instantiated by HA-SW-001 (missed excursion), HA-SW-002 (false accept), HA-SW-013 (silent outage), HA-SW-009 (unsafe recovery), and the whole HAL-1000 module.
HAZ-DATA — Data & decision integrity
Definition | Hazards from loss, corruption, misattribution, or misleading presentation of data used for a decision.
Common sources | Storage, record format, export/API, dashboards, time source, gap-fill/estimation.
Potential hazardous situations | Torn/lost/fabricated record; a gap presented as a clean pass; estimated data shown as measured; misattributed record.
Typical harms | Wrong disposition (information harm); regulatory/data-integrity violation.
Applicable product types | Any data-for-decision product (medical records, GDP logistics, industrial historians).
Lifecycle phases | Operation, record management, export, audit.
Related failure mechanisms | Power-fail torn write, buffer overwrite, missing tamper-evidence, misleading visualization.
Typical control strategies | Inherent/protective (power-fail-safe atomic writes, tamper-evidence, honest gap), information (flag estimated data).
Associated standards | ISO 13485/WHO GDP (ALCOA+), 21 CFR Part 11 (audit).
Example search questions | Can a record be torn/lost/fabricated on power loss? Is a gap ever hidden? Is estimated data flagged?
SCM applicability | Applicable (dominant) — the record underpins every disposition; instantiated by HA-DATA-001 (misattributed/falsified read), HA-DATA-004 (gap-as-clean-pass), HA-DATA-006 (export/record loss), HA-DATA-007 (estimated-as-measured), and HAL-1000/1200 record paths.
HAZ-CONT — Connectivity / control (loss of control authority)
Definition | Hazards from loss or corruption of a control loop or command authority — a device that acts on, or fails to act on, a physical process.
Common sources | Actuator command paths, control loops, remote control links, autonomy stacks.
Potential hazardous situations | Unintended actuation; loss of control; command injection into a control loop.
Typical harms | Physical injury/collision, process upset.
Applicable product types | Robotics, drones, automotive, industrial control, motorised medical.
Lifecycle phases | Operation.
Related failure mechanisms | Control-loop instability, command loss/corruption, actuator fault.
Typical control strategies | Inherent (fail-safe defaults), protective (limits/interlocks, watchdog), redundancy.
Associated standards | ISO 13849, IEC 61508, ISO 10218/13482, ISO 26262.
Example search questions | Does the device command any physical actuator? Can loss of a link cause unsafe motion?
SCM applicability | Not Applicable (confirm) — the SCM is an information/monitoring device; it does not control temperature or actuate anything (HA-CTX-001/004). No control loop exists; confirmed by Systems (EM-ENG-001). No SCM record instantiates this class. (Its connectivity-for-alerting concern is HAZ-COMM, not HAZ-CONT.)
HAZ-COMM — Communications / alerting
Definition | Hazards from loss, delay, or misdirection of communication — where the message (an alert/telemetry) is safety-relevant, not a control command.
Common sources | Cellular/Wi-Fi/BLE links, antenna/RF budget, store-and-forward, notification channels, routing.
Potential hazardous situations | Excursion alert delayed/lost; alert routed to the wrong responder; in-container RF cannot close the link.
Typical harms | No time to act on an excursion → compromised dose administered (information harm).
Applicable product types | Connected/IoT products, telemetry/telehealth, fleet monitoring.
Lifecycle phases | Operation, transport.
Related failure mechanisms | RF attenuation, coverage gap, notification failure, wrong-tenant routing.
Typical control strategies | Inherent (commit-locally, store-and-forward), protective (alert-before-bulk, ack/escalation, dead-letter), information (flag last-known).
Associated standards | 3GPP (LTE-M), IEC 62368-1 (RF), carrier/roaming specs.
Example search questions | Does any excursion depend on a live link to exist? Can the in-container link close? Is an alert acked/escalated?
SCM applicability | Applicable (dominant) — in-container RF is a carried unknown (antenna Evidence-Pending); instantiated by HA-COMM-001/002 (delayed/missing alert), HA-COMM-003 (RF-in-container), HA-COMM-006 (band/region), HA-CLOUD-001 (alert-path drop).
HAZ-CYBER — Cybersecurity-induced safety
Definition | Hazards where a security compromise causes a safety consequence — falsified data, suppressed alarms, unauthorised config/firmware, platform compromise.
Common sources | Telemetry integrity, alarm path, config/OTA authorisation, debug ports, cloud platform, identity/keys.
Potential hazardous situations | Falsified telemetry hides an excursion; alarm silently suppressed; unauthorised threshold change; fleet blinded by ransomware.
Typical harms | Wrong disposition at device or fleet scale (information harm); data-integrity violation.
Applicable product types | All connected products, especially medical/critical-infrastructure.
Lifecycle phases | Operation, update, service, decommission.
Related failure mechanisms | Data injection, replay, auth bypass, left-open debug, key compromise, platform breach.
Typical control strategies | Protective (capture-time integrity/tamper-evidence, non-suppressible alarms, signed OTA, secure boot, RBAC), process (secure decommission/erasure).
Associated standards | IEC 62443, AAMI TIR57, ISO 27001, secure-boot/PKI practice.
Example search questions | Can data be falsified or an alarm suppressed undetectably? Are config/OTA authenticated? Are debug ports locked? Is decommission secure?
SCM applicability | Applicable (dominant) — security↔safety is explicit; instantiated by HA-CYBER-001 (falsified telemetry), HA-CYBER-002 (suppressed alarm), HA-CYBER-005 (OTA), HA-CYBER-006 / HA-SVC-006 (debug/secure-boot), HA-CYBER-008 / HA-SVC-007 (data/key remanence), HA-CYBER-009 / HA-MPE-004 (platform compromise).
HAZ-USE — Use-related / human factors
Definition | Hazards from foreseeable use error, misuse, or use-environment mismatch — including the vulnerable, low-training last-mile user.
Common sources | Verdict display, arming workflow, alerts, provisioning, cold/gloved/low-light/low-literacy use.
Potential hazardous situations | Misread verdict; forgot-to-arm; wrong profile; ignored/muted alert; override of a reject; cold device discarded.
Typical harms | Wrong disposition → compromised/withheld dose (information harm).
Applicable product types | All human-operated products, especially medical/consumer with diverse users.
Lifecycle phases | Install/arm, operation, receipt/disposition, recharge, reuse.
Related failure mechanisms | Perception/legibility error, workflow slip, mode confusion, alarm fatigue, misuse.
Typical control strategies | Inherent (legible, language-independent verdict; pre-arm gate; confirmed recorded override), information (IFU/training), process.
Associated standards | IEC 62366-1, ISO 14971, HE75.
Example search questions | Can the vulnerable user misread the verdict, forget to arm, or mistake an inhibited device for dead? Is an override deliberate and recorded?
SCM applicability | Applicable (dominant) — the last-mile vulnerable user is a context anchor (HA-CTX-005); instantiated by HA-USE-001 (misread verdict), HA-USE-003/004 (config/arm), HA-USE-007 (override), HA-USE-009 (charge-inhibit misread), HA-USE-011 (reuse without recalibration — ties HA-SVC-001).
HAZ-INST — Installation / service / maintenance
Definition | Hazards introduced during installation, servicing, recalibration, reprovisioning, or maintenance — the reuse-loop escapes.
Common sources | Recalibration, firmware/config restoration, seal/assembly restoration, mode hygiene, service handling.
Potential hazardous situations | Miscalibration escapes; wrong firmware/config; seal not restored; debug left on; service-energy injury.
Typical harms | Wrong disposition (information harm); ingress→record fault; technician injury.
Applicable product types | Reusable/serviceable products, installed systems, medical (reprocessed).
Lifecycle phases | Service, recalibration, reprovision, redeploy, decommission.
Related failure mechanisms | Calibration escape, config/firmware error, seal damage, left-open mode.
Typical control strategies | Process (controlled service + post-service BIST) backstopped by design (pre-arm gate, calibration-expiry gate, mode-lock).
Associated standards | ISO 13485/WHO GDP (calibrated-status, service records), IEC 62443 (secure service).
Example search questions | Can a service escape (calibration/config/firmware/seal/mode) ride into the field, and does a gate catch it before redeploy?
SCM applicability | Applicable (dominant) — the reuse loop is unique to the D1 SCM and is a hazard amplifier (HAL-300 §5); instantiated by the whole HAL-1600 module — HA-SVC-001 (miscalibration, headline), HA-SVC-002/005 (firmware/config), HA-SVC-003 (seal), HA-SVC-006 (debug mode).
HAZ-ENV — Environmental (operating environment)
Definition | Hazards from the operating/transport environment — temperature extremes, humidity/condensation, ingress, vibration, altitude — acting on the device.
Common sources | −30…+60 °C range, cold-to-warm cycling, humidity, transport vibration/shock, dust/water.
Potential hazardous situations | Condensation corrodes record electronics; cold degrades the display/battery; ingress from a breached seal.
Typical harms | Record-integrity fault → wrong disposition; loss of function in cold.
Applicable product types | Field/outdoor/transport products, cold-chain, aerospace.
Lifecycle phases | Transport, storage, operation, cleaning.
Related failure mechanisms | Condensation/corrosion, IP failure, cold-brittleness, vibration fatigue.
Typical control strategies | Inherent (IP rating, conformal coat, wide-temp parts), protective (condensation tolerance).
Associated standards | IEC 60529 (IP), IEC 60068 (environmental test), IEC 60721.
Example search questions | Can condensation reach record electronics? Does the device hold function/record across −30…+60 °C and cold-to-warm cycling?
SCM applicability | Applicable (dominant) — extreme-environment device by definition; instantiated by HA-ENV-001 (cold/freeze detection), HA-ENV-002 (cold display), HA-ENV-003 (condensation/ingress→electronics), HA-ENV-005 (vibration).
HAZ-MFG — Manufacturing / process escape
Definition | Hazards from a manufacturing or assembly process escape — a wrong build, calibration/provisioning error, or inadequate screen that ships a latent defect.
Common sources | Assembly, EOL calibration/provisioning, test coverage, workmanship (solder/PCB).
Potential hazardous situations | Miscalibrated unit passes EOL; reversed protection part; damaged seal; wrong identity/config; latent solder defect.
Typical harms | Wrong disposition at fleet scale (information harm); thermal event; ingress.
Applicable product types | All manufactured products.
Lifecycle phases | Procurement, assembly, EOL calibration/provisioning, test.
Related failure mechanisms | Calibration escape, part orientation/rating error, seal damage, coverage gap, latent workmanship.
Typical control strategies | Process (poka-yoke, AOI/ICT, EOL calibration coverage, incoming inspection) backstopped by design/field controls; traces to PFMEA (Stage 11).
Associated standards | ISO 13485/WHO GDP, IPC-A-610/6012, WHO PQS E006.
Example search questions | Can a miscalibrated/mis-built/mis-provisioned unit pass EOL and reach the field? Is EOL coverage adequate on safety-critical functions?
SCM applicability | Applicable (dominant) — fleet-scale latent escapes; instantiated by the whole HAL-1700 module — HA-MFG-001 (EOL calibration escape, headline), HA-MFG-003 (reversed protection), HA-MFG-004 (seal), HA-MFG-007/008 (coverage/latent solder).
HAZ-SUPP — Supply chain
Definition | Hazards from the supply chain — counterfeit/wrong parts, compromised SOUP/components, single-source fragility, tampered goods.
Common sources | Cell/component suppliers, SOUP/third-party software, contract manufacturers, logistics.
Potential hazardous situations | Counterfeit cell built in; component with a known vulnerability; tampered device in transit.
Typical harms | Thermal event; wrong disposition; security compromise.
Applicable product types | All products with external supply.
Lifecycle phases | Procurement/incoming, manufacture.
Related failure mechanisms | Counterfeit, substitution, vulnerable SOUP, tamper.
Typical control strategies | Process (supplier qualification, incoming verification, SBOM/vuln management, tamper-evidence).
Associated standards | ISO 13485 (supplier control), IEC 62443 (supply-chain security), UN 38.3 (cell qualification).
Example search questions | Could a counterfeit/wrong cell or a vulnerable component enter the build? Is incoming verification adequate?
SCM applicability | Applicable — chemistry Evidence-Pending raises counterfeit-cell risk; instantiated by HA-MFG-002 (wrong/counterfeit cell) and HA-CYBER-005 (vulnerable SOUP/OTA), with HA-LIFE-001 (incoming).
HAZ-MISS — Mission / public-health consequence
Definition | Hazards at the mission scale — the aggregate consequence to the vaccine lot, the immunisation programme, and the population, that device-level failures roll up to.
Common sources | Aggregate of the record-integrity chain; systematic/common-mode failures; false-accept/false-reject balance.
Potential hazardous situations | Mass administration of compromised doses; false-reject stock-out; regulatory hold; loss of programme trust.
Typical harms | Population-scale preventable disease (outbreak); mass under-vaccination; programme certification loss.
Applicable product types | Mission-critical/public-health/critical-infrastructure products.
Lifecycle phases | Aggregated over the programme's operational life.
Related failure mechanisms | Common-mode escape, systematic false-accept/reject, data-integrity failure at scale.
Typical control strategies | Inherent/protective (record-integrity chain, independent verdict cross-check, common-mode prevention), process (overall residual & benefit-risk governance, post-market surveillance).
Associated standards | ISO 14971 (overall residual/benefit-risk), WHO-UNICEF EVM, WHO PQS E006.
Example search questions | Can a systematic failure release compromised doses at population scale? Can false rejects cause a stock-out? Does a common-mode escape defeat the record-integrity chain?
SCM applicability | Applicable (dominant) — the whole product exists to protect this mission; instantiated by the whole HAL-1800 module — HA-MPE-002 (population-scale outbreak, S5, the ultimate top event), HA-MPE-001 (lot loss), HA-MPE-003 (stock-out), HA-MPE-006 (programme trust).
HAZ-PROP — Property / financial
Definition | Hazards resulting in property or financial loss without (or alongside) human harm — loss of high-value product, equipment, or facility.
Common sources | Vaccine lot value, the reusable fleet, transport/facility assets.
Potential hazardous situations | Valid lot destroyed; device fleet lost; facility/aircraft damage from a fire.
Typical harms | Loss of a lot/shipment; fleet loss; facility/aircraft loss.
Applicable product types | High-value-goods, logistics, aerospace, industrial.
Lifecycle phases | Operation, transport, disposition.
Related failure mechanisms | Wrong reject, lost record, fire, fleet attrition.
Typical control strategies | Inherent/protective (record integrity to avoid needless loss; fire containment), process.
Associated standards | ISO 14971 (property in harm definition), insurance/GDP loss criteria.
Example search questions | Can a valid high-value lot be destroyed needlessly? Can a fire cause facility/aircraft loss?
SCM applicability | Applicable — the vaccine lot and the fleet are high-value property; instantiated by HA-MPE-001 (lot loss, property face) and the property face of HA-TFC-002 / HA-ENRG-002/003 (fire → facility/aircraft loss).
HAZ-ECOL — Environmental / ecological (product's effect on the environment)
Definition | Hazards from the product's effect on the environment — hazardous-material release, e-waste, and end-of-life impact at scale.
Common sources | Lithium cells, electronics/e-waste, hazardous materials, packaging.
Potential hazardous situations | Improper battery/e-waste disposal; fleet-scale release; waste-stream fire.
Typical harms | Environmental contamination; ecological harm; waste-stream fire.
Applicable product types | Battery/electronic products, especially large fleets.
Lifecycle phases | Decommission, disposal, recycling.
Related failure mechanisms | Improper disposal, missing take-back, non-recyclable design.
Typical control strategies | Information (EOL instructions), process (take-back/recycling), inherent (design-for-disassembly/recyclability).
Associated standards | UN 38.3/IATA (transport of returned cells), WEEE/battery directives, e-waste regulation.
Example search questions | Where do the cell and electronics go at end-of-life? Is there a take-back/recycling path at fleet scale?
SCM applicability | Applicable — a reusable Li-battery fleet reaches EOL at scale; instantiated by HA-SVC-007 (single-device disposal) and HA-MPE-005 / HA-ENV-008 (fleet-scale environmental release).
6. SCM applicability summary
Class | Grade (SCM) | One-line rationale | Instantiating SCM records
HAZ-ELEC | Applicable | SELV device but high-energy cell/charge path + service energy | HA-ELEC-001/006/007, HA-ENRG-004, HA-SVC-004
HAZ-THERM | Applicable (dominant) | D3 battery/energy tension (cold-charge, dead-in-cold, self-heat) | HA-ENRG-001/006/008, HA-TFC-001, HA-MFG-002
HAZ-MECH | Applicable | Enclosure/probe/small-parts + cell-crush path | HA-MECH-001/002/003/009, HA-ENRG-003
HAZ-PRES | Applicable-limited | Cell venting + altitude in air transport only | HA-ENRG-002, HA-TFC-003 (+ UN 38.3 T1)
HAZ-CHEM | Applicable-limited | Electrolyte (fault) + cleaning-agent compatibility only | HA-TFC-005, HA-TFC-003 (offgas)
HAZ-BIO | Applicable-limited | Not patient-contacting; cleaning-residue/missed-cleaning only | HA-USE-010, HA-TFC-005
HAZ-RAD | Applicable-limited | RF/EM only, low IoT power; HAZ-EMC governs interference | HA-ENRG-009
HAZ-ACOU | Not Applicable (confirm) | No acoustic energy source at harmful SPL | — (none)
HAZ-EMC | Applicable | Immunity (no false read) + emissions (aircraft interference) | HA-ELEC-003/004
HAZ-FIRE | Applicable (dominant) | Li-cell fire is the top physical-safety consequence | HA-TFC-002/003, HA-ENRG-002/003, HA-SVC-007, HA-MPE-005
HAZ-FUNC | Applicable (dominant) | Defining information-device class (detection/verdict/safe-state) | HA-SW-001/002/009/013 (all HAL-1000)
HAZ-DATA | Applicable (dominant) | The record underpins every disposition | HA-DATA-001/004/006/007
HAZ-CONT | Not Applicable (confirm) | Monitoring device; controls/actuates nothing | — (none; alerting is HAZ-COMM)
HAZ-COMM | Applicable (dominant) | Alert delivery + in-container RF (antenna Evidence-Pending) | HA-COMM-001/002/003/006, HA-CLOUD-001
HAZ-CYBER | Applicable (dominant) | Security↔safety (falsified data, suppressed alarm, OTA, debug) | HA-CYBER-001/002/005/006/008/009, HA-SVC-006
HAZ-USE | Applicable (dominant) | Vulnerable last-mile user is a context anchor | HA-USE-001/003/004/007/009/011
HAZ-INST | Applicable (dominant) | Reuse loop is a unique D1 hazard amplifier | HA-SVC-001/002/003/005/006 (all HAL-1600)
HAZ-ENV | Applicable (dominant) | −30…+60 °C extreme-environment device | HA-ENV-001/002/003/005
HAZ-MFG | Applicable (dominant) | Fleet-scale latent escapes; traces to PFMEA | HA-MFG-001/003/004/007/008 (all HAL-1700)
HAZ-SUPP | Applicable | Chemistry Evidence-Pending → counterfeit-cell + SOUP risk | HA-MFG-002, HA-CYBER-005
HAZ-MISS | Applicable (dominant) | The product exists to protect this mission | HA-MPE-001/002/003/006 (all HAL-1800)
HAZ-PROP | Applicable | Vaccine lot + reusable fleet are high-value property | HA-MPE-001, HA-TFC-002 / HA-ENRG-002/003 (property face)
HAZ-ECOL | Applicable | Reusable Li-battery fleet reaches EOL at scale | HA-SVC-007, HA-MPE-005, HA-ENV-008
Dominant applicable classes (SCM profile): HAZ-FUNC, HAZ-DATA, HAZ-COMM, HAZ-CYBER, HAZ-USE, HAZ-THERM, HAZ-FIRE, HAZ-ENV, HAZ-MISS (plus the reuse/manufacturing amplifiers HAZ-INST, HAZ-MFG). Applicable-limited: HAZ-PRES (cell venting/altitude), HAZ-CHEM (electrolyte/cleaning), HAZ-BIO (cleaning-residue, not patient-contacting), HAZ-RAD (RF/EM only). Not Applicable (confirm): HAZ-ACOU (no acoustic source), HAZ-CONT (no control/actuation).
7. Traceability notes
Prompt-list role (HAL-000 §4) | this taxonomy is the hazard-identification prompt list; each domain register (HAL-500…HAL-1800) cites the HAZ- class(es) in each record's Hazard ID (library) field.
Applicability is a decision, not an omission | the two Not-Applicable classes (HAZ-ACOU, HAZ-CONT) are confirmed with a rationale and reviewer (System-Safety / EM-ENG-001); the four Applicable-limited classes (HAZ-PRES, HAZ-CHEM, HAZ-BIO, HAZ-RAD) state the narrowed scope.
No scoring here | severity/probability/acceptability live in HAL-2000/2100/2200 and are applied in the domain registers; this module lists only typical harms/controls to guide identification.
Coverage cross-check | every SCM domain record maps to at least one HAZ- class (Hazard ID field), and every Applicable/Applicable-limited class maps to at least one SCM record (section 6 table) — the two-way check that hazard identification is complete.
Downstream | the dominant classes (HAZ-FUNC/DATA/COMM/CYBER/USE/THERM/FIRE) feed the Stage-11 FTA/DFMEA/ SW-FMEA/PFMEA scoping; HAZ-MISS (HA-MPE-002) is the top FTA event of the library; HAZ-MFG traces to the PFMEA.
8. Closure notes
• This module is a library, not a register: sections 5 and 6 hold the class definitions and the SCM applicability decisions; it closes when every SCM domain record's Hazard-ID field cites a class here and every Applicable class is instantiated (both directions verified in section 7's coverage cross-check).
Reusability | the 23 classes are product-agnostic and reusable across the org's target industries (medical, consumer, drone, automotive, industrial, aerospace, robotics); for a new product only the applicability lines are re-decided — the class library itself is stable.
Maintenance | if a design change introduces a new energy/consequence path (e.g., adding an audible alarm would move HAZ-ACOU to Applicable, or a control/actuation feature would move HAZ-CONT to Applicable), the applicability line is re-opened under the ECO/ECN process (HAL-000 §7) and the corresponding domain register gains a record.
• The SCM headline classes — HAZ-FUNC/DATA/COMM/CYBER/USE (information harm) and HAZ-THERM/FIRE (battery) and HAZ-MISS (mission) — are the identification anchors that make this an information-device-plus-reusable- battery-IoT hazard profile, and are the classes that dominate HAL-2900 and the Risk-Management Summary.
HAL-200 (HA-BND) — System Boundary & Exposure Model
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: EM-ENG-001 with System-Safety. Defines the system boundary, the exposure model (how energy, data, and control cross the boundary to reach a person/property/mission), and the required boundary diagrams. Based on the system architecture (S1–S9), DOL, and the fault-management architecture.
1. Purpose
Draw the boundary of "the SCM system" and model exposure pathways — the routes by which a hazard reaches a harmed perspective — so no hazard is missed at an interface and so the dominant information exposure pathway (device → record/verdict/alert → human decision → vaccine → beneficiary) is made explicit.
2. Scope
Physical device (S1–S7), device firmware (S3-FW), connectivity (S5), cloud (S8), mobile/BLE app (S9), accessories (probe, charger, external antenna), and the human/organisational decision layer that converts device information into a disposition action.
3. Applicable standards & references
ISO 14971 (foreseeable sequences, exposure); IEC 62443 / AAMI TIR57 (trust boundaries); architecture files (logical-physical-architecture.md, data-flow-control-flow.md, power-energy-architecture.md, communication-architecture.md). Cited by designation.
4. Boundary register (HA-BND records)
HA-BND-001 — System boundary definition
Inside the boundary | the sealed device (enclosure, cell CMP-006/pack CMP-007, BMS CMP-008, charge path CMP-010, charge-inhibit CMP-012, MCU CMP-015, secure element CMP-016, sensors CMP-019/020/021/022/023, radios CMP-024/025/026, HMI CMP-030/031/032, storage CMP-033, PCB CMP-034), the device firmware (CMP-035…045), the external probe (CMP-020) and USB-C charger interface (CMP-004), and the cloud/mobile services (CMP-046…055).
On the boundary (interfaces) | USB-C charge/data port; probe connector; RF egress (LTE-M/Wi-Fi/BLE/GNSS antennas); OLED/LED/button HMI; BLE provisioning link.
Outside the boundary (context) | the vaccine and its passive shipper; the cellular carrier network; notification recipients; the human decision-maker; third-party USB power sources; the disposal/recycling chain.
Source | Component list, logical-physical-architecture. Status: Baselined.
HA-BND-002 — Energy-flow exposure model (physical harm pathways)
Stored/loose energy that can cross to a person/property | cell chemical/electrical energy (thermal runaway → fire/burn → nearby people, aircraft, cold store); charge-path electrical energy (USB-C source fault → overvoltage/short → thermal); self-heat element thermal energy (hot surface → burn); mechanical energy (drop/shock → cell crush → thermal event; detached small parts → choking). SELV voltages limit shock exposure. Exposure gate: the enclosure, BMS, charge-inhibit, and retention are the barriers on this pathway. Source: power-energy-architecture, DO-SAFE-001/002/010. Status: Baselined — feeds HAL-500/600/700/800.
HA-BND-003 — Data/decision exposure model (the dominant information pathway)
The defining SCM pathway | sensor → detection engine (CMP-039) → verdict engine (CMP-040) → local record/OLED and cloud record/dashboard/alert (CMP-046/048/049/050) → human disposition decision → vaccine accepted/rejected → beneficiary. Harm enters when the information at any node is missing/incorrect/delayed and the human acts on it. Exposure gate: two-sided detection, verdict integrity (suppress false accept on gaps), honest-gap annunciation, alert delivery, receipt QA (external contributing control). Source: data-flow-control-flow, fault-management, DO-SAFE-004/005/006/009. Status: Baselined — the top exposure pathway; feeds HAL-1000/1100/1200/1400/1500.
HA-BND-004 — Control-flow exposure model
Control paths that can create an unsafe/misleading state | charge enable/inhibit (HW-primary, FW defense-in-depth); arming state machine; OTA apply/rollback; safe-state entry; power-state governor (radio gating). A defect or compromise in any control path can either create a physical hazard (charge) or a false-reassurance information state (arming/OTA/safe-state). Exposure gate: HW-independent charge interlock, pre-arm gate, fail-safe A/B OTA, watchdog. Source: state-mode-model, DO-SAFE-007/014. Status: Baselined — feeds HAL-500/1000/1300.
HA-BND-005 — Trust-boundary / attack-surface model
Trust boundaries | device↔cloud (mutually authenticated); device↔mobile/BLE (authenticated pairing); cloud↔responder notification; USB-C data/OTA; debug port (locked in production); RF (jamming/DoS). Each trust boundary is a cyber→safety entry point where an adversary could falsify data, suppress an alarm, alter thresholds, or block comms → wrong disposition. Exposure gate: secure element/root of trust, signed OTA, tamper-evidence, debug-port lockdown, RBAC. Source: DOL-1100 (DO-SEC), TIR57. Status: Baselined — feeds HAL-1300.
HA-BND-006 — Human-interaction map
Interaction points | provisioning/arming (mobile/BLE + button), receipt read (OLED verdict, gloved/ cold/offline), alert receipt/acknowledgement (responder phone/dashboard), depot charge/triage (LED/OLED), disposal. Each is a use-error exposure point where the human is the last barrier — or the failure point — on the information pathway. Source: HAL-100 (HA-CTX-005/006), IEC 62366-1. Status: Baselined — feeds HAL-400/1500.
HA-BND-007 — External-dependency map
Dependencies whose failure crosses into a hazard | cellular coverage/roaming + in-container RF attenuation (delayed/missing alert); cloud availability (alert/record); notification channel delivery; time source (record defensibility); third-party USB charger (charge-input abuse); disposal chain (environmental). Source: HA-CTX-010, communication/deployment architecture. Status: Baselined — feeds HAL-1100/1200/1800.
5. Required boundary diagrams (references)
System context, energy-flow, data-flow, control-flow, trust-boundary, human-interaction, and external-dependency diagrams are maintained in the Stage-08 architecture (data-flow-control-flow.md, power-energy-architecture.md, communication-architecture.md, deployment-architecture.md) and referenced here rather than duplicated. The data-flow and human-interaction diagrams are the two governing artefacts for the information-harm pathway.
6. Traceability / closure notes
The exposure model scopes which perspectives each domain register must consider. HA-BND-003 (information pathway) is the single most important exposure model in this library and is referenced by every S4/S5 software/data/connectivity/use record. Feeds the PHA, Detailed PHA, and the FTA top-event selection (Stage 11).
HAL-2000 (HA-SEV) — Severity
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: Quality/Regulatory (EM-ENG-006) with System-Safety. This module fixes the severity scale used by every HA-<CODE>-NNN record in the library. Severity is fixed by outcome (the credible worst-case harm), independent of probability, per ISO 14971 and CLAUDE.md rule 3 (severity-first Action Priority). Standards cited by designation; revisions to confirm.
1. Purpose
Provide a single, product-specific severity scale covering every consequence dimension the SCM can touch — human/health harm, public-health/mission, property/financial, information/regulatory — so that severity is assigned consistently and can never be traded away against a low probability.
2. Scope
All SCM SKUs, all lifecycle phases, all perspectives (beneficiary/patient, clinic staff, logistics/3PL, depot technician, installer, bystander, organization/programme, mission, environment). The defining SCM subtlety: the device is an information device — its most consequential harm path is informational (a wrong disposition decision), not a direct physical injury. The severity scale therefore treats an undetected/incorrect/delayed excursion decision that leads to a compromised dose administered (or a potent dose destroyed) as a primary, high-severity harm, on par with a physical thermal/fire event.
3. Applicable standards & references
ISO 14971 (risk management; severity fixed by outcome); IEC 62366-1 (use-error harm); IEC 62368-1 (thermal/ electrical injury); WHO PQS E006 / WHO-UNICEF Effective Vaccine Management (public-health consequence of potency loss). Cited by designation.
4. Governance & method
• Severity = credible worst-case harm of the hazardous situation, assessed at the harm end of the risk chain, not the initiating cause. It is not reduced by controls (controls reduce probability; residual severity is only lowered where the control genuinely changes the harm outcome, per HAL-2800).
Multi-dimensional | a record takes the highest applicable severity across the dimensions below, and records the governing dimension in its Rationale.
• Population/scale is explicit. For the SCM, a single wrong "accept" can release a lot of doses to many beneficiaries; scale is captured in the S4/S5 boundary.
• Reversibility, duration, detectability, and ability to intervene are considered but do not downgrade a catastrophic outcome; they inform probability (P2) and detectability instead.
5. Severity scale (S1–S5)
Level | Label | Human / health harm | Public-health / mission (information-device harm) | Property / financial | Information / regulatory
S5 | Catastrophic | Death or permanent disability (e.g., fire/thermal-runaway causing fatal injury; enclosure fire in an occupied aircraft cargo hold) | Population-scale preventable disease: mass administration of potency-lost vaccine across many beneficiaries → outbreak / vaccine-derived harm; or mass stock-out (false-reject cascade) leaving a population unvaccinated during an outbreak | Loss of a national/programme vaccine campaign; multi-lot loss with public-health impact | Systemic falsification of the cold-chain record base → withdrawal of programme certification
S4 | Critical | Serious injury requiring medical treatment (burn from hot surface/thermal event; smoke inhalation) | A lot/shipment of compromised doses administered to a group of beneficiaries → individuals unprotected / preventable disease; or a valid lot wrongly destroyed → local stock-out → unvaccinated individuals | Loss of a full shipment/lot of vaccine; product recall | Major GDP / data-integrity violation; regulatory finding forcing recall or hold
S3 | Serious | Moderate, reversible injury (minor burn, laceration needing first aid) | Wrong disposition of a limited quantity of doses; significant avoidable wastage; a monitored trip released on an incomplete record | Loss of a limited quantity of product; significant unplanned cost | Reportable data-integrity deviation; audit non-conformance
S2 | Minor | Minor injury, no treatment (pinch, sharp-edge scratch, transient discomfort) | Small avoidable wastage; nuisance re-inspection; low-confidence but correctable decision | Minor product/equipment loss; rework | Minor documentation deviation
S1 | Negligible | No injury; nuisance only | No disposition impact; cosmetic | Negligible loss | Cosmetic record annotation
6. Severity anchor for the SCM (governing dimension)
Per UN-SAFE-001 / STN-020 / PRD-SAFE-001, the severity anchor is the disposition-error outcome. Both directions harm the beneficiary:
• False accept (compromised dose administered) → loss of potency → beneficiary not protected → preventable disease. Worst credible case at lot/population scale = S5; typical single-lot = S4.
• False reject (potent dose destroyed) → stock-out → unvaccinated beneficiary → preventable disease. Worst credible case (campaign-scale) = S5; typical = S4. This anchor is why software/data/connectivity/use hazards in this library carry S4–S5 severity even though the device never physically touches a patient — the harm is real and mediated by the disposition decision.
7. Severity register (assignments used across the library)
Harm archetype | Governing dimension | Severity | Example records
Undetected excursion → compromised dose administered (lot/pop.) | Public-health | S5 (worst) / S4 (typical lot) | HA-SW-001, HA-DATA-001, HA-COMM-001/002, HA-CYBER-001, HA-MPE-002
False "accept" on incomplete/faulted record | Public-health | S4 | HA-SW-002, HA-DATA-004, HA-USE-007
Missed freeze excursion (freeze-sensitive vaccine) | Public-health | S4 | HA-SW-001, HA-ENV-001
Delayed/missing excursion alert → no time to act | Public-health | S4 | HA-COMM-001/002, HA-CLOUD-001
False reject → stock-out → unvaccinated | Public-health | S4 (S5 campaign-scale) | HA-USE-009, HA-MPE-003
Battery thermal runaway / fire (transport, clinic) | Human injury | S5 (occupied aircraft) / S4 | HA-ENRG-001/002/003, HA-TFC-002
Cold-charge lithium plating → latent internal short | Human injury | S4 | HA-ENRG-001, HA-ELEC-002
Hot self-heat surface burn | Human injury | S3 | HA-TFC-001
Small parts / probe choking (child bystander) | Human injury | S4 (choking child) | HA-MECH-003
Electric shock (SELV device) | Human injury | S2 | HA-ELEC-001
Lost record in deep cold (dead-in-cold, silent) | Public-health | S4 | HA-ENRG-006, HA-SW-013
Improper Li-battery disposal | Environmental | S3 | HA-SVC-007, HA-MPE-005
8. Traceability / closure notes
Severity feeds HAL-2200 (acceptability, severity-first), every domain register (Severity field), the PHA and Detailed PHA, and the Risk-Management Summary. Severities are Judgment anchored to ISO 14971 and the STN-020 severity anchor; they are not Evidence-Pending (outcome is known) — only probability is Evidence-Pending. High-severity hazardous situations (S4/S5) become FTA top events in Stage 11 (ERFKL).
HAL-2100 (HA-PROB) — Probability
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: System-Safety / Reliability with EM-ENG-006. Defines the decomposed probability model P = P1 × P2 and its evidence bases. Per ISO 14971 and the SCM concept §7, most SCM probabilities are Evidence-Pending (closed by named tests); this module records the basis and uncertainty of every estimate and never asserts a probability as fact.
1. Purpose
Give every HA record a consistent, honest probability estimate decomposed into P1 (initiating cause → hazardous situation) and P2 (hazardous situation → harm), with an explicit basis and uncertainty flag, so that risk prioritisation is transparent and severity-first (CLAUDE.md rule 3).
2. Scope
All SCM hazard records. Because the SCM is an information device, P2 (situation → harm) is often the governing, uncertain factor: whether a missed excursion actually results in a compromised dose being administered depends on downstream human/organisational checks (independent shipping indicators, receipt QA, disposition policy) that are outside the device boundary and are captured as contributing conditions.
3. Applicable standards & references
ISO 14971 (P1×P2 decomposition; probability may be unknown pending evidence); reliability handbook methods (field/lab data where available). Cited by designation.
4. Governance & method — basis taxonomy
Every estimate records one primary basis and an uncertainty flag (Low/Med/High uncertainty):
1. Field data — none yet (pre-launch); post-market surveillance (Stage 17) will refine.
2. Reliability data — component FIT/datasheet (battery, connectors) — partial, Evidence-Pending.
3. Test — closed by a named PV-/VAL- test (e.g., 90-day cold endurance, UN 38.3, detection-logic test).
4. Simulation / analysis — energy budget, RF-in-container link budget (Evidence-Pending).
5. Historical / analogous — WHO-UNICEF freezing-excursion literature; industry Li-ion event rates.
6. Expert judgment — System-Safety judgment where no data exists (flagged High uncertainty).
7. Conservative categorical — assign a deliberately conservative level pending evidence.
8. Unknown pending evidence — probability cannot yet be estimated; carried as a gap, risk driven by severity until closed.
5. Probability levels (P1 and P2 each on a 1–5 categorical scale)
Level | Label | Qualitative frequency (per device over service life / per exposure) | Typical basis
5 | Frequent | Expected to occur repeatedly / routinely | field/test
4 | Probable | Likely to occur several times | test/analogous
3 | Occasional | Likely to occur sometime in the fleet | analogous/judgment
2 | Remote | Unlikely but credible | judgment/conservative
1 | Improbable | So unlikely it may be assumed not to occur, but not impossible | analysis/judgment
6. Combined probability P = P1 × P2 (mapped to five bands)
Multiply the two levels (1–25) and map to a combined probability band used by the risk matrix (HAL-2200):
P1×P2 product | Combined band | Symbol
20–25 | Very High | PVH
12–16 | High | PH
6–10 | Medium | PM
3–5 | Low | PL
1–2 | Very Low | PVL
The product is a prioritisation aid only; where either factor is "unknown pending evidence", the combined band is set conservatively (one band above the point estimate) and the record is flagged Evidence-Pending.
7. Probability posture for the SCM top hazards (basis register)
Hazard archetype | P1 basis | P2 basis | Combined (initial, pre-control) | Uncertainty
Cold-charge lithium plating (D3) | Occasional (depots receive cold-returned devices; Test PV cold-charge) | Occasional→Probable (plating→latent short over cycles; analogous Li-ion) | Medium–High | High (chemistry Evidence-Pending)
Dead-in-cold / energy exhaustion mid-trip (D3) | Occasional (90-day/−30 °C tension; Test endurance) | Probable (silent gap reads as uneventful → wrong accept; judgment) | High | High (budget Evidence-Pending)
Missed excursion — detection logic error | Remote (SW verification) | Probable (undetected → dose administered; judgment) | Medium | Med
Missed freeze excursion | Occasional (historical WHO-UNICEF freezing common) | Probable | High | Med
Delayed/missing alert — in-container RF attenuation | Occasional–Probable (Sim/Test RF-in-container, PRD-VER-005 Evidence-Pending) | Occasional (store-and-forward + receipt QA mitigate) | Medium–High | High (antenna Evidence-Pending)
Battery thermal runaway (protection intact) | Improbable–Remote (analogous cell event rate + BMS) | Occasional (enclosed near people) | Low–Medium | Med
Miscalibration escape at EOL | Remote (Test EOL calibration coverage) | Probable (biased reading → wrong accept) | Medium | Med
Use-error: misread OLED verdict (cold/gloved/low-training) | Occasional (Usability summative, Evidence-Pending) | Occasional | Medium | High
Cyber: falsified telemetry / suppressed alarm | Remote (Pen-test, controls) | Probable (undetected → dose administered) | Medium | Med
8. Traceability / closure notes
Probability feeds HAL-2200 (matrix) and the residual re-assessment (HAL-2800, after controls verified). All initial probabilities are pre-control estimates; residual probability is re-estimated only after the control is defined, implemented, verified, and evidenced. Evidence-Pending probabilities are closed by: 90-day cold endurance & reserve test, cold-charge-inhibit fault-injection, UN 38.3 abuse, detection-logic test, RF-in-container qualification (PRD-VER-005), EOL calibration coverage, and the summative usability study. These named tests become Verification/Validation references (PV-/VAL-) in Stages 13–14.
HAL-2200 (HA-ACC) — Risk Acceptability
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: EM-ENG-006 (Quality/Regulatory) with System-Safety and EM-ENG-001. Defines the risk-acceptability matrix, regions, Action Priority, and decision rules applied to every HA record. Severity-first per ISO 14971 and CLAUDE.md rule 3. Standards by designation.
1. Purpose
Convert each record's (Severity × Combined-Probability) into an acceptability region and an Action Priority (AP), define who may accept residual risk, and state the special rules for catastrophic severity and uncertain probability that dominate this information device.
2. Scope
All SCM hazard records, initial and residual. The matrix is severity-dominant: probability may move a risk within a severity band but cannot, on its own, make a catastrophic outcome broadly acceptable.
3. Applicable standards & references
ISO 14971 (risk acceptability, benefit-risk, overall residual risk); ISO 14971 Annex C concepts (risk matrix as an aid, not a substitute for judgment). Cited by designation.
4. Governance & method
• Combined probability band (PVL…PVH) from HAL-2100 × severity (S1…S5) from HAL-2000 → cell in the matrix.
• Each cell maps to a region: R (Unacceptable), A (Undesirable / ALARP — reduce as far as practicable, benefit-risk to accept), B (Broadly acceptable).
Action Priority | R → High, A → Medium, B → Low (severity-first; see §6 rules).
• The score never hides the harm scenario — the full risk chain is always the primary artefact.
5. Risk-acceptability matrix (Severity × Combined probability)
Severity ↓ / Prob → | PVL (Very Low) | PL (Low) | PM (Medium) | PH (High) | PVH (Very High)
S5 Catastrophic | A (High) | A (High) | R (High) | R (High) | R (High)
S4 Critical | A (Med) | A (High) | R (High) | R (High) | R (High)
S3 Serious | B (Low) | A (Med) | A (Med) | R (High) | R (High)
S2 Minor | B (Low) | B (Low) | A (Med) | A (Med) | R (High)
S1 Negligible | B (Low) | B (Low) | B (Low) | A (Med) | A (Med)
Catastrophic rule (S5): an S5 risk is never classified B (broadly acceptable), regardless of how low the probability. Even at PVL it is A and requires documented benefit-risk justification and the maximum practicable risk reduction (ALARP) plus disclosure — see §6.
6. Decision rules (severity-first)
1. Catastrophic-severity rule. Any S5 hazard is treated as High Action Priority for design purposes until reduced to A with benefit-risk justification; it can never be closed as "broadly acceptable."
2. Uncertain-probability rule. Where P1 or P2 is unknown pending evidence, escalate the combined band one level (conservative) and drive the risk by severity; the risk stays Open / Evidence-Pending until the named test closes the probability. (This governs the D3 battery and RF-in-container top risks.)
3. Additional-reduction rule (ALARP). Any risk in region A must show that further practicable reduction has been applied or justified as not practicable, following the control hierarchy (HAL-2300).
4. Warnings-challenge rule. A risk reduced only by "information for safety" (labels/IFU/training) where inherent or protective control is feasible is not accepted; it is escalated (HAL-2300).
5. Single-fault rule. For S4/S5 physical hazards (battery/charge/thermal), a single fault shall not lead to harm — a protective backstop independent of the primary control is required (HAL-2400).
6. Benefit-risk & overall residual. No individual risk is accepted in isolation; the overall residual risk (HAL-2900) and benefit-risk (Risk-Management Summary) govern release.
7. Risk-acceptance authority
Region / severity | Acceptance authority
B (Broadly acceptable) | Responsible discipline engineer + System-Safety record
A (ALARP), S1–S3 | System-Safety + EM-ENG-006 (Quality/Regulatory)
A (ALARP), S4–S5 | EM-ENG-006 + EM-ENG-001 (Chief Systems) + benefit-risk analysis
R (Unacceptable) | Cannot be accepted — design change required; escalate to programme risk board
Any S5 residual | EM-ENG-001 + EM-ENG-006 sign-off + documented benefit-risk + post-market monitoring
8. Traceability / closure notes
The matrix is applied in every domain register (Initial Risk, Residual Risk fields), aggregated in the PHA, Detailed PHA, and Risk-Management Summary. Release criterion: no R region risk may remain open at release; every A S4/S5 risk carries a benefit-risk statement and residual-risk sign-off; every Evidence-Pending probability has a named closing test. High-severity (S4/S5) hazardous situations are handed to Stage 11 (ERFKL) as FTA top events and drive deeper DFMEA/SW-FMEA.
HAL-2300 (HA-HIER) — Risk-Control Hierarchy applied to the SCM
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Owner: Principal System-Safety / Risk-Management Engineer, with Quality/Regulatory (EM-ENG-006), Human-Factors, and Chief Systems (EM-ENG-001). This module fixes how controls are chosen — the ISO 14971 control hierarchy — and classifies every Risk-Control Requirement (RCR-001…058, HAL-2400) by hierarchy level. It enforces two program rules the SCM cannot be released without: the warnings-challenge rule (information for safety may not be the primary control for the low-training last-mile user where design prevention is feasible) and the single-fault rule for S4/S5 physical hazards. Standards cited by designation; revisions to confirm.
---
1. Purpose
Make the selection order of every SCM risk control explicit and auditable. ISO 14971 requires that risk be reduced in a fixed priority — first by designing the hazard out, then by protective measures, then by information for safety, and only lastly by process/operational measures — and that a lower-priority control is not accepted where a higher-priority one is practicable. This module:
1. explains the four-level hierarchy as applied to an information device whose dominant harm is a wrong disposition decision, not a direct injury;
2. classifies all 58 RCRs by level;
3. applies the warnings-challenge rule — for the vulnerable last-mile user, labels/IFU/training cannot carry a safety risk that inherent or protective design can carry instead — and shows where design carries it (e.g., RCR-043 legible verdict, RCR-045 pre-arm gate, RCR-001 HW charge interlock);
4. records the single-fault rule for the S4/S5 physical (battery/charge/thermal) hazards (RCR-001/002).
2. Scope
All accepted controls in HAL-2400 (RCR-001…058) across every domain register (HAL-500…HAL-1800), every SKU, and the full lifecycle. The hierarchy classification here is the rationale layer behind the Class field already carried in the RCR register; this module is authoritative for hierarchy disputes and for the warnings-challenge and single-fault determinations. Out of scope: the numeric acceptability regions (HAL-2200) and the implementation/verification status (HAL-2500/2600).
3. Applicable standards & references
• ISO 14971 — risk management; the control-option priority (inherent safety by design → protective measures / safeguards → information for safety), and the requirement that further practicable reduction be applied before accepting residual risk (ALARP).
• IEC 62366-1 — usability engineering: use-related risk is reduced by design first; information for safety is the weakest control and must be validated, not assumed effective.
• IEC 62368-1 — the hazard-based safety-engineering safeguard model (basic/supplementary/reinforced safeguards) — the electrical/battery analogue of the protective layer and the single-fault backstop.
• IEC 62304 — software safety: risk-control measures implemented in software are design measures, must be specified as requirements, and cannot be assumed to fail safe without architecture.
• IEC 62443 / AAMI TIR57 — security controls that are safety-relevant enter the same hierarchy (a protective measure that must be independent and non-defeatable).
• UN 38.3 / IATA DGR — transport-safety qualification behind the battery protective/inherent controls.
Upstream | HAL-2000 (severity), HAL-2100 (probability), HAL-2200 (acceptability, §6.4 warnings-challenge, §6.5 single-fault), HAL-2400 (RCR register — the master control list), DOL-1200 (DO-SAFE implementation), CLAUDE.md rule 3 (severity-first) and the org-wide warnings challenge.
4. Governance & method — the four-level hierarchy
Controls are selected top-down; a lower level is accepted only when the higher levels have been applied as far as practicable and the residual still needs reduction (ISO 14971 ALARP).
• Level 1 — Inherent safety by design (ISD). Remove or design out the hazard, or make the safe outcome the default that no fault, config, or user action can undo. Strongest; preferred. SCM examples: prevent charging below 0 °C at the hardware (RCR-001); make the verdict incapable of showing a false "accept" over a gap (RCR-021); make the verdict legible and language-independent so it cannot be misread (RCR-043).
• Level 2 — Protective measures / safeguards (PROT). Add a measure that detects the hazardous condition and acts (interlock, backstop, limiter, safe-state, alarm-with-escalation). Does not remove the hazard but bounds it. SCM examples: independent secondary fuse/OVP (RCR-002); probe-fault detection (RCR-020); escalating acknowledged alerts (RCR-029); pre-arm readiness gate (RCR-045).
• Level 3 — Information for safety (INFO). Labels, IFU, on-device status, training. Weakest; depends on a human perceiving, understanding, and acting correctly. In this library the information controls are design-enforced where they carry real risk — the device forces the honest state to be shown, rather than relying on a label (RCR-022 honest gap, RCR-040 incomplete-monitoring qualifier, RCR-046 charge-inhibit legibility). Pure information (RCR-052 EOL instructions) is used only where no higher control is feasible.
• Level 4 — Process / operational (PROC). Manufacturing, calibration, service, transport, and disposal procedures and their controls. Necessary for latent/production hazards but never the sole barrier for an S4/S5 physical hazard. SCM examples: UN 38.3 transport qualification (RCR-007); EOL calibration coverage (RCR-053); correct-chemistry/part verification (RCR-054).
Method rules applied here:
• Severity-first (CLAUDE.md rule 3, HAL-2200). The higher the severity, the stronger the required level — S4/S5 physical hazards demand Level 1/2 with a single-fault backstop, never Level 3/4 alone.
• Warnings-challenge (HAL-2200 §6.4). A risk reduced only by information for safety, where inherent or protective control is feasible, is not accepted — it is escalated and re-controlled at a higher level. This is decisive for the SCM because the last-mile disposition user is low-training, working cold, gloved, in low light, possibly not in their first language.
• Single-fault (HAL-2200 §6.5). For S4/S5 physical hazards a single fault shall not reach harm; an independent, common-cause-resistant backstop is required.
• Non-defeatable. Level 1/2 controls on the release-critical set (RCR-001, RCR-021/022) must not be bypassable by firmware fault, configuration, or user action.
5. Categories — how the hierarchy maps to the SCM's two harm classes
The SCM has two harm classes, and the hierarchy is applied to each:
1. Physical-energy harm (battery/charge/thermal/mechanical/electrical). Classic ISO 14971 / IEC 62368-1 territory. Controlled by Level 1 (design the energy hazard out / contain it) + Level 2 (independent protective backstop, single-fault rule) + Level 4 (transport & build process). Information (Level 3) is never the primary control for a fire/thermal hazard.
2. Information-integrity harm (missed/incorrect/delayed excursion → wrong disposition → compromised or destroyed dose). The SCM-defining class. Because the harm is mediated by a human decision, the temptation is to control it with labels and training (Level 3). The library rejects that: the excursion decision is carried by inherent design (two-sided detection RCR-019, verdict integrity RCR-021, legible verdict RCR-043, pre-arm gate RCR-045) and protective measures (probe-fault detection RCR-020, escalating alerts RCR-029, alarm-not-suppressible RCR-035), with information controls made design-enforced (RCR-022/040/046) rather than advisory.
6. Control-hierarchy classification register (RCR-001 … RCR-058)
Legend — Level: 1 = inherent safety by design · 2 = protective measure · 3 = information for safety (D-E = design-enforced) · 4 = process/operational. Where an RCR spans levels, the primary (highest, strongest) level governs and secondary levels are noted. "SF" = on the single-fault-tolerant release-critical set. "WC" = a control that carries a risk the warnings-challenge would otherwise have left to information.
6.1 Battery / charge / thermal — physical-energy harm (severity-first, single-fault)
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-001 | 1 ISD | 2, 3 | HW charge-inhibit <0 °C, MCU-independent | SF, WC. Charge safety designed out at HW; FW is defense-in-depth and cannot defeat it. Not left to a "warm before charging" label.
RCR-002 | 2 PROT | 1 | Autonomous OV/UV/OC/OT/short + non-resettable secondary fuse/OVP | SF. Independent secondary backstop terminates the hazard if the primary BMS IC fails (single-fault rule).
RCR-003 | 2 PROT | — | Reject/limit out-of-spec charge source; PD PDO cap | Protective; BMS (RCR-002) is the independent backstop layer.
RCR-004 | 1 ISD | 2 | Cell retained without crush/puncture through drop/vibe/transport | Inherent mechanical protection of the cell; the energy hazard is designed not to be initiated.
RCR-005 | 2 PROT | — | Self-heat may only enable the charge guard, never bypass inhibit or exceed upper limit | Protective, subordinate to RCR-001; bounded so it cannot become a new hazard (see HA-NEW-001).
RCR-006 | 1 ISD | 3 | Energy governor + reserve; honest low-battery/dead-in-cold; 90-day endurance | Inherent (governor/reserve) prevents silent loss of monitoring; the information part (honest indication) is design-enforced, not advisory.
RCR-007 | 4 PROC | — | UN 38.3-tested, IATA-DGR-compliant, transport documentation | Process/qualification behind the physical controls.
RCR-008 | 1 ISD | — | Fire-enclosure contains a single-cell event, low-tox offgas | Inherent containment (IEC 62368-1 fire enclosure).
6.2 Mechanical
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-009 | 1 ISD | — | No accessible sharp/pinch/crush point | Hazard designed out of the enclosure.
RCR-010 | 1 ISD | 3 | Small parts retained/sized; warned if not | Inherent retention preferred; warning only where retention is not feasible (child-bystander choking, S4).
RCR-011 | 1 ISD | 2 | Enclosure/seal maintains ingress + cell containment through life | Inherent + protective (ingress & containment).
RCR-012 | 1 ISD | — | Fasteners/joints resist loosening under vibration/thermal cycling | Inherent robustness.
6.3 Electrical / EMC
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-013 | 1 ISD | — | All accessible circuits SELV; no hazardous voltage in single-fault | Inherent — shock hazard designed out.
RCR-014 | 2 PROT | — | EMC immunity (no false reading/verdict) + emissions (aircraft/medical) | Protective; immunity design so RF cannot corrupt the decision.
RCR-015 | 2 PROT | — | Conformal coat + seal protect record electronics across −30↔+60 °C | Protective barrier for record-integrity electronics.
RCR-016 | 4 PROC | — | AOI + EOL screening; no latent single-point escape on safety nets | Process screening of latent defects.
6.4 Environmental
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-017 | 1 ISD | — | Operate + retain record −30…+60 °C incl. 90-day endurance | Inherent environmental capability.
RCR-018 | 2 PROT | — | Tolerate condensation from cold↔warm cycling | Protective barrier against condensation.
6.5 Software / detection / verdict / record / time — information-integrity harm
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-019 | 1 ISD | — | Two-sided detection, freeze reliability = heat, no undetected path | WC. Detection designed so no excursion goes undetected — not a "check the vaccine" instruction.
RCR-020 | 2 PROT | — | Probe absent/open/short/wrong & implausible reads detected, never a valid temp | WC. Protective fault detection; the false-accept path is closed by design, not by telling the user to place the probe well.
RCR-021 | 1 ISD | — | Verdict suppresses false "accept" on any gap/fault; device↔cloud consistent | SF-equivalent, WC, non-defeatable. The most important information-harm inherent control; must not be bypassable by FW/config/user.
RCR-022 | 3 INFO (D-E) | — | Monitoring gap made visible; incomplete record never shows a clean pass | WC. Design-enforced honesty — the device forces the honest state; not a label.
RCR-023 | 2 PROT | — | Defined safe state on fault, watchdog-backed logged recovery | Protective safe-state; logging prevents masking (see HA-NEW-005).
RCR-024 | 2 PROT | — | Power-fail-safe atomic writes; no torn/lost/fabricated record | Protective record-integrity barrier.
RCR-025 | 2 PROT | — | Trusted monotonic UTC; drift bounded and disclosed when untrusted | Protective + design-enforced disclosure.
RCR-026 | 2 PROT | 4 | Calibration integrity-protected; expired/invalid blocks arming | Protective (pre-arm gate) + process (interval).
RCR-027 | 2 PROT | — | Availability self-check flags "not monitoring" on outage | Protective silent-outage prevention.
6.6 Connectivity / cloud — alert delivery
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-028 | 1 ISD | — | Every excursion committed locally, independent of connectivity | Inherent — no excursion depends on a live link to exist.
RCR-029 | 2 PROT | — | Alert-ahead-of-bulk, ack-required, escalate if unacknowledged | WC. Protective escalation carries the "responder must notice" risk, not a training reminder.
RCR-030 | 2 PROT | — | Store-and-forward on link loss; never record loss; alert-first on reconnect | Protective degraded-mode.
RCR-031 | 1 ISD | — | RF closes in-container budget, or IP-sealed external-antenna fallback | Inherent link capability (RF-in-container, Evidence-Pending).
RCR-032 | 2 PROT | — | Cloud never drops a safety event; dead-letter/retry; reconciliation | Protective delivery guarantee.
RCR-033 | 2 PROT | — | Authoritative device↔trip↔owner↔responder binding | Protective routing correctness.
6.7 Cybersecurity-induced safety
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-034 | 2 PROT | — | Capture-time authenticity/integrity; falsified data detectable/rejected | Protective; falsification cannot silently pass.
RCR-035 | 2 PROT | — | Alert path not silently suppressible; suppression itself raises an event | Protective + design-enforced detectability.
RCR-036 | 2 PROT | — | Threshold/config changes authenticated, authorised, bounded, audited | Protective change-control.
RCR-037 | 2 PROT | — | OTA signed, rollback-protected, fail-safe A/B, never bricks; windows disclosed | Protective update integrity (S5 via HA-CYBER-005).
RCR-038 | 2 PROT | — | Secure boot to HW root of trust; debug ports locked; tamper-evident | Protective platform integrity.
RCR-039 | 2 PROT | 4 | Cloud resists compromise (RBAC/SSO, monitoring, tested DR) | Protective + process; guards fleet-wide S5 (HA-CYBER-009).
6.8 Data & decision
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-040 | 3 INFO (D-E) | — | Disposition surface shows incomplete-monitoring qualifier; never a clean pass on a gap | WC. Design-enforced decision-support honesty (ties RCR-021/022).
RCR-041 | 3 INFO (D-E) | — | Estimated/interpolated/last-known data flagged, never presented as measured | WC. Design-enforced labelling of estimate vs measurement.
RCR-042 | 2 PROT | — | Export/API preserves attribution, time, gaps, tamper-evidence | Protective export integrity.
6.9 Use-related / human factors — where the warnings-challenge is decisive
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-043 | 1 ISD | — | Verdict legible & language-independent (icon+colour+word), gloved, low-light, −30 °C | WC. The verdict is designed to be readable by the vulnerable user — not "trained to interpret." Carries HA-USE-001/008.
RCR-044 | 1 ISD | — | Verdict-override requires deliberate, confirmed, recorded action | WC. Accidental acceptance designed out; override is hard and audited, not discouraged by a note. Carries HA-USE-007.
RCR-045 | 2 PROT | — | Pre-arm readiness gate blocks arming unless SoC/cal/FW/profile/probe all pass | WC, non-defeatable. The "is it ready?" judgement is taken off the low-training user and enforced by the device. Carries HA-USE-003/004.
RCR-046 | 3 INFO (D-E) | — | Charge-inhibited/fault/unready states legible with cause; cold ≠ dead | WC. Design-enforced legibility so a cold charge-inhibited device is not discarded. Carries HA-USE-009.
RCR-047 | 2 PROT | — | Alerting tuned to validated false-alarm/missed targets; consolidation | Protective alarm-management (avoids fatigue muting a real excursion).
RCR-048 | 2 PROT | — | Provisioning validates trip profile against stability library | Protective mismatch detection.
6.10 Service / maintenance
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-049 | 2 PROT | 4 | Recalibration interval enforced by pre-arm expiry gate | Protective enforcement (via RCR-026) + process interval.
RCR-050 | 4 PROC | 2 | Service restores seal/config/FW; post-service BIST + pre-arm gate | Process with a protective verification gate.
RCR-051 | 2 PROT | 4 | Debug/dev modes locked before deployment; unexited mode detected | Protective + process.
RCR-052 | 3 INFO | 4 | Safe EOL instructions (battery/data/key) + design support | Information + process — used where no higher control removes the disposal hazard.
6.11 Manufacturing / supply-chain
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-053 | 4 PROC | — | EOL calibration coverage/accuracy + birth record; miscalibrated unit fails EOL | Process screening of the reuse-amplified miscalibration hazard.
RCR-054 | 4 PROC | 2 | Correct chemistry + correct-rated/oriented protection parts (poka-yoke + AOI) | Process (with poka-yoke, an inherent-flavoured error-proofing).
RCR-055 | 4 PROC | — | Unique cryptographic identity + serial; duplicates prevented | Process provisioning integrity.
RCR-056 | 4 PROC | — | Correct regional/firmware config provisioned + verified at EOL | Process configuration verification.
6.12 Overarching governance
RCR | Primary level | Also | Control (short) | Hierarchy rationale
RCR-057 | 4 PROC | 1/2 | Record-integrity chain preserves mission evidence (aggregate of RCR-019…024/042) | Governance aggregate of the inherent/protective record controls.
RCR-058 | 4 PROC | — | Overall residual + benefit-risk evaluated before release (HAL-2900) | Governance gate.
Distribution (primary level): Level 1 (ISD) — RCR-001, 004, 006, 008, 009, 010, 011, 012, 013, 017, 019, 021, 028, 031, 043, 044 (16); Level 2 (PROT) — RCR-002, 003, 005, 014, 015, 018, 020, 023, 024, 025, 026, 027, 029, 030, 032, 033, 034, 035, 036, 037, 038, 039, 042, 045, 047, 048, 049, 051 (28); Level 3 (INFO, all design-enforced except RCR-052) — RCR-022, 040, 041, 046, 052 (5); Level 4 (PROC) — RCR-007, 016, 050, 053, 054, 055, 056, 057, 058 (9). The portfolio is inherent/protective-dominant (44 of 58) — the correct posture for a severity-first, low-training-user information device.
7. Warnings-challenge — where inherent/protective design carries the last-mile risk
The last-mile disposition user is the weakest link in the harm chain and the least able to compensate: low training, working at −30 °C, gloved, in low light, under stock-out/production pressure, sometimes reading a non-primary language. Per HAL-2200 §6.4 and CLAUDE.md, information for safety cannot be the primary control for these use-related S4 hazards where design can carry the risk. The library discharges the challenge as follows:
Use hazard (would default to a warning/training) | Warnings-challenge outcome — the higher-level control that carries it
Misread verdict (cold/gloved/low-literacy/low-light/language) — HA-USE-001/008, HA-ENV-002 | RCR-043 (Level 1) — verdict is icon+colour+word, language-independent, readable gloved/cold/low-light within the responsiveness budget. Not "train the user to read the OLED."
Override a "reject" under pressure (automation bias) — HA-USE-007 | RCR-044 (Level 1) — override requires a deliberate, confirmed, recorded action; accidental acceptance designed out. Not "advise against overriding."
Forgot-to-arm / void trip believed monitored — HA-USE-004; wrong profile — HA-USE-003 | RCR-045 (Level 2, non-defeatable) — pre-arm readiness gate blocks arming unless SoC/cal/FW/profile/probe all pass; un-armed/void state unmistakable. Not "remember to arm."
Cold charge-inhibited device mistaken for dead → discarded — HA-USE-009 | RCR-046 (Level 3 design-enforced) + RCR-001 (Level 1) — inhibited state legible with cause; the charge safety itself is inherent HW. Not "don't throw away cold devices."
Ignored/muted alert (alarm fatigue) — HA-USE-002 | RCR-047 + RCR-029 (Level 2) — alerting tuned to validated false-alarm/missed targets with consolidation; alerts escalate if unacknowledged. Not "pay attention to alerts."
Gap misread as an uneventful trip — HA-DATA-004, HA-USE-004 | RCR-021 (Level 1) + RCR-022/040 (Level 3 design-enforced) — the verdict cannot show a clean pass over a gap; the surface forces the incomplete-monitoring qualifier. Not "check for gaps."
Probe left in air / against wall — HA-USE-005 | RCR-020 (Level 2) — an unrepresentative/absent probe read is detected and never scored as a valid product-zone temperature. Not "place the probe correctly."
Residual reliance on information (accepted): RCR-052 (EOL disposal) remains a Level-3/4 control because no inherent/protective measure removes the end-of-life disposal hazard for a battery product; it is accepted with design support (battery removal, key/data erasure) and is not a disposition-decision safety control. All other information controls are design-enforced (RCR-022/040/041/046) — the device forces the honest state; the human is not relied upon to remember or interpret.
8. Single-fault rule — S4/S5 physical hazards
Per HAL-2200 §6.5 and HAL-2400 §4, for the S4/S5 physical (battery/charge/thermal) hazards a single fault shall not lead to harm; an independent, common-cause-resistant backstop is required.
Physical S4/S5 hazard | Primary (Level 1/2) | Independent backstop | Common-cause resistance
Cold-charge plating → latent short (HA-ENRG-001, S4) | RCR-001 MCU-independent HW charge-inhibit on cell-coupled sensing | FW defense-in-depth (subordinate) + RCR-046 legibility | HW interlock is independent of FW; inhibit persists in any FW/comms fault and cannot be defeated by FW.
Cell OV/OC/OT/short → thermal runaway (HA-ENRG-002, S5) | RCR-002 autonomous BMS protection | Non-resettable secondary fuse/OVP | Secondary fuse/OVP is independent of the BMS IC — hazard terminated if the primary IC fails. Verified by protection fault-injection with the BMS forced failed.
Mechanical crush → cell short (HA-ENRG-003 / HA-MECH-009, S5) | RCR-004 cell retention | RCR-008 fire-enclosure containment | Retention (mechanical) and containment (enclosure) are different physical mechanisms — no shared failure.
Out-of-spec charge source (HA-ENRG-005/HA-ELEC-002, S4) | RCR-003 input protection | RCR-002 BMS backstop | Input front-end and BMS are two independent layers; both must fail to reach the cell.
Note: the information-integrity S4/S5 hazards are not "single-fault physical" in the IEC 62368-1 sense, but the library applies an analogous non-defeatability rule to their inherent controls — RCR-021/022 (verdict integrity / honest gap) must not be bypassable by FW fault, configuration, or user action, and the device and cloud verdicts cross-check (RCR-021) so a single computation error cannot silently release a false accept.
9. Traceability notes
• Each RCR's hierarchy level here refines the Class field in HAL-2400; any change to a control's level is an ECO against both modules.
• Every Level-1/2 control traces control → RCR (HAL-2400) → DO-SAFE / DO- implementation (DOL-1200) → PV-/VAL- verification (Stages 13–14, Evidence-Pending); the implementation review is HAL-2500, verification method HAL-2600.
• The warnings-challenge outcomes here are the acceptance basis for the use-related residuals in HAL-2800 and are validated by the summative usability study (VAL-, Evidence-Pending).
• The single-fault determinations feed Stage 11 (ERFKL): HA-ENRG-001/002/003, HA-MECH-009 and the S5 records become FTA top events; the independence claims (RCR-001 vs FW, RCR-002 fuse vs BMS) are the FTA common-cause/independence assumptions to be proven.
10. Residual-risk / closure notes
• The hierarchy is inherent/protective-dominant (44/58 controls at Level 1/2) — the required posture for a severity-first information device with a vulnerable last-mile user. No S4/S5 safety risk is carried by information alone; every warnings-challenge use hazard has a Level-1/2 carrier (§7).
Open at baseline | the two D3 headline controls have inherent designs whose effectiveness is Evidence-Pending — RCR-001 (chemistry / cold-charge boundary) and RCR-006/017 (90-day cold endurance) — and RCR-031 (RF-in-container). Their hierarchy placement is correct (Level 1); their verification is the open item (HAL-2600).
Release gate contribution | this module is satisfied when (a) every S4/S5 hazard is controlled at Level 1/2, (b) every physical S4/S5 hazard has an independent single-fault backstop, (c) every warnings-challenge use hazard has a Level-1/2 or design-enforced carrier, and (d) no safety risk is accepted on information alone where a higher level is practicable. All four are met in design; verification is carried in HAL-2600 and residual acceptance in HAL-2800/2900.
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HAL-2400 (HA-RCR) — Risk-Control Requirements
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: System-Safety with EM-ENG-001. The authoritative Risk-Control Requirement register. Every accepted control from the hazard registers becomes a formal, testable RCR-###, allocated to a subsystem, implemented in a controlled design output (DO-…), and verified (PV-…). This register feeds Stage 12 (Risk-Control Requirements) and re-baselines the preliminary PRD-SAFE requirements. Per HAL-2300 the control hierarchy is applied (inherent → protective → information → process); warnings-only controls are challenged where design prevention is feasible.
1. Purpose
Convert controls into single-fault-tolerant, non-defeatable, testable requirements with clear allocation, acceptance criteria, and implementation/verification links — so no risk is "controlled" on a spreadsheet alone.
2. Scope
All accepted controls across HAL-500…HAL-1800. Each RCR lists: control class, allocation (subsystem/CMP), the DO-SAFE / DO- implementation reference (DOL-1200 and disciplines), the verification method, and the hazards it controls.
3. RCR register (RCR-001 … RCR-058)
Legend — Class: ISD = inherent safety by design · PROT = protective measure · INFO = information for safety · PROC = process/operational. Impl = Implementation Reference (design output). Verif = verification method (PV-…/VAL-…, Evidence-Pending until Stages 13–14).
Battery / charge / thermal (severity-first physical safety)
RCR | Class | Requirement (testable) | Allocation / CMP | Impl (DO) | Verif | Controls hazards
RCR-001 | ISD+PROT | Charging shall be prevented below 0 °C (or cell datasheet limit) by an MCU-independent HW interlock on cell-coupled sensing; inhibit persists in any FW/comms fault; FW is defense-in-depth and cannot defeat it. | S2 / CMP-010/012 | DO-SAFE-001 → DO-EE-007+DO-FW-005 | Cold-charge fault-injection sweep | HA-ENRG-001, HA-ELEC-002, HA-USE-009
RCR-002 | PROT | Autonomous OV/UV/OC/OT/short protection with a non-resettable secondary backstop (fuse/OVP); single-fault-safe (hazard terminated if primary IC fails). | S2 / CMP-008/009 | DO-SAFE-002 → DO-EE-005+DO-EE-010 | Protection fault-injection (DO-TST-009) | HA-ENRG-002, HA-TFC-002, HA-ELEC-006
RCR-003 | PROT | Device shall reject/limit an out-of-spec charge source (reverse-insertion, OVP, surge, short); USB-PD PDO capped. | S2 / CMP-014 | DO-SAFE-003 → DO-EE-009+DO-EE-010 | Reverse/OVP/short test | HA-ENRG-005, HA-ELEC-002
RCR-004 | ISD+PROT | Cell shall be retained without crush/puncture through drop/vibration/transport; thermal/electrical clearance maintained. | S1/S2 / CMP-007 | DO-SAFE-010 → DO-MECH-012+DO-EE-004 | Drop/vibe + UN 38.3 | HA-ENRG-003, HA-MECH-009, HA-TFC-002
RCR-005 | PROT | Self-heat (if provided) shall only enable the charge guard by warming the cell; it shall not bypass inhibit or exceed the safe upper thermal limit. | S2 / CMP-013 | DO-SAFE-001 (self-heat) → DO-EE-007 | Self-heat bound + inhibit-integrity test | HA-ENRG-001, HA-TFC-001
RCR-006 | ISD+INFO | An energy governor + reserve shall prevent silent loss of monitoring: reserve preserves record-commit and a low-battery alert; honest low-battery/dead-in-cold indication; 90-day cold endurance met. | S2/S3 / CMP-037/011 | DO-SAFE-012 → DO-FW-004+DO-EE-008 | 90-day cold endurance (PRD-VER-004) | HA-ENRG-006, HA-SW-013
RCR-007 | PROC | The device and battery shall be UN 38.3-tested and IATA-DGR-compliant, with transport documentation provided. | S2 / CMP-006/007 | DO-EE-015, DO-REG-006 | UN 38.3 test | HA-ENRG-003, HA-TFC-002, HA-MPE-005
RCR-008 | ISD | Enclosure and internal materials shall contain a single-cell thermal event without external flame propagation and with low-toxicity offgassing consistent with IEC 62368-1 fire enclosure. | S1/S2 / CMP-001/034 | DO-MECH-005, DO-PCB-010 | Fire-enclosure / abuse test | HA-TFC-002/003
Mechanical
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-009 | ISD | No accessible hazardous sharp edge, pinch, or crush point in normal handling/service. | S1 / CMP-001 | DO-MECH-002/003 | Edge/pinch inspection (IEC 62368-1) | HA-MECH-001/002
RCR-010 | ISD+INFO | Small parts (probe tip, connector caps, any lanyard) shall be retained or sized/warned to prevent choking/strangulation of a child bystander. | S1 / CMP-005/020 | DO-MECH-006, DO-LBL warning | Small-parts (16 CFR 1501-style) + retention test | HA-MECH-003
RCR-011 | ISD+PROT | Enclosure + seals shall maintain ingress rating and cell containment through the specified drop/vibration/transport and reuse life. | S1 / CMP-001/002 | DO-SAFE-010/011 → DO-MECH-005/012 | Drop/vibe + IP test | HA-MECH-004/009, HA-ENV-003
RCR-012 | ISD | Fasteners/joints shall resist loosening under vibration/thermal cycling over reuse life (no detachment). | S1 / CMP-001/034 | DO-MECH-004 | Vibration + thermal-cycle inspection | HA-MECH-006
Electrical / EMC
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-013 | ISD | All accessible circuits shall be SELV; no accessible hazardous voltage in normal or single-fault. | S2 / CMP-034 | DO-EE (rails), DO-PCB | IEC 62368-1 electrical safety | HA-ELEC-001
RCR-014 | PROT | The device shall meet EMC immunity (ESD/RF/surge/EFT per IEC 61000) such that no interference produces a false reading, false verdict, or missed excursion; and shall not interfere with aircraft/medical equipment (emissions). | S3/S5 / CMP-034/024 | DO-EE (EMC), DO-PCB-010 | EMC immunity + emissions test | HA-ELEC-003/004
RCR-015 | PROT | Conformal coat + seal shall protect record-integrity electronics from condensation/leakage/corrosion across −30↔+60 °C cycling over reuse life. | S1/S3 / CMP-002/034 | DO-SAFE-011 → DO-MECH-005+DO-PCB-010 | IP + cold-cycle + cleanliness | HA-ELEC-006/007, HA-ENV-003, HA-TFC-006
RCR-016 | PROC | Latent solder/PCB/component defects shall be screened by AOI + EOL test coverage; no latent single-point failure escapes on safety-critical nets. | S3 / CMP-034 | DO-TST (AOI/ICT), DO-PCB | AOI/ICT coverage review | HA-ELEC-008, HA-MFG-008
Environmental
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-017 | ISD | The device shall operate and retain the record across −30 °C…+60 °C, including the 90-day cold endurance and reserve. | SYS / CMP-006/015/033 | DO-SAFE-012, DO-EE, DO-FW | IEC 60068 + 90-day endurance | HA-ENV-001/002, HA-ENRG-006
RCR-018 | PROT | The device shall tolerate condensation from cold-to-warm cycling without record-integrity failure. | S1/S3 / CMP-002/034 | DO-SAFE-011 | Condensation cycle test | HA-ENV-003, HA-TFC-006
Software / detection / verdict / record / time
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-019 | ISD | Two-sided (heat and freeze) product-stability excursion detection with freeze reliability equal to heat; no undetected-excursion path. | S3 / CMP-039 | DO-SAFE-004 → DO-FW-006+DO-EE-013 | Detection-logic test (profiled excursions) | HA-SW-001, HA-ENV-001, HA-DATA-001
RCR-020 | PROT | Absent/open/short/wrong probe and implausible sensor reads shall be detected and flagged, never read as a valid product-zone temperature. | S4 / CMP-020/038 | DO-SAFE-005 → DO-FW-007+DO-EE-013 | Probe-fault-injection test | HA-SW-008, HA-USE-005, HA-DATA-001
RCR-021 | ISD | The verdict shall suppress a false "accept" on any data gap/sensor fault/incomplete monitoring; device and cloud verdicts consistent. | S4 / CMP-040/049 | DO-SAFE-006 → DO-FW-008+DO-CLD-006 | Verdict-consistency + gap-handling test | HA-SW-002/007, HA-DATA-004, HA-USE-007
RCR-022 | INFO(design-enforced) | Any monitoring gap/outage/lapse shall be made visible in the record and on-device status (honest gap); an incomplete record never shows an unqualified clean pass. | S4 / CMP-040/041 | DO-SAFE-006/008 | Gap-exposure test | HA-SW-013, HA-DATA-004, HA-USE-004
RCR-023 | PROT | On any unsafe/misleading fault the device shall enter a defined safe state (preserve record, hold cell safe, honest status) with watchdog-backed logged recovery. | S3/S4 / CMP-018/037/045 | DO-SAFE-007 → DO-FW-003+DO-FW-013 | Fault-injection + watchdog test | HA-SW-009, HA-ELEC-008, HA-ENRG-006
RCR-024 | PROT | Storage shall be power-fail-safe (atomic writes, brown-out barrier); no torn/lost/fabricated record; no overwrite of unsent data; tamper-evident. | S4 / CMP-033/041 | DO-SAFE-008 → DO-FW-010+DO-EE-002 | Power-fail-during-write test | HA-SW-013, HA-ELEC-008, HA-DATA-006
RCR-025 | PROT | The device shall maintain trusted monotonic UTC time across sleep/brown-out/service; drift shall be bounded and disclosed when time is untrusted. | S3 / CMP-017 | DO-EE (RTC), DO-FW | Clock-drift/discipline test | HA-SW-005, HA-DATA-003
RCR-026 | PROC+PROT | Calibration coefficients shall be integrity-protected; expired/invalid calibration shall block arming (pre-arm gate). | S4 / CMP-038 | DO-FW-007, DO-SAFE-007 | Calibration-expiry gate test | HA-SW-014, HA-MFG-001, HA-SVC-001
RCR-027 | PROT | An availability self-check shall verify the whole sensing→logging chain and flag the device not monitoring on an outage (silent-outage prevention). | S4 / CMP-045 | DO-SAFE-006/007 → DO-FW | Self-check outage test | HA-SW-013, HA-ENRG-006
Connectivity / cloud (alert delivery)
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-028 | ISD | Every excursion event shall be committed locally independent of connectivity, so no excursion depends on a live link to exist. | S4 / CMP-041/042 | DO-SAFE-009 → DO-FW-011 | Offline-excursion test | HA-COMM-001/002
RCR-029 | PROT | Excursion alerts shall be sent ahead of bulk, require acknowledgement, and escalate if unacknowledged; cloud-side consolidation avoids fatigue. | S4 / CMP-042/048 | DO-SAFE-009 → DO-CLD-007/008 | Alert-latency + escalation test | HA-COMM-001, HA-CLOUD-001, HA-USE-002
RCR-030 | PROT | On connectivity loss the device shall degrade to store-and-forward buffering, never record loss, and deliver alert-first on reconnection. | S4 / CMP-041/042 | DO-SAFE-009 | Store-and-forward test | HA-COMM-001/003
RCR-031 | ISD | The RF link (integral antenna) shall close the in-container budget across markets, or an IP-sealed external-antenna fallback shall be provided; RF-in-container performance qualified. | S4 / CMP-028/029 | DO-IF (antenna), DO-EE | RF-in-container qualification (PRD-VER-005) | HA-COMM-003
RCR-032 | PROT | The cloud shall not drop a safety event: dead-letter/retry on failure, never silent discard; end-to-end delivery reconciliation. | S4 / CMP-046/048 | DO-CLD-007 (SRS-CLOUD-023) | Fault-injection on ingest/notify | HA-CLOUD-001/002/004
RCR-033 | PROT | Device↔trip↔owner↔responder binding shall be authoritative and correct; alerts routed to the right responder; wrong-tenant binding prevented. | S4 / CMP-047 | DO-CLD (registry) | Association/routing test | HA-CLOUD-003, HA-DATA-002
Cybersecurity-induced safety
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-034 | PROT | Telemetry and the record shall carry capture-time authenticity/integrity (tamper-evidence); falsified/injected data shall be detectable and rejected. | S4 / CMP-016/044/049 | DO-SEC (integrity), DO-SAFE-008 | Integrity/pen-test | HA-CYBER-001, HA-DATA-001
RCR-035 | PROT | The alarm/alert path shall not be silently suppressible; suppression/blocking shall itself raise a detectable event and escalate. | S4 / CMP-048/042 | DO-SEC, DO-SAFE-009 | Alarm-suppression pen-test | HA-CYBER-002, HA-COMM-002
RCR-036 | PROT | Excursion thresholds/config changes shall be authenticated, authorised, bounded, and audited; unauthorised change prevented. | S4 / CMP-044/053/054 | DO-SEC, DO-CLD (audit) | Config-tamper test | HA-CYBER-003, HA-USE-003
RCR-037 | PROT | OTA shall be signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; never bricks a device; unmonitored windows disclosed. | S4 / CMP-035/043/051 | DO-SAFE-014 → DO-FW-014/015+DO-CLD-010 | OTA rollback + version-skew test | HA-SW-010, HA-CYBER-005, HA-CLOUD-007
RCR-038 | PROT | Secure boot against a hardware root of trust; production debug ports locked/disabled; tamper-evident enclosure. | S3 / CMP-016/044 | DO-SEC-006, DO-FW | Secure-boot + debug-lock pen-test | HA-CYBER-006/008, HA-SVC-006
RCR-039 | PROT+PROC | The cloud shall resist compromise (RBAC/SSO, monitoring of auth/integrity/backlog events, tested backup/restore) so a platform incident cannot blind the fleet or corrupt records silently. | S8 / CMP-053/054 | DO-CLD, DO-SEC | Cloud security review + DR test | HA-CYBER-009, HA-CLOUD-004/006
Data & decision
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-040 | INFO(design-enforced) | The disposition surface shall show the incomplete-monitoring qualifier and never present a gap/fault as a clean pass (ties RCR-021/022). | S4 / CMP-040/050/055 | DO-SAFE-006, DO-HF-002 | Decision-support usability test | HA-DATA-004, HA-USE-007
RCR-041 | INFO(design-enforced) | Estimated/interpolated/last-known data (e.g., GNSS, gap-fill) shall be flagged as such, never presented as measured/authoritative. | S3 / CMP-023/050 | DO-CLD, DO-HF | Visualization review | HA-DATA-007, HA-COMM-003
RCR-042 | PROT | Export/API shall preserve attribution, time, gaps, and tamper-evidence in a documented format; no silent data loss. | S3 / CMP-052/054 | DO-CLD (API/audit) | Export-integrity test | HA-DATA-006, HA-MPE-004
Use-related / human factors
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-043 | ISD | The accept/reject/quarantine verdict shall be legible and language-independent (icon+colour+word), readable gloved, in low light, and at −30 °C within the responsiveness budget. | S6 / CMP-030/032 | DO-SAFE-013, DO-HF-002/005 | Cold-display + summative usability | HA-USE-001/008, HA-ENV-002
RCR-044 | ISD | A verdict-override shall require a deliberate, confirmed, recorded action; accidental acceptance of a compromised dose shall be prevented. | S3/S6 / CMP-040/055 | DO-HF, DO-CLD (audit) | Override-workflow usability test | HA-USE-007, HA-IUE-007
RCR-045 | PROT | The pre-arm readiness gate shall block arming unless SoC-for-trip, valid/unexpired calibration, valid firmware/profile, and probe-present all pass; an un-armed/void state shall be unmistakable. | S3/S4 / CMP-045/031 | DO-SAFE-007, DO-FW-032 | Pre-arm gate + arming/void test | HA-USE-003/004, HA-SW-003, HA-SVC-005
RCR-046 | INFO(design-enforced) | Charge-inhibited/fault/unready states shall be made legible with cause (LED/OLED); a cold, charge-inhibited device shall not be mistaken for a dead/discardable device. | S6 / CMP-030/032 | DO-SAFE-013 → DO-HF-005+DO-EE-008 | State-legibility test | HA-USE-009, HA-ENRG-001
RCR-047 | PROT | Alerting shall be tuned to validated false-alarm/missed-excursion targets that avoid alarm fatigue muting a real excursion; consolidation applied. | S4 / CMP-039/048 | DO-SAFE-004, DO-CLD-003 | Alarm-management validation | HA-USE-002, HA-SW-006
RCR-048 | PROT | Provisioning shall validate the selected trip profile against the product-stability library and flag an implausible/mismatched profile. | S4 / CMP-055/047 | DO-APP, DO-CLD | Profile-validation test | HA-USE-003, HA-SW-003
Service / maintenance
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-049 | PROT | Recalibration interval shall be enforced by the pre-arm expiry gate (RCR-026); an out-of-calibration device cannot be armed. | S3 / CMP-045 | DO-FW, DO-SVC | Interval-enforcement test | HA-SVC-001, HA-MFG-001
RCR-050 | PROC | Service shall restore seal, configuration, and firmware; a post-service BIST + pre-arm gate shall verify integrity before redeploy. | S1/S3 / CMP-045 | DO-SVC, DO-TST | Post-service BIST test | HA-SVC-003/005
RCR-051 | PROT | Diagnostic/developer/debug modes shall be locked/disabled before field deployment; an unexited service mode is detected. | S3 / CMP-044 | DO-SEC, DO-SVC | Mode-lock verification | HA-SVC-006, HA-CYBER-006
RCR-052 | INFO+PROC | Safe end-of-life instructions (battery removal/recycling, data/key erasure) shall be provided and supported by design. | Env / CMP-006/044 | DO-SVC, DO-LBL, DO-SEC | Disposal-procedure review | HA-SVC-007, HA-CYBER-008, HA-MPE-005
Manufacturing / supply-chain
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-053 | PROC | EOL calibration shall meet defined coverage and accuracy and record a birth record; a miscalibrated unit cannot pass EOL. | S4 / CMP-019/020 | DO-TST-... , DO-MFG | EOL calibration coverage audit | HA-MFG-001, HA-SVC-001
RCR-054 | PROC | Correct cell chemistry and correct-rated/oriented protection components shall be verified (poka-yoke + AOI); counterfeit/wrong parts prevented. | S2 / CMP-006/008/014 | DO-MFG, DO-TST | Incoming + AOI verification | HA-MFG-002/003, HA-ENRG-002
RCR-055 | PROC | Each device shall carry a unique cryptographic identity and serial; duplicate identity/serial prevented. | S3 / CMP-016 | DO-MFG (provisioning), DO-SEC | Identity-uniqueness check | HA-MFG-005, HA-CYBER-006
RCR-056 | PROC | Correct regional/firmware configuration shall be provisioned and verified at EOL (band plan, market profile). | S3 / CMP-027/035 | DO-MFG, DO-FW | Config-verification test | HA-MFG-006, HA-COMM-006
Overarching governance
RCR | Class | Requirement | Allocation | Impl | Verif | Controls
RCR-057 | PROC | The record-integrity chain (detection→verdict→record→export) shall preserve mission evidence such that a wrong disposition/lot-loss cannot arise from data loss or corruption (aggregate of RCR-019…024/042). | SYS | DO-SAFE-004…008 | Traceability + integration test | HA-MPE-001/002/004
RCR-058 | PROC | Overall residual risk and benefit-risk shall be evaluated per the ISO 14971 risk-management plan before release (HAL-2900). | SYS | Risk-Management File | Risk-management review | PRD-SAFE-011, HA-ORES-001
4. RCR discipline (per HAL-2400 standard)
• Every RCR is testable, allocated, and linked to a controlled design output and a verification.
• Single-fault tolerance applies to S4/S5 physical hazards (RCR-001/002 charge/battery) — an independent protective backstop is required and shall not share a common-cause with the primary control.
Not defeatable | RCR-001 (charge-inhibit) and RCR-021/022 (verdict integrity / honest gap) must not be bypassable by FW fault, config, or user action.
Independence & common-cause | the HW charge interlock (RCR-001) is independent of the FW; the secondary fuse (RCR-002) is independent of the BMS IC; device and cloud verdicts (RCR-021) cross-check.
5. Traceability / closure notes
This register feeds Stage 12 (Risk-Control Requirements) and re-baselines the preliminary PRD-SAFE-001…012. Each RCR's Implementation Reference is a DO-SAFE / DO- output (DOL-1200 + disciplines); each Verification Reference becomes a PV-…/VAL-… in Stages 13–14 (currently Evidence-Pending). New risks introduced by these controls are analysed in HAL-2700; residual risk after each is evaluated in HAL-2800/2900. High-severity controls (RCR-001/002/019/021/022/028/029) are the release-critical set.
HAL-2500 (HA-CIR) — Risk-Control Implementation Review
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-2500, record code HA-CIR · Owner: System-Safety (lead) with the discipline owners named per control, and Quality/Regulatory (EM-ENG-006). This module takes every accepted risk control from the HAL-2400 RCR register and reviews whether it is actually a control — not a spreadsheet promise. Per ISO 14971 §7 a risk-control measure must be implemented, verified, and free of unacceptable new risk; this module answers the first of those (implementation completeness and quality); HAL-2600 answers verification and HAL-2700 answers new risk. It is not an FMEA and not the RCR register itself: each record is a structured implementation audit of one control (or a tightly coupled control group). Standards cited by designation; revisions to confirm.
---
1. Purpose
Confirm, for each key SCM risk control, that it has been converted to a testable requirement (RCR), allocated to a subsystem, and implemented in a controlled design output (DO-…), and then interrogate the quality of that implementation against the attributes that decide whether a control genuinely reduces risk: independence, common-cause resistance, non-defeatability, dependency posture (connectivity / cloud / human response), fail-safe behaviour, observability/testability, and preservation into production. A control that is bypassable, common-cause-coupled to the hazard it guards, or silently dependent on a link that the operating environment (in-container, deep cold, offline) removes is not an adequate control and is flagged here for redesign — before verification (HAL-2600) is even attempted.
2. Scope
The release-critical control set and the controls the hazard registers most depend on, drawn from the HAL-2400 RCR register (RCR-001 … RCR-058) and their DO-SAFE / DO- implementations (DOL-1200 and disciplines):
• Charge-inhibit below 0 °C (RCR-001 / DO-SAFE-001) — the D3 physical top control.
• Cell protection + secondary backstop / thermal containment (RCR-002/004/008 / DO-SAFE-002/010, DO-MECH-005).
• Excursion-detection logic integrity (RCR-019 / DO-SAFE-004) — the software headline control.
• Verdict integrity + honest-gap (RCR-021/022/040 / DO-SAFE-006, DO-HF-002) — no-false-reassurance.
• Sensor/probe plausibility (RCR-020 / DO-SAFE-005).
• Alarm generation + delivery (RCR-028/029/030 / DO-SAFE-009).
• Calibration integrity + pre-arm gate (RCR-026/045 / DO-FW-007, DO-SAFE-007).
• Record-integrity / power-fail-safe storage (RCR-024 / DO-SAFE-008).
• Data authenticity / capture-time integrity (RCR-034 / DO-SEC).
• Secure, never-brick OTA (RCR-037 / DO-SAFE-014).
• Energy-reserve / loss-of-monitoring governor (RCR-006 / DO-SAFE-012).
• State legibility (charge-inhibited ≠ dead) (RCR-046 / DO-SAFE-013).
Lower-severity controls (mechanical edge/pinch, EMC, material selection, EOL screening) are implementation- reviewed in aggregate in §7; they are individually straightforward and carry no independence/defeatability tension.
3. Applicable standards & references
• ISO 14971 §6/§7 — risk-control implementation, verification of implementation, and the requirement that each control be traceable and effective.
• IEC 62304 — the safety-relevant firmware controls (detection, verdict, plausibility, storage, OTA) are treated at the highest software safety class (Class-C-equivalent); "implemented in a controlled DO" means an IEC 62304-controlled software item with detailed design and unit/integration verification.
• IEC 62368-1 / UN 38.3 / IATA DGR — the physical charge/battery/thermal controls (independence, single-fault tolerance, containment).
Upstream | HAL-2300 (control hierarchy — a control's class is checked here), HAL-2400 (the RCR register being audited), DOL-1200 (DO-SAFE implementations), fault-management-diagnostics.md and state-mode-model.md (fail-safe / safe-state architecture), the design-output register (DO-… allocations).
Evidence-honesty note | implementation quality is assessed against the design intent recorded in the DO-SAFE outputs and architecture; because the DOs are Draft/Preliminary and V&V is pending (Stages 13–14), every "adequate as designed" judgement here is Design-adequate / Evidence-Pending — the claim is that the design provides the attribute, and the named PV-/VAL- (HAL-2600) is what will prove it in silicon/firmware. No control is asserted as "verified"; that word is reserved for a closed PV-.
4. Governance & method — the implementation-review checklist
Each HA-CIR record answers a fixed 11-point checklist; every point is Yes / No / Partial / N-A with a one-line justification, and any No/Partial on a release-critical control raises a finding carried to the Risk-Management Summary:
1. Requirement? Converted to a testable RCR-### (HAL-2400)?
2. Allocated? Assigned to a specific subsystem/component (CMP-###)?
3. Implemented in a controlled DO? A released/controlled design output owns it (DO-SAFE-…/DO-)?
4. Independent? Independent of the element whose failure it guards (esp. FW-independent for physical safety)?
5. Common-cause resistant? Free of a shared failure with the primary path (shared clock, rail, chemistry, firmware image, calibration store, cloud tenant)?
6. Non-defeatable? Cannot be bypassed by FW fault, config, user action, or (for safety) a single fault?
7. Dependency posture? Does it depend on connectivity, the cloud, or a human response to function? (A safety control must not depend on the very link the environment removes.)
8. Fails safe? On its own failure, does the system go to a safe/honest state rather than a false-clean one?
9. Observable? Can the control's health/actuation be observed (self-check, status, diagnostic)?
10. Testable? Is there a defined way to verify it (fault-injection for safety controls) — the PV- in HAL-2600?
11. Production-preservable? Will EOL test, provisioning, and change-control keep it present and correct on every shipped/serviced unit (no silent regression)?
5. Categories & guiding questions
Physical-energy controls (charge/battery/thermal) are judged hardest on independence + single-fault + non-defeatability; information controls (detection/verdict/gap/data) on non-defeatability + fail-safe + observability; delivery controls (alert/cloud) on dependency posture + fail-safe (degrade-not-lose); lifecycle controls (calibration/OTA/config) on production-preservability + common-cause.
6. Control-Implementation-Review Register (HA-CIR records)
HA-CIR-001 — Charge-inhibit below 0 °C (RCR-001 / DO-SAFE-001) ★ RELEASE-CRITICAL
Control reviewed | MCU-independent hardware interlock that prevents charging below 0 °C (or cell-datasheet limit), with FW as defense-in-depth. Controls HA-ENRG-001 (cold-charge plating → latent short → runaway), HA-ELEC-002, HA-USE-009.
Class (HAL-2300) | ISD + PROT (inherent + protective) — the strongest tier.
• 1 Requirement? Yes — RCR-001 (testable: "charging prevented below 0 °C by an MCU-independent HW interlock; inhibit persists in any FW/comms fault").
• 2 Allocated? Yes — CMP-010 (charger) / CMP-012 (charge-path switch) with cell-coupled temperature sensing; S2.
• 3 Implemented in a controlled DO? Yes — DO-SAFE-001 → DO-EE-007 (HW interlock) + DO-FW-005 (FW guard). DO is Draft/Preliminary (reconcile-with-HAL flagged), so design-adequate / Evidence-Pending.
4 Independent? Yes (design intent) — the interlock senses cell temperature and gates the charge path in hardware, independent of the MCU/firmware; this is the load-bearing property. Finding | the temperature sensor the interlock trusts must not be the same element/channel used by the detection engine (avoid a shared-sensor common cause) — verify sensor independence in PV-ENRG-001.
• 5 Common-cause resistant? Partial — independent of FW, but shares the cell temperature measurement domain and the charge rail; a common-mode cold-sensor bias could in principle mis-read temperature. Mitigated by conservative threshold margin + the secondary protection IC (RCR-002) catching an out-of-window charge. Carried as a finding for fault-injection.
• 6 Non-defeatable? Yes (design intent) — HAL-2400 §4 mandates FW cannot defeat it; a self-heat feature may only enable charging by warming the cell (RCR-005), never bypass the inhibit. To be proven by the inhibit-integrity + self-heat-bound test.
• 7 Dependency posture? None — purely local hardware; no connectivity/cloud/human dependency. Correct for a physical safety control.
• 8 Fails safe? Yes — failure mode is inhibit charging (device may arrive under-charged → dead-in-cold information path, HA-USE-009 / HA-ENRG-006), which is a safe direction for the fire hazard; the residual is a monitoring-availability risk handled by RCR-006/046, not a safety risk.
• 9 Observable? Yes — charge-inhibited state + cause made legible (RCR-046 / DO-SAFE-013), so a cold, inhibited device is not mistaken for dead.
• 10 Testable? Yes — cold-charge fault-injection sweep (PV-ENRG-001, HAL-2600); fault-injection is mandatory for this S-class control.
11 Production-preservable? Yes (with finding) — correct charger/protection parts and orientation verified by poka-yoke + AOI (RCR-054); finding | the interlock threshold and sensor calibration must be covered by EOL test (RCR-053) so a mis-built interlock cannot escape.
Implementation verdict | Adequate as designed / Evidence-Pending. The single most important physical control; independence and non-defeatability are correctly specified. Open findings: sensor independence from the detection channel, cold-sensor common-mode margin, EOL coverage of the threshold.
Owner | Battery/Power Safety with System-Safety · Review History: 2026-07-19 v1.0 initial (ORCH-017)
HA-CIR-002 — Cell protection + secondary backstop + thermal containment (RCR-002/004/008) ★ RELEASE-CRITICAL
Control reviewed | Autonomous OV/UV/OC/OT/short protection with a non-resettable secondary backstop (fuse/OVP) (RCR-002); cell retention without crush/puncture (RCR-004); fire-enclosure containment of a single-cell event (RCR-008). Controls HA-ENRG-002/003, HA-TFC-002 (occupied-aircraft fire S5), HA-TFC-003.
Class | PROT + ISD (protection + retention + inherent fire enclosure).
• 1 Requirement? Yes — RCR-002/004/008, each testable.
• 2 Allocated? Yes — CMP-008/009 (BMS + fuse), CMP-007 (pack/retention), CMP-001/034 (enclosure/fire barrier).
• 3 Controlled DO? Yes — DO-SAFE-002 → DO-EE-005 (BMS) + DO-EE-010 (fuse); DO-SAFE-010 → DO-MECH-012 + DO-EE-004; DO-MECH-005 + DO-PCB-010 (fire enclosure). Evidence-Pending.
• 4 Independent? Yes — the secondary fuse/OVP is independent of the primary protection IC (RCR-002 explicitly single-fault-safe); the fire enclosure is independent of both cell protection layers (a containment layer that assumes the electrical layers already failed). This is the correct defense-in-depth for an S5.
5 Common-cause resistant? Yes (design intent) — fuse is non-resettable and mechanically/thermally distinct from the IC; enclosure containment does not share a failure with the electronics. Finding | confirm the protection IC and fuse do not share a single sense resistor / current path that could open-circuit both senses.
• 6 Non-defeatable / single-fault-safe? Yes — HAL-2200 §6.5 single-fault rule is satisfied: no single fault (IC failure, FW fault, one mechanical shock) reaches the S5 outcome because an independent backstop or containment layer remains.
• 7 Dependency posture? None — fully autonomous local hardware; correct.
• 8 Fails safe? Yes — protection trips to a de-energised/limited state; the fuse fails open (safe); the enclosure contains rather than propagates.
• 9 Observable? Partial — a tripped protection / blown fuse renders the device inert; state legibility (RCR-046) and depot triage detect a non-functional unit, but a latent weakened fuse is not directly observable in field — accepted, screened at EOL/service.
• 10 Testable? Yes — protection fault-injection (PV-ENRG-002 / DO-TST-009), UN 38.3 abuse suite, single-cell containment / fire-enclosure abuse (PV-TFC-002). Mandatory fault-injection.
• 11 Production-preservable? Yes — correct cell chemistry + correct-rated/oriented protection parts verified by incoming + AOI (RCR-054); UN 38.3 is a transport prerequisite per lot/design change.
Implementation verdict | Adequate as designed / Evidence-Pending. Proper independent, single-fault-safe, non-defeatable defense-in-depth for the S5 fire path. Open findings: shared-sense-path check (IC↔fuse), enclosure containment proof.
Owner | Thermal/Battery-Safety + Mechanical + System-Safety · Review History: 2026-07-19 v1.0 initial
HA-CIR-003 — Excursion-detection logic integrity (RCR-019 / DO-SAFE-004) ★ RELEASE-CRITICAL
Control reviewed | Two-sided (heat and freeze) excursion detection with freeze reliability equal to heat, no undetected-excursion path. Controls HA-SW-001 (missed excursion, headline), HA-ENV-001 (freeze miss), HA-DATA-001.
Class | ISD (inherent — no undetected-excursion path by construction) + PROC (IEC 62304 verification).
• 1 Requirement? Yes — RCR-019.
• 2 Allocated? Yes — CMP-039 (detection engine), fed by CMP-038 (sensor drivers).
• 3 Controlled DO? Yes — DO-SAFE-004 → DO-FW-006 + DO-EE-013; Class-C-equivalent IEC 62304 item. Evidence-Pending.
4 Independent? Partial (by design, needs a caveat) — detection is a single firmware function; its independent cross-check is the cloud recompute (RCR-021, DO-CLD-006) computing a verdict from the same quality-annotated record by a separate code path. That cross-check is the independence layer, but it is cloud-and-connectivity-dependent (see point 7). On-device, detection has no independent twin. Finding | the device must not silently depend on the cloud recompute to catch a detection defect; the on-device logic must be correct-by-verification (PV-SW-001), with the cloud as a second check, not the only one.
• 5 Common-cause resistant? Partial — device and cloud detection share the product-stability profile and the calibration coefficients; a wrong profile (HA-SW-003) or biased calibration (HA-SW-004) is a common cause that defeats both. Mitigated by profile validation (RCR-048) and calibration integrity (RCR-026), which are separate controls reviewed in HA-CIR-007. This coupling is the reason those controls are also release-relevant.
• 6 Non-defeatable? Yes (design intent) — "no undetected-excursion path" is an inherent requirement; freeze logic must be as reliable as heat. Proven by the profiled-excursion test with explicit equal freeze/heat coverage.
• 7 Dependency posture? Detection: none (local). Its independent cross-check: cloud + connectivity-dependent. The primary detection is correctly local and offline-capable; only the second-opinion recompute needs the link. Acceptable provided on-device detection is treated as authoritative and verified, not as a draft the cloud fixes.
• 8 Fails safe? Yes (design intent) — a detection fault must not produce a false clean pass; paired with verdict suppression (RCR-021) so an uncertain/incomplete evaluation degrades to a qualified/reject, not an accept.
• 9 Observable? Partial — a missed excursion has no signature (stated in HA-SW-001); the control's health is observed indirectly via device↔cloud verdict divergence and the profiled-excursion test, not in field. This is inherent to the hazard and is why verification rigor (not field detection) is the defence.
• 10 Testable? Yes — detection-logic test on a library of profiled excursions (heat/freeze/boundary/ramp/dwell) + static/code analysis + unit/integration per IEC 62304 (PV-SW-001).
11 Production-preservable? Yes (with finding) — the logic ships in firmware; finding | OTA (RCR-037) and depot firmware load (HA-SVC-002) can deploy wrong detection logic — the detection control's field integrity depends on the OTA/firmware-provenance control (HA-CIR-010) and config verification (RCR-056). Version-skew between device and cloud logic must be prevented.
Implementation verdict | Adequate as designed / Evidence-Pending — with two structural findings: (a) do not let the cloud recompute mask an unverified on-device logic; (b) profile + calibration are common-cause inputs that must be independently controlled. Both are already addressed by named controls; the finding is to hold their verification as prerequisites to accepting HA-SW-001 residual.
Owner | Embedded/Software Safety + Cloud + System-Safety · Review History: 2026-07-19 v1.0 initial
HA-CIR-004 — Verdict integrity + honest-gap (RCR-021/022/040 / DO-SAFE-006, DO-HF-002) ★ RELEASE-CRITICAL
Control reviewed | Verdict suppresses a false "accept" on any gap/sensor-fault/incomplete monitoring; any gap made visible; the disposition surface never presents a gap/fault as a clean pass; device and cloud verdicts consistent. Controls HA-SW-002 (false accept), HA-DATA-004 (automation-bias headline), HA-SW-007, HA-SW-013.
Class | ISD (make the false-accept-over-gap state unrepresentable) + design-enforced INFO (qualifier always shown).
• 1 Requirement? Yes — RCR-021/022/040. 2 Allocated? Yes — CMP-040 (verdict), CMP-049 (cloud recompute), CMP-030/050/055 (surfaces). 3 Controlled DO? Yes — DO-SAFE-006 → DO-FW-008 + DO-CLD-006 + DO-HF-002. Evidence-Pending.
• 4 Independent? Yes — the two verdict computations (device CMP-040 and cloud CMP-049) are independent code paths that must agree; divergence is itself detectable (a genuine independent cross-check, unlike detection which is single-path on-device).
• 5 Common-cause resistant? Partial — same product-stability profile feeds both (shared with HA-CIR-003); mitigated by profile validation. The gap/quality flags are a shared input — if storage writes a wrong "no gap" flag (HA-SW-013 / RCR-024), both verdicts trust it. So verdict integrity is only as good as record integrity (HA-CIR-008) — a stated dependency, not a defect.
• 6 Non-defeatable? Yes — HAL-2400 §4 marks the honest-gap / no-false-reassurance invariant as non-defeatable by FW fault, config, or user action; an override to accept against an honest reject requires a deliberate, recorded action (RCR-044).
• 7 Dependency posture? Device verdict: local (edge read on OLED works offline). Cloud verdict: connectivity-dependent second check. Correct posture — the operator can get an honest verdict at receipt with zero connectivity (edge read, AD-5).
• 8 Fails safe? Yes — the defining behaviour: on any doubt (gap, fault, divergence) the verdict is not a clean accept; it qualifies or rejects. This is the module's core fail-safe.
• 9 Observable? Yes — device↔cloud divergence is monitored; the qualifier is always visible; availability self-check (RCR-027) flags a non-monitoring device.
• 10 Testable? Yes — verdict-consistency + gap-handling test (PV-SW-002 / PV-DATA-004): inject gapped/faulted/incomplete records, confirm no unqualified accept and device=cloud; summative usability (VAL-USE-004) proves the human reads it correctly.
• 11 Production-preservable? Yes — logic in firmware + cloud; version-skew controlled with OTA (HA-CIR-010).
Implementation verdict | Adequate as designed / Evidence-Pending. The strongest information control — genuine independence (device vs cloud), non-defeatable, fail-safe, offline-capable. Dependency on record-integrity flags (HA-CIR-008) is explicit and controlled.
Owner | Embedded/Software Safety + Human-Factors + Cloud · Review History: 2026-07-19 v1.0 initial
HA-CIR-005 — Sensor / probe plausibility (RCR-020 / DO-SAFE-005)
Control reviewed | Absent/open/short/wrong probe and implausible reads are detected and flagged, never read as a valid product-zone temperature. Controls HA-SW-008 (probe reads room temp → false accept, headline), HA-USE-005/006, HA-DATA-001.
Class | PROT (fault detection) + ISD (never read invalid as valid, by construction).
1–3 | RCR-020; CMP-020 (probe) + CMP-038 (drivers/plausibility); DO-SAFE-005 → DO-FW-007 + DO-EE-013 (HW presence/open/short detect). Evidence-Pending.
• 4 Independent? Yes — pairs hardware presence/open/short detection (DO-EE-013) with firmware plausibility (range/rate/stuck/int-vs-ext divergence); two mechanisms, so a single-layer miss does not read a floating input as valid.
5 Common-cause resistant? Partial — an in-range but wrong value (probe in air not product, HA-USE-005; room temp inside acceptance window) can defeat both HW and FW checks because the signal is electrically valid. This is the known residual: plausibility cannot detect a physically-mis-placed but electrically-healthy probe — mitigated by pre-arm probe-present gate (RCR-045), internal-vs-external divergence, and use-procedure/HF (HA-USE-005). Finding | carry HA-USE-005 (placement) as the un-closable-by-plausibility remainder.
• 6 Non-defeatable? Yes — a detected probe fault blocks a clean pass and the pre-arm gate requires probe-present.
• 7 Dependency posture? None — local. 8 Fails safe? Yes — a flagged probe fault drives qualify/reject, never a clean accept. 9 Observable? Yes — fault flagged in record + status. 10 Testable? Yes — probe-fault-injection (PV-SW-008). 11 Production-preservable? Yes — front-end + firmware; EOL functional check of presence/open/short detect.
Implementation verdict | Adequate as designed / Evidence-Pending, with the explicit remainder that placement error (HA-USE-005) is a human-factors control, not a plausibility control.
Owner | Embedded/Software Safety + Metrology + Human-Factors · Review History: 2026-07-19 v1.0 initial
HA-CIR-006 — Alarm generation + delivery (RCR-028/029/030 / DO-SAFE-009) ★ RELEASE-CRITICAL
Control reviewed | Every excursion committed locally independent of connectivity (RCR-028); alerts sent ahead of bulk, ack-required, escalate if unacknowledged (RCR-029); connectivity loss degrades to store-and-forward, never record loss, alert-first on reconnection (RCR-030). Controls HA-SW-006 (alarm not generated), HA-COMM-001/002 (delayed/missing alert), HA-COMM-003, HA-CLOUD-001.
Class | ISD (local commit) + PROT (ack/escalate/degrade).
1–3 | RCR-028/029/030; CMP-041/042 (log + comms FW), CMP-048 (cloud notify); DO-SAFE-009 → DO-FW-011 + DO-CLD-007/008. Evidence-Pending.
• 4 Independent? Yes — the existence of the excursion (local commit) is independent of the delivery of the alert (network/cloud); this is the load-bearing split (HAL-1100 §1). A delivery failure cannot erase the excursion.
5 Common-cause resistant? Partial — device-side generation and delivery share the comms firmware (CMP-042); a comms-agent defect could both fail to emit and fail to flag the failed dispatch. Mitigated by ack/escalation (a never-received alert self-detects) and end-to-end reconciliation (SYS-176). Finding | the failed-dispatch condition must itself be logged locally so a comms-agent defect is observable at receipt.
• 6 Non-defeatable? Yes (design intent) — the alert path shall not be silently suppressible; suppression/loss raises a detectable event (RCR-035, cyber twin).
• 7 Dependency posture? This is the critical one: delivery is connectivity- AND cloud-dependent by nature. The control is designed around that: it converts a delivery failure from data-loss into latency, and provides an offline compensating detection (on-receipt OLED edge read). The residual (finite in-transit dark window, RF-in-container HA-COMM-003) cannot be removed by this control — it is bounded, not eliminated, and is the carried connectivity residual.
• 8 Fails safe? Yes — on link loss it buffers and delivers alert-first; it never assumes sent=received.
• 9 Observable? Yes — un-acked backlog visible to the fleet; connectivity diagnostics (SYS-113).
• 10 Testable? Yes — alert-latency + store-and-forward test (PV-SW-006 / PV-COMM-001), delivery-path fault-injection (PV-COMM-002), RF-in-container qualification (PRD-VER-005).
• 11 Production-preservable? Yes — firmware + cloud config; regional config verified at EOL (RCR-056).
Implementation verdict | Adequate as designed / Evidence-Pending — with the explicit, un-removable dependency: live alerting depends on a link the operating environment can remove; the control bounds (not eliminates) the consequence via local-commit + edge-read + escalation. HA-COMM-003 (RF-in-container) stays Open until PRD-VER-005.
Owner | Connectivity + Cloud + Firmware + System-Safety · Review History: 2026-07-19 v1.0 initial
HA-CIR-007 — Calibration integrity + pre-arm gate (RCR-026/045 / DO-FW-007, DO-SAFE-007)
Control reviewed | Calibration coefficients integrity-protected; expired/invalid calibration blocks arming; the pre-arm readiness gate blocks arming unless SoC-for-trip, valid/unexpired calibration, valid firmware/profile, and probe-present all pass. Controls HA-SW-004 (stale/corrupt calibration), HA-SW-014, HA-SVC-001/MFG-001 (calibration escape), HA-USE-003/004, HA-SVC-005.
Class | PROT (integrity + expiry) + ISD (unrepresentable armed-with-invalid-calibration state) + PROC (calibration lifecycle).
1–3 | RCR-026/045; CMP-038 (calibration), CMP-045 (readiness gate), CMP-033/041 (NV store); DO-FW-007 + DO-SAFE-007 (→ DO-FW-032). Evidence-Pending.
4 Independent? Partial — the gate that enforces validity runs in the same firmware whose calibration it checks; independence is provided by an integrity primitive (signature/CRC) over the coefficient block so corruption is detected rather than trusted. Finding | the integrity key/anchor must not be defeatable by the same fault that corrupts the block.
• 5 Common-cause resistant? Partial — calibration is a common-cause input to detection AND verdict (see HA-CIR-003/004); a systematic fleet-wide calibration bias (HA-DATA-001 S5-worst) defeats everything downstream. This is the common-cause of the information chain — mitigated by EOL calibration coverage + birth record (RCR-053) and interval enforcement, but it is the reason a fleet-wide calibration escape is an S5 in HAL-2900.
• 6 Non-defeatable? Yes — expired/invalid calibration blocks arming; a device cannot be silently armed on bad calibration.
• 7 Dependency posture? None at runtime (local gate). Calibration provisioning depends on depot/manufacturing process (HA-SVC-001/HA-MFG-001) — a process dependency, controlled by RCR-049/053.
• 8 Fails safe? Yes — failure blocks arming (safe: an un-armed device is unmistakable, not a falsely-monitoring one).
• 9 Observable? Yes — un-armed/void state made unmistakable (RCR-045); calibration status shown.
• 10 Testable? Yes — calibration-integrity (corrupt-block rejection) + expiry-gate test (PV-SW-004), EOL calibration coverage audit (PV-MFG-001).
• 11 Production-preservable? Partial → finding. This is the hardest preservation case: reuse + depot recalibration (D1) means calibration is re-established many times over life; a depot calibration escape (HA-SVC-001) reintroduces the bias the control exists to prevent. Preservation depends on the depot process control + interval gate + post-service BIST (RCR-049/050). Carried as a standing finding to Service/Manufacturing.
Implementation verdict | Adequate as designed / Evidence-Pending — calibration is the information chain's dominant common-cause; its lifecycle (EOL + depot recal + interval gate) is a release-relevant process, not a one-time design fact.
Owner | Metrology + Embedded/Software Safety + Service/Manufacturing · Review History: 2026-07-19 v1.0 initial
HA-CIR-008 — Record-integrity / power-fail-safe storage (RCR-024 / DO-SAFE-008)
Control reviewed | Power-fail-safe storage (atomic writes, brown-out barrier); no torn/lost/fabricated record; no overwrite of unsent data; tamper-evident. Controls HA-SW-013 (silent gap), HA-DATA-006 (silent data loss), HA-ELEC-008/010, HA-ENRG-006.
Class | PROT (atomic/barrier) + design-enforced INFO (a loss surfaces as an honest gap).
1–3 | RCR-024; CMP-041/033 (storage/NV); DO-SAFE-008 → DO-FW-010 + DO-EE-002 (rail interlock) + DO-SEC-006 (tamper-evidence). Evidence-Pending.
• 4 Independent? Yes — the brown-out barrier is a hardware rail interlock (DO-EE-002) independent of firmware timing; it prevents a write from starting when the rail cannot complete it. HW + FW two-layer.
5 Common-cause resistant? Partial — cold-induced voltage sag (D3) is a common stressor for both the write and the barrier; margin over the −30 °C brown-out profile is the mitigation, proven by test. Finding | validate the barrier threshold at the coldest, lowest-SoC corner.
• 6 Non-defeatable? Yes — a lost write becomes an honest gap (RCR-022), never a silent absence or a fabricated sample.
• 7 Dependency posture? None — local. 8 Fails safe? Yes — the defining property: any interruption yields a gap-flagged record, which the verdict then refuses to clean-pass (ties HA-CIR-004).
• 9 Observable? Yes — gap visible in record + status; tamper-evidence detectable. 10 Testable? Yes — power-fail-during-write test over the cold profile (PV-SW-013 / PV-DATA-006).
• 11 Production-preservable? Yes — firmware + rail HW; brown-out threshold set in hardware, screened at EOL.
Implementation verdict | Adequate as designed / Evidence-Pending. Good HW/FW independence; the honest-gap fail-safe is what makes verdict integrity (HA-CIR-004) trustworthy. Open finding: cold-corner barrier margin.
Owner | Embedded/Software Safety + Electrical · Review History: 2026-07-19 v1.0 initial
HA-CIR-009 — Data authenticity / capture-time integrity (RCR-034 / DO-SEC)
Control reviewed | Telemetry and the record carry capture-time authenticity/integrity (tamper-evidence); falsified/injected data detectable and rejected. Controls HA-CYBER-001 (falsified data, headline), HA-DATA-001, HA-COMM-004 (transport integrity), HA-CYBER-004 (replay).
Class | PROT (cryptographic integrity).
1–3 | RCR-034; CMP-016 (secure element/identity), CMP-044 (crypto), CMP-049 (cloud verify); DO-SEC (integrity) + DO-SAFE-008. Evidence-Pending.
• 4 Independent? Yes — integrity is anchored at capture time in the device secure element (CMP-016), independent of the transport and cloud; the cloud verifies against the device root of trust, so a compromised link or cloud node cannot forge a valid record (the anchor is the device, not the wire — SYS-132).
5 Common-cause resistant? Partial → finding. The whole scheme rests on the device key / secure element; key compromise or remanence (HA-CYBER-008) is a common cause that defeats authenticity. Mitigated by secure boot + debug-lock (RCR-038) and disposal key-erasure (RCR-052). Finding | the key-management lifecycle (provisioning uniqueness RCR-055, disposal erasure) is the common-cause boundary of this control and must be verified end-to-end.
• 6 Non-defeatable? Yes (design intent) — injected/replayed/tampered data fails verification and is rejected; replay resisted by capture-time + sequence binding.
• 7 Dependency posture? None for integrity generation (local at capture); verification is cloud-side but the anchor is the device. Correct — a falsified record is detectable even offline at the edge read because the device signs its own record.
• 8 Fails safe? Yes — a record that fails integrity is treated as untrusted/rejected, not accepted.
• 9 Observable? Yes — integrity failures surfaced and monitored (SRS-CLOUD-024). 10 Testable? Yes — integrity/pen-test (PV-CYBER-001), reorder/duplicate/corrupt injection (PV-COMM-004). Independent security verification required (SC control).
• 11 Production-preservable? Yes (with finding) — unique cryptographic identity provisioned + verified at EOL (RCR-055); debug ports locked (RCR-038). Preservation = key lifecycle discipline (the finding above).
Implementation verdict | Adequate as designed / Evidence-Pending. Correct device-anchored integrity. The control's boundary is the key lifecycle (provisioning→operation→disposal), which must be independently security-verified.
Owner | Product-Security + Cloud + System-Safety · Review History: 2026-07-19 v1.0 initial
HA-CIR-010 — Secure, never-brick OTA (RCR-037 / DO-SAFE-014)
Control reviewed | OTA signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; never bricks a device; unmonitored windows disclosed. Controls HA-SW-010 (unsafe OTA), HA-CYBER-005 (malicious OTA, headline), HA-CLOUD-007 (uncoordinated rollout), HA-COMM-005 (cert lifecycle).
Class | PROT (signed A/B + rollback) + design-enforced INFO (disclose unmonitored window).
1–3 | RCR-037; CMP-035/043 (bootloader/OTA agent), CMP-051 (campaign), CMP-016 (root of trust); DO-SAFE-014 → DO-FW-014/015 + DO-CLD-010. Evidence-Pending.
• 4 Independent? Yes — A/B slots with an immutable bootloader verifying a signature against a hardware root of trust (CMP-016); the update mechanism is independent of the running image and rolls back on failed health-check.
5 Common-cause resistant? Partial → finding (fleet common-cause). OTA is the fleet-wide common-cause vector: a bad-but-signed image, or an uncoordinated rollout, can deploy wrong detection logic (HA-CIR-003) or blind many devices at once. Mitigated by staged/canary rollout + automatic rollback + version-skew control (RCR-037 + HA-CLOUD-007). Finding | rollout policy (canary %, health gate, halt-on-regression) is a release-critical operational control, not just a device feature.
• 6 Non-defeatable? Yes — unsigned/rolled-back images rejected; secure boot enforced (RCR-038).
• 7 Dependency posture? Cloud + connectivity-dependent (campaign) but device-safe on its own — a device that never receives an OTA keeps running the last good image; an interrupted OTA rolls back. So the safety does not depend on the link; only the update does.
• 8 Fails safe? Yes — failed apply → automatic rollback to last good; trip data preserved; never bricks.
• 9 Observable? Yes — version state + unmonitored-window disclosure; campaign health monitored.
• 10 Testable? Yes — OTA rollback + version-skew test (PV-SW-010), OTA pen-test (PV-CYBER-005). Independent security verification (SC).
• 11 Production-preservable? Yes — signing keys + rollout policy under change control; secure boot verified at EOL.
Implementation verdict | Adequate as designed / Evidence-Pending. Device-side OTA safety is sound (A/B, rollback, secure boot). The residual is the operational rollout policy as a fleet common-cause — held as a release-critical process finding.
Owner | Firmware + Product-Security + Cloud · Review History: 2026-07-19 v1.0 initial
HA-CIR-011 — Energy-reserve / loss-of-monitoring governor (RCR-006 / DO-SAFE-012)
Control reviewed | Energy governor + reserve prevents silent loss of monitoring; reserve preserves record-commit + a low-battery alert; honest low-battery/dead-in-cold indication; 90-day cold endurance met. Controls HA-ENRG-006 (dead-in-cold, headline), HA-ENRG-008, HA-SW-013.
1–3 | RCR-006; CMP-037 (governor) / CMP-011 (fuel gauge); DO-SAFE-012 → DO-FW-004 + DO-EE-008. Evidence-Pending.
• 4 Independent? Partial — the governor is firmware; the reserve is a physical energy margin. The honest "dead-in-cold" indication depends on the device still having enough energy to say it is dying — a reserve floor is the design answer.
• 5 Common-cause resistant? Partial → key tension (D3). The −30 °C / 90-day / capacity-fade combination (D3, KEY RISK) is the common stressor; a mis-estimated cold capacity defeats both endurance and the reserve. This is Evidence-Pending on the battery datasheet + 90-day cold endurance test — the largest energy uncertainty.
• 6 Non-defeatable? Yes — reserve is protected; a low-battery state cannot be silently ignored (honest indication + alert).
• 7 Dependency posture? Low-battery alert delivery is connectivity-dependent, but the record-commit on reserve and the on-device honest indication are local; a silent gap becomes an honest gap even offline.
• 8 Fails safe? Yes — exhaustion produces an honest gap + low-battery indication, not a silent uneventful-looking trip.
• 9 Observable? Yes — honest low-battery/dead-in-cold indication (ties RCR-046). 10 Testable? Yes — 90-day cold endurance (PRD-VER-004) + reserve-floor + honest-indication test.
• 11 Production-preservable? Partial — depends on cell chemistry sourced (RCR-054) and capacity-fade over reuse (HA-ENRG-008) staying within the budget over service life; interval/health checks at depot.
Implementation verdict | Design-adequate / Evidence-Pending — dominated by the D3 cold-endurance unknown. The honest-gap behaviour is sound; the endurance margin is a top open item closed only by the 90-day cold test.
Owner | Battery/Power + Embedded/Software Safety · Review History: 2026-07-19 v1.0 initial
HA-CIR-012 — State legibility: charge-inhibited ≠ dead (RCR-046 / DO-SAFE-013)
Control reviewed | Charge-inhibited/fault/unready states made legible with cause (LED/OLED); a cold, charge-inhibited device not mistaken for dead/discardable. Controls HA-USE-009 (inhibit mistaken for dead → stock-out), ties HA-ENRG-001/006.
1–3 | RCR-046; CMP-030/032 (OLED/LED); DO-SAFE-013 → DO-HF-005 + DO-EE-008. Evidence-Pending.
4 Independent? N-A (an information/HMI control). 5 Common-cause? shares the HMI; if the OLED is dead in extreme cold (HA-ENV-002) the LED is the fallback legibility channel — finding | cause must be legible on the low-power LED alone when the OLED is cold-slow.
• 6 Non-defeatable? N-A. 7 Dependency posture? None — local HMI. 8 Fails safe? Yes — a cold, inhibited device shows why, so it is not discarded (a false-discard is the failure this prevents). 9 Observable? self-evidently (it is the observability). 10 Testable? cold-display + state-legibility test (PV-USE / cold). 11 Production-preservable? Yes — HMI firmware; EOL display/LED check.
Implementation verdict | Adequate as designed / Evidence-Pending, with the cold-OLED fallback finding (LED must carry cause legibly at −30 °C).
Owner | Human-Factors + Electrical · Review History: 2026-07-19 v1.0 initial
7. Aggregate review of lower-severity controls
The remaining RCRs are implementation-reviewed in aggregate; none carries an independence/defeatability tension and all are Design-adequate / Evidence-Pending:
Mechanical safety (RCR-009/010/011/012) | edge/pinch, small-parts retention, ingress/containment, fastener retention — inherent-by-design, local, fail-safe (a cracked enclosure is inspectable), testable (edge/pinch inspection, drop/vibe/IP), production-preserved by mould/tool control + inspection.
Electrical/EMC (RCR-013/014/015/016) | SELV, EMC immunity/emissions, conformal-coat protection, latent- defect screening — protective/inherent, testable (IEC 62368-1/61000, AOI/ICT), preserved by EOL coverage.
Environmental (RCR-017/018) | −30…+60 °C operation + condensation tolerance — inherent (materials/coat), testable (IEC 60068), the cold-endurance part ties the D3 unknown (HA-CIR-011).
Time integrity (RCR-025) | bounded/disclosed drift — protective + design-enforced info, local, fail-safe (untrusted time disclosed), testable (clock-drift test).
Cyber cluster (RCR-035/036/038/039) | non-suppressible alarm, authorised config change, secure boot/debug- lock, cloud resilience — protective, require independent security verification (SC), testable by pen-test.
Association/binding (RCR-033/055) | authoritative device↔trip↔owner binding + unique identity — inherent, testable (association/routing + identity-uniqueness), production-preserved by provisioning verification.
Data-decision info controls (RCR-040/041/042) | honest surface, estimate-flagging, integrity-preserving export — design-enforced info (survive the warnings-challenge), validated by summative usability.
Human-factors (RCR-043/044/045/047/048) | legible verdict, recorded override, pre-arm gate, alarm tuning, profile validation — inherent/protective, validated by summative usability; the pre-arm gate (RCR-045) is reviewed under HA-CIR-007.
Service/manufacturing (RCR-049…056) | calibration interval, service restore/BIST, mode-lock, disposal, EOL calibration/parts/identity/config — process controls; their common thread is production/service preservation (checklist point 11), which HA-CIR-007/009/010 identify as the recurring risk.
Governance (RCR-057/058) | record-integrity aggregate + overall benefit-risk — meta-controls closed in HAL-2900 / Risk-Management Summary.
8. Findings summary (carried to Risk-Management Summary)
1. Sensor independence (HA-CIR-001/003): the temperature channel trusted by the charge interlock and by the detection engine must be independently verifiable; avoid a shared-sensor common cause.
2. Calibration as dominant common-cause (HA-CIR-003/004/007): detection + verdict + charge-margin all trust calibration; a fleet-wide calibration escape is an S5 driver (HAL-2900). Hold EOL + depot-recal + interval-gate verification as prerequisites to accepting the information-chain residuals.
3. Cloud recompute must not mask unverified on-device logic (HA-CIR-003): on-device detection/verdict must be correct-by-verification, with the cloud as a second check, not the only one.
4. Delivery dependency is un-removable (HA-CIR-006): live alerting depends on a link the environment can remove; the consequence is bounded (local-commit + edge-read + escalation), not eliminated — HA-COMM-003 stays Open until PRD-VER-005.
5. OTA rollout policy is a fleet common-cause (HA-CIR-010): canary/health-gate/halt-on-regression is a release-critical operational control.
6. Key lifecycle bounds data authenticity (HA-CIR-009): provisioning-uniqueness + disposal-erasure must be security-verified end-to-end.
7. D3 cold-endurance unknown (HA-CIR-011): the reserve/governor is design-adequate but the endurance margin is Evidence-Pending on the battery datasheet + 90-day cold test.
8. Production/service preservation (HA-CIR-007 and §7): reuse + depot service can silently regress calibration, seal, firmware, and config; post-service BIST + pre-arm gate are the preservation net.
All controls reviewed are Design-adequate / Evidence-Pending; none is judged inadequate-as-designed. The findings are prerequisites and watch-items, not redesign demands. Verification of each control is specified in HAL-2600; new risks introduced by these controls in HAL-2700; residual risk after control in HAL-2800/2900.
HAL-2600 (HA-RCV) — Risk-Control Verification
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-2600, record code HA-RCV · Owner: Quality/Regulatory (EM-ENG-006) with System-Safety and the discipline verification owners. This module defines how each accepted risk control will be shown to work — the verification method, the evidence type, the acceptance criterion, and the planned PV-… (product verification, Stage 13) or VAL-… (validation, Stage 14) reference. Per ISO 14971 §7, every risk-control measure must be verified for implementation and for effectiveness; HAL-2500 confirmed implementation, this module plans the effectiveness verification. It is the bridge from the RCR register (HAL-2400) to the Product Verification Library (Stage 13) and Validation Library (Stage 14). Everything here is planned / Evidence-Pending: no PV-/VAL- has yet been executed. Standards cited by designation; revisions to confirm.
---
1. Purpose
For every risk control, specify the method (analysis / inspection / test / fault-injection / demonstration / summative usability), the evidence type it produces, and the acceptance criterion that closes the control's Evidence-Pending probability in the hazard registers — and bind it to a named PV-/VAL- in Stages 13–14. Two governing rules are made explicit here: (a) safety controls require fault-injection (proving the control acts when the fault occurs, not merely that it exists in nominal operation); and (b) safety-critical (SC) and security controls require independent verification — verified by someone other than the implementer, and for security by adversarial test (pen-test), not developer self-test.
2. Scope
Verification planning for the release-critical control set (HA-CIR-001…012) and the full RCR register (RCR-001…058), mapped to the PV-/VAL- identifiers already seeded across the domain modules (PV-ENRG-###, PV-TFC-###, PV-SW-###, PV-COMM-###, PV-CLOUD-###, PV-CYBER-###, PV-DATA-###, PV-USE-###, PV-MFG-###, and the product-level PRD-VER-004/005), plus VAL-USE-### summative usability. This module does not execute verification (that is Stage 13/14) — it defines the plan the hazard analysis requires, so that Stage 13 cannot close without covering every safety control.
3. Applicable standards & references
• ISO 14971 §7 — verification of risk-control implementation and effectiveness.
• IEC 62304 §5.5–5.7 — software unit/integration/system verification; the SC firmware controls (detection/verdict/plausibility/storage/OTA) require documented unit + integration + system tests plus static/code analysis at Class-C-equivalent rigor.
• IEC 62366-1 §5.9 — summative (validation) usability for the decision-support and legibility controls.
• IEC 62368-1 / IEC 62133-2 / UN 38.3 / IATA DGR — the abuse/fault tests for the physical charge/battery/ thermal controls (fault-injection is intrinsic to these suites).
• IEC 60068 / IEC 60529 / IEC 61000 — environmental, ingress, and EMC verification.
Upstream | HAL-2400 (RCR register + Verif column), HAL-2500 (implementation review + findings), the domain modules (which each name their PV-/VAL-), the Verification standard (ORCH-014) and Validation standard (ORCH-015) that will own the executed evidence.
Evidence-honesty note | every entry below is a plan; acceptance criteria are stated as targets, not results. Where a control's probability is currently escalated for uncertainty (D3 battery, RF-in-container), the corresponding PV-/PRD-VER- is the named closing test and the control stays Open until it reports.
4. Governance & method — verification classes
Fault-injection (FI) | required for every safety control — inject the fault the control guards and confirm the control acts and the system reaches the safe/honest state. (Nominal-operation tests are necessary but not sufficient for a safety control.)
Independent verification (IV) | required for SC controls (S4/S5 contributors) and all security controls — performed by a party independent of the implementer; security controls additionally require adversarial verification (pen-test).
Evidence types | Test report (bench/chamber/abuse), FI log, static/code-analysis report, chamber/endurance data, EMC/abuse certificate, usability study report, security assessment/pen-test report, audit record.
Closure | a control's Evidence-Pending residual (HAL-2800) closes only when its PV-/VAL- passes to its acceptance criterion; SC controls additionally need the IV sign-off.
5. Verification Register (HA-RCV records)
Legend — Method: FI = fault-injection · T = test · A = analysis · I = inspection · D = demonstration · SU = summative usability. IV = independent verification required (SC/security). Status: all Evidence-Pending (Stage 13/14).
Physical: charge / battery / thermal (single-fault-safe, fault-injection mandatory)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-001 | Charge-inhibit <0 °C (RCR-001) | FI + T | Cold-charge FI log; interlock-integrity report | No charge current below 0 °C across the full cold-sensor + SoC + FW-fault matrix; inhibit persists in every injected FW/comms fault; self-heat cannot bypass | PV-ENRG-001 | Yes
HA-RCV-002 | Cell protection + secondary backstop (RCR-002) | FI + T | Protection FI log (DO-TST-009); UN 38.3 cert | OV/UV/OC/OT/short each terminated; single-fault (primary IC failed) still safe via independent fuse/OVP; UN 38.3 T1–T8 pass | PV-ENRG-002 + UN 38.3 | Yes
HA-RCV-003 | Cell retention (RCR-004) | T | Drop/vibe + UN 38.3 report | No crush/puncture through drop/vibration/transport; clearances maintained | PV-MECH-004 + UN 38.3 | Yes
HA-RCV-004 | Fire-enclosure containment (RCR-008) | FI(abuse) + T | Single-cell abuse/containment report; offgas characterisation | Single-cell thermal event contained with no external flame propagation; offgas low-toxicity characterised | PV-TFC-002 (+ PV-TFC-003 offgas) | Yes
HA-RCV-005 | Charge-input abuse protection (RCR-003) | FI + T | Reverse/OVP/surge/short test report | Out-of-spec source rejected/limited; safe under reverse-insertion/OVP/short; PD PDO capped | PV-ELEC-002 | Yes
HA-RCV-006 | Self-heat bound (RCR-005) | FI + T | Touch-temp + inhibit-integrity report | Self-heat only enables the guard; touchable surface within IEC 62368-1 burn limits; cannot bypass inhibit | PV-TFC-001 | Yes
HA-RCV-007 | Energy reserve / no silent loss (RCR-006) | T(endurance) + FI | 90-day cold endurance data; reserve-floor FI log | 90-day cold endurance met; reserve preserves record-commit + low-battery alert; exhaustion yields honest gap not silent trip | PRD-VER-004 | Yes
Software / detection / verdict / record / time (Class-C-equivalent, FI + static analysis, IV for SC)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-008 | Two-sided detection integrity (RCR-019) | FI + T + A | Profiled-excursion test; static-analysis report; unit/integration | Every profiled excursion (heat/freeze/boundary/ramp/dwell) detected; freeze coverage explicitly equal to heat; no undetected-excursion path | PV-SW-001 (+ VAL-SW-001) | Yes
HA-RCV-009 | Verdict integrity / honest gap (RCR-021/022) | FI + T | Gap-handling FI log; device↔cloud consistency report | No unqualified accept over any injected gap/fault; device=cloud verdict; incomplete record never a clean pass | PV-SW-002 (+ PV-DATA-004) | Yes
HA-RCV-010 | Sensor/probe plausibility (RCR-020) | FI + T | Probe-fault-injection log | Absent/open/short/wrong probe + implausible reads flagged, never read as valid product-zone temp | PV-SW-008 | Yes
HA-RCV-011 | Profile validation + pre-arm gate (RCR-048/045) | FI + T | Profile-validation + pre-arm blocks-arming log | Mismatched/implausible profile rejected; arming blocked unless SoC/calibration/firmware/profile/probe all pass | PV-SW-003 | Yes
HA-RCV-012 | Calibration integrity + expiry gate (RCR-026) | FI + T | Corrupt-block rejection + expiry-gate log | Corrupt/stale calibration rejected; expired/invalid calibration blocks arming | PV-SW-004 | Yes
HA-RCV-013 | Time integrity (RCR-025) | T + FI | Clock-drift/discipline over cold profile; untrusted-time disclosure | Drift bounded over 90-day cold; monotonic across brown-out/sleep; untrusted time disclosed | PV-SW-005 (+ PV-DATA-003) | Yes
HA-RCV-014 | Alarm generation + escalation (RCR-029) | FI + T | Alert-latency + escalation-on-no-ack log | Alert generated + sent ahead of bulk; escalates if unacknowledged; lost hand-off/full-queue/task-starvation recovered | PV-SW-006 (+ VAL-SW-006) | Yes
HA-RCV-015 | Safe-state / watchdog (RCR-023) | FI + T | Fault-injection + watchdog-recovery log | Any injected unsafe fault → defined safe state; record + cell held safe; logged recovery | PV-SW-009 | Yes
HA-RCV-016 | Power-fail-safe storage (RCR-024) | FI + T | Power-fail-during-write log over cold profile | No torn/lost/fabricated record; no overwrite of unsent data; loss surfaces as honest gap; tamper-evident | PV-SW-013 (+ PV-DATA-006) | Yes
HA-RCV-017 | Availability self-check (RCR-027) | FI + T | Self-check outage log | Whole sensing→logging chain self-checked; device flagged not-monitoring on outage | PV-SW-011/012 | Yes
HA-RCV-018 | Verdict-display integrity (RCR-021/043) | FI + T + SU | Display-integrity FI; summative usability | Displayed verdict = computed verdict (no inversion/stale/mis-association); legible gloved/cold/low-light | PV-SW-007 + VAL-USE-001 | Yes
Connectivity / cloud (delivery — degrade-not-lose; FI on the delivery path)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-019 | Local excursion commit (RCR-028) | FI + T | Offline-excursion + store-and-forward log | Every excursion committed locally with no bearer; delivered alert-first on reconnection; zero record loss | PV-COMM-001 | Yes
HA-RCV-020 | Alert delivery reconciliation (RCR-029/032) | FI + T | Delivery-path FI (drop sessions/acks) | No silently dropped safety event; retry + reconciliation + escalation recover every alert; no un-acked eviction | PV-COMM-002 (+ PV-CLOUD-001) | Yes
HA-RCV-021 | RF-in-container link / antenna (RCR-031) | T(qual) + A | RF-in-container qualification report | In-container attenuation, LTE-M/Wi-Fi/BLE margin, per-trip delivery-success, GNSS behaviour quantified; integral-vs-external antenna decided | PRD-VER-005 (+ PV-COMM-003) | Yes
HA-RCV-022 | Ordered/idempotent ingest (RCR-030/032) | FI + T | Reorder/duplicate/corrupt injection log | Exactly-once, in-order, integrity-checked, reconciled cloud state | PV-COMM-004 | Yes
HA-RCV-023 | Session-enabler robustness (RCR-025/032) | FI + T | Cert-expiry/time-loss/DNS-failure FI + monitoring | Graceful degrade; disclosed drift; backlog surfaced; fleet monitoring alerts on common-cause | PV-COMM-005 | Yes
HA-RCV-024 | Roaming/regional config (RCR-056/031) | T + A | Roaming/coverage characterisation; config-verification | Coverage continuity across representative lanes; correct regional band-plan verified at EOL | PV-COMM-006 | Yes
HA-RCV-025 | Authoritative device↔trip binding (RCR-033/055) | FI + T | Association/routing + identity-uniqueness log | Correct binding across re-arm cycles; wrong-tenant/duplicate rejected; alerts routed correctly | PV-COMM-008 (= PV-CLOUD-003) | Yes
HA-RCV-026 | Cloud never-drop + resilience (RCR-032/039) | FI + T | Ingest/notify FI; DR/backup-restore test | No dropped safety event; tested backup/restore; auth/integrity/backlog monitored | PV-CLOUD-001/002 + DR test | Yes
Cybersecurity-induced safety (independent adversarial verification mandatory)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-027 | Capture-time data authenticity (RCR-034) | FI + pen-test | Integrity/pen-test report | Falsified/injected/replayed data detected and rejected; device-anchored integrity holds against a compromised link/cloud | PV-CYBER-001 | Yes (adversarial)
HA-RCV-028 | Non-suppressible alarm path (RCR-035) | pen-test + FI | Alarm-suppression pen-test | Suppression/blocking raises a detectable event and escalates; cannot be silently muted | PV-CYBER-002 | Yes (adversarial)
HA-RCV-029 | Authorised config change (RCR-036) | pen-test + T | Config-tamper test | Threshold/config change authenticated, authorised, bounded, audited; unauthorised change prevented | PV-CYBER-003 | Yes (adversarial)
HA-RCV-030 | Secure OTA (RCR-037) | FI + pen-test | OTA rollback + version-skew + pen-test | Signed, rollback-protected A/B; never bricks; malicious/unsigned image rejected; staged rollout + halt-on-regression | PV-CYBER-005 (+ PV-SW-010) | Yes (adversarial)
HA-RCV-031 | Secure boot / debug-lock (RCR-038) | pen-test | Secure-boot + debug-lock pen-test | Boot verified against HW root of trust; production debug ports locked/disabled; tamper-evident | PV-CYBER-006 | Yes (adversarial)
HA-RCV-032 | Cloud security + DR (RCR-039) | T + review | Cloud security review + DR test | RBAC/SSO; auth/integrity/backlog monitored; tested backup/restore; compromise cannot silently blind/corrupt | PV-CYBER-009 | Yes
Data & decision / human factors (design-enforced info + summative usability)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-033 | Honest disposition surface (RCR-040) | SU + I | Gap-exposure review + summative usability | Surface never shows a gap/fault as a clean pass; low-training user does not read an incomplete record as clean | PV-DATA-004 + VAL-USE-004 | Yes
HA-RCV-034 | Estimate-provenance flagging (RCR-041) | I + SU | Visualisation/provenance review | Estimated/interpolated/last-known flagged, never shown as measured; users distinguish estimate from measurement | PV-DATA-007 + VAL-USE-005 | Yes
HA-RCV-035 | Integrity-preserving export (RCR-042) | T | Export-integrity test | Export preserves attribution/time/gaps/tamper-evidence; no silent loss | PV-DATA-006 (export) | Yes
HA-RCV-036 | Legible verdict (RCR-043) | SU + T(cold) | Cold-display + summative usability | Verdict legible/language-independent, gloved/low-light/−30 °C; icon+colour+word, not colour-only | PV-USE-001 + VAL-USE-001 | Yes
HA-RCV-037 | Recorded override (RCR-044) | SU + T | Override-workflow usability | Override deliberate, confirmed, recorded; accidental accept of a compromised dose prevented | VAL-USE-007 | Yes
HA-RCV-038 | Alarm-fatigue management (RCR-047) | T + SU | Alarm-management validation | False-alarm/missed-excursion targets met; consolidation prevents fatigue muting a real excursion | PV-USE-002 + VAL-USE-002 | Yes
Service / manufacturing (production/service preservation — HA-CIR §11)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-039 | EOL calibration coverage + birth record (RCR-053) | T + audit | EOL calibration coverage audit | Defined coverage/accuracy; miscalibrated unit cannot pass EOL; birth record captured | PV-MFG-001 | Yes
HA-RCV-040 | Correct chemistry/parts (RCR-054) | I + T | Incoming + AOI verification | Correct cell chemistry + correct-rated/oriented protection parts; counterfeit/wrong parts prevented | PV-MFG-002/003 | Yes
HA-RCV-041 | Identity uniqueness + regional config (RCR-055/056) | T | Identity-uniqueness + config-verification | Unique crypto identity/serial; correct regional/firmware config verified at EOL | PV-MFG-005/006 | Yes
HA-RCV-042 | Post-service BIST + interval gate (RCR-049/050/051) | T + FI | Post-service BIST; interval-enforcement; mode-lock | Out-of-cal device un-armable; seal/config/firmware restored + BIST-verified; debug modes locked before deploy | PV-SVC-001/003/006 | Yes
HA-RCV-043 | Latent-defect screening (RCR-016) | T | AOI/ICT coverage review | No latent single-point defect escapes on safety-critical nets | PV-MFG-008 | Yes
Mechanical / electrical / environmental (standards-based test)
HA-RCV | Control (RCR) | Method | Evidence type | Acceptance criterion (target) | PV-/VAL- | IV
HA-RCV-044 | Edge/pinch/small-parts (RCR-009/010) | I + T | Edge/pinch inspection; small-parts test | No accessible hazardous edge/pinch; small parts retained/sized/warned | PV-MECH-001/003 | No
HA-RCV-045 | Ingress/condensation protection (RCR-011/015/018) | T | IP + cold-cycle + insulation-resistance | Seal + conformal coat protect record electronics across −30↔+60 °C condensation cycling over reuse life | PV-TFC-006 + PV-ENV-003 | Yes
HA-RCV-046 | SELV + EMC (RCR-013/014) | T | IEC 62368-1 + IEC 61000 report | No accessible hazardous voltage; EMC immunity (no false read/verdict/miss); emissions safe near aircraft/medical | PV-ELEC-001/003/004 | Yes
HA-RCV-047 | Environmental range (RCR-017) | T(chamber) | IEC 60068 + endurance | Operate/retain record −30…+60 °C incl. 90-day cold endurance + reserve | PV-ENV-001 + PRD-VER-004 | Yes
HA-RCV-048 | UN 38.3 / IATA transport (RCR-007) | T(abuse) | UN 38.3 report; DGR documentation | UN 38.3 pass; IATA-DGR-compliant with transport documentation | UN 38.3 + DO-REG-006 | Yes
6. Verification-planning notes (per HAL-2600 standard)
• Fault-injection is required, not optional, for every safety control (HA-RCV-001/002/004/008/009/010/012/ 014/016/019/020/027/028/030). A nominal-operation pass does not close a safety control; the control must be shown to act at the moment the fault occurs and to leave the system safe/honest.
• Independent verification (IV) is marked Yes for every S4/S5 contributor and every security control; the security controls (HA-RCV-027…032) additionally require adversarial (pen-test) verification by an independent assessor — developer self-test is insufficient (SC/security rule).
• The Evidence-Pending residuals in HAL-2800 close only when the paired PV-/VAL- passes to its acceptance criterion (SC controls also need IV sign-off). The carried-Open items — D3 cold endurance (HA-RCV-007 / PRD-VER-004) and RF-in-container (HA-RCV-021 / PRD-VER-005) — are the two verifications that gate the largest uncertain probabilities and therefore the release recommendation (Risk-Management Summary).
• Summative usability (IEC 62366-1) is the verification of last resort for the human-mediated information controls (HA-RCV-018/033/034/036/037/038): the decision-support surfaces are only "verified" when a representative low-training last-mile user, gloved and in low light at cold, reads them correctly.
7. Traceability / closure notes
• Every PV-/VAL- named here already exists as an Evidence-Pending reference in a domain module (HAL-500… 1800); this module consolidates them into a verification plan so Stage 13 (PVL, ORCH-014) and Stage 14 (PVaL, ORCH-015) inherit a complete, safety-driven coverage matrix and cannot close leaving a safety control unverified.
Feeds Stage 13/14 | each HA-RCV row becomes a planned PVL/PVaL protocol; the executed evidence returns to close the residual in HAL-2800 and the benefit-risk in HAL-2900 / Risk-Management Summary.
Feeds Stage 11 (ERFKL) | the S4/S5 hazardous situations behind these controls are FTA top events; the fault-injection cases here seed the FTA/FMEA fault sets.
Release gate | no release without every safety control's PV- passed + IV sign-off; the two carried-Open qualification tests (PRD-VER-004/005) are explicit conditions of the Conditionally-Ready recommendation.
HAL-2700 (HA-NEW) — New Risks Introduced by Risk Controls
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-2700, record code HA-NEW · Owner: System-Safety (lead) with the discipline owners of each control, and Quality/Regulatory (EM-ENG-006). ISO 14971 §7.4 requires that each risk-control measure be reviewed for new risks it introduces or existing risks it makes worse, and that those be fed back into the risk analysis. This module collects, chains, scores, and controls the new risks introduced by the SCM's risk controls — the "the cure has side-effects" pass. Each record uses the same risk-chain discipline as the domain modules; new risks that reach a real harm are scored severity-first and given their own control (often already present as an existing RCR). Standards cited by designation; revisions to confirm.
---
1. Purpose
Make explicit that the SCM's safety controls are not free: charge-inhibit can leave a device dead-on- arrival; a self-heat-before-charge feature adds an energy/thermal source; encryption/integrity adds latency and a key-lifecycle dependency; automatic alerting + escalation adds alarm load and fatigue; a redundant/ independent second sensor or cross-check adds a disagreement-logic problem; conservative fail-safe (suppress-accept, honest-gap, flag-any-doubt) trades safety for availability/wastage/stock-out; OTA and cloud cross-checks add fleet-wide common-cause and connectivity dependence. Each is identified, chained to its worst credible outcome, scored, and shown to be controlled — or flagged as an accepted trade with a benefit- risk line in HAL-2900.
2. Scope
New risks arising from the controls reviewed in HAL-2500 (HA-CIR) and the full RCR register (HAL-2400). Grouped by the control that introduces them. Several new risks are already owned as domain records (e.g., the charge-inhibit → dead-in-cold information path is HA-ENRG-006 / HA-USE-009); this module cross-references rather than duplicates, and adds the ones not already carried. The dominant systemic trade — safety conservatism raising false-reject/wastage/stock-out — is the recurring theme and ties HA-MPE-003.
3. Applicable standards & references
• ISO 14971 §7.4 — new-risk / risk-increase review of every control; feedback into the analysis.
• IEC 62366-1 — alarm fatigue and automation-related new risks (the human side of over-alerting).
• IEC 62368-1 / IEC 62133-2 — the self-heat element as a new thermal/energy source.
Upstream | HAL-2400 (RCRs), HAL-2500 (HA-CIR findings), HAL-2000/2100/2200 (S/P/acceptability), HAL-2900 (where the accepted trades are aggregated into benefit-risk), and the domain modules that already own several of these new risks (HA-ENRG-006, HA-USE-002/009, HA-MPE-003, HA-TFC-001).
Evidence-honesty note | the magnitude of each trade (how much wastage a conservative verdict causes, how much latency encryption adds, whether self-heat is even adopted) is design-/data-dependent and Evidence-Pending; scored here on severity with judgment-based probability and a named closing activity.
4. Governance & method
• Every record names the originating control (RCR / HA-CIR), the new hazard, its risk chain, a severity-first score, the control on the new risk (usually an existing RCR — a control balancing a control), and a disposition: Controlled, Accepted trade (benefit-risk), or Cross-referenced to domain record.
• The safety-vs-availability trade is treated honestly: the SCM's conservative direction (favour reject/ qualify over false accept) is the correct safety choice, but its new risk (wastage/stock-out → HA-MPE-003 → unvaccinated population) is itself a public-health harm and is balanced, not ignored (alarm tuning RCR-047, false-alarm targets, ALARP both ways).
5. New-Risk Register (HA-NEW records)
HA-NEW-001 — Charge-inhibit below 0 °C → device under-charged / dead-on-arrival if left cold
Originating control | RCR-001 / HA-CIR-001 (charge-inhibit below 0 °C).
New hazard | A device that cannot charge in a cold depot arrives at deployment under-charged or flat → cannot complete the 90-day trip, or is discarded as "dead."
Risk chain | Cold depot → charge inhibited (correctly, for fire safety) → SoC not restored → device armed under-charged (or not deployable) → energy exhaustion mid-trip → silent loss of monitoring (or stock-out from a device pulled as dead).
Harm/Loss | Loss of monitoring on a trip → wrong disposition (S4 via HA-ENRG-006); or a serviceable device discarded → stock-out (S3 via HA-USE-009).
Severity | S4 (loss-of-monitoring worst case) — governing dimension public-health.
Probability (P1×P2) | P1=2 (Remote — depot warm-before-charge process + self-heat option + pre-arm SoC gate) × P2=3 (Occasional — an under-charged arming reaches a gap only if the reserve/governor also under-performs; the pre-arm SoC-for-trip gate blocks most) ⇒ PM — Basis: judgment; Medium uncertainty, Evidence-Pending on the 90-day cold budget (PRD-VER-004).
Control on the new risk | RCR-045 (pre-arm SoC-for-trip gate blocks arming under-charged), RCR-006 (energy reserve + honest low-battery), RCR-046 (charge-inhibited ≠ dead legibility so it is not discarded), depot warm-before-charge process (+ optional self-heat RCR-005).
Disposition | Controlled + cross-referenced to HA-ENRG-006 (silent gap) and HA-USE-009 (discarded-as-dead). The charge-inhibit is non-negotiable for fire safety; its availability cost is managed, not removed. Carried in HAL-2900 as a safety-vs-availability trade.
Owner | Battery/Power + Human-Factors · Review History: 2026-07-19 v1.0 initial (ORCH-017)
HA-NEW-002 — Self-heat-before-charge → new hot surface + new energy/thermal source
Originating control | RCR-005 / HA-CIR-001 (self-heat enables cold charging by warming the cell — a mitigation for HA-NEW-001).
New hazard | The self-heat element is a new touchable hot surface (burn) and a new energy path that, if it exceeded its bound or ran without control, could over-warm the cell.
Risk chain | Self-heat activates to warm a cold cell → element/surface heats → (a) technician contact → burn; or (b) runaway/unbounded heating → cell over-temperature → ties the thermal set.
Harm/Loss | Moderate reversible burn (S3) — already owned as HA-TFC-001; the over-warm path is bounded into the charge-safety envelope.
Severity | S3 (burn) — cross-referenced.
Probability (P1×P2) | P1=2 × P2=2 ⇒ PL (per HA-TFC-001), Evidence-Pending on self-heat feasibility (CMP-013).
Control on the new risk | RCR-005 (self-heat only enables the guard, cannot bypass inhibit, bounded below the safe thermal/burn limit), RCR-046 (state legibility), touch-temperature limits (IEC 62368-1).
Disposition | Cross-referenced to HA-TFC-001; conditional on self-heat adoption (if CMP-013 not adopted, this record is N-A and cold charging relies on warm-then-charge process alone, strengthening HA-NEW-001). A clear example of one control (self-heat) introducing the risk another control (touch-temp bound) must catch.
Owner | Thermal/Battery-Safety · Review History: 2026-07-19 v1.0 initial
HA-NEW-003 — Cryptographic integrity/encryption → added latency, energy cost, and key-lifecycle dependency
Originating control | RCR-034/037/038 / HA-CIR-009/010 (capture-time authenticity, secure OTA, secure boot).
New hazard | Signing/encryption adds compute + radio latency and energy draw (tension with the D3 budget and alert-latency), and the whole authenticity scheme becomes dependent on the device key / secure element — a new single point whose compromise or remanence (HA-CYBER-008) defeats integrity fleet-wide.
Risk chain | (a) Crypto/TLS overhead → longer sessions / more energy → marginally later alert or shorter endurance; (b) key compromise/remanence → forged-but-valid records → false accept, or repudiation.
Harm/Loss | (a) Bounded added alert latency (S3, ties HA-COMM-007) + energy cost (D3 budget); (b) key compromise → falsified record → compromised dose (S4, owned by HA-CYBER-001/008).
Severity | S4 (key-compromise path) — governing dimension public-health/security.
Probability (P1×P2) | P1=2 (Remote — hardware crypto/secure element sized for the budget; key lifecycle controlled) × P2=3 (Occasional — a key compromise reaches harm readily but requires defeating secure boot/debug-lock) ⇒ PM.
Control on the new risk | hardware-accelerated crypto sized into the power/latency budget (energy↔latency trade tracked vs PRD-VER-004/005); RCR-055 (unique identity), RCR-038 (secure boot + debug-lock), RCR-052 (disposal key-erasure) bound the key lifecycle — the HA-CIR-009 finding.
Disposition | Controlled (latency/energy = accepted trade; key-lifecycle = controlled with finding). The integrity benefit vastly outweighs the latency/energy cost for an evidentiary device. Carried in HAL-2900 (energy budget + key-lifecycle security).
Owner | Product-Security + Connectivity + Power · Review History: 2026-07-19 v1.0 initial
HA-NEW-004 — Automatic alerting + escalation → alarm load → alarm fatigue → a real excursion muted
Originating control | RCR-028/029/032 / HA-CIR-006 (alert-first, ack-required, escalate-if-unacknowledged, cloud notify).
New hazard | Alert-first prioritisation, escalation chains, and conservative detection generate notification volume; a fatigued responder mutes/ignores alerts and misses the one real excursion.
Risk chain | Many alerts (incl. false/low-value) → responder desensitised / mutes channel → real excursion alert ignored → no intervention → compromised dose.
Harm/Loss | Real excursion not acted on → compromised dose administered (S4) — owned as HA-USE-002 (alarm-fatigue headline).
Severity | S4 — public-health; cross-referenced.
Probability (P1×P2) | P1=3 (Occasional — alert volume is a real operational reality, especially in outbreak/congestion) × P2=3 (Occasional — a muted channel reaches harm, but consolidation + escalation-to-a-different-person mitigate) ⇒ PM.
Control on the new risk | RCR-047 (alerting tuned to validated false-alarm/missed-excursion targets; consolidation avoids fatigue), RCR-029 (cloud-side consolidation; escalation to a different responder rather than repeat-spamming the same one), alarm-management validation (VAL-USE-002).
Disposition | Controlled (accepted trade, tuned). The alerting benefit requires disciplined tuning; over-alerting is itself a hazard. This is the archetypal control-introduced risk — more alerting is not automatically more safety. Carried in HAL-2900 and HA-USE-002.
Owner | Human-Factors + Connectivity/Cloud · Review History: 2026-07-19 v1.0 initial
HA-NEW-005 — Redundant / independent second sensor or device↔cloud cross-check → disagreement logic
Originating control | RCR-020 (int-vs-ext divergence), RCR-021 (device↔cloud verdict consistency) / HA-CIR-003/004/005.
New hazard | Adding an independent second measurement or an independent recompute creates a disagreement problem: when the two disagree, which is trusted? A naive "trust either / trust the benign one" rule can reintroduce a false accept; a "trust neither" rule can flood rejects.
Risk chain | Two channels/verdicts disagree → disagreement-resolution logic picks the wrong (benign) one → false accept; or over-conservatively rejects → wastage.
Harm/Loss | False accept (S4) if disagreement resolves optimistically; or wastage/stock-out (S3, ties HA-MPE-003) if it resolves over-conservatively.
Severity | S4 — public-health.
Probability (P1×P2) | P1=2 (Remote — disagreement-resolution is a defined, safety-biased rule: any disagreement → qualify/reject, never a clean accept) × P2=3 (Occasional — a subtle disagreement-logic defect is invisible, like the parent detection defect) ⇒ PM.
Control on the new risk | disagreement resolution is safety-biased by construction — divergence never yields an unqualified accept (extends RCR-021/022); the resolution rule is itself verified (part of PV-SW-002/008); internal-vs-external divergence flags a probe fault (RCR-020) rather than silently averaging.
Disposition | Controlled. Redundancy is only a benefit if the disagreement rule is safety-biased and verified; an un-verified voter is a new single point. Feeds HAL-2600 (the disagreement cases are explicit FI cases) and the FTA (Stage 11).
Owner | Embedded/Software Safety · Review History: 2026-07-19 v1.0 initial
HA-NEW-006 — Conservative fail-safe (suppress-accept / honest-gap / flag-any-doubt) → false-reject → wastage → stock-out
Originating control | RCR-019/020/021/022/024/040 / HA-CIR-003/004/005/008 (the entire no-false-reassurance direction) + RCR-045 (pre-arm gate blocks a doubtful arming).
New hazard | The SCM is deliberately biased to reject/qualify/void on any doubt (gap, marginal value, fault, invalid calibration, un-armed). This prevents false accepts but destroys potent product and blocks serviceable devices — a false-reject/availability harm.
Risk chain | Marginal/incomplete/uncertain trip → conservative reject/quarantine/void → potent lot destroyed or device not deployed → stock-out → unvaccinated population → preventable disease.
Harm/Loss | Potent product destroyed / stock-out → unvaccinated beneficiary (S4) — owned as HA-MPE-003 (false-reject stock-out) and ties HA-USE-009.
Severity | S4 — public-health (the other direction of the disposition harm; a stock-out is a real harm, not a free safety win).
Probability (P1×P2) | P1=3 (Occasional — conservative flagging is by design frequent) × P2=2 (Remote — most conservative flags trigger a review, not automatic destruction; human + cloud recompute can re-adjudicate a marginal case) ⇒ PM.
Control on the new risk | RCR-047 (false-alarm/false-reject targets validated — ALARP in both directions), RCR-044 (a deliberate, recorded, informed override path exists so a genuinely-good marginal lot is not blindly destroyed), cloud recompute re-adjudication, honest qualifier (RCR-040) so the human sees why and can make a proportionate call.
Disposition | Accepted trade (benefit-risk), balanced. The conservative direction is the correct safety choice, but the wastage/stock-out harm is real and public-health-significant; it is bounded by validated false-reject targets and an audited override, not by weakening the safety bias. This is the central benefit-risk tension of the whole device — carried prominently in HAL-2900 and the Risk-Management Summary.
Owner | System-Safety + Human-Factors + EM-ENG-006 · Review History: 2026-07-19 v1.0 initial
HA-NEW-007 — Pre-arm readiness gate → legitimate arming blocked → availability/stock-out
Originating control | RCR-045/026/048 / HA-CIR-007 (pre-arm gate blocks arming unless SoC/calibration/firmware/profile/probe all pass).
New hazard | An over-strict or falsely-tripped gate blocks a legitimate deployment (e.g., calibration marginally within interval read as expired; a valid profile read as mismatched), leaving no monitored device for a shipment.
Risk chain | Gate falsely blocks arming → no device for the trip → shipment ships unmonitored or not at all → lost monitoring or stock-out.
Harm/Loss | Lost monitoring / stock-out (S3–S4, ties HA-MPE-003 / HA-ENRG-006).
Severity | S3 (availability, typically caught before an unmonitored shipment ships) escalating to S4 if it silently forces an unmonitored ship.
Probability (P1×P2) | P1=2 (Remote — gate thresholds set with margin; clear cause shown) × P2=2 (Remote — a blocked arming is visible and prompts a fix/spare, rarely reaching an unmonitored ship) ⇒ PL.
Control on the new risk | gate thresholds set with justified margin (not hair-trigger); cause shown legibly (RCR-046) so the depot can remedy (recharge, recalibrate, reload) rather than abandon; spare-device logistics; the gate is fail-safe toward not-arming (the safe direction).
Disposition | Controlled (accepted trade). A gate that occasionally blocks a good arming is far safer than one that passes a bad one; kept ALARP by margin + legible cause + service support.
Owner | Human-Factors + Service · Review History: 2026-07-19 v1.0 initial
HA-NEW-008 — OTA + cloud cross-check → fleet-wide common-cause and connectivity/cloud dependence
Originating control | RCR-037 (OTA), RCR-021 (cloud recompute), RCR-032/039 (cloud delivery/resilience) / HA-CIR-003/006/010.
New hazard | Controls that use the cloud or OTA introduce fleet-wide common-cause (one bad signed image, one rules-engine misconfig, one cloud outage affects the whole fleet at once) and make some assurance connectivity/cloud-dependent — the very dependency the offline-first architecture tries to avoid.
Risk chain | Bad OTA / rules misconfig / cloud outage → many devices deploy wrong logic, or are blinded, or lose the cross-check simultaneously → multiple wrong dispositions across a programme.
Harm/Loss | Fleet-scale wrong disposition / blinding (S4, S5 at population scale) — owned by HA-CLOUD-002/007/008, HA-CYBER-005/009, HA-COMM-005.
Severity | S4/S5 — public-health at scale.
Probability (P1×P2) | P1=2 (Remote — staged/canary rollout, config change-control, cloud redundancy, offline-first so a cloud outage does not lose data) × P2=3 (Occasional — a common-cause reaches many devices before it is noticed) ⇒ PM, escalated for the S5-worst common-cause.
Control on the new risk | staged/canary OTA + automatic rollback + halt-on-regression (HA-CIR-010 finding), config change-control with staged rollout (avoid a single fleet-wide flip), cloud monitoring of auth/integrity/backlog (RCR-032/039), and the offline-first invariant (device safety never depends on the cloud — a cloud outage delays, never loses; on-device detection/verdict/commit stand alone).
Disposition | Controlled (with standing findings). The cloud/OTA benefits (fleet management, second-opinion recompute, patchability) are real but purchased with common-cause exposure; kept acceptable only by rollout discipline + offline-first. Carried prominently in HAL-2900 (common-cause + systemic connectivity dependence).
Owner | Cloud + Firmware + Product-Security + System-Safety · Review History: 2026-07-19 v1.0 initial
HA-NEW-009 — External-antenna fallback → new sealed connector (ingress) + per-market variant (config error)
Originating control | RCR-031 / HA-CIR-006 (external-antenna fallback if the integral antenna cannot close the in-container budget).
New hazard | An IP-sealed external-antenna connector is a new ingress path and mechanical seal risk; per-market antenna/band variants add a provisioning/config-error path.
Risk chain | External connector seal degrades → ingress → record-electronics fault (ties HA-ENV-003 / HA-TFC-006); or wrong-region variant provisioned → wrong-band transmit / coverage loss (ties HA-MFG-006 / HA-COMM-006).
Harm/Loss | Record-integrity fault → wrong disposition (S3–S4); or coverage/compliance error (S3–S4).
Severity | S4 worst (record-integrity/coverage → wrong disposition).
Probability (P1×P2) | P1=2 × P2=2 ⇒ PL — conditional on the external-antenna variant being adopted (Evidence-Pending on PRD-VER-005 / the integral-vs-external decision).
Control on the new risk | RCR-011/015 (sealed-connector ingress protection to the same IP/condensation standard), RCR-056 (correct regional config provisioned + verified at EOL), configuration poka-yoke.
Disposition | Controlled + conditional. Only materialises if the integral antenna fails the in-container budget; flagged to Mechanical (seal) and Regulatory (per-market). Cross-referenced HA-COMM-003/006, HA-MFG-006.
Owner | Connectivity + Mechanical + Regulatory · Review History: 2026-07-19 v1.0 initial
HA-NEW-010 — Fire-containment enclosure + low-toxicity materials → mass/cost + material-property constraints
Originating control | RCR-008 / HA-CIR-002 (fire enclosure contains a single-cell event; low-toxicity offgas materials).
New hazard | A robust fire-containment enclosure adds mass/cost and constrains material choice; a directed-venting path could conflict with the ingress seal (RCR-011); low-toxicity selection may trade against other material properties (strength, cold-toughness).
Risk chain | Heavier/constrained enclosure → handling ergonomics (HA-MECH-010) / cold-brittleness of a lower-strength material → crack on drop → ingress (HA-MECH-004); vent-vs-seal conflict → ingress path.
Harm/Loss | Handling injury (S1–S2) or record-integrity fault via ingress (S3) — low severity, contained.
Severity | S3 worst (ingress via vent/seal conflict).
Probability (P1×P2) | P1=2 × P2=2 ⇒ PL.
Control on the new risk | integrated design of vent-path + seal (a vent that reseals / directs offgas without a standing ingress hole), cold-toughness in material selection (RCR-017 range + drop/vibe), ergonomic handling (RCR-009/HA-MECH-010).
Disposition | Controlled (design-integration trade). The S5 containment benefit dominates the mass/cost/ingress trade; managed by integrated mechanical design. Cross-referenced HA-TFC-002/003, HA-MECH-004/010.
Owner | Mechanical + Thermal/Battery-Safety · Review History: 2026-07-19 v1.0 initial
HA-NEW-011 — Frequent time re-discipline + safety-plane short backoff → extra radio energy (D3 budget)
Originating control | RCR-025 (bounded/re-disciplined time), RCR-029/030 (safety-plane alert-first, tuned-shorter backoff) / HA-CIR-006.
New hazard | Re-disciplining the clock and running a shorter backoff for the safety plane spend extra radio energy — a tension with the D3 90-day cold budget.
Risk chain | More radio activity → higher energy draw → earlier exhaustion → loss of monitoring (ties HA-ENRG-006).
Harm/Loss | Earlier energy exhaustion → loss of monitoring (S4 via HA-ENRG-006) — but this is a budget effect, bounded by the governor.
Severity | S4 (loss-of-monitoring) but low-probability contribution.
Probability (P1×P2) | P1=2 × P2=2 ⇒ PL — the energy governor (RCR-006) allocates and bounds radio energy; safety-plane activity is prioritised within the budget.
Control on the new risk | energy governor budgets radio activity (RCR-006); safety-plane backoff tuned for latency within the budget (not unlimited); re-discipline opportunistic (piggybacked on existing sessions). The energy↔latency trade is quantified against PRD-VER-004.
Disposition | Controlled (accepted trade, budget-bounded). A minor, bounded energy cost for time defensibility and alert timeliness. Cross-referenced HA-ENRG-006, HA-COMM-007.
Owner | Firmware + Power + Connectivity · Review History: 2026-07-19 v1.0 initial
HA-NEW-012 — Aggressive / over-sensitive detection → false-reject + alarm load
Originating control | RCR-019 (two-sided detection with freeze equal to heat; no undetected-excursion path) / HA-CIR-003.
New hazard | Tuning detection to guarantee capture of every real excursion raises the false-positive (false-reject) rate and adds to alarm load.
Risk chain | Over-sensitive thresholds → benign fluctuations declared excursions → false rejects (wastage, ties HA-NEW-006/HA-MPE-003) + more alerts (ties HA-NEW-004/HA-USE-002).
Harm/Loss | Wastage/stock-out (S4 via HA-MPE-003) + alarm fatigue (S4 via HA-USE-002).
Severity | S4 — the same two trades as HA-NEW-004/006, sourced at the detection tuning.
Probability (P1×P2) | P1=3 (Occasional — a mis-tuned threshold is plausible) × P2=2 (Remote — validated detection targets + product-stability profiles bound false-positive rate) ⇒ PM.
Control on the new risk | detection tuned to the product-stability profile (thresholds are the vaccine's real stability limits, not arbitrary), RCR-047 validated false-alarm/missed-excursion targets balancing sensitivity vs specificity, profile validation (RCR-048).
Disposition | Controlled (accepted trade, profile-anchored). Sensitivity is anchored to real product stability, not maximised blindly; balanced against wastage/fatigue by validated targets. Cross-referenced HA-NEW-004/006, HA-USE-002, HA-MPE-003.
Owner | Embedded/Software Safety + Human-Factors · Review History: 2026-07-19 v1.0 initial
6. Cross-cutting observations
• The dominant control-introduced risk is the safety-vs-availability trade (HA-NEW-006/007/012): almost every information safety control favours reject/qualify/void, which converts avoided false-accepts into wastage/stock-out — itself a public-health harm (HA-MPE-003, unvaccinated population). The correct posture is ALARP in both directions: keep the safety bias, but validate false-reject targets and provide an audited, informed override — never weaken the safety bias to reduce wastage.
• The second theme is fleet-wide common-cause and connectivity/cloud dependence (HA-NEW-003/008): OTA, cloud recompute, and cryptographic key/lifecycle add power that is purchased with fleet-scale exposure; managed by rollout discipline, offline-first, and key-lifecycle security.
• The third theme is one control needing another (HA-NEW-001↔002, HA-NEW-005, HA-NEW-010): charge-inhibit needs self-heat/warm-before-charge; self-heat needs a touch-temp bound; redundancy needs a safety-biased voter; containment needs a vent-vs-seal integration. No control is evaluated in isolation.
7. Traceability / closure notes
• Every HA-NEW record either cross-references an existing domain record that already owns the harm (HA-ENRG-006, HA-USE-002/009, HA-MPE-003, HA-TFC-001, HA-CLOUD-002/007/008, HA-CYBER-001/005/008/009, HA-MFG-006, HA-COMM-006) or is controlled by an existing balancing RCR (RCR-045/046/047/044/055/038/052). No new uncontrolled hazard is introduced by the control set.
• Accepted trades (HA-NEW-004/006/007/011/012 and the latency/energy parts of HA-NEW-003) are aggregated into the benefit-risk statement in HAL-2900 and the Risk-Management Summary; each carries a validated target (false-alarm/false-reject, latency, energy budget) as its closure condition.
• Standing findings to other stages: OTA rollout policy (Cloud/Firmware), key lifecycle (Product-Security), external-antenna seal + per-market config (Mechanical/Regulatory), D3 energy budget for crypto/radio (Power). These feed Stage 12 (Risk-Control Requirements) and the domain design outputs.
• The disagreement-logic and self-heat cases seed Stage 11 (ERFKL) FTA/FMEA fault sets; the balancing controls' verification is planned in HAL-2600.
HAL-2800 (HA-RES) — Individual Residual Risk
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-2800, record code HA-RES · Owner: EM-ENG-006 (Quality/Regulatory) + System-Safety, with EM-ENG-001 for S4/S5 sign-off. ISO 14971 §7.4/§8 requires that the residual risk of each individual hazard be estimated after its risk controls are implemented, and its acceptability decided. This module records, for each top SCM hazard, the residual severity and probability after controls, the effectiveness of those controls, the evidence status (predominantly Evidence-Pending until Stages 13–14 V&V), the residual acceptability classification, the acceptance authority (HAL-2200 §7), and the disclosure / post-market monitoring each residual requires. Overall (aggregate) residual risk is evaluated in HAL-2900. Standards cited by designation; revisions to confirm.
---
1. Purpose
Give each top hazard an honest, per-hazard residual-risk verdict: what remains after the controls, how confident we are that the controls work (control-effectiveness), what evidence still owes (Evidence-Pending vs Verified), where the residual sits on the acceptability matrix, who may accept it, and what disclosure and field monitoring the residual obliges. Residual severity is normally unchanged (the outcome, if it ever occurred, is the same); controls act on probability and on detectability/exposure. Where a control's verification is still pending, the residual is stated as projected and the hazard remains Controlled- pending-verification — not closed.
2. Scope
The ~30 top hazards across the domain modules — every S4/S5 record plus the significant S3 record-integrity and availability contributors — drawn from HAL-500…HAL-1800 and consolidated here with a uniform residual verdict. The full ~120-record inventory is not repeated; the lower-severity S1–S2 records reduce to region B (broadly acceptable) via their controls and are noted in aggregate (§7). The residual severity/probability bands and regions follow HAL-2000/2100/2200.
3. Applicable standards & references
• ISO 14971 §7.4 / §8 — individual residual-risk estimation and acceptability; overall residual and benefit-risk (§8) is HAL-2900.
• HAL-2200 — acceptability matrix, regions R/A/B, Action Priority, decision rules (catastrophic-severity, uncertain-probability, single-fault, warnings-challenge), acceptance authority (§7).
• HAL-2500/2600/2700 — control implementation (effectiveness input), verification plan (evidence status), and new-risk trades (which qualify a residual as an "accepted trade").
Evidence-honesty note | every residual probability is projected, Evidence-Pending until its named PV-/VAL- (HAL-2600) passes; residual severity is stated unchanged unless a control genuinely reduces the worst-case outcome. Control-effectiveness is a judgement of the designed control quality (HAL-2500), explicitly labelled design-adequate / to-be-verified — never asserted as demonstrated.
4. Governance & method — the residual verdict fields
Each HA-RES record states: Hazard (ref) · Initial risk · Controls (RCR) · Control-effectiveness (High/Moderate, design-adequate/to-be-verified) · Residual severity · Residual probability (projected) · Residual region / AP · Evidence status (Evidence-Pending / Partial / Verified) · Classification (Controlled-pending-verification / Open / Broadly acceptable) · Acceptance authority · Disclosure / monitoring · Closing evidence (PV-/VAL-).
5. Individual Residual-Risk Register (HA-RES records)
HA-RES-001 — Missed excursion (detection logic error) — HA-SW-001 ★
Initial | S4 (S5 worst) × PM ⇒ R (High). Controls: RCR-019 (two-sided detection, freeze=heat, no undetected-excursion path) + RCR-047.
Control-effectiveness | High, design-adequate/to-be-verified — inherent "no undetected-excursion path" + IEC 62304 Class-C rigor; residual is P2-dominated (harm remains human-mediated). HA-CIR-003 findings apply (cloud must not mask on-device logic; calibration/profile common-cause).
Residual | S4 (unchanged) × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-SW-001 profiled-excursion w/ equal freeze/heat + static analysis + VAL-SW-001). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 (S4, region A) + benefit-risk.
Disclosure / monitoring | disclose detection limits in IFU; post-market monitor missed-excursion field signals (returns, complaints, receipt-QA discrepancies); FTA top event (Stage 11).
Closing evidence | PV-SW-001, VAL-SW-001.
HA-RES-002 — False accept over incomplete/gapped record — HA-SW-002 / HA-DATA-004 ★
Initial | S4 × PM ⇒ R (High). Controls: RCR-021 (suppress false accept, device=cloud) + RCR-022 (gap visible) + RCR-040 (honest surface).
Control-effectiveness | High — false-accept-over-gap made unrepresentable by construction; genuine device↔cloud independence; non-defeatable invariant (HAL-2400 §4).
Residual | S4 (unchanged) × PVL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-SW-002 / PV-DATA-004 gap-handling + VAL-USE-004 summative usability). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | honest-gap qualifier is user-facing (by design); monitor override rates and incomplete-record dispositions; FTA top event.
Closing evidence | PV-SW-002, PV-DATA-004, VAL-USE-004.
HA-RES-003 — Probe-fault read as valid product-zone temp — HA-SW-008 ★
Initial | S4 × PM ⇒ R (High). Controls: RCR-020 (HW presence/open/short + FW plausibility) + RCR-045 (pre-arm probe-present).
Control-effectiveness | High for electrical faults; Moderate for in-range-but-wrong (placement). The plausibility control cannot detect a physically-misplaced healthy probe (HA-CIR-005 remainder → HA-USE-005).
Residual | S4 × PL (projected) ⇒ A (Medium); the placement remainder is carried under HA-RES-014 (HA-USE-005).
Evidence status | Evidence-Pending (PV-SW-008 probe-fault-injection). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | IFU on probe placement; monitor probe-fault flag rates; FTA top event.
Closing evidence | PV-SW-008.
HA-RES-004 — Silent gap / dead-in-cold not logged — HA-SW-013 / HA-ENRG-006 ★
Initial | S4 × PM ⇒ R (High). Controls: RCR-024 (power-fail-safe, loss→honest gap) + RCR-006 (energy reserve/governor) + RCR-027 (availability self-check) + RCR-022.
Control-effectiveness | High for the honest-gap behaviour (design-adequate); Moderate on the energy-margin (D3 cold-endurance is the open unknown, HA-CIR-011).
Residual | S4 × PL (projected) ⇒ A (Medium) — contingent on the 90-day cold endurance test.
Evidence status | Partial/Evidence-Pending — honest-gap closes on PV-SW-013; the energy margin is Open on PRD-VER-004 (D3). Classification: Controlled-pending-verification (energy margin Open).
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk; the D3 margin is a named release condition.
Disclosure / monitoring | honest low-battery/dead-in-cold indication user-facing; monitor field battery/endurance and gap rates.
Closing evidence | PV-SW-013, PRD-VER-004 (carried Open).
HA-RES-005 — Cold-charge plating → latent short → thermal runaway — HA-ENRG-001 ★ (D3)
Initial | S4/S5 × (uncertain) ⇒ R (High) (single-fault + catastrophic rules). Controls: RCR-001 (MCU-independent charge-inhibit <0 °C) + RCR-002 (independent protection backstop) + RCR-005 (bounded self-heat).
Control-effectiveness | High — independent, non-defeatable, fail-safe (inhibit), single-fault-safe with the secondary backstop (HA-CIR-001/002).
Residual | S4 (unchanged; S5 if it reached fire) × PVL (projected) ⇒ A (Medium) (S4×PVL) — the fire consequence is the HA-RES-006 S5 line.
Evidence status | Evidence-Pending (PV-ENRG-001 cold-charge fault-injection; sensor-independence finding). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk (S4/S5, single-fault).
Disclosure / monitoring | transport declaration; monitor charge-inhibit/field thermal events; FTA top event.
Closing evidence | PV-ENRG-001 (+ UN 38.3).
HA-RES-006 — Thermal-runaway propagation → fire in occupied aircraft hold — HA-TFC-002 ★★ (S5 anchor)
Initial | S5 × (any) ⇒ R (High) (catastrophic rule — never broadly acceptable). Controls: RCR-002 (protection) + RCR-004 (retention) prevent initiation; RCR-008 (fire enclosure) contains; RCR-007 (UN 38.3 / IATA-DGR).
Control-effectiveness | High — independent defense-in-depth (prevent + contain), single-fault-safe (HA-CIR-002).
Residual | S5 (unchanged) × PVL (projected) ⇒ A (High\) — per HAL-2200 §6.1 an S5 stays region A even at PVL and can never be broadly acceptable.
Evidence status | Evidence-Pending (PV-TFC-002 single-cell containment/fire-enclosure abuse + UN 38.3). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-001 + EM-ENG-006 sign-off + documented benefit-risk + post-market monitoring (any S5 residual, HAL-2200 §7).
Disclosure / monitoring | mandatory transport disclosure (UN 38.3 / IATA-DGR documentation); mandatory post-market thermal-event monitoring; FTA top event; the SCM's defining physical S5 residual.
Closing evidence | PV-TFC-002, UN 38.3.
HA-RES-007 — Delayed excursion alert (offline latency) — HA-COMM-001 ★
Initial | S4 × PM ⇒ R (High). Controls: RCR-028 (local commit) + RCR-029 (alert-first/ack/escalate) + RCR-030 (store-and-forward) + on-receipt edge read.
Control-effectiveness | High on data-loss (converted to latency); the finite in-transit dark window cannot be removed (HA-CIR-006).
Residual | S4 (unchanged) × PL (projected) ⇒ A (Medium) — residual is the actionable-window risk, not data loss.
Evidence status | Evidence-Pending (PV-COMM-001 latency + store-and-forward; interacts with PRD-VER-005). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | disclose that live alerting is connectivity-bound; on-receipt edge read is the compensating check; monitor alert-latency distributions per lane; FTA top event.
Closing evidence | PV-COMM-001, PRD-VER-005.
HA-RES-008 — Missing excursion alert (never delivered) — HA-COMM-002
Initial | S4 × PM ⇒ R (High). Controls: RCR-028 + RCR-029 (ack/escalate) + RCR-035 (non-suppressible) + SYS-176 reconciliation.
Control-effectiveness | High — end-to-end reconciliation converts silent loss into detectable backlog.
Residual | S4 × PVL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-COMM-002 delivery-path fault-injection). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | monitor un-acked backlog fleet-wide; FTA top event.
Closing evidence | PV-COMM-002.
HA-RES-009 — In-container RF attenuation → no live link — HA-COMM-003 ★ (carried Open)
Initial | S4 × PH (uncertain-probability escalation) ⇒ R (High). Controls: RCR-031 (close in-container budget or IP-sealed external antenna) + RCR-030 (store-and-forward bounds the consequence).
Control-effectiveness | Consequence bounded by store-and-forward (High); link margin itself unknown (to-be-quantified).
Residual | S4 (unchanged) × PM→PL only after PRD-VER-005 quantifies margin and fixes the antenna strategy ⇒ target A (Medium).
Evidence status | Open / Evidence-Pending — the single largest open connectivity item; cannot close until PRD-VER-005 reports and the integral-vs-external decision (CMP-028 vs CMP-029) is made. Classification: Open.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk; named release condition.
Disclosure / monitoring | disclose in-container link limitation; RF/antenna + Regulatory neighbour flags; monitor per-trip delivery-success.
Closing evidence | PRD-VER-005 (+ PV-COMM-003).
HA-RES-010 — Falsified/injected telemetry (security→safety) — HA-CYBER-001 ★
Initial | S4 × PM ⇒ R (High). Controls: RCR-034 (capture-time device-anchored authenticity) + RCR-038 (secure boot/debug-lock) + RCR-055 (unique identity).
Control-effectiveness | High — device-anchored integrity holds against a compromised link/cloud; bounded by the key lifecycle (HA-CIR-009 finding).
Residual | S4 × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending — requires independent adversarial verification (PV-CYBER-001 pen-test). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk; security IV sign-off.
Disclosure / monitoring | monitor integrity-failure events; coordinated vulnerability disclosure; FTA/threat-model (Stage 11).
Closing evidence | PV-CYBER-001 (independent pen-test).
HA-RES-011 — Suppressed / blocked alarm (attacker) — HA-CYBER-002
Initial | S4 × PM ⇒ R (High). Controls: RCR-035 (alarm path not silently suppressible — suppression raises a detectable event) + RCR-029 (ack/escalate).
Control-effectiveness | High — suppression self-detects via missing-ack escalation.
Residual | S4 × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-CYBER-002 alarm-suppression pen-test, IV). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk; security IV.
Disclosure / monitoring | monitor alarm-suppression/anomaly events; FTA top event.
Closing evidence | PV-CYBER-002.
HA-RES-012 — Malicious/compromised OTA → fleet unsafe/blind — HA-CYBER-005 / HA-SW-010 ★
Initial | S4 (S5 fleet-scale) × PM ⇒ R (High). Controls: RCR-037 (signed, rollback-protected A/B, never-brick, staged rollout) + RCR-038 (secure boot).
Control-effectiveness | High device-side; the fleet-rollout common-cause is the residual (HA-CIR-010 / HA-NEW-008 finding — canary + halt-on-regression).
Residual | S4 (S5 worst) × PL (projected) ⇒ A (Medium/High for the S5-worst).
Evidence status | Evidence-Pending (PV-CYBER-005 OTA pen-test + PV-SW-010 rollback/version-skew, IV). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-001 + EM-ENG-006 (S5-worst) + benefit-risk; security IV.
Disclosure / monitoring | disclose OTA/unmonitored-window; monitor campaign health + version distribution; FTA top event.
Closing evidence | PV-CYBER-005, PV-SW-010.
HA-RES-013 — Cloud compromise / outage → fleet blind + records altered — HA-CYBER-009 / HA-CLOUD-008 ★
Initial | S4/S8 (fleet/business) × PM ⇒ R (High). Controls: RCR-039 (cloud resilience, RBAC/SSO, monitoring, tested backup/restore) + RCR-032 (never-drop) + the offline-first invariant (device safety independent of cloud).
Control-effectiveness | High for data survival (offline-first + local commit); Moderate for the fleet-blinding availability during an outage.
Residual | S4 × PL (projected) ⇒ A (Medium) — records survive on-device; the residual is delayed/degraded fleet visibility.
Evidence status | Evidence-Pending (cloud security review + DR test; PV-CYBER-009). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | monitor cloud auth/integrity/backlog + availability; DR-tested; incident response.
Closing evidence | PV-CYBER-009 + DR test.
HA-RES-014 — Use-error cluster: misread verdict / muted alert / mis-armed / override / probe placement — HA-USE-001/002/004/007/005 ★
Initial | S4 × PM (each) ⇒ R (High). Controls: RCR-043 (legible verdict) + RCR-047 (alarm tuning) + RCR-045 (pre-arm gate) + RCR-044 (recorded override) + RCR-040 (honest surface); probe-placement partly HF/IFU (HA-CIR-005 remainder).
Control-effectiveness | Moderate-to-High, human-mediated — these residuals turn on the human and are closed only by summative usability; the last-mile low-training/vulnerable user (HA-CTX-005) keeps P2 elevated.
Residual | S4 × PL (projected) ⇒ A (Medium); probe-placement (HA-USE-005) remains the plausibility-uncatchable remainder, ALARP via HF + IFU.
Evidence status | Evidence-Pending, usability-gated (VAL-USE-001/002/004/007). Classification: Controlled-pending-validation.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | IFU/training; monitor override rates, muted-alert rates, mis-arm/void rates, probe-placement errors; FTA top events.
Closing evidence | VAL-USE-001/002/004/007.
HA-RES-015 — Wrong/stale calibration or config → biased detection — HA-SW-004 / HA-SVC-001 / HA-MFG-001 ★
Initial | S4 (S5 fleet-wide calibration bias) × PM ⇒ R (High). Controls: RCR-026 (integrity + expiry gate) + RCR-053 (EOL calibration coverage + birth record) + RCR-049 (interval enforcement).
Control-effectiveness | High per-device; the fleet-wide systematic calibration bias is the S5-worst common-cause (HA-CIR-007 finding).
Residual | S4 (S5 worst) × PL (projected) ⇒ A (Medium/High for the S5-worst).
Evidence status | Evidence-Pending (PV-SW-004 + PV-MFG-001 EOL coverage; depot recal process). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-001 + EM-ENG-006 (S5-worst) + benefit-risk.
Disclosure / monitoring | monitor calibration-drift trends + EOL/depot calibration escapes fleet-wide; FTA top event.
Closing evidence | PV-SW-004, PV-MFG-001.
HA-RES-016 — Wrong subject/trip attribution / mis-routing — HA-DATA-002 / HA-COMM-008 / HA-CLOUD-003
Initial | S4 × PM ⇒ R (High). Controls: RCR-033 (authoritative binding) + RCR-055 (unique identity).
Control-effectiveness | High — mis-attribution made unrepresentable by authoritative binding.
Residual | S4 × PVL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-COMM-008 / PV-CLOUD-003 association/routing). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | monitor binding/routing anomalies; provisioning validation.
Closing evidence | PV-COMM-008 (= PV-CLOUD-003).
HA-RES-017 — Cloud notify / rules-engine failure or misconfig — HA-CLOUD-001 / HA-CLOUD-002
Initial | S4 × PM ⇒ R (High). Controls: RCR-032 (never-drop, dead-letter/retry, reconciliation) + RCR-036 (authorised, bounded, audited config change) + device-side commit (RCR-028).
Control-effectiveness | High — device holds truth; cloud rules cross-check the device verdict, config change-controlled.
Residual | S4 × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-CLOUD-001/002 + config-tamper). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | monitor notify success + rules-config changes; staged config rollout.
Closing evidence | PV-CLOUD-001/002.
HA-RES-018 — Record/DB loss (device or cloud) → missing evidence — HA-DATA-006 / HA-CLOUD-004/006
Initial | S4 × PM ⇒ R (High). Controls: RCR-024 (power-fail-safe, loss→honest gap) + RCR-032 (cloud never-drop) + RCR-042 (integrity-preserving export) + tested backup/restore (RCR-039).
Control-effectiveness | High — loss surfaces as an honest gap, never a silent absence.
Residual | S4 × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (PV-DATA-006 + DR test). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | monitor loss/gap rates + DR test cadence; ties record-integrity aggregate RCR-057.
Closing evidence | PV-DATA-006, DR test.
HA-RES-019 — Vaccine-lot loss / false-reject stock-out → unvaccinated population — HA-MPE-001 / HA-MPE-003 ★
Initial | S4 × PM ⇒ R (High). Controls: the opposite direction of the safety controls — RCR-047 (validated false-reject targets), RCR-044 (audited informed override), cloud re-adjudication, honest qualifier (RCR-040).
Control-effectiveness | Moderate — this is the accepted-trade residual (HA-NEW-006); it is reduced by validated false-reject targets, not by weakening the safety bias.
Residual | S4 × PL (projected) ⇒ A (Medium) — an accepted public-health trade, balanced against the false-accept residuals.
Evidence status | Evidence-Pending (VAL-USE-002 alarm/false-reject targets + field wastage data). Classification: Controlled-pending-validation (accepted trade).
Acceptance authority | EM-ENG-001 + EM-ENG-006 + benefit-risk (both-directions ALARP).
Disclosure / monitoring | monitor false-reject/wastage/stock-out rates as a first-class field metric; a rising false-reject rate is as reportable as a missed excursion.
Closing evidence | VAL-USE-002 + post-market wastage monitoring.
HA-RES-020 — GDP / data-integrity violation at scale — HA-MPE-004
Initial | S4 (regulatory/programme) × PM ⇒ R (High). Controls: RCR-042 (integrity-preserving export/audit) + RCR-057 (record-integrity aggregate) + time/attribution controls (RCR-025/033).
Control-effectiveness | High — ALCOA+ record integrity by construction.
Residual | S4 × PL (projected) ⇒ A (Medium).
Evidence status | Evidence-Pending (integration + export-integrity + audit). Classification: Controlled-pending-verification.
Acceptance authority | EM-ENG-006 + EM-ENG-001 + benefit-risk.
Disclosure / monitoring | audit-readiness; monitor data-integrity deviations.
Closing evidence | PV-DATA-006 (export), traceability integration test.
HA-RES-021 — Condensation / ingress / cleaning → record-electronics fault — HA-TFC-005/006 / HA-ENV-003/004 / HA-ELEC-006/007
Initial | S3 × PM ⇒ A (Medium). Controls: RCR-015/018 (conformal coat + seal across −30↔+60 °C cycling) + RCR-020/021 (plausibility/verdict catch residual wet-fault reads).
Control-effectiveness | Moderate-High — barrier + screening; reduces to a reliability precursor.
Residual | S3 × PL (projected) ⇒ A (Medium) / trending B.
Evidence status | Evidence-Pending (PV-TFC-006 condensation cycle + IP + insulation-resistance). Classification: Controlled-pending-verification.
Acceptance authority | System-Safety + EM-ENG-006 (S3, region A).
Disclosure / monitoring | cleaning-agent IFU; monitor ingress/BIST failures over reuse life.
Closing evidence | PV-TFC-006, PV-ENV-003.
HA-RES-022 — Time / units / location / estimate misrepresentation — HA-SW-005 / HA-DATA-003/007
Initial | S3 × PM ⇒ A (Medium). Controls: RCR-025 (bounded/disclosed drift) + RCR-041 (estimate flagged) + RCR-042 (export provenance).
Control-effectiveness | High — untrusted time/estimates disclosed by construction.
Residual | S3 × PL (projected) ⇒ A (Medium) — accepted with disclosure.
Evidence status | Evidence-Pending (PV-SW-005 / PV-DATA-003/007). Classification: Controlled-pending-verification.
Acceptance authority | System-Safety + EM-ENG-006.
Disclosure / monitoring | provenance/units disclosed on the surface; monitor untrusted-time incidence.
Closing evidence | PV-SW-005, PV-DATA-003.
6. Residual summary (per-hazard)
• No individual residual remains in region R as projected — every R-region hazard is reduced to region A (or, for S5, region A-High which can never be broadly acceptable) assuming the named verification passes. This is the central caveat: the residual verdicts are projected, Evidence-Pending, not demonstrated.
• Two hazards are carried Open pending their qualification tests: HA-RES-009 (RF-in-container, PRD-VER-005) and the energy-margin part of HA-RES-004 (D3 cold endurance, PRD-VER-004). These gate the release recommendation.
• The S5 residual (HA-RES-006, aircraft-hold fire) stays region-A-High permanently by rule and carries the most stringent acceptance (EM-ENG-001+006 + benefit-risk + mandatory disclosure + post-market monitoring).
• Two residuals are accepted public-health trades (HA-RES-019 false-reject/stock-out; the availability parts of HA-RES-004/014) — balanced against the false-accept residuals, not treated as free.
• Security residuals (HA-RES-010/011/012/013) additionally require independent adversarial verification before their projected residual can be accepted.
7. Lower-severity residuals (aggregate)
The S1–S2 records (mechanical edge/pinch, hot enclosure surface, RF exposure, minor material degradation, transport vibration, dust/UV) reduce to region B (broadly acceptable, Low AP) via their inherent/protective controls (RCR-009/010/011/012/013/014/017), accepted by the responsible discipline engineer + System-Safety record, with standards-based verification (IEC 62368-1, IEC 60068, IEC 61000). They carry no disclosure obligation beyond normal labelling and no dedicated post-market metric.
8. Traceability / closure notes
• Each HA-RES record's residual feeds HAL-2900 (overall/aggregate residual + benefit-risk) and the Risk-Management Summary (release recommendation), and its acceptance authority follows HAL-2200 §7.
• No hazard is closed here. Every S3–S5 residual is Controlled/Open-pending-verification; closure transfers to Stages 13–14 when the named PV-/VAL- passes (SC/security also need IV). The two carried-Open items (PRD-VER-004/005) are explicit conditions of the Conditionally-Ready recommendation.
• Disclosure/monitoring obligations collected here become the post-market surveillance plan (Stage 17) and the IFU/labelling inputs; the S5 and security residuals carry mandatory monitoring.
• FTA top events (the S4/S5 records flagged ★) hand to Stage 11 (ERFKL) for deeper fault-tree/FMEA decomposition.
HAL-2900 (HA-ORES) — Overall Residual Risk & Benefit-Risk
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Owner: EM-ENG-001 (Chief Systems) + EM-ENG-006 (Quality/Regulatory) with System-Safety. ISO 14971 §8 requires that, after all individual residual risks are judged acceptable, the overall residual risk of the device be evaluated as a whole — because risks that are individually acceptable can, in aggregate, in combination, or through common causes, be unacceptable — and that an overall benefit-risk determination conclude whether the medical/ public-health benefit outweighs the overall residual risk. This module performs that whole-device evaluation for the SCM and states the benefit-risk conclusion that feeds the Risk-Management Summary and the release recommendation. Standards cited by designation; revisions to confirm.
---
1. Purpose
Step back from the per-hazard view (HAL-2800) and ask the ISO 14971 §8 questions the individual records cannot: Do many individually-acceptable moderate (region-A) risks accumulate into an unacceptable whole? Can multiple hazards converge on the same beneficiary through one pathway (the undetected/mis-handled excursion)? Are there common causes that defeat several controls at once (firmware, clock, calibration, OTA, cloud, key)? Is the device systemically dependent on connectivity/cloud in a way no single record captures? What is the rare catastrophic contribution (battery fire in an aircraft)? How much uncertainty sits under the whole picture (D3, RF-in-container, WHO PQS accuracy target)? And, finally: does the benefit outweigh the overall residual risk for a public-health vaccine-integrity device?
2. Scope
The whole SCM as released (all SKUs, full lifecycle, multi-market). Inputs: every individual residual (HAL-2800), the control set (HAL-2400), the implementation review + findings (HAL-2500), the verification plan (HAL-2600), and the control-introduced trades (HAL-2700). Output: the overall residual-risk statement, the benefit-risk conclusion, the disclosure/monitoring obligations, and the conditions on release — consumed by the Risk-Management Summary and Stage 11/12/17.
3. Applicable standards & references
• ISO 14971 §8 — overall residual-risk evaluation and overall benefit-risk analysis; §9 production and post-production information as an input to keeping the conclusion valid.
• WHO PQS (cold-chain monitor performance) / WHO-UNICEF EVM — the public-health benefit frame and the accuracy/performance targets against which benefit is judged (target/revision Evidence-Pending, D4).
• EU GDP / ALCOA+ — the evidentiary-record benefit and its data-integrity obligations.
• HAL-2200 §6 — the catastrophic-severity, uncertain-probability, single-fault, and benefit-risk decision rules that govern the whole-device conclusion.
Evidence-honesty note | the benefit-risk conclusion is drawn on projected, Evidence-Pending residuals (V&V is not yet executed). The conclusion is therefore explicitly conditional — it states what must be true (which tests must pass) for the projected acceptability to become demonstrated acceptability. No probability or performance number is asserted as fact; the WHO PQS accuracy target, the D3 cold endurance, and the RF-in-container margin are named as the load-bearing unknowns.
4. Governance & method — how "overall" is assessed
Overall residual risk is assessed along six lenses (§5.1–§5.6), each of which can make an individually-acceptable set unacceptable in aggregate: accumulation, convergence-on-one-beneficiary, common-cause, systemic dependence, rare-catastrophic, and uncertainty. Each lens yields a judgement and, where needed, a cross-cutting control or condition. The benefit-risk conclusion (§6) weighs the resulting whole- device residual against the device's public-health benefit, following HAL-2200 §6.6 (no individual risk accepted in isolation; overall residual governs release).
5. Overall residual-risk evaluation
5.1 Accumulation of moderate (region-A) risks
After controls, almost every top hazard sits in region A (ALARP), not region B — the SCM does not have a handful of moderate risks, it has ~20 moderate residuals, most of them S4, most of them Evidence-Pending. Individually each is acceptable with benefit-risk; in aggregate the device is a dense field of S4 region-A residuals, which is itself a signal that this is a high-consequence information device whose safety rests on many controls all working. The aggregate judgement:
• The accumulation is acceptable only because the residuals share a small number of well-defined defences (honest-gap/no-false-reassurance, local-commit, independent physical backstops) rather than being twenty independent weak spots. The record-integrity aggregate control (RCR-057) exists precisely to assert that the detection→verdict→record→export chain as a whole preserves mission evidence, not just each node.
Cross-cutting condition | the release decision must treat the S4 region-A set as a portfolio — a material slippage in any of the shared defences (e.g., the honest-gap invariant) degrades many residuals at once, so those shared invariants are release-critical and their verification is non-negotiable.
5.2 Convergence: multiple hazards, one beneficiary, one pathway
The SCM's defining feature is that ~15 distinct hazards all converge on the same harm: an undetected / incorrect / delayed excursion → a spoiled vaccine administered as if good → loss of potency → the beneficiary is not protected → preventable disease. Detection error (HA-SW-001), false accept (HA-SW-002), probe fault (HA-SW-008), silent gap (HA-SW-013/HA-ENRG-006), biased calibration (HA-SW-004/HA-DATA-001), wrong profile (HA-SW-003), delayed/missing/undelivered alert (HA-COMM-001/002/003, HA-CLOUD-001), falsified/suppressed data (HA-CYBER-001/002), mis-attribution (HA-DATA-002), misread verdict / muted alert / override (HA-USE-001/002/007), and misleading visualisation (HA-DATA-005) all reach the same beneficiary through the same disposition decision. The aggregate judgement:
• Because they converge, the beneficiary's total probability of an undetected-excursion harm is the union of many moderate probabilities, not any single one — the whole is larger than each part. This is the single most important overall-residual finding.
• This is contained by the fact that the pathway has two independent choke points that every path must pass: (a) the honest verdict / no-false-reassurance node (a spoiled or unverifiable trip cannot present as a clean accept — RCR-021/022/040), and (b) the human disposition decision supported by a legible, honest surface (RCR-043/044). The device's safety architecture is deliberately built so that however the upstream path fails, it converges on "not a clean accept" rather than "a false clean accept." The overall residual is acceptable only if those two choke points hold — making their verification (PV-SW-002, VAL-USE-004) the most load-bearing in the whole programme.
Cross-cutting condition | the convergence must be re-checked at Stage 11 (ERFKL) with a fault tree whose top event is "spoiled dose administered as good" and whose OR-gate collects all these paths, to confirm no single upstream failure defeats both choke points.
5.3 Common causes (defeat several controls at once)
Several elements are common causes whose single failure degrades multiple residuals simultaneously — the aggregate risk they carry exceeds their per-record appearance:
• Firmware — detection, verdict, plausibility, storage, time, alerting all run in one firmware image; a systemic firmware defect (or a bad OTA, HA-CYBER-005/HA-SW-010) is a common cause across the entire information chain. Contained by: IEC 62304 Class-C rigor, the honest-gap invariant surviving a fault (safe-state RCR-023), independent cloud recompute, and staged OTA with rollback.
• Clock/time — mis-time corrupts detection windows, record defensibility, and TLS sessions at once (HA-SW-005/HA-DATA-003/HA-COMM-005). Contained by: bounded/disclosed drift (RCR-025).
• Calibration — a systematic fleet-wide calibration bias defeats detection and verdict on every device simultaneously (HA-DATA-001 S5-worst) — the information chain's dominant common cause (HA-CIR-007). Contained by: EOL calibration coverage + birth record + interval gate; but this is the largest common- cause residual and its verification (PV-MFG-001) + depot process control is a named condition.
• Cloud tenant / rules config — one misconfig or compromise affects many devices (HA-CLOUD-002/008, HA-CYBER-009). Contained by: config change-control, offline-first, cloud monitoring, DR.
• Cryptographic key / secure element — key compromise/remanence defeats authenticity fleet-wide (HA-CYBER-008). Contained by: key lifecycle (provisioning-uniqueness + disposal-erasure), secure boot.
Overall judgement | the common causes are real and are the reason the device is not merely "many independent small risks." Each has a designed containment, but calibration and OTA/cloud are the two common causes whose containment is a process/operational control (EOL + depot recal; canary rollout), not a one-time design fact — so they need standing post-market monitoring, not just a pre-release test.
5.4 Systemic connectivity / cloud dependence
The SCM is architected offline-first precisely so that device safety (detect, commit, verdict, edge-read) does not depend on connectivity or the cloud — a deliberate decision (D2) that makes the systemic- dependence risk bounded. But two capabilities remain genuinely dependent: live/timely alerting (connectivity-bound, HA-COMM-001/003) and the independent cloud recompute cross-check (cloud-bound). The aggregate judgement:
• The safety floor is local: even with zero connectivity for a whole trip, the excursion exists in the record and is readable at receipt (edge read) — so a total connectivity failure degrades the device to a logger with on-receipt verdict, not to an unsafe state. This is the property that makes systemic dependence acceptable.
• The residual systemic dependence is timeliness: without a link, the window to intervene mid-transit (re-ice/segregate) is lost. Whether that window is even usable is lane-dependent and Evidence-Pending (HA-COMM-001 assumptions + PRD-VER-005). The overall residual therefore carries an honest statement that the SCM's real-time-alerting benefit is conditional on connectivity that the operating environment (in-container, deep cold, cross-border) frequently removes — the device is most valuable as an auditable record + on-receipt verdict, with real-time alerting as a best-effort overlay.
5.5 Rare catastrophic contribution (battery fire in aircraft cargo)
One residual stands apart in kind: HA-TFC-002 — a lithium-cell thermal-runaway fire in an occupied aircraft cargo hold (S5, potential loss of aircraft/lives). It is not part of the vaccine-disposition pathway; it is a rare, physically-catastrophic contribution from carrying a rechargeable battery on aircraft. The aggregate judgement:
• Reduced to region-A-High at PVL by independent defense-in-depth (prevent: RCR-002/004; contain: RCR-008 fire enclosure; qualify: UN 38.3/IATA-DGR), it can never be broadly acceptable (HAL-2200 §6.1) and carries the most stringent acceptance and mandatory post-market thermal-event monitoring + transport disclosure.
• It is the one residual whose severity dominates absolutely: no aggregation with the information risks is meaningful (a different harm, a different population), so it is held as a standalone benefit-risk line — the public-health benefit of the device must independently justify carrying a qualified lithium battery on aircraft, which UN 38.3 / IATA-DGR compliance is the accepted basis for.
5.6 Uncertainty under the whole picture
The overall residual sits on three load-bearing unknowns, each of which could move multiple residuals:
• D3 — low-temperature battery: −30 °C / 90-day endurance + charge safety + capacity fade over reuse. Drives loss-of-monitoring (HA-ENRG-006), dead-on-arrival (HA-NEW-001), and the cold-charge fire precursor (HA-ENRG-001). Closed by the 90-day cold endurance test (PRD-VER-004) + cold-charge FI + the battery datasheet. The single largest energy uncertainty.
• RF-in-container link margin + antenna decision. Drives all delayed/missing-alert residuals. Closed by PRD-VER-005. The single largest connectivity uncertainty; carried Open.
• WHO PQS accuracy / performance target (and US classification). The benefit side depends on the device meeting the recognised cold-chain-monitor accuracy target; the exact PQS category/spec revision is Evidence-Pending (D4) and must be confirmed in the regulatory stage — the benefit claim is provisional until the target is fixed and met.
Overall judgement | the whole-device residual is not yet demonstrated — it is a projected acceptability resting on these three unknowns plus the full V&V. This is stated honestly in the benefit-risk conclusion rather than papered over.
6. Overall benefit-risk determination
The benefit. The SCM serves a public-health function of high value: it protects the potency of vaccines in the distribution chain by detecting temperature excursions, alerting responders, and providing an auditable cold-chain record for release decisions. Its benefit is both directions of the disposition: preventing a spoiled dose being administered (a beneficiary unknowingly unprotected → preventable disease) and preventing the needless destruction of potent product (stock-out → an unvaccinated population → preventable disease). Against the counterfactual (no monitoring, or a dumb single-use indicator), a correct, honest, auditable monitor materially improves both patient safety and vaccine availability at population scale.
The overall residual. After controls (as projected, Evidence-Pending), the device carries: a dense field of S4 region-A information residuals that converge on one beneficiary through one pathway, contained by two independent choke points (honest verdict + supported human decision); a small set of common causes (firmware, calibration, OTA/cloud, key, clock) with designed containment, two of which (calibration, OTA/cloud) rely on standing operational control; a bounded systemic connectivity dependence for timely alerting, with a local safety floor; a rare catastrophic aircraft-fire residual held to region-A-High by independent defense-in-depth and UN 38.3; and an accepted public-health trade (conservative safety bias → some wastage/false-reject), balanced by validated false-reject targets.
The determination. On the projected residuals, the public-health benefit of the SCM outweighs its overall residual risk, conditional on the following being demonstrated — i.e., this is a positive but conditional benefit-risk conclusion, not an unconditional one:
1. The two convergence choke points are verified — honest verdict / no-false-reassurance (PV-SW-002, PV-DATA-004) and supported human decision (VAL-USE-004, VAL-USE-001).
2. The physical S5 (aircraft-hold fire) is verified contained (PV-TFC-002 + UN 38.3) and carries its mandatory disclosure + monitoring.
3. The two carried-Open unknowns close acceptably: D3 cold endurance (PRD-VER-004) and RF-in-container margin (PRD-VER-005).
4. The dominant common causes are verified and put under post-market monitoring: fleet-wide calibration (PV-MFG-001 + depot recal control) and OTA/cloud rollout discipline (staged/canary + halt-on-regression).
5. The security controls pass independent adversarial verification (PV-CYBER-001/002/005/006).
6. The benefit target is fixed and met: the WHO PQS accuracy/performance target (and US classification) is confirmed (D4) and the device meets it.
7. Post-market surveillance monitors both harm directions — missed-excursion signals and false-reject/ wastage/stock-out rates — as first-class metrics, feeding the ISO 14971 §9 loop to keep this conclusion valid.
If any of 1–6 fails to close acceptably, the benefit-risk conclusion is not satisfied and release is withheld or restricted (e.g., no air-transport claim without UN 38.3; no real-time-alert claim on lanes where PRD-VER-005 shows the link cannot close, defaulting to the auditable-record + on-receipt-verdict benefit).
7. Overall residual-risk statement (for the Risk-Management Summary)
After application of the risk controls, the Smart Cold Chain Monitor's overall residual risk is judged
acceptable in relation to the public-health benefit of protecting vaccine potency and availability,
conditional on completion of the planned verification and validation (Stages 13–14) — in particular the
honest-verdict and decision-support verifications (the two pathway choke points), the aircraft-hold fire
containment (UN 38.3 + fire-enclosure abuse), the D3 cold-endurance and RF-in-container qualifications (the
two carried-Open unknowns), independent security verification, and confirmation that the device meets the
recognised WHO PQS accuracy target. The S5 aircraft-fire residual can never be broadly acceptable and is
accepted only with EM-ENG-001+006 benefit-risk sign-off, transport disclosure, and mandatory post-market
thermal-event monitoring. The conservative safety bias's false-reject/wastage residual is an accepted,
monitored public-health trade. Overall recommendation: Conditionally Ready — release pending V&V (see
Risk-Management Summary).
8. Traceability / closure notes
• This overall evaluation is authored under RCR-058 (overall residual risk + benefit-risk evaluated per the ISO 14971 risk-management plan before release) and consumes every individual residual (HAL-2800), the control set (HAL-2400), and the trades (HAL-2700).
• Feeds the Risk-Management Summary (the formal release recommendation and open-risk list) and, via it, Stage 12 (Risk-Control Requirements re-baseline), Stage 13–14 (the V&V that must close the conditions), and Stage 17 (post-market surveillance plan carrying the disclosure/monitoring obligations).
Feeds Stage 11 (ERFKL) | the §5.2 convergence fault tree ("spoiled dose administered as good") and the §5.3 common-cause list are the top events / common-cause seeds for the fault-tree and FMEA work.
Keeps valid via ISO 14971 §9 | the benefit-risk conclusion is explicitly conditional and monitored; production and post-production information (field wastage, missed-excursion signals, thermal events, security incidents) feeds back to re-affirm or revise this conclusion.
HAL-300 (HA-LIFE) — Lifecycle Hazard Analysis
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: System-Safety with Manufacturing, Service, and Logistics engineers. Sweeps every lifecycle phase of the reusable (D1) SCM to ensure no phase-specific hazard is missed, and points each to the domain register that carries the detailed record. Based on D1 (reusable), UN-1200 (service), PRD-1500/1600 (mfg/service), and the lifecycle sweep.
1. Purpose
Enumerate the lifecycle phases and, for each, the hazard exposure introduced or amplified, so the library covers the whole life (CLAUDE.md rule 6), not just nominal operation — with particular attention to the reuse loop (return → clean → recharge → recalibrate → reprovision → redeploy), which is unique to D1.
2. Scope
Procurement → manufacturing/assembly/provisioning → packaging → transport → storage → installation/arming → normal/degraded/fault operation → receipt/disposition → cleaning → recharge → maintenance/calibration → software update → cyber incident/recovery → decommission → disposal/recycling.
3. Applicable standards & references
ISO 14971 (lifecycle risk); IEC 62366-1 (use phases); UN 38.3 / IATA DGR (transport); WHO EVM (storage/ handling); IEC 62304 (software lifecycle/OTA). Cited by designation.
4. Lifecycle register (HA-LIFE records — phase → dominant hazard → domain link)
ID | Lifecycle phase | Dominant hazard exposure (SCM) | Perspective | Governing domain record(s)
HA-LIFE-001 | Procurement / incoming | Wrong/counterfeit low-temp cell; wrong-rated protection part; SOUP with vuln | Mfg / mission | HA-MFG-002, HA-ENRG-002, HA-CYBER-005
HA-LIFE-002 | Manufacturing / assembly | Reversed protection component; damaged seal; latent solder/PCB damage | Mfg / beneficiary | HA-MFG-003/004/008, HA-ELEC-007
HA-LIFE-003 | Calibration / provisioning (EOL) | Miscalibration escape; wrong crypto identity/duplicate serial; wrong regional/firmware config | Mfg / beneficiary | HA-MFG-001/005/006, HA-DATA-003
HA-LIFE-004 | Packaging / labeling | Missing UN 38.3/IATA marks; wrong IFU/language; small-parts warning absent | Logistics / bystander | HA-MFG-006, HA-USE-008, HA-MECH-003
HA-LIFE-005 | Transport (air/ground) | Li-battery thermal event in transit; crush/shock → cell damage; vibration fatigue | Air-cargo crew / property | HA-ENRG-003, HA-TFC-002, HA-ENV-005, HA-MECH-009
HA-LIFE-006 | Storage (depot, unarmed) | Deep discharge / self-discharge → dead device; mistaken-for-monitoring | Technician / mission | HA-ENRG-006, HA-USE-004
HA-LIFE-007 | Installation / arming | Forgot-to-arm; wrong profile; probe mis-placed → void/undetected trip | Dispatch / beneficiary | HA-USE-003/004/005, HA-SW-003
HA-LIFE-008 | Normal operation (monitoring) | Missed/incorrect excursion; delayed alert; dead-in-cold | Beneficiary | HA-SW-001, HA-COMM-001, HA-ENRG-006
HA-LIFE-009 | Degraded operation (low energy/comms) | Silent loss of monitoring; alert deferred; graceful-degradation misjudged | Beneficiary | HA-ENRG-006, HA-COMM-001, HA-SW-013
HA-LIFE-010 | Fault operation / recovery | FW hang → false reassurance; torn record on brown-out; unsafe recovery | Beneficiary | HA-SW-009/013, HA-ELEC-008
HA-LIFE-011 | Receipt / disposition | Misread verdict; ignored alert; override of reject; false accept on gap | Last-mile nurse / beneficiary | HA-USE-001/002/007, HA-SW-002
HA-LIFE-012 | Cleaning (between trips) | Cleaning-agent ingress/incompatibility; missed cleaning → contamination | Technician / beneficiary | HA-TFC-005, HA-ENV-004, HA-USE-010
HA-LIFE-013 | Recharge (depot, post-cold) | Cold-charge lithium plating; charge-inhibit mistaken for dead; non-compliant source | Technician / beneficiary | HA-ENRG-001, HA-USE-009, HA-ELEC-002
HA-LIFE-014 | Maintenance / recalibration | Recalibration miss/escape; config not restored; seal not restored | Technician / beneficiary | HA-SVC-001/003/005, HA-USE-011
HA-LIFE-015 | Software update (OTA) | Bricked/unmonitored fleet; version skew; unsafe rollback; unmonitored window | Fleet / beneficiary | HA-SW-010, HA-CLOUD-007, HA-CYBER-005
HA-LIFE-016 | Cyber incident / recovery | Falsified data / suppressed alarm / blocked comms → wrong disposition; ransomware → fleet blind | Beneficiary / mission | HA-CYBER-001/002/007/009
HA-LIFE-017 | Decommission / return logistics | Damaged cell shipped; residual data/keys; wrong retire decision (SoH) | Technician / environment | HA-SVC-006/007, HA-CYBER-008
HA-LIFE-018 | Disposal / recycling | Improper Li-battery disposal → fire/environmental; data remanence | Environment / bystander | HA-SVC-007, HA-MPE-005
5. Reuse-loop emphasis (D1)
The reuse loop (HA-LIFE-012→014) is the phase set unique to the reusable decision and is a hazard amplifier: cold-returned devices concentrate the cold-charge hazard (HA-ENRG-001); reuse without recalibration concentrates the miscalibration→false-accept hazard (HA-SVC-001); reuse without cleaning concentrates contamination (HA-TFC-005); reprovisioning concentrates wrong-config (HA-USE-003). Each is gated by the pre-arm readiness check (SYS-154 / SRS-DIAG-003) — the last-defence control before redeploy.
6. Traceability / closure notes
The lifecycle sweep guarantees phase coverage and routes each phase to its domain record. Feeds the Applicability Register, the PHA (phase view), and the Detailed PHA. No phase is left without at least one governing hazard record; the reuse-loop phases each trace to a pre-arm-gate control (RCR-079).
HAL-400 (HA-IUE) — Intended Use & Human Exposure Questionnaire
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: Human-Factors / System-Safety with EM-ENG-006. The structured intended-use & human-exposure questionnaire (ISO 14971 Annex A / IEC 62366-1) — every question is answered or justified N/A (a blank is not acceptable). Answers seed the use-related (HAL-1500), software (HAL-1000), and data (HAL-1400) registers.
1. Purpose
Systematically interrogate the SCM's intended use and human-exposure characteristics to surface hazards that a component-first analysis would miss — especially the information-harm and use-error pathways that dominate this device.
2. Scope
The full user set (HA-CTX-005), all interaction points (HA-BND-006), and the disposition decision the device supports.
3. Applicable standards & references
ISO 14971 Annex A (questions to identify characteristics related to safety); IEC 62366-1 (use specification, use scenarios, use errors). Cited by designation.
4. Intended-use & human-exposure questionnaire (HA-IUE records)
ID | Question | Answer for the SCM | Hazard seeded
HA-IUE-001 | What is the intended purpose and is safety dependent on correct information? | Cold-chain monitoring supporting accept/reject disposition. Yes — safety depends directly on correct/timely information (information device). | HA-SW-001/002, HA-DATA-001
HA-IUE-002 | Does a delayed critical action cause harm? | Yes — a delayed/missing excursion alert removes the window to act (segregate/reject/re-ice), so a compromised dose may be administered. | HA-COMM-001, HA-CLOUD-001
HA-IUE-003 | Can a harmful decision result from incorrect information? | Yes — the central hazard. A false "accept" → compromised dose administered; a false "reject" → potent dose wasted → stock-out. | HA-SW-002, HA-DATA-001, HA-USE-007
HA-IUE-004 | Is a vulnerable population exposed? | Yes — beneficiaries are often infants/children/pregnant women receiving vaccines; a child bystander may access a discarded device (small parts/battery). | HA-CTX-006, HA-MECH-003, HA-MPE-002
HA-IUE-005 | Is there physical dependence / patient contact? | No direct patient contact — the harm is mediated by the vaccine and the disposition decision. Not a therapeutic device. N/A for direct-contact hazards. | N/A (physical contact)
HA-IUE-006 | Is user intervention required on failure, and is failure detectable? | Yes — on excursion/fault the user must segregate/reject/escalate. Detectability is the crux: silent failures (dead-in-cold, FW hang, missed excursion) are the danger; addressed by honest-gap + self-check. | HA-SW-013, HA-ENRG-006, HA-COMM-001
HA-IUE-007 | Is there a safe override/bypass, and can it be misused? | The user may override a verdict (accept a "reject" or reject an "accept") — a foreseeable, high-consequence use error (automation bias / production pressure). Override must be deliberate and recorded. | HA-USE-007
HA-IUE-008 | What misuse exposure exists? | Forgot-to-arm, wrong profile, probe mis-placed, ignored/muted alert, cold-device discarded, non-compliant charge, field opening — see HA-CTX-003. | HA-USE-003/004/005/009, HA-ENRG-001
HA-IUE-009 | Are there training gaps? | Yes — the last-mile nurse is minimally trained, may never have seen the device before, and reads it once, offline, gloved, in cold. Training cannot be relied on as the primary control (warnings-challenge rule). | HA-USE-001/005
HA-IUE-010 | Are there accessibility constraints? | Yes — gloved actuation, low light, deep cold (OLED response), low literacy, non-primary language → verdict must be legible and language-independent (icon/colour). | HA-USE-001/008, HA-ENV-002
HA-IUE-011 | Can alarms be misunderstood or missed? | Yes — a misread OLED verdict, a muted/ignored alert (alarm fatigue), or a charge-inhibit indication mistaken for a dead device. | HA-USE-001/002/009
HA-IUE-012 | Is there automation bias (over-trust in the device)? | Yes — users may trust a "clean" record without independent shipping indicators, so a false reassurance (gap shown as healthy, or unmonitored trip believed monitored) is especially dangerous. | HA-DATA-004, HA-USE-004, HA-SW-002
HA-IUE-013 | Can the device be used unmonitored while believed monitored? | Yes — forgot-to-arm, wrong profile, dead-in-cold, or a void trip. This is a top hazard (false reassurance). Controlled by the pre-arm gate + honest status. | HA-USE-004, HA-SW-003
HA-IUE-014 | Does the device create physical hazards during handling/charging/transport? | Yes — Li-cell thermal/fire, cold-charge plating, hot self-heat surface, small-parts choking, USB-C abuse. Physical hazards of a reusable Li-ion cellular device. | HA-ENRG-001/002/003, HA-TFC-001, HA-MECH-003
HA-IUE-015 | Is disposal/end-of-life a hazard? | Yes — improper Li-battery disposal (fire/environmental) and data/key remanence. | HA-SVC-007, HA-CYBER-008, HA-MPE-005
5. Traceability / closure notes
Every question is answered (none blank; HA-IUE-005 justified N/A for direct patient contact). The questionnaire confirms the information-harm + use-error pathways dominate and that training/warnings cannot be the primary control for the last-mile user (warnings-challenge rule, HAL-2300). Seeds HAL-1000 (software), HAL-1400 (data), and HAL-1500 (use-related); feeds the summative usability study (Validation).
HAL-500 (HA-ENRG) — Energy Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-500, record code HA-ENRG · Owner: Principal Battery/Power Engineer, with System-Safety (lead) and Quality/Regulatory (EM-ENG-006). This module analyses stored- and delivered-energy hazards — the lithium cell's chemical/electrical energy, the charge path, and the energy budget — following the ISO 14971 harm-and-exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain (initiating cause → hazardous situation → harm), scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM's energy — the stored electrochemical energy of the rechargeable cell, the electrical energy of the charge path, and the finite energy budget that keeps the device monitoring — can cross the system boundary (HA-BND-002) to harm a person, property, or the mission. Two of these are carried as program TOP hazards per HAL-000 §7:
• HA-ENRG-001 — cold-charge lithium plating (the D3 low-temperature-chemistry tension), and
• HA-ENRG-006 — dead-in-cold silent loss of monitoring (the D3 90-day/−30 °C endurance tension).
2. Scope
All SCM SKUs (reusable, multi-trip). In-scope energy elements: rechargeable low-temp cell (CMP-006), pack (CMP-007), BMS/protection AFE (CMP-008), secondary fuse (CMP-009), charge-path/PMIC (CMP-010), fuel-gauge (CMP-011), cell-temperature sensor (CMP-012), self-heat element (CMP-013), charge-input protection front-end (CMP-014), USB-C port (CMP-004), energy governor firmware (CMP-037), and the RF radios as radiated-energy sources (CMP-024/025/026). Lifecycle phases covered: recharge at depot, armed transport (air/ground), receipt, long unarmed storage, service/handling, and reuse over service life (HAL-300). Physical-injury energy paths that terminate in a fire/thermal consequence are cross-referenced to HAL-800 (HA-TFC); mechanical initiating causes (drop/crush) are shared with HAL-600 (HA-MECH); charge-source electrical faults are shared with HAL-700 (HA-ELEC). The information harm of energy exhaustion (HA-ENRG-006/008) uses the HA-BND-003 information pathway.
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• IEC 62368-1 — electrically-caused fire/thermal/energy injury; fire enclosure; safeguards.
• IEC 62133-2 — safety of rechargeable lithium cells/batteries (charge limits, temperature windows, protection).
• UN 38.3 (UN Manual of Tests and Criteria, Part III, 38.3) — lithium-battery transport safety (T1–T8: altitude, thermal, vibration, shock, external short, impact/crush, overcharge, forced discharge).
• IATA DGR — dangerous-goods rules for lithium batteries in air transport (relevant to HA-TFC-002).
• IEC 61000 — EMC (radiated-energy immunity/emissions; used for HA-ENRG-009).
• IEC 60068 — environmental test methods (cold/thermal-cycle; endurance context).
Upstream | D3 (energy/chemistry tension, concept §7); PRD-SAFE; UN-900 (safety needs); DOL-1200 (DO-SAFE controls); HAL-2400 (RCR register); power-energy-architecture.md; component-list.md.
Chemistry note (evidence honesty) | the final low-temperature cell chemistry (CMP-006) is Evidence-Pending. Standard Li-ion cannot be charged below 0 °C without lithium plating (IEC 62133-2 charge-temperature window; general Li-ion electrochemistry); cold usable capacity is derated. No chemistry, capacity, C-rate, or datasheet figure is asserted as fact here — each is carried Evidence-Pending with the named closing test.
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A.
• Severity is fixed by credible worst-case outcome (HAL-2000), never traded against probability.
• Probability is decomposed P1×P2 with an explicit basis and uncertainty flag (HAL-2100); where the D3 chemistry/budget unknowns govern, the band is set conservatively (one level up) and the record is held Evidence-Pending with a named PV-/VAL- test (HAL-2200 §6.2 uncertain-probability rule).
Single-fault rule (HAL-2200 §6.5) | for the S4/S5 battery/charge hazards a single fault shall not reach harm — an independent protective backstop is required (fuse independent of BMS IC; HW charge-interlock independent of FW). This is enforced through RCR-001/002.
Control hierarchy (HAL-2300) | inherent-by-design → protective → information → process. Warnings-only control of a physical energy hazard is challenged (HAL-2200 §6.4).
• Residual risk is re-estimated only after a control is defined, implemented (DO-…), and verified (PV-…); residual probabilities below are projected, Evidence-Pending until the named tests close.
5. Categories & guiding questions (energy prompt list)
1. Stored electrochemical energy (cell): Can the cell enter thermal runaway (overcharge, over-temp, internal short, external short, mechanical damage, latent plating)? Is a single-fault backstop present?
2. Charge energy (path & source): Can charging occur outside the safe temperature/voltage/current window? Can a non-compliant external source inject over-voltage, reverse polarity, surge, or short?
3. Energy budget (availability): Can the device run out of energy — especially in deep cold — such that monitoring stops silently and the gap is misread as an uneventful trip?
4. Degradation over reuse: Does capacity fade, deep discharge, or cold cycling erode the budget or damage the cell over the reusable service life?
5. Handling energy (service): Does stored charge expose a technician to arc/short/burn or ESD damage?
6. Radiated energy: Do the RF radios expose people/equipment (SAR/emissions)?
7. Parasitic/resistive energy: Can loose connections or vibration create localised resistive heating?
6. Hazard Analysis Register (HA-ENRG records)
HA-ENRG-001 — Cold-charge lithium plating → latent internal short → thermal runaway/fire ★ TOP HAZARD (D3)
Hazard ID (library) | HAZ-THERM / HAZ-FIRE (electrochemical stored-energy release)
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Recharge (depot) after cold-return · System State: Charging · Operational Mode: Depot turnaround / recharge
Perspective | Depot technician and any bystander at the charging bench; downstream: mission (a cell that later fails in transit)
Source | D3 (low-temp chemistry tension, concept §7); PRD-SAFE; analogous Li-ion electrochemistry literature; IEC 62133-2 charge-temperature window
Initiating Cause/Event | A device returned cold from a −30 °C trip is placed on charge while the cell is still below 0 °C (cell-coupled temperature, not ambient), before it has warmed.
Contributing Conditions | Standard Li-ion cannot accept charge below ~0 °C without metallic lithium plating on the anode (physics); high-throughput depot wants fast turnaround; ambient may read above 0 °C while the cell core is still cold; final chemistry Evidence-Pending (CMP-006).
Foreseeable Sequence of Events | Cold cell put on charge → lithium plates as dendrites instead of intercalating → repeated cold-charge cycles grow dendrites → dendrite penetrates separator → latent internal short forms and lies dormant → later a normal charge/discharge or a warm-cell event tips it into self-heating.
Hazard | Stored electrochemical energy of a compromised cell (internal short).
Hazardous Situation | A device with a latent internal short is charging or in service near a person (depot bench, or later inside a shipper).
Sequence to Harm | Internal short → localised self-heating → separator meltdown → thermal runaway → vent/flame → burn / smoke inhalation to the technician, or fire in the shipper/transport.
Harm/Loss | Serious burn requiring treatment; loss of the device and potentially the vaccine lot; in transport, escalates to HA-TFC-002.
Severity | S4 (Critical) — governing dimension human injury (serious burn/thermal event needing treatment), per HAL-2000 §7 ("cold-charge lithium plating → latent internal short = S4"). Worst-case in an occupied space escalates to S5 via HA-TFC-002, but the direct cold-charge outcome anchors at S4.
Probability (P1×P2) | P1=3 (Occasional — depots routinely receive cold-returned devices; cold-charge is a foreseeable normal event) × P2=3→4 (Occasional→Probable — plating→latent short accrues over cycles; analogous Li-ion) ⇒ product 9–12 ⇒ Medium–High band, escalated to High (PH) under the uncertain-probability rule — Basis: analogous electrochemistry + judgment; High uncertainty, Evidence-Pending — closed by cold-charge fault-injection sweep (PV-ENRG-001) and cold cycle-life plating characterisation on the selected chemistry (PV-ENRG-008).
Detectability/Warning | Latent short is not directly detectable by the user before it self-heats; the control is to prevent the charge, not detect the damage. Charge-inhibited state must be legible so the cold device is not misread as dead (RCR-046).
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control (R region cannot be accepted; HAL-2200 §6). Single-fault rule applies.
Risk-Control Strategy | 1 (inherent) + 2 (protective) — prevent charge below 0 °C by design, backstopped independently; information for legibility.
Risk-Control Requirement | RCR-001 — charging prevented below 0 °C by an MCU-independent HW interlock on cell-coupled sensing; inhibit persists in any FW/comms fault; FW is defense-in-depth and cannot defeat it. (+ RCR-005 self-heat may only enable the guard by warming the cell, never bypass inhibit; + RCR-046 charge-inhibited state legible.)
Implementation Reference | DO-SAFE-001 (→ DO-EE-007 HW interlock + DO-FW-005 FW defense-in-depth) · CMP-010 (charge path), CMP-012 (cell-temp NTC), CMP-013 (self-heat, if provided)
Verification Reference | PV-ENRG-001 (Evidence-Pending) — cold-charge-inhibit fault-injection sweep across the 0 °C boundary with FW forced-faulty (proves inhibit persists HW-only); PV-ENRG-008 (Evidence-Pending) — cold cycle-life plating characterisation on selected chemistry
Validation Reference | VAL-ENRG-001 (Evidence-Pending) — depot-turnaround field-representative validation that cold devices are safely gated and correctly indicated (ties usability of RCR-046)
Residual Probability | projected Very Low (PVL) after HW interlock verified (plating pathway removed at source) · Residual Severity: S4 (unchanged — outcome unchanged if it ever occurred) · Residual Risk: AP Medium (region A, S4×PVL) — ALARP, benefit-risk sign-off (EM-ENG-006 + EM-ENG-001)
New Risks Introduced | Self-heat element (RCR-005) introduces a hot-surface burn path → HA-TFC-001, and a charge-inhibit that could be misread as a dead device → HA-USE-009 (controlled by RCR-046).
Overall Residual Contribution | After control, small residual; feeds HAL-2900 as a D3 headline residual carried with benefit-risk. FTA top event candidate (Stage 11).
Status | Evidence-Pending (chemistry; probability closed by PV-ENRG-001/008) — Controlled-pending-verification
Owner | Battery/Power Engineer
Rationale/Evidence | IEC 62133-2 (charge-temperature window); Li-ion plating electrochemistry (analogous); D3; RCR-001; DO-SAFE-001. Severity per HAL-2000 §7.
Assumptions | Cell-coupled temperature sensing (CMP-012) resolves the 0 °C boundary with margin; final chemistry retains the standard <0 °C charge prohibition (revisit if an LTO-class chemistry that tolerates cold charge is selected — would lower P1).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-002 — Cell OV/OC/OT/internal-short → thermal runaway/fire
Hazard ID (library) | HAZ-THERM / HAZ-FIRE
Product/Config | SCM all SKUs · Lifecycle Phase: Any (charge, transport, storage) · System State: Charging or discharging · Operational Mode: Any powered
Perspective | Any person near the device — depot technician, air-cargo crew (occupied hold), clinic staff; mission and property
Source | PRD-SAFE; IEC 62368-1; IEC 62133-2; UN 38.3 (T7 overcharge); analogous Li-ion event rates
Initiating Cause/Event | A cell fault or abuse — overvoltage on charge, overcurrent, over-temperature, or a manufacturing-latent internal short — drives the cell outside its safe operating window.
Contributing Conditions | Sealed enclosure concentrates heat; device may be in an occupied aircraft hold; primary protection IC (BMS) could itself fail; wrong-rated protection part (see HA-MFG/RCR-054).
Foreseeable Sequence of Events | Fault condition → cell voltage/current/temperature exceeds limit → exothermic decomposition begins → self-heating accelerates → separator collapse → thermal runaway → vent, flame, offgas.
Hazard | Stored electrochemical energy release (thermal runaway).
Hazardous Situation | A cell in runaway inside the sealed device in an occupied space (hold/depot/clinic).
Sequence to Harm | Runaway → fire/heat/toxic offgas → burn, smoke inhalation, or (occupied aircraft) catastrophic fire.
Harm/Loss | Serious injury (S4) up to death/fatal fire in an occupied hold (S5); loss of device, lot, and potentially aircraft/facility.
Severity | S5 (Catastrophic) worst-case (fire in an occupied aircraft cargo hold — HAL-2000 §5/§7) / S4 typical (depot/clinic burn). Take the highest applicable: S5.
Probability (P1×P2) | P1=1→2 (Improbable–Remote — protection IC + fuse; analogous cell-event rates are low) × P2=3 (Occasional — enclosed near people once initiated) ⇒ product 3–6 ⇒ Low–Medium band (PL–PM) — Basis: analogous Li-ion event rate + designed protection; Medium uncertainty. Closed by protection fault-injection and UN 38.3 (PV-ENRG-002).
Detectability/Warning | Onset may give thermal/swelling/venting warning seconds before flame; BMS over-temp trip is the designed early action. Not reliably human-detectable in a sealed hold.
Initial Risk | AP High (region R — S5 × PM band under conservative escalation; and S5 catastrophic-severity rule forces High regardless)
Risk Acceptability | Unacceptable without control; S5 can never be "broadly acceptable" (HAL-2200 §6.1). Single-fault rule mandatory.
Risk-Control Strategy | 2 (protective) + 1 (inherent fire enclosure)
Risk-Control Requirement | RCR-002 — autonomous OV/UV/OC/OT/short protection with a non-resettable secondary backstop (fuse/OVP), single-fault-safe (hazard terminated if primary IC fails). (+ RCR-008 enclosure contains a single-cell event without external flame; + RCR-054 correct protection parts verified.)
Implementation Reference | DO-SAFE-002 (→ DO-EE-005 BMS + DO-EE-010 fuse) · DO-SAFE-010 (retention) · CMP-008 (BMS), CMP-009 (fuse)
Verification Reference | PV-ENRG-002 (Evidence-Pending) — protection fault-injection (DO-TST-009: force BMS IC failed, confirm fuse/OVP terminates) + UN 38.3 abuse suite (T5 external short, T7 overcharge)
Validation Reference | N/A (physical abuse verification carries the assurance; no separate use-validation)
Residual Probability | projected Very Low (PVL) after dual-independent protection verified · Residual Severity: S5 (unchanged) · Residual Risk: AP High→Medium (region A, S5×PVL) — per §6.1 an S5 residual stays region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off + post-market monitoring
New Risks Introduced | Non-resettable fuse creates a nuisance dead-device path (not a safety hazard) — handled as reliability, not HAL.
Overall Residual Contribution | S5 residual → mandatory benefit-risk line in HAL-2900; FTA top event (Stage 11).
Status | Controlled-pending-verification (Evidence-Pending probability)
Owner | Battery/Power Engineer
Rationale/Evidence | IEC 62133-2, IEC 62368-1 (fire enclosure), UN 38.3; RCR-002/008. Severity per HAL-2000 §5/§7.
Assumptions | Secondary fuse/OVP is genuinely independent (no common-cause) from the BMS IC (RCR-002 independence clause); enclosure fire-containment claim to be proven by abuse test (ties HA-TFC-002).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-003 — Mechanical crush/puncture of cell (drop/transport) → internal short → thermal event
Hazard ID (library) | HAZ-THERM / HAZ-FIRE (mechanically initiated)
Product/Config | SCM all SKUs · Lifecycle Phase: Transport / handling · System State: Any · Operational Mode: Armed transport or handling
Perspective | Air-cargo crew, logistics handler, depot technician; mission/property
Source | UN 38.3 (T4 shock, T6 impact/crush); PRD-SAFE; ties HA-MECH-009
Initiating Cause/Event | Drop, crushing load in transit, or impact deforms the enclosure enough to crush or puncture the cell.
Contributing Conditions | Air/ground transport shock/vibration; stacking loads; enclosure/retention margin; cell soft-pouch vs hard-case format (Evidence-Pending, CMP-007).
Foreseeable Sequence of Events | Impact/crush → cell deformation/puncture → separator breach → internal short → self-heating → thermal event.
Hazard | Stored electrochemical energy release via mechanical damage.
Hazardous Situation | A mechanically damaged cell in a device in transport or being handled near people.
Sequence to Harm | Internal short → thermal runaway → fire/burn/offgas → injury; in an occupied hold escalates to HA-TFC-002 (S5).
Harm/Loss | Serious burn / thermal event (S4); catastrophic in occupied hold (S5).
Severity | S4 (Critical) typical / S5 worst (occupied aircraft) — take S5 as credible worst-case (HAL-2000 §7 lists HA-ENRG-003 among S5-capable battery events).
Probability (P1×P2) | P1=2 (Remote — retention + enclosure designed to UN 38.3; but transport abuse is real) × P2=3 (Occasional once cell is breached) ⇒ product 6 ⇒ Medium band (PM) — Basis: UN 38.3 design intent + judgment; Medium uncertainty. Closed by drop/vibration + UN 38.3 (PV-ENRG-003).
Detectability/Warning | Post-drop damage may be invisible externally; no reliable user warning of a breached cell.
Initial Risk | AP High (region R — S5 × PM; catastrophic-severity rule)
Risk Acceptability | Unacceptable without control; single-fault + inherent retention required.
Risk-Control Strategy | 1 (inherent — retain cell without crush/puncture) + 2 (protective — enclosure containment)
Risk-Control Requirement | RCR-004 — cell retained without crush/puncture through drop/vibration/transport; thermal/electrical clearance maintained. (+ RCR-007 UN 38.3-tested & IATA-compliant; + RCR-008 enclosure containment.)
Implementation Reference | DO-SAFE-010 (→ DO-MECH-012 retention + DO-EE-004 pack) · CMP-007 (pack), CMP-001 (enclosure)
Verification Reference | PV-ENRG-003 (Evidence-Pending) — drop/vibration to profile + UN 38.3 T4/T6 (shock, impact/crush)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after retention + UN 38.3 verified · Residual Severity: S5 (unchanged) · Residual Risk: AP High→Medium (region A, S5×PVL) with benefit-risk sign-off
New Risks Introduced | None new (shares controls with HA-MECH-009, HA-TFC-002).
Overall Residual Contribution | S5 residual carried in HAL-2900 benefit-risk; FTA top event. Ties HA-MECH-009 (mechanical) and HA-TFC-002 (transport fire).
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Battery/Power Engineer (with Mechanical)
Rationale/Evidence | UN 38.3 T4/T6; IEC 62368-1; RCR-004/007/008. Severity per HAL-2000 §7.
Assumptions | Final cell format/retention closes the crush margin; transport profile bounded by IATA DGR conditions.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-004 — Stored charge during service handling → arc/short/burn to technician
Hazard ID (library) | HAZ-THERM (electrical energy, low-voltage high-current)
Product/Config | SCM all SKUs · Lifecycle Phase: Service / maintenance (battery replacement, board rework) · System State: Powered-down but energy present · Operational Mode: Service
Perspective | Depot / repair technician
Source | PRD-SAFE; IEC 62368-1; judgment (Li-ion pack service)
Initiating Cause/Event | During service a tool or conductor bridges the charged cell/pack terminals, or a lead is shorted, releasing high current.
Contributing Conditions | Cell is a SELV but high-energy source (large short-circuit current); service may occur without a battery-safe procedure; sealed enclosure implies field service is out-of-scope (HA-CTX-004) but depot service is in-scope.
Foreseeable Sequence of Events | Technician opens device → tool bridges terminals → arc / high short-circuit current → localised heating/spark → burn or startle → possible cell damage (feeds HA-ENRG-002).
Hazard | Stored electrical energy (short-circuit current / arc).
Hazardous Situation | Technician handling an energised pack with exposed conductors.
Sequence to Harm | Arc/short → burn to hand / eye flash / startle → minor-to-moderate injury.
Harm/Loss | Moderate reversible burn requiring first aid (S3).
Severity | S3 (Serious) — human injury, moderate reversible burn (HAL-2000 §5, S3 anchor).
Probability (P1×P2) | P1=2 (Remote — controlled depot service with procedure) × P2=2 (Remote — arc→treatable burn) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62368-1; Medium uncertainty.
Detectability/Warning | Energised state indicated (LED); procedure and secondary fuse limit sustained current.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP — reduce as far as practicable via protective + process controls.
Risk-Control Strategy | 2 (protective — fuse limits energy) + 4 (process — battery-safe service procedure)
Risk-Control Requirement | RCR-002 (secondary fuse limits fault current) + RCR-013 (SELV: no accessible hazardous voltage). Service procedure per DO-SVC (RCR-050 post-service BIST).
Implementation Reference | DO-SAFE-002 (fuse) · CMP-008/009 · DO-SVC (service procedure)
Verification Reference | PV-ENRG-004 (Evidence-Pending) — service-short fault-injection (confirm fuse limits current); procedure review
Validation Reference | VAL-ENRG-004 (Evidence-Pending) — service-procedure usability with representative technicians
Residual Probability | projected Very Low (PVL) with fuse + procedure · Residual Severity: S3 · Residual Risk: AP Low→Medium (region A/B boundary, S3×PVL → B Low)
New Risks Introduced | None.
Overall Residual Contribution | Small; occupational-safety residual noted in HAL-2900.
Status | Controlled-pending-verification
Owner | Battery/Power Engineer (with Service)
Rationale/Evidence | IEC 62368-1 energy-source classification; RCR-002/013. Severity per HAL-2000 §5.
Assumptions | Field service is prohibited (HA-CTX-004); only trained depot service opens the device.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-005 — Out-of-spec external charge source (OVP/reverse/surge from third-party USB) → overcharge → thermal event
Hazard ID (library) | HAZ-THERM / HAZ-FIRE (charge-input electrical energy)
Product/Config | SCM all SKUs · Lifecycle Phase: Recharge (depot or field, third-party charger) · System State: Charging · Operational Mode: Recharge
Perspective | Person charging the device (technician / field user); property
Source | PRD-SAFE; USB-C/USB-PD; IEC 62368-1; ties HA-ELEC-002; HA-CTX-010 (third-party chargers)
Initiating Cause/Event | A non-compliant or faulty third-party USB source presents over-voltage, reversed polarity, surge, or a short to the USB-C input.
Contributing Conditions | Users foreseeably charge from any available USB source (HA-CTX-003 misuse); no guarantee the source is compliant; enclosure concentrates heat.
Foreseeable Sequence of Events | Bad source connected → over-voltage/surge reaches charge path → PMIC/charge control stressed → cell driven toward overcharge → over-temperature → thermal event.
Hazard | Uncontrolled charge-input electrical energy → cell overcharge.
Hazardous Situation | Device charging from a hazardous source near a person.
Sequence to Harm | Overcharge → over-temperature/venting → burn / fire → injury.
Harm/Loss | Serious burn / thermal event (S4); property loss.
Severity | S4 (Critical) — human injury (thermal), consistent with HA-ELEC-002 (HAL-2000 §7 cold-charge/electrical S4 grouping).
Probability (P1×P2) | P1=3 (Occasional — third-party sources are common, some non-compliant) × P2=2 (Remote — input protection + BMS backstop must both fail to reach cell) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + USB-PD field reality; Medium uncertainty. Closed by reverse/OVP/short test (PV-ENRG-005).
Detectability/Warning | Device should reject/limit the source and indicate a charge fault (RCR-046 legibility); protection is designed to act before the cell is stressed.
Initial Risk | AP High (region R — S4 × PM)
Risk Acceptability | Unacceptable without control; single-fault rule (input protection + BMS are independent layers).
Risk-Control Strategy | 2 (protective — reject/limit bad source) with BMS (RCR-002) as independent backstop
Risk-Control Requirement | RCR-003 — device rejects/limits an out-of-spec charge source (reverse-insertion, OVP, surge, short); USB-PD PDO capped. (+ RCR-002 BMS backstop.)
Implementation Reference | DO-SAFE-003 (→ DO-EE-009 USB-C sink + DO-EE-010) · CMP-014 (input protection front-end), CMP-004 (USB-C port)
Verification Reference | PV-ENRG-005 (Evidence-Pending) — reverse/OVP/surge/short input test; USB-PD PDO-cap test
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after input protection verified · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | None (protection acts inertly).
Overall Residual Contribution | Small; ties HA-ELEC-002 (shared control). Feeds HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending)
Owner | Battery/Power Engineer (with Electrical)
Rationale/Evidence | IEC 62368-1; USB-PD spec; RCR-002/003. Severity per HAL-2000 §7.
Assumptions | Two independent layers (input front-end + BMS) with no common-cause; PD cap value Evidence-Pending (DO-SAFE-003 note).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-006 — Energy exhaustion / DEAD-IN-COLD mid-trip → SILENT loss of monitoring → gap reads as uneventful → wrong disposition ★ TOP HAZARD (D3, information harm)
Hazard ID (library) | HAZ-THERM (energy depletion) → information-integrity harm (HA-BND-003 pathway)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport (deep cold) · System State: Monitoring/logging · Operational Mode: Armed / logging
Perspective | Beneficiary/patient (ultimate harmed party via a compromised dose); disposition decision-maker (last-mile clinic); mission/programme
Source | D3 (90-day/−30 °C endurance tension, concept §7); PRD-SAFE/REL; HA-BND-003; HAL-2000 §7 ("lost record in deep cold, silent = S4")
Initiating Cause/Event | The cell's usable energy is exhausted mid-trip — cold derates capacity (physics), the 90-day/−30 °C budget is in genuine tension (D3), and the device powers down before the trip ends.
Contributing Conditions | Cold capacity derating (Evidence-Pending, CMP-006); long trips near the endurance limit; radios drawing more than budget in poor coverage; gauge inaccuracy in cold (CMP-011). The dangerous property is that the gap is silent — a dead device logs nothing, and an absence of excursion records can be misread as "no excursion occurred."
Foreseeable Sequence of Events | Cold + long trip → usable energy exhausted → device powers down mid-trip → an excursion occurs during the dark window and is not recorded → device (if it revives) or cloud shows a clean-looking record with a gap → human reads the gap as an uneventful trip → accepts a compromised dose.
Hazard | Loss of monitoring energy (availability failure) creating a silent information gap.
Hazardous Situation | A compromised dose at point-of-care believed acceptable because monitoring silently stopped.
Sequence to Harm | Silent gap → false "accept" → potency-lost dose administered → beneficiary not protected → preventable disease (or, if the gap triggers a false reject, stock-out).
Harm/Loss | A lot/shipment of compromised doses administered to a group of beneficiaries (S4; S5 at population scale).
Severity | S4 (Critical) typical / S5 worst (lot/population scale) — governing dimension public-health (HAL-2000 §6 severity anchor; §7 "lost record in deep cold = S4"). Take S4 as the record anchor, noting S5 credible worst-case.
Probability (P1×P2) | P1=3 (Occasional — the 90-day/−30 °C combination is in genuine tension; cold devices near budget limit) × P2=4 (Probable — a silent gap is readily misread as uneventful; downstream receipt QA is an external, imperfect check) ⇒ product 12 ⇒ High band (PH) — Basis: judgment + D3 energy tension; High uncertainty, Evidence-Pending — closed by the 90-day cold endurance & reserve test (PV-ENRG-006 / PRD-VER-004).
Detectability/Warning | The control makes the failure honest: reserve preserves record-commit and a low-battery alert; dead-in-cold state must be indicated, and any monitoring gap must be exposed so the verdict cannot show a clean pass over a gap (RCR-022/027). Without the control, the failure is silent — the core danger.
Initial Risk | AP High (region R — S4 × PH)
Risk Acceptability | Unacceptable without control; drives design by severity while probability is Evidence-Pending.
Risk-Control Strategy | 1 (inherent — energy governor + reserve) + 3 (information for safety — honest gap, low-battery alert, never a clean pass over a gap)
Risk-Control Requirement | RCR-006 — energy governor + reserve prevents silent loss of monitoring; reserve preserves record-commit and a low-battery alert; honest low-battery/dead-in-cold indication; 90-day cold endurance met. (+ RCR-027 availability self-check flags "not monitoring"; + RCR-022 honest gap in verdict/record; + RCR-021 verdict suppresses false accept on a gap.)
Implementation Reference | DO-SAFE-012 (→ DO-FW-004 governor + DO-EE-008 gauge) · CMP-037 (energy governor), CMP-011 (fuel gauge), CMP-033 (storage — power-fail-safe commit)
Verification Reference | PV-ENRG-006 (Evidence-Pending) — 90-day cold endurance & reserve test (PRD-VER-004) at −30 °C with representative duty cycle; reserve-preserves-commit-and-alert test
Validation Reference | VAL-ENRG-006 (Evidence-Pending) — decision-support validation that a gap is never read as a clean pass (ties RCR-040 usability)
Residual Probability | projected Low (PL) after governor/reserve verified and gap-exposure enforced (residual driven by the honest-gap information control, not by eliminating exhaustion) · Residual Severity: S4 (unchanged) · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP with benefit-risk sign-off; stays Evidence-Pending until PV-ENRG-006 closes the endurance budget
New Risks Introduced | A larger cell/reserve to meet endurance adds stored energy → marginally increases HA-ENRG-002/003 thermal exposure (energy↔safety trade, noted for HAL-2700).
Overall Residual Contribution | D3 headline residual — the endurance budget is the single most load-bearing open item; carried explicitly in HAL-2900 benefit-risk. FTA top event (Stage 11); ties HA-SW-013 (silent outage) and HA-USE-004 (gap misread).
Status | Evidence-Pending (budget) — Open until PV-ENRG-006/PRD-VER-004 closes
Owner | Battery/Power Engineer (with Firmware)
Rationale/Evidence | D3; cold-capacity derating physics; HAL-2000 §6/§7; RCR-006/027; DO-SAFE-012. CMP-037 is "the single most load-bearing FW component for the D3 budget" (component-list).
Assumptions | Final chemistry's cold usable capacity meets the SYS-087/090 budget with reserve margin (Evidence-Pending); receipt QA is a contributing (external) check, not a controlled barrier.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-007 — Deep discharge in long unarmed storage → cell damage/venting on later recharge
Hazard ID (library) | HAZ-THERM (over-discharge → damaged cell → thermal on recharge)
Product/Config | SCM all SKUs · Lifecycle Phase: Long unarmed storage between deployments · System State: Deep sleep / off · Operational Mode: Idle storage
Perspective | Depot technician who later recharges; mission (fleet reliability)
Source | PRD-SAFE; IEC 62133-2 (over-discharge); judgment
Initiating Cause/Event | A device sits in storage long enough that self-discharge + parasitic draw takes the cell into deep over-discharge below the safe floor.
Contributing Conditions | Reusable fleet with idle periods; no periodic top-up policy; UV protection cutoff may not fully isolate parasitic loads (Evidence-Pending).
Foreseeable Sequence of Events | Long storage → cell over-discharged (copper dissolution / internal damage physics) → later placed on charge → damaged cell heats/gasses/vents on recharge.
Hazard | Over-discharge-damaged cell → thermal/venting energy on recharge.
Hazardous Situation | A technician recharging an over-discharged (internally damaged) cell.
Sequence to Harm | Recharge of damaged cell → gassing/venting/heat → minor burn or device damage.
Harm/Loss | Moderate reversible injury / device loss (S3).
Severity | S3 (Serious) — human injury moderate; also property loss.
Probability (P1×P2) | P1=2 (Remote — UV cutoff + storage policy) × P2=2 (Remote — damaged-cell recharge → venting) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62133-2; Medium uncertainty.
Detectability/Warning | Gauge/BMS should flag an over-discharged cell and block/limit recharge; charge-fault indication (RCR-046).
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP
Risk-Control Strategy | 2 (protective — UV cutoff + block recharge of a damaged cell) + 4 (process — storage top-up policy)
Risk-Control Requirement | RCR-006 (energy governor / low-battery management incl. storage reserve) + RCR-054 (correct cell + protection; over-discharge protection thresholds per datasheet). Storage top-up procedure per DO-SVC.
Implementation Reference | DO-SAFE-012 (governor) · CMP-008 (UV protection), CMP-037 (governor), CMP-011 (gauge)
Verification Reference | PV-ENRG-007 (Evidence-Pending) — over-discharge → controlled-recharge safety test; storage self-discharge characterisation
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) with UV cutoff + block-damaged-recharge · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small; fleet-reliability + minor-safety residual in HAL-2900.
Status | Controlled-pending-verification
Owner | Battery/Power Engineer
Rationale/Evidence | IEC 62133-2 over-discharge; RCR-006/054. Severity per HAL-2000 §5.
Assumptions | A storage top-up / max-shelf policy exists; UV cutoff isolates parasitic loads below the safe floor.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-008 — Capacity fade over reuse cycles → premature end-of-trip energy loss → loss of monitoring
Hazard ID (library) | HAZ-THERM (degradation) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs (reusable, multi-trip) · Lifecycle Phase: Late-life reuse · System State: Monitoring/logging · Operational Mode: Armed / logging
Perspective | Beneficiary/patient; disposition decision-maker; mission
Source | PRD-SAFE/REL; HA-CTX-008 (cycle-life Evidence-Pending); ties HA-ENRG-006
Initiating Cause/Event | After many charge/discharge cycles (worsened by cold cycling), the cell's usable capacity fades below the level needed to complete a full trip.
Contributing Conditions | Reusable over multi-year life (D1); cold cycling accelerates fade (physics); cycle-life is Evidence-Pending (HA-CTX-008); the pre-arm readiness gate must account for SoH, not just SoC.
Foreseeable Sequence of Events | Repeated cycles → capacity fades → an armed old device is dispatched with insufficient reserve → energy exhausted before trip end → silent monitoring gap → wrong disposition (as HA-ENRG-006).
Hazard | State-of-health degradation eroding the monitoring energy budget.
Hazardous Situation | An aged, faded-capacity device armed for a trip it cannot complete.
Sequence to Harm | Premature exhaustion → silent gap → false accept → compromised dose administered.
Harm/Loss | A lot of compromised doses administered (S4).
Severity | S4 (Critical) — public-health (same anchor as HA-ENRG-006).
Probability (P1×P2) | P1=3 (Occasional — fade is certain over life; the question is whether the readiness gate catches it) × P2=3 (Occasional — SoH-aware pre-arm gate reduces the reach-to-harm) ⇒ product 9 ⇒ Medium band (PM), escalated to High under uncertainty — Basis: judgment + Evidence-Pending cold cycle-life; High uncertainty — closed by cold cycle-life test (PV-ENRG-008).
Detectability/Warning | SoH tracked by fuel gauge (CMP-011); pre-arm readiness gate blocks arming without sufficient SoC-for-trip (RCR-045) — this is the primary defence.
Initial Risk | AP High (region R — S4 × PM/PH)
Risk Acceptability | Unacceptable without control
Risk-Control Strategy | 1 (inherent — SoH-aware readiness gate) + 4 (process — recalibration/EOL retirement)
Risk-Control Requirement | RCR-006 (energy governor + reserve, 90-day endurance incl. aged margin) + RCR-049 (recalibration/service interval enforcement) + RCR-045 (pre-arm gate blocks arming unless SoC-for-trip passes).
Implementation Reference | DO-SAFE-012 (governor/reserve) · DO-SAFE-007 (pre-arm gate) · CMP-011 (gauge SoH), CMP-037 (governor)
Verification Reference | PV-ENRG-008 (Evidence-Pending) — cold cycle-life test to end-of-life capacity; SoH-gate blocks-arming test
Validation Reference | VAL-ENRG-008 (Evidence-Pending) — fleet retirement-policy validation
Residual Probability | projected Low (PL) after SoH-aware gate verified · Residual Severity: S4 · Residual Risk: AP High→Medium (region A, S4×PL) — ALARP; Evidence-Pending until cycle-life known
New Risks Introduced | Aggressive retirement increases fleet cost / disposal volume (env, HA-MPE-005) — not a safety hazard.
Overall Residual Contribution | Ties HA-ENRG-006 as a D3 endurance residual; carried in HAL-2900. Cold cycle-life is a named Evidence-Pending closure.
Status | Evidence-Pending (cold cycle-life) — Controlled-pending-verification
Owner | Battery/Power Engineer
Rationale/Evidence | Li-ion capacity-fade physics; HA-CTX-008; RCR-006/045/049. Severity per HAL-2000 §6.
Assumptions | Fuel-gauge SoH estimate is accurate enough at cold to drive the pre-arm gate (gauge accuracy Evidence-Pending, CMP-011).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-009 — RF / electromagnetic energy exposure near people/equipment (SAR / emissions)
Hazard ID (library) | HAZ-THERM/RADN (radiated electromagnetic energy)
Product/Config | SCM all SKUs · Lifecycle Phase: Armed transport / operation · System State: Transmitting (LTE-M/Wi-Fi/BLE/GNSS) · Operational Mode: Upload/alert window
Perspective | People near the device; nearby aircraft/medical equipment (emissions)
Source | IEC 61000 (EMC); RF-exposure limits (SAR/MPE regulatory); ties HA-ELEC-003/004
Initiating Cause/Event | The radios emit RF energy during transmit windows.
Contributing Conditions | Low-power IoT radios; device is inside an insulated shipper (attenuates external field); transmit is duty-cycled by the energy governor (CMP-037).
Foreseeable Sequence of Events | Radio transmits → RF field present → (if above limits) exposure to a nearby person or interference to equipment.
Hazard | Radiated RF/electromagnetic energy.
Hazardous Situation | A person or sensitive equipment within the near-field during transmit.
Sequence to Harm | RF exposure above limit → tissue heating (SAR) — highly unlikely at IoT power levels; or emissions interfere with aircraft/medical equipment.
Harm/Loss | Minor / negligible for SAR at these power levels; equipment interference addressed under EMC.
Severity | S2 (Minor) for the direct RF-exposure path (low-power radios, duty-cycled; well below harm thresholds). (Emissions-into-aircraft interference is the higher-severity concern and is owned by HA-ELEC-004 / RCR-014, S5-capable — cross-referenced, not double-counted here.)
Probability (P1×P2) | P1=2 (Remote — compliant low-power radios) × P2=1 (Improbable — SAR harm at IoT levels) ⇒ product 2 ⇒ Very Low band (PVL) — Basis: RF-power analysis + regulatory limits; Low uncertainty. Closed by emissions/SAR compliance (PV-ENRG-009).
Detectability/Warning | N/A to user; managed by compliance design + duty-cycling.
Initial Risk | AP Low (region B — S2 × PVL)
Risk Acceptability | Broadly acceptable for the SAR path (ALARP via compliant design); emissions-into-equipment handled by RCR-014.
Risk-Control Strategy | 1 (inherent — compliant low-power radios) + 2 (protective — EMC emissions control)
Risk-Control Requirement | RCR-014 — meet EMC emissions so as not to interfere with aircraft/medical equipment; RF exposure within regulatory limits.
Implementation Reference | DO-EE (EMC), DO-PCB-010 · CMP-024/025/026 (radios), CMP-028 (antenna)
Verification Reference | PV-ENRG-009 (Evidence-Pending) — RF-exposure (SAR/MPE) assessment + EMC emissions test (with HA-ELEC-004)
Validation Reference | N/A
Residual Probability | Very Low (PVL) · Residual Severity: S2 · Residual Risk: AP Low (region B)
New Risks Introduced | None.
Overall Residual Contribution | Negligible (direct RF path); the emissions-into-aircraft path is the material one, carried under HA-ELEC-004.
Status | Controlled-pending-verification
Owner | Battery/Power Engineer (defers emissions to Electrical/EMC, EM-ENG-003)
Rationale/Evidence | IEC 61000; RF-exposure limits; RCR-014. Severity per HAL-2000 §5.
Assumptions | Radios operate at typical IoT power; SAR is well below limits at the enclosure surface.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-010 — Charge overcurrent to a hot cell (upper temperature limit) → gassing
Hazard ID (library) | HAZ-THERM (charge outside upper temperature window)
Product/Config | SCM all SKUs · Lifecycle Phase: Recharge in a hot environment (+60 °C ambient) · System State: Charging · Operational Mode: Recharge
Perspective | Person charging; property
Source | IEC 62133-2 (charge temperature window has an upper limit too); PRD-SAFE; judgment
Initiating Cause/Event | The device is charged while the cell is above its safe upper charge-temperature limit (hot ambient, or self-heating on top of charging).
Contributing Conditions | +60 °C operating ceiling; self-heat element (CMP-013) adds heat; charge-inhibit logic must guard the upper boundary as well as the 0 °C lower boundary.
Foreseeable Sequence of Events | Hot cell put on charge → charging above upper temp limit → accelerated side reactions / gassing → cell stress (precursor to venting).
Hazard | Charging outside the upper safe temperature window → gassing/degradation.
Hazardous Situation | A hot cell being charged.
Sequence to Harm | Overheat/gassing → venting → minor burn or accelerated degradation (feeds HA-ENRG-002 in the worst case).
Harm/Loss | Moderate reversible injury / cell degradation (S3).
Severity | S3 (Serious) — human injury moderate; degradation.
Probability (P1×P2) | P1=2 (Remote — temp-qualified charge enable both bounds) × P2=2 (Remote) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62133-2; Medium uncertainty.
Detectability/Warning | Cell-temp NTC (CMP-012) qualifies charge enable at both boundaries; charge-fault indication.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP
Risk-Control Strategy | 2 (protective — temperature-qualified charge enable, both bounds)
Risk-Control Requirement | RCR-002 (OT protection) + RCR-005 (self-heat cannot exceed the safe upper thermal limit) — charge enable qualified by cell temperature at the upper as well as the lower boundary.
Implementation Reference | DO-SAFE-002 (OT) / DO-SAFE-001 (self-heat bound) · CMP-012 (cell-temp NTC), CMP-010 (charger enable), CMP-013 (self-heat)
Verification Reference | PV-ENRG-010 (Evidence-Pending) — upper-temperature charge-inhibit test; self-heat upper-bound test
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small; ties HA-TFC-008 (hot-ambient charge surface).
Status | Controlled-pending-verification
Owner | Battery/Power Engineer
Rationale/Evidence | IEC 62133-2 (upper charge-temperature limit); RCR-002/005. Severity per HAL-2000 §5.
Assumptions | Charge-enable logic guards both temperature bounds; self-heat cannot be active while ambient is already hot (interlock).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-011 — ESD / stored-charge event on service handling → component damage
Hazard ID (library) | HAZ-THERM/RADN (electrostatic discharge energy)
Product/Config | SCM all SKUs · Lifecycle Phase: Service / manufacturing handling · System State: Powered-down, board exposed · Operational Mode: Service / assembly
Perspective | Technician (no injury); mission (a latently-damaged safety-critical net could escape → false reading)
Source | IEC 61000-4-2 (ESD); IPC/ESD handling; ties HA-ELEC-008
Initiating Cause/Event | ESD from a technician/tool discharges into exposed circuitry during service or assembly.
Contributing Conditions | Board exposed during service; no or degraded ESD controls; the damaged part may be on a safety-critical net (sensor/verdict path) and fail latently.
Foreseeable Sequence of Events | ESD event → semiconductor damage → part fails now or degrades latently → later a wrong reading or missed excursion (feeds HA-ELEC-008 information path).
Hazard | Electrostatic discharge energy into sensitive components.
Hazardous Situation | A latently ESD-damaged safety-critical component in a fielded device.
Sequence to Harm | Latent damage → wrong reading / missed excursion → wrong disposition (information harm) — but caught by screening + plausibility checks.
Harm/Loss | Minor directly (no injury); the escalation to information harm is owned by HA-ELEC-008. This record scores the direct handling event.
Severity | S2 (Minor) — no injury; direct component-damage/nuisance. (Escalation to S4 information harm is carried by HA-ELEC-008/RCR-016/020, not double-counted.)
Probability (P1×P2) | P1=2 (Remote — ESD-controlled service/assembly) × P2=2 (Remote — screening catches most escapes) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IPC/ESD practice; Medium uncertainty.
Detectability/Warning | AOI/ICT + EOL test + sensor plausibility checks screen latent damage.
Initial Risk | AP Low→Medium (region B/A — S2 × PL → B Low)
Risk Acceptability | Broadly acceptable / ALARP via handling + screening.
Risk-Control Strategy | 4 (process — ESD handling) + 2 (protective — on-board ESD protection) + screening
Risk-Control Requirement | RCR-013/014 (ESD immunity + protection) + RCR-016 (latent defect screening AOI/ICT/EOL, no latent single-point escape on safety-critical nets).
Implementation Reference | DO-EE (ESD protection), DO-PCB · DO-TST (AOI/ICT) · CMP-034 (PCB), CMP-016 (secure element handling)
Verification Reference | PV-ENRG-011 (Evidence-Pending) — ESD immunity test (IEC 61000-4-2); AOI/ICT/EOL coverage review
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) · Residual Severity: S2 · Residual Risk: AP Low (region B)
New Risks Introduced | None.
Overall Residual Contribution | Negligible directly; the information-harm escalation is carried by HA-ELEC-008.
Status | Controlled-pending-verification
Owner | Battery/Power Engineer (defers to Electrical/EM-ENG-003 for ESD immunity)
Rationale/Evidence | IEC 61000-4-2; IPC/ESD; RCR-014/016. Severity per HAL-2000 §5.
Assumptions | ESD-controlled environment for all service/assembly; screening coverage adequate on safety-critical nets.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ENRG-012 — Vibration / loose-connection resistive heating
Hazard ID (library) | HAZ-THERM (parasitic resistive heating)
Product/Config | SCM all SKUs · Lifecycle Phase: Transport (vibration) over reuse life · System State: Powered / charging · Operational Mode: Any
Perspective | Person near the device; property
Source | IEC 60068 (vibration); IEC 62368-1; ties HA-MECH-006; judgment
Initiating Cause/Event | Vibration/thermal cycling loosens a power-carrying connection (pack interconnect, connector, solder joint), raising contact resistance.
Contributing Conditions | Reusable multi-trip transport vibration; thermal cycling −30↔+60 °C; a high-current node (charge/pack) is the concern.
Foreseeable Sequence of Events | Vibration → connection loosens → contact resistance rises → localised I²R heating at the joint → local hot spot (precursor; could feed a cell thermal event in the worst case).
Hazard | Parasitic resistive (I²R) heating at a degraded high-current joint.
Hazardous Situation | A hot resistive joint inside a powered device.
Sequence to Harm | Local overheating → discoloration/odour → minor burn if touched, or degradation; worst case escalates to HA-ENRG-002.
Harm/Loss | Minor injury / degradation (S2).
Severity | S2 (Minor) — minor; a precursor rather than a direct serious-injury path (escalation owned by HA-ENRG-002).
Probability (P1×P2) | P1=2 (Remote — vibration-resistant joints per RCR-012) × P2=2 (Remote) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 60068; Medium uncertainty.
Detectability/Warning | Over-temp protection (RCR-002) trips on gross heating; no direct user warning of a marginal joint.
Initial Risk | AP Low→Medium (region B/A — S2 × PL → B Low)
Risk Acceptability | Broadly acceptable / ALARP
Risk-Control Strategy | 1 (inherent — vibration-resistant joints) + 2 (protective — OT trip)
Risk-Control Requirement | RCR-012 (fasteners/joints resist loosening under vibration/thermal cycling) + RCR-014 (as it relates to interconnect integrity) + RCR-002 (OT backstop).
Implementation Reference | DO-MECH-004 (joints) · DO-EE-004 (pack interconnect) · CMP-007 (pack), CMP-034 (PCB)
Verification Reference | PV-ENRG-012 (Evidence-Pending) — vibration + thermal-cycle inspection (joint integrity, contact-resistance drift)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) · Residual Severity: S2 · Residual Risk: AP Low (region B)
New Risks Introduced | None.
Overall Residual Contribution | Negligible; monitored as a precursor to HA-ENRG-002.
Status | Controlled-pending-verification
Owner | Battery/Power Engineer (with Mechanical)
Rationale/Evidence | IEC 60068 (vibration/thermal cycling); RCR-012/002. Severity per HAL-2000 §5.
Assumptions | High-current joints are the material concern; signal joints are non-safety here.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale, severity anchor §6/§7), HAL-2100 (P1×P2, basis taxonomy, Evidence-Pending posture §7), and HAL-2200 (matrix, regions R/A/B, AP, severity-first + uncertain-probability + single-fault rules).
Controls | RCR-001…006, RCR-012/013/014, RCR-045/046/049/054 (HAL-2400) are the energy-module controls; RCR-001/002 are on the release-critical single-fault-tolerant set.
Implementation | DO-SAFE-001/002/003/010/012/013 (DOL-1200) are the Implementation References; each points to concrete DO-EE/DO-FW/DO-MECH outputs and to CMP-004/007/008/009/010/011/012/013/014/037.
Cross-module ties | HA-ENRG-002/003 → HA-TFC-002 (transport fire) and HA-TFC-003 (offgas); HA-ENRG-001 → HA-TFC-001 (self-heat surface) + HA-USE-009 (charge-inhibit misread); HA-ENRG-005 → HA-ELEC-002; HA-ENRG-006/008 → HA-SW-013 (silent outage) + HA-USE-004 (gap misread) via the HA-BND-003 information pathway; HA-ENRG-009 emissions → HA-ELEC-004; HA-ENRG-011 → HA-ELEC-008; HA-ENRG-012 → HA-MECH-006.
Downstream | S4/S5 records (HA-ENRG-001/002/003/005/006/008) become FTA top events in Stage 11 (ERFKL) and drive deeper battery/charge DFMEA; all PV-/VAL- references are Evidence-Pending and feed Stages 13–14. The D3 items (HA-ENRG-001/006/008) are the carried headline unknowns per HAL-000 §7.
8. Residual-risk / closure notes
Two carried TOP hazards (D3) | HA-ENRG-001 (cold-charge plating) and HA-ENRG-006 (dead-in-cold silent loss of monitoring). Both are severity-driven while probability is Evidence-Pending; neither can be closed until its named test reports — cold-charge fault-injection + cold cycle-life (PV-ENRG-001/008), and the 90-day cold endurance & reserve test (PV-ENRG-006 / PRD-VER-004).
• S5 residuals (HA-ENRG-002/003) can never be classified broadly acceptable (HAL-2200 §6.1); after the dual-independent protection and UN 38.3 evidence they sit in region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off and post-market monitoring.
Release gate (HAL-2200 §8) | no R region energy risk may remain open at release; every remaining region-A S4/S5 energy risk carries a benefit-risk statement and residual-risk sign-off, and every Evidence-Pending probability has a named closing test listed above.
Overall contribution | the energy module's residual is dominated by the two D3 items; they are the principal battery/thermal inputs to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
New-risk watch (HAL-2700) | the energy↔safety trade (larger cell/reserve for endurance increases stored energy and thermal exposure) and the self-heat surface (HA-TFC-001) are tracked as control-introduced risks.
HAL-600 (HA-MECH) — Mechanical Hazard Analysis Register
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: Principal Mechanical Engineer with System-Safety (lead), EM-ENG-001 (Chief Systems), and EM-ENG-006 (Quality/Regulatory). This module is the mechanical domain register (HAZ-MECH). It is a harm-and-exposure-first hazard analysis per ISO 14971, not an FMEA: every record is authored as a complete risk chain (initiating cause → hazard → hazardous situation → sequence to harm → harm), severity fixed by outcome, probability decomposed P1×P2 with an honest basis, and risk scored severity-first (CLAUDE.md rule 3). Standards cited by designation; revisions to confirm. Baseline v1.0 (2026-07-18), pre-DR.
1. Purpose
Identify, characterise, and control the mechanical hazards of the SCM across its full reusable lifecycle — sharp edges, pinch/crush, small parts and choking, enclosure/seal breakage, probe-lead handling, fastener integrity, retention in the shipper, port/vent blockage, battery mechanical retention, and handling ergonomics. Because the SCM is an information device (HA-CTX-001), several mechanical hazards do not end in a direct physical injury but in loss/corruption of the monitoring record or an unrepresentative reading → a wrong disposition decision → beneficiary harm (HA-BND-003). Those records carry S4 severity even though the mechanical event itself is minor, and are traced to the same record-integrity controls as the software domain.
2. Scope
All SCM SKUs (reusable multi-trip device), the external probe + lead (CMP-020), the connector/caps (CMP-005), the mount/retention feature (CMP-003), the sealed enclosure and seals (CMP-001/002), fasteners/joints, the USB-C port body (CMP-004), and the battery pack mechanical retention (CMP-007). Lifecycle phases in scope: manufacture/assembly, provisioning, deployment/transport (air/ground), receipt/read, depot service/recharge, and end-of-life/disposal (where a child bystander may access a discarded device — HA-CTX-006). Out of scope: the passive shipper's thermal performance (context only) and the therapeutic product (object of the disposition decision only). Battery thermal-runaway chemistry is analysed in HAL-500 (HA-ENRG); this register covers only the mechanical initiating path to a cell event (HA-MECH-009, ties HA-ENRG-003).
3. Applicable standards & references
• ISO 14971 — risk management; risk chain, severity-by-outcome, P1×P2, acceptability (HAL-2000/2100/2200).
• IEC 62368-1 — mechanical hazards (sharp edges, corners, pinch/crush points), enclosure integrity, fire enclosure for a contained single-cell event.
• 16 CFR 1501 (small-parts concept) / ISO 8124-1 small-parts cylinder concept — choking-hazard sizing for parts accessible to a child bystander (concept cited; SCM is not a toy but a discarded device is foreseeably child-accessible — HA-CTX-003/006).
• IEC 60068-2-6 / -2-27 / -2-64 — vibration, shock, and random-vibration test methods (fastener loosening, retention, cell mechanical loading in transport).
• IEC 60529 — ingress protection (enclosure crack/seal breach → ingress path; detailed in HAL-900).
• UN 38.3 — lithium-battery transport abuse (mechanical: shock, vibration, crush) — ties HAL-500.
Upstream | HAL-000/100/200 (governance, context, boundary), HAL-2000/2100/2200/2400 (scales, RCRs), DOL-1200 (DO-SAFE), component-list.md (CMP), fault-management-diagnostics.md (F-STO, honest-gap invariant).
4. Governance & method
Method | ISO 14971 harm-and-exposure-first. Each record uses the 34-field structure (HAL-000 §9); every field is answered or justified N/A. Hazard prompts come from the mechanical energy-flow exposure model (HA-BND-002) and the lifecycle sweep (HAL-300); the taxonomy suggests, it never auto-declares.
• Severity is fixed by the credible worst-case outcome (HAL-2000); it is not reduced by controls unless the control changes the harm outcome. Wrong-disposition mechanical paths inherit the S4 anchor (HAL-2000 §6).
• Probability is P1 (cause → hazardous situation) × P2 (situation → harm), each 1–5, mapped to bands (HAL-2100 §6). Basis and uncertainty are recorded; unknowns are Evidence-Pending with a named closing test — never invented. Where a factor is unknown, the band is escalated one level (HAL-2200 §6.2).
• Risk is scored severity-first on the HAL-2200 matrix (region R/A/B → AP High/Med/Low). The S5 catastrophic rule and the single-fault rule (S4/S5 physical hazards) apply to HA-MECH-009.
• Control hierarchy (HAL-2300): inherent design → protective → information → process; warnings-only controls are challenged where design prevention is feasible (HAL-2200 §6.4).
• Residual risk is reassessed only after the control is defined (RCR), implemented (DO-…), and verified (PV-…) — never on a spreadsheet action (HAL-000 §8; HAL-2800).
5. Hazard categories & guiding questions (mechanical)
1. Sharp edges / burrs / corners — Can any accessible surface cut or abrade a handler in normal use or service? (HA-MECH-001)
2. Pinch / crush points — Do the battery door, seams, or serviceable joints trap a finger during open/close? (HA-MECH-002)
3. Small parts / strangulation — Can a probe tip, connector cap, or lanyard detach and be swallowed by a child bystander near a discarded device? (HA-MECH-003)
4. Enclosure / seal breakage — Does a drop crack the housing or defeat the seal, opening an ingress path to record-integrity electronics? (HA-MECH-004)
5. Probe-lead handling — Can the lead be entangled, strained, cut, or dislodged so the probe reads an unrepresentative temperature? (HA-MECH-005)
6. Fastener / joint integrity — Do fasteners loosen under vibration and thermal cycling over reuse life, detaching or rattling parts? (HA-MECH-006)
7. Retention in shipper — Does the mount hold the device and probe in the product zone through transport? (HA-MECH-007)
8. Port / vent blockage — Can packaging block a port or vent and trap moisture? (HA-MECH-008)
9. Battery mechanical retention — Is the cell retained without crush/puncture through drop, shock, and vibration? (HA-MECH-009)
10. Handling ergonomics — Can placement/service cause a strain or handling injury? (HA-MECH-010)
6. Hazard Analysis Register (HA-MECH records)
HA-MECH-001 — Accessible sharp edge / burr → laceration during handling
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Deployment / receipt / depot service · System State: Any (powered or off) · Operational Mode: Manual handling
Perspective | Clinic staff, logistics/3PL handler, depot technician (installer/bystander)
Source | HA-BND-002 (mechanical energy exposure); IEC 62368-1 mechanical-hazard prompt; lifecycle sweep HAL-300
Initiating Cause/Event | Un-deburred moulding parting line, trimmed rib, exposed PCB/heatsink edge, or a chipped enclosure edge from prior drop
Contributing Conditions | Gloved-vs-bare handling; repeated grasping over reuse life; cold-embrittled polymer at −30 °C reducing edge toughness
Foreseeable Sequence of Events | Sharp/burred edge left accessible after manufacture or created by field damage → user grips or wipes the device → skin contact under force
Hazard | Accessible sharp edge / burr (concentrated mechanical energy)
Hazardous Situation | A person's hand is in contact with an unguarded sharp edge during grip, insertion, or cleaning
Sequence to Harm | Contact under grip force → skin laceration/abrasion
Harm/Loss | Minor cut/abrasion, no treatment or first-aid only
Severity | S2 (Minor) — human-injury dimension; superficial, reversible, self-treating per HAL-2000
Probability (P1×P2) | P1=2 (Remote — edges controlled by moulding/finishing but a burr/chip is credible) × P2=2 (Remote — contact under enough force to break skin) ⇒ 4, PL (Low) — Basis: Expert judgment + IEC 62368-1 edge criteria; Low–Med uncertainty; confirmed by edge inspection
Detectability/Warning | High — visible/tactile at incoming inspection and to the user; no hidden hazard
Initial Risk | Low (B) — S2 × PL
Risk Acceptability | Broadly acceptable (B); ALARP still applied (edges are cheap to remove by design)
Risk-Control Strategy | 1 inherent (design out the edge)
Risk-Control Requirement | RCR-009 — No accessible hazardous sharp edge, pinch, or crush point in normal handling/service
Implementation Reference | DO-MECH-002/003 (enclosure edge/finish) · CMP-001
Verification Reference | PV-MECH-001 (Evidence-Pending) — edge/burr accessibility inspection per IEC 62368-1
Validation Reference | N/A (design-inspection hazard; no use validation required)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S2 (unchanged) · Residual Risk: Low (B) (Evidence-Pending PV-MECH-001)
New Risks Introduced | None
Overall Residual Contribution | Negligible to overall residual risk
Status | Open — Evidence-Pending (edge inspection)
Owner | Mechanical Engineer
Rationale/Evidence | IEC 62368-1 treats accessible sharp edges as a design-controllable mechanical hazard; radius/finish requirements remove it inherently. Judgment; low consequence.
Assumptions | Enclosure polymer and finish specified to a defined edge radius; no metal fins exposed in normal grip zones.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-002 — Pinch/crush at battery door or seam during service
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Depot service / recharge / battery service · System State: Off or charging · Operational Mode: Service (open/close)
Perspective | Depot technician, service handler
Source | HA-BND-002; IEC 62368-1 pinch/crush prompt; HA-CTX-003 (field/service misuse)
Initiating Cause/Event | A serviceable joint (battery door, split-line seam, hinged cover) closes on a finger during open/close
Contributing Conditions | Gloved handling reducing tactile feedback; spring-loaded or press-fit closure; cold-stiffened seals raising closing force
Foreseeable Sequence of Events | Technician opens the serviceable feature → finger placed in the closing gap → feature snaps/closes → finger pinched
Hazard | Pinch/crush point at a moving mechanical joint (stored/applied mechanical energy)
Hazardous Situation | A finger is inside the closing gap of the door/seam during actuation
Sequence to Harm | Closure force on soft tissue → pinch/bruise
Harm/Loss | Minor pinch injury/bruise, no treatment
Severity | S2 (Minor) — human-injury dimension; transient, reversible
Probability (P1×P2) | P1=2 (Remote — service is infrequent; gap geometry can be designed safe) × P2=2 (Remote — enough force to injure) ⇒ 4, PL (Low) — Basis: Judgment + IEC 62368-1 pinch criteria; Low uncertainty
Detectability/Warning | Medium — gap is visible; closing force is felt; no concealed mechanism
Initial Risk | Low (B) — S2 × PL
Risk Acceptability | Broadly acceptable (B); ALARP applied via gap/force design
Risk-Control Strategy | 1 inherent (safe gap geometry / limited closing force)
Risk-Control Requirement | RCR-009 — no accessible pinch/crush point in normal handling/service
Implementation Reference | DO-MECH-002/003 (serviceable-feature geometry) · CMP-001
Verification Reference | PV-MECH-002 (Evidence-Pending) — pinch-point inspection / closing-force measurement (IEC 62368-1)
Validation Reference | N/A
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S2 · Residual Risk: Low (B) (Evidence-Pending PV-MECH-002)
New Risks Introduced | None
Overall Residual Contribution | Negligible
Status | Open — Evidence-Pending
Owner | Mechanical Engineer
Rationale/Evidence | IEC 62368-1 pinch-point clearances design the hazard out; low severity/probability. Judgment.
Assumptions | Battery service is depot-only (not field, per HA-CTX-004); closure geometry provides finger clearance or low force.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-003 — Detachable small part (probe tip / connector cap / lanyard) → choking or strangulation of a child bystander
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs (probe CMP-020, connector cap CMP-005, any lanyard) · Lifecycle Phase: End-of-life / disposal (discarded device) · System State: Off (discarded) · Operational Mode: N/A (uncontrolled disposal)
Perspective | Child bystander (HA-CTX-006) — the vulnerable non-user
Source | HA-CTX-003/006 (foreseeable misuse — discarding; child access); 16 CFR 1501 small-parts concept
Initiating Cause/Event | A device is discarded rather than returned to the reuse loop; a small component (probe tip, connector cap, lanyard) detaches or is already loose
Contributing Conditions | Small-parts-sized components; a lanyard long enough to loop a neck; a curious child able to reach discarded waste; no return discipline at last mile
Foreseeable Sequence of Events | Device discarded → child finds it → detaches/mouths a small part or loops a lanyard → airway obstruction or ligature
Hazard | Detachable small part / flexible lanyard (mechanical airway/ligature hazard)
Hazardous Situation | A small part is in a child's mouth, or a lanyard is around a child's neck
Sequence to Harm | Airway obstruction (choking) or neck constriction (strangulation) → hypoxia
Harm/Loss | Serious injury requiring medical treatment; credible worst case is asphyxial injury/death of a child
Severity | S4 (Critical) — human-injury dimension; child-choking anchor (HAL-2000 §7). Worst credible outcome (asphyxia) is at the S4/S5 boundary; assigned S4 per the SCM severity register, treated with High AP.
Probability (P1×P2) | P1=2 (Remote — a device reaching a child depends on failed return discipline + disposal + child access) × P2=3 (Occasional — given access, mouthing/looping is a known child behaviour) ⇒ 6, PM (Medium) — Basis: Historical/analogous child small-parts incident literature + judgment; Med uncertainty (return-rate Evidence-Pending)
Detectability/Warning | Low from the child's side (no supervision assumed); design/label are the only barriers
Initial Risk | High (R) — S4 × PM (severity-first; child-safety)
Risk Acceptability | Unacceptable (R) as-is — design change required; cannot be accepted on labelling alone (warnings-challenge rule, HAL-2200 §6.4)
Risk-Control Strategy | 1 inherent (retain/size parts) + 3 information (disposal warning) — inherent leads
Risk-Control Requirement | RCR-010 — small parts (probe tip, connector caps, any lanyard) shall be retained or sized/warned to prevent choking/strangulation of a child bystander
Implementation Reference | DO-MECH-006 (captive/retained parts, no free lanyard or break-away lanyard) + DO-LBL (disposal/keep-from-children warning) · CMP-005 / CMP-020
Verification Reference | PV-MECH-003 (Evidence-Pending) — small-parts cylinder test (16 CFR 1501-style) + part-retention pull test + lanyard break-away/length check
Validation Reference | VAL-MECH-003 (Evidence-Pending) — disposal-instruction comprehension (ties HA-SVC end-of-life)
Residual Probability | P1=1 (retained/oversized parts, break-away lanyard) × P2=2 ⇒ PVL · Residual Severity: S4 (outcome unchanged if a part still detaches) · Residual Risk: A (Med) — S4 × PVL, ALARP with benefit-risk and disposal instruction (Evidence-Pending PV-MECH-003)
New Risks Introduced | Captive probe tip could complicate legitimate probe replacement — managed by depot-only service (RCR-050)
Overall Residual Contribution | Moderate — a headline child-safety record; remains on the residual-risk summary until retention verified
Status | Open — Evidence-Pending (small-parts + retention tests) · release-critical child-safety record
Owner | Mechanical Engineer
Rationale/Evidence | 16 CFR 1501 small-parts concept and child-choking epidemiology ground the severity and the inherent-first control; a discarded reusable device is foreseeably child-accessible (HA-CTX-003). Warnings alone are not accepted where retention is feasible.
Assumptions | No safety-required free lanyard; if a lanyard is provided it is break-away and length-limited; probe tip and caps are captive or larger than the small-parts cylinder.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-004 — Enclosure crack/breakage on drop → ingress → record-electronics fault → data loss → wrong disposition
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Deployment / transport / handling · System State: Armed (monitoring) · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path, HA-BND-003)
Source | HA-BND-002/003; drop in the lifecycle sweep; F-STO fault class (fault-management §2)
Initiating Cause/Event | Device dropped or crushed in handling/transport → enclosure cracks or a seal joint opens
Contributing Conditions | Cold-embrittled polymer at −30 °C; repeated drops over reuse life; ingress path then admits condensation/liquid to the PCB (ties HA-ENV-004)
Foreseeable Sequence of Events | Drop → enclosure crack/seal breach → moisture/contaminant ingress → record-integrity electronics fault (corrosion/short/reset) → samples lost or record corrupted → gap in the trip record
Hazard | Loss of enclosure/seal integrity (ingress path to record electronics)
Hazardous Situation | An armed device is silently logging with a compromised record path during a monitored trip
Sequence to Harm | Record gap/corruption → if the honest-gap invariant fails, incomplete record reads as a clean pass → limited quantity of product wrongly dispositioned
Harm/Loss | Wrong disposition of a limited quantity of doses / a monitored trip released on an incomplete record
Severity | S3 (Serious) — public-health dimension, limited quantity (HAL-2000 §5); bounded because a single device covers one shipment
Probability (P1×P2) | P1=3 (Occasional — drops are routine in distribution) × P2=3 (Occasional — crack→ingress→electronics fault→undetected gap; the honest-gap controls in RCR-022/024 are the mitigating downstream check) ⇒ 9, PM (Medium) — Basis: Judgment + reliability analogous (transport drop rates); Med uncertainty; closed by drop + IP test
Detectability/Warning | Medium — a visible crack is detectable at receipt; a hairline seal breach is not, but the record gap is annunciated if RCR-022 holds
Initial Risk | Medium (A) — S3 × PM
Risk Acceptability | Undesirable/ALARP (A) — reduce as far as practicable via robust enclosure + honest-gap record
Risk-Control Strategy | 1 inherent (rugged enclosure/seal) + 2 protective (honest-gap record so a fault can't read as a pass)
Risk-Control Requirement | RCR-011 — enclosure + seals maintain ingress rating and cell containment through drop/vibration/transport and reuse life (supported by RCR-024/022 record integrity for the harm-end backstop)
Implementation Reference | DO-SAFE-011 (ingress/condensation) → DO-MECH-005/012 (enclosure/seal) · CMP-001/002; harm-end backstop DO-SAFE-008 (record integrity)
Verification Reference | PV-MECH-004 (Evidence-Pending) — drop test (IEC 60068-2-31/-2-27) + post-drop IP re-test (IEC 60529)
Validation Reference | N/A (verification-closed; harm-end validated under HA-DATA/HA-SW)
Residual Probability | P1=2 × P2=2 ⇒ PL · Residual Severity: S3 · Residual Risk: A (Med) trending B — after rugged-enclosure verification + honest-gap record (Evidence-Pending PV-MECH-004)
New Risks Introduced | A more rugged (thicker) enclosure can raise RF attenuation in-container — flag to Electrical/RF (RCR-031, HA-COMM-003) for antenna trade
Overall Residual Contribution | Low–moderate; contributes to the record-integrity residual set
Status | Open — Evidence-Pending (drop + IP)
Owner | Mechanical Engineer
Rationale/Evidence | Drop-induced ingress is a classic reuse-life failure; severity is bounded to limited quantity because the honest-gap invariant (fault-management §6) prevents a silent clean pass. Judgment + standard test methods.
Assumptions | RCR-022/024 (honest gap, power-fail-safe record) are implemented so a corrupted record cannot present as an unqualified accept (cross-domain dependency on HA-SW/HA-DATA).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-005 — Probe lead entanglement/strain → probe damaged/dislodged → unrepresentative reading → wrong disposition
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs (external probe + lead CMP-020) · Lifecycle Phase: Provisioning / transport / receipt · System State: Armed · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-003; HA-CTX-009 (probe safety-relevant); F-SNS fault class
Initiating Cause/Event | Lead snagged, pulled, crushed by shipper contents, or repeatedly flexed → probe conductor damage or the probe tip pulled out of the product zone
Contributing Conditions | Cold-stiffened lead insulation (brittle at −30 °C); strain at the connector or tip; no strain relief; reuse-life flex fatigue
Foreseeable Sequence of Events | Lead strained/entangled → probe damaged (intermittent/offset) or dislodged from product → probe now measures a non-product-zone temperature or a drifted value → record contains an unrepresentative product temperature
Hazard | Damaged/dislodged product-zone probe (unrepresentative measurement)
Hazardous Situation | An armed device logs an unrepresentative product-zone temperature believed to be valid
Sequence to Harm | Wrong/unrepresentative reading → if not flagged by RCR-020, verdict computed on bad data → wrong accept/reject of the shipment
Harm/Loss | Wrong disposition of a shipment (compromised dose administered, or potent dose destroyed)
Severity | S4 (Critical) — public-health dimension via wrong disposition (HAL-2000 §6 anchor)
Probability (P1×P2) | P1=3 (Occasional — lead strain/entanglement is a common field event) × P2=3 (Occasional — undetected wrong reading reaches a wrong disposition; RCR-020 probe-fault detection is the governing downstream check) ⇒ 9, PM (Medium) — Basis: Judgment + F-SNS; Med uncertainty; closed by probe-fault-injection + lead flex/pull test
Detectability/Warning | Medium — a fully open/short probe is detectable (RCR-020); a partially damaged probe giving a plausible-but-wrong reading is the harder case, mitigated by plausibility/inter-channel checks
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via strain relief + probe-fault detection
Risk-Control Strategy | 1 inherent (strain relief/robust lead) + 2 protective (probe-fault/plausibility detection so a bad reading is never treated as valid)
Risk-Control Requirement | RCR-011 (mechanical robustness of lead/retention) and RCR-020 (absent/open/short/wrong/implausible probe reads detected and never read as a valid product-zone temperature)
Implementation Reference | DO-MECH (strain relief, lead spec) + DO-SAFE-005 → DO-FW-007/DO-EE-013 (probe-fault detection) · CMP-020/CMP-005
Verification Reference | PV-MECH-005 (Evidence-Pending) — lead flex/pull/cold-brittleness test; and probe-fault-injection test (per RCR-020)
Validation Reference | VAL-MECH-005 (Evidence-Pending) — receipt-read usability confirming a probe-fault flag blocks a clean accept (ties HA-USE-005)
Residual Probability | P1=2 × P2=2 ⇒ PL · Residual Severity: S4 · Residual Risk: A (Med) — after strain relief + probe-fault detection verified (Evidence-Pending)
New Risks Introduced | Over-stiff strain relief could reduce probe placement flexibility — manage in DO-MECH usability
Overall Residual Contribution | Moderate — part of the product-zone-accuracy residual set; hand to Stage 11 as an FTA contributor to "unrepresentative reading"
Status | Open — Evidence-Pending (mechanical + probe-fault tests)
Owner | Mechanical Engineer (with Firmware/Sensing for RCR-020)
Rationale/Evidence | Product-zone accuracy is safety-critical (HA-CTX-009); a damaged lead is a foreseeable mechanical initiator of the information-harm path. The partial-damage/plausible-wrong-reading case is why RCR-020 plausibility is required, not just open/short detection.
Assumptions | Probe is independently identifiable and continuity-monitored (CMP-020); inter-channel plausibility (internal vs probe) is available (fault-management §3).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-006 — Fastener loosening under vibration/thermal cycling → part detachment/rattle
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Transport / reuse life · System State: Any · Operational Mode: Autonomous
Perspective | Depot technician / handler; secondary information-integrity effect
Source | HA-BND-002; IEC 60068 vibration; reuse-life sweep
Initiating Cause/Event | Threaded fasteners/snap joints loosen under transport random vibration and −30↔+60 °C thermal cycling (differential CTE)
Contributing Conditions | No thread-lock/retention; dissimilar-material CTE mismatch; many reuse cycles; cold-relaxed press fits
Foreseeable Sequence of Events | Repeated vibration + thermal cycling → fastener backs out → internal part loosens → rattle, intermittent internal contact, or a detached fragment inside the enclosure
Hazard | Loosened fastener / detached internal part (loose mechanical element)
Hazardous Situation | A loose part or backed-out fastener inside the sealed device during operation
Sequence to Harm | Loose part → at worst an intermittent electrical contact/short or a rattling annoyance; a detached conductive fragment could bridge a net (ties HA-ELEC-008)
Harm/Loss | Minor — nuisance rattle; at worst a repairable intermittent fault caught at depot
Severity | S2 (Minor) — property/nuisance; escalates only if it causes a record-electronics intermittent (then handled under HA-ELEC-008)
Probability (P1×P2) | P1=2 (Remote — with proper retention) × P2=2 (Remote — loosening→harm) ⇒ 4, PL (Low) — Basis: Judgment + IEC 60068 methods; Low–Med uncertainty; closed by vibration + thermal-cycle test
Detectability/Warning | Medium — rattle is audible; internal loosening detectable at depot BIST/inspection
Initial Risk | Low (B) — S2 × PL
Risk Acceptability | Broadly acceptable (B); ALARP via fastener retention
Risk-Control Strategy | 1 inherent (fastener retention / joint design)
Risk-Control Requirement | RCR-012 — fasteners/joints resist loosening under vibration/thermal cycling over reuse life (no detachment)
Implementation Reference | DO-MECH-004 (fastening/thread-lock/retention) · CMP-001/034
Verification Reference | PV-MECH-006 (Evidence-Pending) — random vibration (IEC 60068-2-64) + thermal-cycle (IEC 60068-2-14) followed by torque-retention/rattle inspection
Validation Reference | N/A
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S2 · Residual Risk: Low (B) (Evidence-Pending PV-MECH-006)
New Risks Introduced | None
Overall Residual Contribution | Negligible directly; feeds HA-ELEC-008 as a possible initiator
Status | Open — Evidence-Pending
Owner | Mechanical Engineer
Rationale/Evidence | Vibration/thermal loosening is a standard transport failure controlled by thread-lock and joint design; low severity. Judgment + standard test methods.
Assumptions | Fasteners specified with retention appropriate to CTE mismatch; no safety-critical net relies on an un-retained joint.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-007 — Retention feature failure → device shifts in shipper → probe leaves product zone → unrepresentative temp → wrong disposition
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs (mount/retention CMP-003) · Lifecycle Phase: Provisioning / transport · System State: Armed · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-003; HA-CTX-003 (probe outside product zone is named foreseeable misuse); F-SNS
Initiating Cause/Event | Mount/retention feature fails to hold the device/probe → device and probe migrate during transport
Contributing Conditions | Vibration/shock in transit; cold-relaxed retention; loading/unloading disturbance; misplacement at provisioning
Foreseeable Sequence of Events | Retention fails → device/probe shifts → probe tip moves out of the product zone (into air, against the shipper wall) → measured temperature no longer represents the product → record shows an unrepresentative product temperature
Hazard | Loss of device/probe position (unrepresentative measurement location)
Hazardous Situation | An armed device logs a temperature from outside the product zone, believed representative
Sequence to Harm | Unrepresentative temperature → verdict on non-product data → wrong accept (missed real excursion) or wrong reject
Harm/Loss | Wrong disposition of a shipment (compromised dose administered, or potent dose destroyed)
Severity | S4 (Critical) — public-health dimension via wrong disposition
Probability (P1×P2) | P1=3 (Occasional — retention slip / misplacement is credible and named foreseeable misuse) × P2=3 (Occasional — out-of-zone reading reaching a wrong disposition; pre-arm probe-present + plausibility are the mitigating checks) ⇒ 9, PM (Medium) — Basis: Judgment + HA-CTX-003 + F-SNS; Med uncertainty; closed by vibration retention test + inter-channel plausibility
Detectability/Warning | Medium — pre-arm gate confirms probe present (RCR-045) but not correct placement; inter-channel divergence (internal vs probe) gives partial detection
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via robust retention + placement guidance + plausibility
Risk-Control Strategy | 1 inherent (robust retention) + 2 protective (pre-arm + plausibility) + 3 information (placement guidance)
Risk-Control Requirement | RCR-011 (retention robustness) and RCR-045 (pre-arm readiness gate — probe-present) with plausibility from RCR-020
Implementation Reference | DO-MECH-012 / DO-MECH (retention feature) + DO-SAFE-007/DO-FW-032 (pre-arm gate) · CMP-003/CMP-020
Verification Reference | PV-MECH-007 (Evidence-Pending) — retention vibration/shock test (IEC 60068-2-64/-2-27); pre-arm probe-present test (per RCR-045)
Validation Reference | VAL-MECH-007 (Evidence-Pending) — provisioning/placement usability (probe-in-product-zone) — ties HA-USE-003/005
Residual Probability | P1=2 × P2=2 ⇒ PL · Residual Severity: S4 · Residual Risk: A (Med) — after retention + pre-arm + placement guidance verified (Evidence-Pending)
New Risks Introduced | A more aggressive retention feature could complicate glove-serviceable placement — manage in DO-MECH/DO-HF
Overall Residual Contribution | Moderate — a product-zone-representativeness contributor; FTA input to "unrepresentative reading"
Status | Open — Evidence-Pending
Owner | Mechanical Engineer (with Human-Factors for placement guidance)
Rationale/Evidence | "Probe outside the product zone" is explicitly named foreseeable misuse (HA-CTX-003); retention failure is its mechanical initiator. Placement correctness is only partially machine-detectable, so retention + guidance + plausibility are all required.
Assumptions | Pre-arm gate detects probe presence (RCR-045); inter-channel plausibility available; the product zone is defined in the provisioning workflow.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-008 — Blocked port/vent by packaging → moisture trap
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Transport / storage · System State: Any · Operational Mode: Autonomous
Perspective | Property/device; minor information-integrity effect
Source | HA-BND-002; HA-ENV condensation path (ties HA-ENV-003)
Initiating Cause/Event | Packaging, tape, or shipper wadding covers a pressure-equalisation vent or drainage feature
Contributing Conditions | Cold-to-warm condensation cycling; a sealed pocket holding moisture against the enclosure/port
Foreseeable Sequence of Events | Vent/port blocked → local moisture cannot equalise/drain → moisture trapped against the enclosure or USB-C port → localised dampness (feeds the ingress/condensation path, HA-ENV-003)
Hazard | Blocked vent/port creating a moisture trap
Hazardous Situation | Trapped moisture at a port/vent during thermal cycling
Sequence to Harm | Moisture trap → contributes to condensation ingress (handled in HA-ENV-003); on its own, cosmetic/nuisance
Harm/Loss | Negligible on its own — nuisance dampness/cosmetic; escalates only as a contributor to HA-ENV-003
Severity | S1 (Negligible) — no disposition impact by itself; contributory only
Probability (P1×P2) | P1=2 (Remote — vent placement can avoid blockage-prone zones) × P2=2 (Remote — trap→any harm) ⇒ 4, PL (Low) — Basis: Judgment; Low uncertainty
Detectability/Warning | Low visually; not annunciated; managed by design placement
Initial Risk | Low (B) — S1 × PL
Risk Acceptability | Broadly acceptable (B)
Risk-Control Strategy | 1 inherent (vent/port placement + sealing so no critical vent is blockage-dependent)
Risk-Control Requirement | RCR-011 — enclosure/seals maintain ingress rating (vent/port design part of ingress control)
Implementation Reference | DO-MECH-005 (seal/vent/port placement) · CMP-001/002/004
Verification Reference | PV-MECH-008 (Evidence-Pending) — port/vent placement review + condensation-cycle contribution check (ties PV-ENV-003)
Validation Reference | N/A
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S1 · Residual Risk: Low (B) (Evidence-Pending)
New Risks Introduced | None
Overall Residual Contribution | Negligible; contributes to HA-ENV-003
Status | Open — Evidence-Pending
Owner | Mechanical Engineer
Rationale/Evidence | A blocked vent is only a hazard as a moisture-trap contributor; the enclosure is sealed (IP-rated), so the direct effect is negligible. Judgment.
Assumptions | No safety function depends on an unobstructed external vent; ingress rating is maintained sealed (CMP-001).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-009 — Battery retention failure under shock → cell crush/puncture → thermal event
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs (battery pack CMP-007, retention) · Lifecycle Phase: Transport / drop / reuse life · System State: Any (worst: in occupied aircraft cargo) · Operational Mode: Autonomous / transit
Perspective | People nearby (handler, aircraft crew/passengers), property/mission (HA-CTX-006); ties HA-ENRG-003
Source | HA-BND-002 (energy exposure); UN 38.3 mechanical abuse; DO-SAFE-010; F-CHG/F-NRG upstream cell hazard
Initiating Cause/Event | Battery mechanical retention fails under drop/shock/vibration → cell moves, is crushed against structure, or is punctured by an internal feature
Contributing Conditions | High-energy Li-ion cell; cold-embrittled retention at −30 °C; repeated transport shock over reuse life; loss of clearance to sharp internal features
Foreseeable Sequence of Events | Retention fails → cell crush/puncture → internal short → self-heating → thermal runaway → fire/venting inside the shipper, near people or in an aircraft cargo hold
Hazard | Mechanically-induced internal short of a lithium cell (stored chemical/electrical energy)
Hazardous Situation | A crushed/punctured cell heating toward runaway inside an enclosure near people
Sequence to Harm | Thermal runaway → fire/burn/smoke → serious burn injury; worst credible case an enclosure fire in an occupied aircraft cargo hold → fatal outcome
Harm/Loss | Serious burn/smoke injury (S4); worst credible catastrophic fire (S5)
Severity | S5 (Catastrophic) worst case (occupied-aircraft fire, HAL-2000 §7) / S4 typical — governed by the human-injury/fire dimension. Catastrophic rule applies (HAL-2200 §6.1); single-fault rule applies (HAL-2200 §6.5).
Probability (P1×P2) | P1=2 (Remote — crush/puncture under shock with proper retention + clearance) × P2=3 (Occasional — crush → thermal event, enclosed near people) ⇒ 6, PM (Medium) — Basis: Analogous Li-ion mechanical-abuse event rates + judgment; Med uncertainty; closed by drop/vibe + UN 38.3 crush
Detectability/Warning | Low pre-event (internal); BMS/thermal cues give little warning once a mechanical short initiates — this is why prevention (retention) dominates
Initial Risk | High (R) — S5 × PM (and per the catastrophic rule an S5 hazard is High AP for design purposes regardless of probability)
Risk Acceptability | Unacceptable (R) as-is — design change / single-fault-tolerant retention required; can never be closed "broadly acceptable" (S5)
Risk-Control Strategy | 1 inherent (crush/puncture-proof retention + clearance) + 2 protective (fire enclosure containment, BMS/fuse backstop from HA-ENRG)
Risk-Control Requirement | RCR-004 — cell retained without crush/puncture through drop/vibration/transport; thermal/electrical clearance maintained (single-fault-tolerant with the containment/protection of RCR-002/RCR-008)
Implementation Reference | DO-SAFE-010 → DO-MECH-012 (retention) + DO-EE-004 (pack) · CMP-007; containment DO-MECH-005 (fire enclosure, RCR-008)
Verification Reference | PV-MECH-009 (Evidence-Pending) — drop (IEC 60068-2-31) + random vibration (IEC 60068-2-64) + UN 38.3 shock/vibration/crush; post-test cell integrity/clearance inspection
Validation Reference | N/A (abuse-verification-closed; overall thermal-event validation under HA-ENRG)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S5 (outcome unchanged if runaway still initiates) · Residual Risk: A (High\) — S5 × PVL per the catastrophic rule; requires benefit-risk + EM-ENG-001/EM-ENG-006 sign-off + post-market monitoring (Evidence-Pending PV-MECH-009 + UN 38.3)
New Risks Introduced | Rigid crush-proof retention adds mass/thermal mass — flag to Environmental/Thermal (cold endurance) and Electrical (pack integration)
Overall Residual Contribution | High — a release-critical S5 physical-safety record; carried on the overall residual-risk summary and handed to Stage 11 as an FTA top event (jointly with HA-ENRG-003)
Status | Open — Evidence-Pending (drop/vibe + UN 38.3) · release-critical, single-fault-tolerant
Owner | Mechanical Engineer (with Battery/Power for HA-ENRG-003 cell/thermal)
Rationale/Evidence | UN 38.3 and Li-ion mechanical-abuse literature ground both the S5 fire outcome and the retention control. The mechanical initiator (retention/clearance) is analysed here; the cell-level thermal-runaway propagation and BMS/fuse backstop are analysed in HA-ENRG-003 — the two records share the RCR-004 retention control and must be reconciled.
Assumptions | Fire-enclosure containment (RCR-008) and autonomous BMS/fuse protection (RCR-002) are present as the independent protective backstop (single-fault tolerance); cell chemistry is Evidence-Pending (D3) which does not change the S5 severity.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-MECH-010 — Awkward/heavy handling → ergonomic/handling injury during placement or service
Hazard ID (library) | HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Provisioning / placement / depot service · System State: Any · Operational Mode: Manual handling
Perspective | Clinic staff, logistics handler, depot technician
Source | HA-BND-002; lifecycle handling sweep; ergonomic prompt
Initiating Cause/Event | Awkward grip, unexpected mass, or an awkward reach when placing/removing the device deep inside a shipper
Contributing Conditions | Repetitive placement across many shipments; cold/gloved handling; reaching into a deep insulated container
Foreseeable Sequence of Events | Repeated/awkward handling → minor musculoskeletal strain or a dropped-device knock
Hazard | Ergonomic handling load (manual-handling strain)
Hazardous Situation | A handler reaching/gripping the device in an awkward posture
Sequence to Harm | Awkward load → transient strain/discomfort
Harm/Loss | Negligible — transient discomfort, no treatment
Severity | S1 (Negligible) — human-injury dimension; the device is small/light (a data logger), so ergonomic load is inherently low
Probability (P1×P2) | P1=2 (Remote) × P2=2 (Remote) ⇒ 4, PL (Low) — Basis: Judgment (small, hand-held mass); Low uncertainty
Detectability/Warning | High — handler feels awkwardness directly
Initial Risk | Low (B) — S1 × PL
Risk Acceptability | Broadly acceptable (B); ALARP via grip/size design
Risk-Control Strategy | 1 inherent (size/grip/weight — glove-friendly form factor)
Risk-Control Requirement | RCR-009 — no hazardous mechanical handling condition in normal handling/service (extended to ergonomic grip)
Implementation Reference | DO-MECH-002/003 (form factor, grip, glove-serviceable retention DO-MECH-012) · CMP-001/003
Verification Reference | PV-MECH-010 (Evidence-Pending) — handling/grip ergonomic review (gloved, cold)
Validation Reference | VAL-MECH-010 (Evidence-Pending) — placement/service usability (ties HA-USE)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S1 · Residual Risk: Low (B) (Evidence-Pending)
New Risks Introduced | None
Overall Residual Contribution | Negligible
Status | Open — Evidence-Pending
Owner | Mechanical Engineer (with Human-Factors)
Rationale/Evidence | As a small battery-powered logger the ergonomic load is inherently minor; grip/size design and glove-serviceability keep it negligible. Judgment.
Assumptions | Device mass and form factor are hand-held; no two-handed heavy lift; retention is glove-serviceable (CMP-003).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
7. Traceability notes
Upstream | severity per HAL-2000; probability model per HAL-2100; acceptability/matrix per HAL-2200; RCRs per HAL-2400 (RCR-004/009/010/011/012/020/045). Context/boundary per HAL-100/200 (HA-BND-002 energy-flow; HA-BND-003 information pathway; HA-CTX-003/006 misuse and child bystander).
Implementation | DO-SAFE-010 (battery retention), DO-SAFE-011 (ingress), DO-SAFE-005 (probe fault), DO-SAFE-007 (safe state/pre-arm), DO-MECH-002/003/004/005/006/012, DO-LBL (disposal warning) per DOL-1200.
Components | CMP-001 (enclosure), CMP-002 (seals), CMP-003 (retention), CMP-004 (USB-C port), CMP-005 (probe connector/cap), CMP-007 (battery pack), CMP-020 (external probe), CMP-034 (PCB).
Cross-domain cross-links (flag to named engineers)
• HA-MECH-009 ↔ HA-ENRG-003 — mechanical retention initiator ↔ cell thermal-runaway; shared RCR-004; reconcile with Battery/Power (HA-ENRG); joint Stage-11 FTA top event.
• HA-MECH-004 ↔ HA-ENV-003/004 & HA-ELEC-006/008 & HA-SW-013/HA-DATA — enclosure crack → ingress → record-electronics fault → data loss; shared RCR-011/024/022; reconcile with Environmental, Electrical, Firmware/Data.
• HA-MECH-005/007 ↔ HA-SW-008 / HA-USE-005 / HA-DATA-001 — unrepresentative reading → wrong disposition; shared RCR-020/045; reconcile with Firmware/Sensing and Human-Factors.
• HA-MECH-004 → HA-COMM-003 — a thicker/more rugged enclosure raises in-container RF attenuation; flag the antenna trade to Electrical/RF (RCR-031).
• HA-MECH-006 → HA-ELEC-008 — a detached conductive fragment as an initiator of a record-electronics intermittent.
Downstream | S4/S5 records (HA-MECH-003/005/007/009) become Stage 11 (ERFKL) FTA top-event / DFMEA inputs; RCRs feed Stage 12; PV-MECH-### become Stage 13 verification.
8. Residual-risk / closure notes
No open R at release | HA-MECH-003 (child choking), HA-MECH-005, HA-MECH-007 (all initial R) reduce to A (Med) only after their RCRs are implemented and verified; HA-MECH-009 reduces to A (High\) and, being S5, can never be "broadly acceptable" — it requires benefit-risk justification and EM-ENG-001 + EM-ENG-006 sign-off plus post-market monitoring (HAL-2200 §7).
Warnings-challenge | HA-MECH-003 is not accepted on the disposal label alone — inherent retention/ sizing (RCR-010) leads; the label is defence-in-depth.
Single-fault tolerance | HA-MECH-009 requires retention (RCR-004) and an independent containment/ protection backstop (RCR-008/002) with no common cause.
Evidence-Pending closures | PV-MECH-001…010 (edge/pinch inspection, small-parts + retention, drop + IP, lead flex/pull + probe-fault-injection, vibration + thermal-cycle, retention vibration, UN 38.3 crush, handling ergonomics). Cell chemistry (D3) is Evidence-Pending but does not change the S5 severity of HA-MECH-009.
Overall contribution | the mechanical domain contributes one release-critical S5 physical record (HA-MECH-009), one release-critical child-safety record (HA-MECH-003), and three product-zone/record- integrity contributors (HA-MECH-004/005/007) to the overall residual-risk summary (HAL-2900). The remaining records are broadly acceptable (B) subject to inspection verification.
Status | all records Open — Evidence-Pending; module closes when every S4/S5 record has a verified RCR/DO/PV, no R remains open, and HAL-2900 gives a benefit-risk-based release recommendation.
HAL-700 (HA-ELEC) — Electrical Hazard Analysis Register
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: Principal Electrical/Hardware Engineer with System-Safety (lead), EM-ENG-001 (Chief Systems), and EM-ENG-006 (Quality/Regulatory). This module is the electrical domain register (HAZ-ELEC). It is a harm-and-exposure-first hazard analysis per ISO 14971, not an FMEA: every record is a complete risk chain, severity fixed by outcome, probability decomposed P1×P2 with an honest basis, and risk scored severity-first (CLAUDE.md rule 3). Standards cited by designation; revisions to confirm. Baseline v1.0 (2026-07-18), pre-DR.
1. Purpose
Identify, characterise, and control the electrical / electronic / EMC hazards of the SCM across its reusable lifecycle — accessible-voltage shock, USB-C charge-input abuse, ESD, EMC susceptibility and emissions, connector leakage/corrosion, insulation/creepage under condensation, latent PCB/solder defects, connector/probe mis-mating, and cold-voltage-sag brown-out during a record write. The SCM is an information device (HA-CTX-001); accordingly most electrical hazards here do not end in electric shock (the device is SELV) but in corruption/loss of the monitoring record or a false reading/verdict → a wrong disposition decision → beneficiary harm (HA-BND-003). Those records carry S4 severity and are traced to the record- integrity and EMC-immunity controls. One record ties directly to a physical thermal event (HA-ELEC-002 ↔ HA-ENRG-005).
2. Scope
All SCM SKUs. Components in scope: USB-C sink/port and input-protection front-end (CMP-004/010/014), rails/ PMIC (CMP-010), MCU + brown-out supervisor (CMP-015), sensors and signal chains (CMP-019/020/021/022), probe connector (CMP-005), radios (CMP-024/025/026), NV storage (CMP-033), PCB + conformal coat (CMP-034), and the firmware components that arbitrate charge, record commit, and safe state (CMP-037/040/041, brown-out barrier). Lifecycle phases: manufacture/EOL screening, provisioning, deployment/transport, receipt, and depot recharge. Out of scope: battery cell chemistry/thermal-runaway propagation (HAL-500 HA-ENRG) except where an electrical fault is the initiator (HA-ELEC-002); software logic defects (HAL-1000 HA-SW) except where an electrical event is the initiator (HA-ELEC-003/008/010). SELV means no mains-voltage hazard exists in the device boundary (HA-BND-001).
3. Applicable standards & references
• ISO 14971 — risk management (scales/matrix per HAL-2000/2100/2200).
• IEC 62368-1 — electrical safety, SELV classification, energy sources (ES1), insulation/creepage/ clearance, capacitor discharge.
• IEC 61000 series — EMC: -4-2 (ESD), -4-3 (radiated RF immunity), -4-4 (EFT/burst), -4-5 (surge), -4-6 (conducted RF), and emissions (CISPR 32-class) for interference with nearby equipment.
• USB Type-C / USB-PD — charge-input contract, PDO negotiation, reverse-insertion/OVP behaviour.
• DO-160 (RTCA, by designation) — airborne-equipment EMC/emissions context for aircraft-hold operation.
• IEC 60529 — ingress (condensation → creepage/leakage path; detailed in HAL-900).
Upstream | HAL-000/100/200; HAL-2000/2100/2200/2400 (RCR-003/013/014/015/016/020/024); DOL-1200 (DO-SAFE-003/008/011); component-list.md; fault-management-diagnostics.md (F-CHG, F-BRN, F-STO, F-TMP).
4. Governance & method
Method | ISO 14971 harm-and-exposure-first, 34-field record (HAL-000 §9); every field answered or N/A. Prompts from the energy-flow (HA-BND-002), information (HA-BND-003), and trust-boundary (HA-BND-005) exposure models.
• Severity fixed by outcome (HAL-2000); electrical faults that reach a wrong disposition inherit the S4 anchor; the aircraft-emissions record considers the nearby-equipment perspective (HA-CTX-006).
• Probability P1×P2 (HAL-2100) with recorded basis/uncertainty; unknowns Evidence-Pending with a named test; uncertain factors escalate the band one level (HAL-2200 §6.2) — this governs HA-ELEC-010 (cold-voltage sag, D3, Evidence-Pending).
• Risk severity-first on the HAL-2200 matrix; single-fault rule applies to HA-ELEC-002 (charge-path → thermal).
• Control hierarchy (HAL-2300): inherent → protective → information → process; warnings-only challenged.
• Residual reassessed only after RCR defined, implemented (DO-…), and verified (PV-…).
5. Hazard categories & guiding questions (electrical)
1. Accessible voltage / shock — Can a person contact a hazardous voltage in normal or single-fault? (device is SELV — HA-ELEC-001)
2. Charge-input abuse — Can a non-compliant/faulty USB-C source damage the charge path and cause a thermal event? (HA-ELEC-002)
3. ESD — Can an ESD strike upset the MCU/sensor and produce a false reading/verdict? (HA-ELEC-003)
4. EMC susceptibility — Can radiated/conducted RF, EFT, or surge in transit corrupt a reading and cause a missed/false excursion? (HA-ELEC-004)
5. Emissions — Can the device interfere with aircraft or medical equipment nearby? (HA-ELEC-005)
6. Connector leakage/corrosion — Can a connector go intermittent and lose data? (HA-ELEC-006)
7. Insulation/creepage — Can condensation bridge creepage and short a net? (HA-ELEC-007)
8. Latent PCB/solder defect — Can a latent defect cause an intermittent record-electronics failure and a silent gap? (HA-ELEC-008)
9. Connector/probe mis-mating — Can a probe be connected to the wrong channel or reversed? (HA-ELEC-009)
10. Brown-out during NV write — Can cold voltage sag tear a record write? (HA-ELEC-010)
6. Hazard Analysis Register (HA-ELEC records)
HA-ELEC-001 — Electric shock from accessible voltage (SELV device)
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs · Lifecycle Phase: Any (handling / charging) · System State: Charging or operating · Operational Mode: Manual handling
Perspective | Any handler (clinic, logistics, depot)
Source | HA-BND-002; IEC 62368-1 energy-source classification
Initiating Cause/Event | Contact with an accessible conductor while the device is powered/charging
Contributing Conditions | USB-C input present; internal boost rails; a single-fault that could raise an accessible node
Foreseeable Sequence of Events | User touches an accessible contact (USB-C shell, probe contact) during handling/charging → current path through the body
Hazard | Accessible electrical energy (voltage/current)
Hazardous Situation | A person contacts an accessible conductive part while energised
Sequence to Harm | Body current → perception/shock
Harm/Loss | At most a minor perceptible shock; the device operates from a low-voltage cell/USB-C SELV source with no hazardous-voltage node
Severity | S2 (Minor) — human-injury dimension; SELV limits the outcome (HAL-2000 §7, electric-shock anchor S2)
Probability (P1×P2) | P1=1 (Improbable — all accessible circuits are SELV, no hazardous voltage in normal or single-fault per RCR-013) × P2=2 (Remote — contact to a perceptible effect) ⇒ 2, PVL (Very Low) — Basis: IEC 62368-1 SELV classification + judgment; Low uncertainty
Detectability/Warning | N/A — hazard designed out (SELV); no annunciation needed
Initial Risk | Low (B) — S2 × PVL
Risk Acceptability | Broadly acceptable (B); ALARP inherent (SELV)
Risk-Control Strategy | 1 inherent (SELV design, no accessible hazardous voltage)
Risk-Control Requirement | RCR-013 — all accessible circuits SELV; no accessible hazardous voltage in normal or single-fault
Implementation Reference | DO-EE (rail architecture) + DO-PCB · CMP-034
Verification Reference | PV-ELEC-001 (Evidence-Pending) — IEC 62368-1 accessible-voltage/energy-source measurement (normal + single-fault)
Validation Reference | N/A
Residual Probability | P1=1 × P2=1 ⇒ PVL · Residual Severity: S2 · Residual Risk: Low (B) (Evidence-Pending PV-ELEC-001)
New Risks Introduced | None
Overall Residual Contribution | Negligible
Status | Open — Evidence-Pending
Owner | Electrical Engineer
Rationale/Evidence | The SCM is a SELV, cell/USB-C-powered device; IEC 62368-1 classifies its accessible circuits as ES1, so shock harm is inherently minor and improbable. Evidence (standard classification) + judgment.
Assumptions | No mains connection; USB-C source is SELV; no internal high-voltage node is made accessible under single fault.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-002 — USB-C reverse-insertion / OVP / surge → charge-path damage → thermal event
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (USB-C sink CMP-004/010, input front-end CMP-014) · Lifecycle Phase: Depot recharge · System State: Charging · Operational Mode: Charge
Perspective | Depot technician / people nearby; ties HA-ENRG-005 (charge-path → cell thermal)
Source | HA-BND-002; HA-CTX-003 (charging from a non-compliant USB source is named foreseeable misuse); F-CHG
Initiating Cause/Event | A non-compliant, faulty, or malicious USB-C source delivers reverse polarity, over-voltage, surge, or a short into the charge input
Contributing Conditions | Third-party chargers (HA-CTX-010); depot charging of many devices; USB-PD mis-negotiation; cable/adapter faults
Foreseeable Sequence of Events | Out-of-spec source connected → over-voltage/surge/short reaches the PMIC/charge path → charge-path component damage/over-current → local heating → in worst case cell overcharge/heating (ties HA-ENRG-005)
Hazard | Uncontrolled electrical energy from an out-of-spec charge source
Hazardous Situation | The charge path is stressed beyond rating during charging, near a person
Sequence to Harm | Charge-path damage / over-current → localised thermal event → at worst cell overcharge → burn/smoke injury
Harm/Loss | Serious burn/smoke injury (charge-path/cell thermal event)
Severity | S4 (Critical) — human-injury dimension (thermal); at the S4/S5 boundary if it drives a cell runaway (then governed jointly with HA-ENRG). Cold-charge lithium-plating latent-short path is the S4 anchor (HAL-2000 §7).
Probability (P1×P2) | P1=3 (Occasional — non-compliant chargers are common in the field) × P2=2 (Remote — reaching an actual thermal event given input protection) ⇒ 6, PM (Medium) — Basis: Analogous USB-C abuse + judgment; Med uncertainty; closed by reverse/OVP/surge/short test
Detectability/Warning | Medium — protection trips are detectable (fault annunciation, RCR-046); the input event itself is not user-visible
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via input protection + PD cap; single-fault rule applies
Risk-Control Strategy | 2 protective (reject/limit out-of-spec source, secondary backstop)
Risk-Control Requirement | RCR-003 — device rejects/limits an out-of-spec charge source (reverse-insertion, OVP, surge, short); USB-PD PDO capped (with RCR-002 secondary backstop and RCR-001 cold-charge inhibit as independent lines)
Implementation Reference | DO-SAFE-003 → DO-EE-009 (USB-C sink) + DO-EE-010 (protection) · CMP-014/CMP-004/CMP-010
Verification Reference | PV-ELEC-002 (Evidence-Pending) — reverse-insertion / OVP / surge (IEC 61000-4-5) / short test on the charge input
Validation Reference | N/A (verification-closed; thermal-event validation under HA-ENRG)
Residual Probability | P1=2 × P2=1 ⇒ PVL · Residual Severity: S4 · Residual Risk: A (Med) — after input protection + PD cap + secondary backstop verified (Evidence-Pending PV-ELEC-002)
New Risks Introduced | An aggressive OVP clamp could nuisance-trip on marginal-but-valid sources — manage in DO-EE tolerance
Overall Residual Contribution | Moderate — physical-safety record; carried with HA-ENRG-005 on the residual summary; FTA contributor to charge-path thermal event
Status | Open — Evidence-Pending · release-critical (charge safety)
Owner | Electrical Engineer (with Battery/Power for HA-ENRG-005)
Rationale/Evidence | Third-party USB sources are an explicit external dependency and foreseeable misuse (HA-CTX-003/010); input abuse is a known charge-path thermal initiator. Input protection + PD cap + independent secondary backstop give single-fault tolerance.
Assumptions | RCR-002 secondary fuse/OVP and RCR-001 cold-charge inhibit are independent of the input front-end (no common cause); PD PDO is capped in FW+HW.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-003 — ESD event → MCU/sensor upset → false reading/verdict → wrong disposition
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs · Lifecycle Phase: Handling / provisioning / receipt · System State: Armed / interactive · Operational Mode: Logging / user read
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-002/003; IEC 61000-4-2; F-SNS/F-SW
Initiating Cause/Event | An ESD discharge to an accessible port (USB-C, probe connector, button, display bezel) during handling
Contributing Conditions | Dry/cold low-humidity environments (high static); gloved handling; repeated touch at receipt read
Foreseeable Sequence of Events | ESD strike → transient upsets the MCU (reset/latch) or corrupts a sensor sample/ADC → a false reading or a mis-computed verdict is logged/displayed → user acts on it
Hazard | Electrostatic discharge transient into signal/control electronics
Hazardous Situation | An armed device produces or displays a corrupted reading/verdict believed valid after an ESD event
Sequence to Harm | ESD upset → false reading/verdict → if not caught by plausibility/watchdog/verdict-integrity → wrong disposition
Harm/Loss | Wrong disposition of a shipment (compromised dose administered, or potent dose destroyed)
Severity | S4 (Critical) — public-health dimension via wrong disposition
Probability (P1×P2) | P1=3 (Occasional — ESD in dry/cold handling is common) × P2=2 (Remote — an ESD upset reaching an undetected wrong verdict, given watchdog + plausibility + verdict integrity) ⇒ 6, PM (Medium) — Basis: Analogous EMC field experience + judgment; Med uncertainty; closed by ESD immunity test
Detectability/Warning | Medium — a reset/watchdog recovery is logged (honest-gap); a silent sample corruption relies on plausibility/inter-channel checks
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via ESD immunity + honest-gap recovery
Risk-Control Strategy | 1 inherent (ESD-hardened I/O layout/TVS) + 2 protective (immunity + watchdog/plausibility so an upset can't read as a valid pass)
Risk-Control Requirement | RCR-014 — EMC immunity (incl. ESD per IEC 61000-4-2) such that no interference produces a false reading, false verdict, or missed excursion (with RCR-023 safe-state and RCR-020/021 plausibility/verdict integrity as harm-end backstops)
Implementation Reference | DO-EE (ESD protection/layout, TVS) + DO-PCB-010 (evidence/coat) · CMP-034; harm-end DO-SAFE-006/007
Verification Reference | PV-ELEC-003 (Evidence-Pending) — ESD immunity test (IEC 61000-4-2, contact + air discharge)
Validation Reference | N/A (harm-end validated under HA-SW/HA-DATA)
Residual Probability | P1=2 × P2=1 ⇒ PVL · Residual Severity: S4 · Residual Risk: A (Med) — after ESD immunity + watchdog/plausibility verified (Evidence-Pending)
New Risks Introduced | None significant
Overall Residual Contribution | Moderate — EMC-immunity residual set; FTA contributor to "false reading/verdict"
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Firmware for watchdog/plausibility)
Rationale/Evidence | ESD is a routine handling stress, elevated in dry/cold cold-chain environments; IEC 61000-4-2 immunity plus honest-gap recovery prevent it from reaching a silent wrong verdict.
Assumptions | Watchdog logs recoveries (RCR-023); plausibility/verdict integrity (RCR-020/021) are implemented so an upset cannot present a clean pass.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-004 — EMC susceptibility (RF / EFT / surge in transit) → sensor error → missed/false excursion → wrong disposition
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs · Lifecycle Phase: Transport (air/ground) · System State: Armed · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-002/003; IEC 61000-4-3/-4-4/-4-6; F-SNS
Initiating Cause/Event | Radiated RF (nearby transmitters, own radios), EFT/burst, or conducted RF couples into the sensor signal chain during transit
Contributing Conditions | Long probe lead acting as an antenna; high-power RF environments (aircraft, industrial); the device's own LTE-M/Wi-Fi bursts coupling into sensing
Foreseeable Sequence of Events | RF/EFT coupling → sensor ADC error or biased reading → detection engine sees a wrong temperature → a real excursion is missed, or a false excursion is raised → wrong verdict logged
Hazard | Electromagnetic interference into the measurement chain
Hazardous Situation | An armed device logs erroneous temperatures during transit believed valid
Sequence to Harm | Sensor error → missed excursion (false accept) or false excursion (false reject) → wrong disposition
Harm/Loss | Wrong disposition of a shipment (compromised dose administered, or potent dose destroyed)
Severity | S4 (Critical) — public-health dimension via wrong disposition
Probability (P1×P2) | P1=2 (Remote — with EMC-immune design and filtered front-ends) × P2=3 (Occasional — an undetected sensor error reaching a wrong disposition) ⇒ 6, PM (Medium) — Basis: IEC 61000 + analogous EMC design + judgment; Med uncertainty; closed by RF/EFT/surge immunity test
Detectability/Warning | Medium — implausible/rate-of-change flags and inter-channel plausibility catch gross errors; a small coherent bias near a threshold is the harder case
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via EMC immunity + plausibility
Risk-Control Strategy | 1 inherent (filtering/shielding/layout) + 2 protective (immunity + plausibility)
Risk-Control Requirement | RCR-014 — EMC immunity (RF/EFT/surge per IEC 61000) such that no interference produces a false reading, false verdict, or missed excursion (with RCR-020 plausibility backstop)
Implementation Reference | DO-EE (EMC filtering/shielding, probe front-end) + DO-PCB-010 · CMP-034/CMP-024
Verification Reference | PV-ELEC-004 (Evidence-Pending) — radiated RF immunity (IEC 61000-4-3), EFT/burst (-4-4), conducted RF (-4-6), surge (-4-5)
Validation Reference | N/A (harm-end under HA-SW/HA-DATA)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S4 · Residual Risk: A (Med) — after EMC immunity + plausibility verified (Evidence-Pending)
New Risks Introduced | Added shielding can raise mass and in-container RF attenuation — flag to RF/antenna (RCR-031, HA-COMM-003)
Overall Residual Contribution | Moderate — EMC-immunity residual set; FTA contributor to "missed/false excursion"
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Firmware/Sensing for plausibility)
Rationale/Evidence | The measurement chain (especially a long probe lead) is an EMC-susceptible antenna in RF-rich transport; IEC 61000 immunity plus plausibility keep interference from becoming a silent wrong verdict.
Assumptions | Probe front-end is filtered; own-radio bursts are managed by layout/timing; plausibility (RCR-020) is implemented.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-005 — Device emissions interfere with aircraft / medical equipment nearby
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (radios CMP-024/025/026) · Lifecycle Phase: Air transport / clinic proximity · System State: Transmitting · Operational Mode: Upload / alert
Perspective | Nearby equipment and its dependents (aircraft systems, medical equipment) — HA-CTX-006
Source | HA-BND-007 (external dependency / nearby equipment); CISPR/DO-160 emissions
Initiating Cause/Event | The device's intentional (LTE-M/Wi-Fi/BLE) or unintentional emissions exceed limits or transmit at a prohibited time/place
Contributing Conditions | Operation inside an aircraft cargo hold; proximity to sensitive medical equipment in a clinic; radios not gated to permitted windows/bands
Foreseeable Sequence of Events | Device transmits/emits → coupling into nearby aircraft avionics or medical equipment → interference with that equipment's function
Hazard | Electromagnetic emissions from the device
Hazardous Situation | Sensitive equipment operating within coupling range of the device's emissions
Sequence to Harm | Interference → degraded function of nearby equipment → indirect harm (moderate, reversible in credible cases; safety-relevant only where the affected equipment is safety-critical)
Harm/Loss | Moderate — interference with nearby equipment; credible cases are reversible/transient (equipment tolerance + regulatory margins)
Severity | S3 (Serious) — property/nearby-equipment dimension; bounded by emissions compliance and equipment immunity margins
Probability (P1×P2) | P1=2 (Remote — with emissions compliance + band/time gating) × P2=2 (Remote — emissions actually degrading compliant nearby equipment) ⇒ 4, PL (Low) — Basis: CISPR/DO-160 compliance + judgment; Med uncertainty; closed by emissions test
Detectability/Warning | Low from the device side; managed by design compliance and operating-mode gating
Initial Risk | Medium (A) — S3 × PL
Risk Acceptability | Undesirable/ALARP (A) — reduce via emissions compliance + band-plan/airplane-mode gating
Risk-Control Strategy | 1 inherent (compliant emissions design) + 4 process (band/time gating, transport mode)
Risk-Control Requirement | RCR-014 — device shall not interfere with aircraft/medical equipment (emissions side of the EMC requirement)
Implementation Reference | DO-EE (EMC emissions) + DO-PCB-010 + DO-FW (radio gating / transport mode) · CMP-024/025/026/034
Verification Reference | PV-ELEC-005 (Evidence-Pending) — radiated/conducted emissions (CISPR 32-class) + DO-160-class check for aircraft-hold operation
Validation Reference | N/A
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S3 · Residual Risk: B (Low) — after emissions compliance + gating verified (Evidence-Pending PV-ELEC-005)
New Risks Introduced | Aggressive radio gating could delay alert delivery — flag to Connectivity (HA-COMM, RCR-029/030) for the alert-latency trade
Overall Residual Contribution | Low
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Firmware/Connectivity for radio gating)
Rationale/Evidence | Emissions compliance (CISPR/DO-160) plus operating-mode gating bound the interference risk; regulatory limits and equipment immunity margins make credible harm moderate and reversible.
Assumptions | Radios are power-gated and band-plan-correct (CMP-042/RCR-056); transport/airplane operating mode available; carriage in aircraft holds is per applicable airline/regulatory approval.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-006 — Connector leakage/corrosion → intermittent contact → data loss
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (probe connector CMP-005, USB-C CMP-004) · Lifecycle Phase: Reuse life / transport · System State: Armed · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-003; HA-ENV condensation (ties HA-ENV-003); F-STO/F-SNS
Initiating Cause/Event | Moisture/condensation and reuse-cycle wear cause leakage current or corrosion at a connector contact → intermittent contact
Contributing Conditions | Cold-to-warm condensation cycling; salt/humidity over reuse life; many insertion cycles at the probe connector
Foreseeable Sequence of Events | Corrosion/leakage at connector → intermittent probe/sensor or storage-related signal → dropped samples or a record gap
Hazard | Intermittent electrical contact from connector leakage/corrosion
Hazardous Situation | An armed device intermittently loses a channel, producing a record gap believed complete
Sequence to Harm | Intermittent → data loss/gap → if not annunciated → limited-quantity wrong disposition on an incomplete record
Harm/Loss | Wrong disposition of a limited quantity / a trip released on an incomplete record
Severity | S3 (Serious) — public-health dimension, limited quantity (single device/shipment)
Probability (P1×P2) | P1=3 (Occasional — connector corrosion is a known reuse-life failure) × P2=2 (Remote — reaching a wrong disposition given honest-gap annunciation) ⇒ 6, PM (Medium) — Basis: Reliability analogous + judgment; Med uncertainty; closed by cold-cycle + cleanliness + insertion-cycle test
Detectability/Warning | Medium — probe-continuity/plausibility (RCR-020) and honest-gap flags catch a lost channel; a marginal intermittent is harder
Initial Risk | Medium (A) — S3 × PM
Risk Acceptability | Undesirable/ALARP (A) — reduce via sealing/coating + continuity monitoring
Risk-Control Strategy | 1 inherent (sealed/plated connectors, conformal coat) + 2 protective (continuity/honest-gap)
Risk-Control Requirement | RCR-015 — conformal coat + seal protect record-integrity electronics from condensation/leakage/corrosion across −30↔+60 °C over reuse life (with RCR-020/022 continuity + honest-gap backstop)
Implementation Reference | DO-SAFE-011 → DO-MECH-005 (seal) + DO-PCB-010 (conformal coat) · CMP-002/005/034
Verification Reference | PV-ELEC-006 (Evidence-Pending) — IP + cold-cycle + cleanliness/SIR test + connector insertion-cycle endurance
Validation Reference | N/A (harm-end under HA-DATA)
Residual Probability | P1=2 × P2=1 ⇒ PVL · Residual Severity: S3 · Residual Risk: B (Low) — after sealing/coat + continuity monitoring verified (Evidence-Pending PV-ELEC-006)
New Risks Introduced | Conformal coat over a connector could impair mating if mis-masked — manage in DO-PCB keep-outs (CMP-034)
Overall Residual Contribution | Low — record-integrity residual set
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Mechanical for sealing, Firmware for continuity)
Rationale/Evidence | Connector corrosion/leakage under condensation is a classic reuse-life data-loss initiator; sealing + conformal coat + continuity monitoring + honest gap keep it from a silent clean pass.
Assumptions | Probe connector detects partial-seat/loose (CMP-005); conformal-coat keep-outs preserve mating; honest-gap (RCR-022) implemented.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-007 — Insulation/creepage breakdown under condensation → short
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (PCB CMP-034) · Lifecycle Phase: Cold-to-warm cycling / reuse life · System State: Any · Operational Mode: Autonomous
Perspective | Property/device + information-integrity (record electronics)
Source | HA-BND-002; IEC 62368-1 creepage/clearance; HA-ENV condensation (ties HA-ENV-003)
Initiating Cause/Event | Condensation bridges a creepage gap between conductors → leakage or a short
Contributing Conditions | −30↔+60 °C condensation cycling; contamination lowering surface insulation resistance; tight creepage on a dense PCB
Foreseeable Sequence of Events | Condensation forms on the PCB → bridges creepage between nets → leakage/short → localised malfunction or record-electronics fault
Hazard | Reduced insulation / creepage breakdown (condensation-bridged conductors)
Hazardous Situation | Conductors bridged by condensation during operation
Sequence to Harm | Short/leakage → record-electronics malfunction (data loss/gap) or a supply fault → limited-quantity wrong disposition; SELV means no shock harm
Harm/Loss | Wrong disposition of a limited quantity / repairable electronics fault
Severity | S3 (Serious) — public-health (limited quantity) / property; no shock harm (SELV)
Probability (P1×P2) | P1=2 (Remote — with creepage design + conformal coat) × P2=2 (Remote — bridged conductors reaching a harmful fault) ⇒ 4, PL (Low) — Basis: IEC 62368-1 creepage + judgment; Med uncertainty; closed by condensation-cycle test
Detectability/Warning | Medium — a resulting fault is annunciated (honest-gap); the transient bridge itself is not directly observed
Initial Risk | Medium (A) — S3 × PL
Risk Acceptability | Undesirable/ALARP (A) — reduce via creepage design + conformal coat
Risk-Control Strategy | 1 inherent (creepage/clearance + conformal coat)
Risk-Control Requirement | RCR-015 — conformal coat + seal protect record-integrity electronics across condensation cycling over reuse life
Implementation Reference | DO-SAFE-011 → DO-PCB-010 (creepage/clearance + conformal coat) · CMP-034
Verification Reference | PV-ELEC-007 (Evidence-Pending) — condensation-cycle test + insulation-resistance/SIR + creepage design review (IEC 62368-1)
Validation Reference | N/A
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S3 · Residual Risk: B (Low) — after creepage + conformal coat verified (Evidence-Pending PV-ELEC-007)
New Risks Introduced | None
Overall Residual Contribution | Low — condensation/record-integrity residual set (shared with HA-ENV-003)
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Mechanical/PCB)
Rationale/Evidence | Condensation-bridged creepage is a standard cold-chain electronics failure; adequate creepage plus conformal coat control it. SELV removes the shock dimension. Evidence (IEC 62368-1) + judgment.
Assumptions | Conformal coat covers safety/record-critical nets with correct keep-outs (CMP-034); creepage sized for pollution degree of a condensing environment.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-008 — Latent PCB/solder defect → intermittent record-electronics failure → silent gap
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (PCB/assembly CMP-034) · Lifecycle Phase: Field (post-manufacture, over reuse life) · System State: Armed · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path); ties HA-SW-013
Source | HA-BND-003; F-STO/F-SW; manufacturing-defect prompt (ties HA-MFG-008)
Initiating Cause/Event | A latent solder/PCB/component defect (cold joint, void, cracked pad, marginal part) escapes EOL screening and manifests intermittently under thermal/vibration stress in the field
Contributing Conditions | −30↔+60 °C thermal cycling flexing a marginal joint; transport vibration; the defect being on a record-integrity net (storage/MCU/sensor)
Foreseeable Sequence of Events | Latent defect → intermittent open/short under stress → record-electronics glitch (reset, dropped sample, storage write failure) → a silent gap in the trip record
Hazard | Latent electrical defect on a record-integrity net
Hazardous Situation | An armed device intermittently fails to record, producing a silent gap believed to be a complete monitored trip
Sequence to Harm | Intermittent record failure → silent gap → if honest-gap fails → wrong accept on an incomplete record
Harm/Loss | Wrong disposition (compromised dose administered) — a silent-gap false-negative
Severity | S4 (Critical) — public-health dimension; silent-gap false negative is the "dead-in-cold"-class harm (HAL-2000 §7)
Probability (P1×P2) | P1=2 (Remote — AOI/ICT/EOL screening catches most; a latent escape is the residual) × P2=3 (Occasional — a record-net intermittent reaching a wrong accept, mitigated by honest-gap/self-check) ⇒ 6, PM (Medium) — Basis: Reliability analogous (solder-defect escape rates) + judgment; Med uncertainty; closed by AOI/ICT coverage review + thermal-cycle/HALT
Detectability/Warning | Medium — POST/BIST, availability self-check (RCR-027), and honest-gap catch a manifested fault; a rare intermittent is the harder case
Initial Risk | High (R) — S4 × PM
Risk Acceptability | Unacceptable (R) as-is — reduce via screening + self-check + honest gap
Risk-Control Strategy | 4 process (AOI/ICT/EOL screening) + 2 protective (self-check + honest-gap so a fault can't read as a pass)
Risk-Control Requirement | RCR-016 — latent solder/PCB/component defects screened by AOI + EOL coverage; no latent single-point failure on safety-critical nets (with RCR-024 power-fail-safe record + RCR-027 availability self-check backstop)
Implementation Reference | DO-TST (AOI/ICT) + DO-PCB (design-for-test) · CMP-034; harm-end DO-SAFE-008 (record) + DO-FW (self-check)
Verification Reference | PV-ELEC-008 (Evidence-Pending) — AOI/ICT coverage review + thermal-cycle/HALT screening + self-check outage test
Validation Reference | N/A (harm-end under HA-SW-013/HA-DATA)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S4 · Residual Risk: A (Med) — after screening + self-check + honest-gap verified (Evidence-Pending)
New Risks Introduced | None
Overall Residual Contribution | Moderate — record-integrity residual set; FTA contributor to "silent gap / dead-in-cold" (with HA-SW-013, HA-ELEC-010)
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Manufacturing for screening, Firmware for self-check)
Rationale/Evidence | A latent defect on a record-integrity net is a silent-gap false-negative source (the SCM's most dangerous failure class per fault-management §1); screening reduces the escape, and self-check + honest gap prevent a silent clean pass. Reliability analogous + judgment.
Assumptions | Safety-critical nets are defined for coverage (RCR-016); availability self-check (RCR-027) and honest gap (RCR-022) are implemented.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-009 — Connector mismatch / reverse probe connection → wrong channel/reading
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (probe connector CMP-005) · Lifecycle Phase: Provisioning / probe service · System State: Pre-arm / armed · Operational Mode: Setup
Perspective | Programme/mission + beneficiary (information-harm path)
Source | HA-BND-003; HA-CTX-009 (probe safety-relevant); F-SNS
Initiating Cause/Event | Probe connected to the wrong channel, reversed, or a mismatched probe used → the wrong or a miscalibrated channel is read as the product-zone temperature
Contributing Conditions | Multiple similar connectors; no keying/identification; field probe swap; different probe types
Foreseeable Sequence of Events | Wrong/reversed/mismatched probe connected → device treats a non-product or miscalibrated channel as the product-zone temperature → record contains a wrong reading
Hazard | Mis-mated / mismatched probe channel (wrong measurement source)
Hazardous Situation | An armed device logs the wrong channel as the product-zone temperature believed valid
Sequence to Harm | Wrong channel/reading → verdict on wrong data → wrong disposition (bounded because probe-ID + plausibility catch gross cases; a same-type mismatch is the residual)
Harm/Loss | Wrong disposition of a shipment; bounded to S3 by probe identification + pre-arm checks
Severity | S3 (Serious) — public-health dimension; bounded by probe-identification/plausibility (a fully wrong channel is detectable; the residual is a same-type offset)
Probability (P1×P2) | P1=2 (Remote — connector keying + probe-ID limit mis-mating) × P2=3 (Occasional — a mismatch that reaches a wrong reading given detection) ⇒ 6, PM (Medium) — Basis: Judgment + HA-CTX-009; Med uncertainty; closed by probe-fault-injection + keying review
Detectability/Warning | Medium — probe-presence/identity + plausibility (RCR-020) detect wrong/absent probe; same-type miscalibration is the harder residual
Initial Risk | Medium (A) — S3 × PM
Risk Acceptability | Undesirable/ALARP (A) — reduce via keying/identification + plausibility
Risk-Control Strategy | 1 inherent (keyed/uniquely identifiable connector) + 2 protective (probe-ID + plausibility)
Risk-Control Requirement | RCR-020 — absent/open/short/wrong probe and implausible reads detected and never read as a valid product-zone temperature (connector keying supports inherent prevention)
Implementation Reference | DO-SAFE-005 → DO-FW-007 (probe-ID/plausibility) + DO-EE-013 + connector keying (DO-MECH/DO-EE) · CMP-005/CMP-020
Verification Reference | PV-ELEC-009 (Evidence-Pending) — probe-fault-injection (wrong/reversed/mismatched) + connector-keying review
Validation Reference | VAL-ELEC-009 (Evidence-Pending) — probe-connection usability (ties HA-USE-005)
Residual Probability | P1=1 × P2=2 ⇒ PVL · Residual Severity: S3 · Residual Risk: B (Low) — after keying + probe-ID + plausibility verified (Evidence-Pending)
New Risks Introduced | Keying that is too restrictive could block legitimate probe variants — manage in DO-MECH/DO-EE
Overall Residual Contribution | Low — product-zone-accuracy residual set
Status | Open — Evidence-Pending
Owner | Electrical Engineer (with Firmware/Sensing and Mechanical for keying)
Rationale/Evidence | Probe identification and connector keying make a wrong-channel connection largely inherently preventable and machine-detectable; the residual same-type offset is bounded, giving S3. Judgment + HA-CTX-009.
Assumptions | Probe is independently identifiable (CMP-020); connector is keyed; plausibility available.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-ELEC-010 — Brown-out during NV write (cold voltage sag, D3) → torn record
Hazard ID (library) | HAZ-ELEC
Product/Config | SCM all SKUs (MCU brown-out supervisor CMP-015, NV storage CMP-033) · Lifecycle Phase: Deployment in deep cold · System State: Armed, writing to NV · Operational Mode: Autonomous logging
Perspective | Programme/mission + beneficiary (information-harm path); ties HA-SW-013
Source | HA-BND-002/003; D3 (cold energy tension); F-BRN/F-STO; fault-management §3 (brown-out barrier)
Initiating Cause/Event | At −30 °C the cell's internal resistance rises and voltage sags under load; a load transient (radio TX, write) collapses the rail below the operating threshold during an NV write
Contributing Conditions | D3 low-temperature chemistry Evidence-Pending (cold voltage-sag behaviour unknown); high write/TX load coincidence; end-of-trip low SoC compounding the sag
Foreseeable Sequence of Events | Cold voltage sag during an NV write → brown-out mid-write → torn/partial record entry or reset → a corrupted/lost record segment → a gap or a torn record
Hazard | Supply brown-out during a non-atomic record write (cold voltage collapse)
Hazardous Situation | An armed device tears a record write in deep cold, leaving a corrupted/lost segment believed complete
Sequence to Harm | Torn/lost record → if not caught by the atomic-write barrier / honest gap → wrong accept on an incomplete record
Harm/Loss | Wrong disposition (compromised dose administered) — a cold-driven silent-gap false-negative
Severity | S4 (Critical) — public-health dimension; torn-record/dead-in-cold false-negative anchor (HAL-2000 §7)
Probability (P1×P2) | P1=3 (Occasional — cold voltage sag is credible and D3 is Evidence-Pending) × P2=3 (Occasional — torn record reaching a wrong accept, mitigated by atomic-write barrier) ⇒ 9 → escalated to PH (High) per the uncertain-probability rule (HAL-2200 §6.2; D3 Evidence-Pending) — Basis: Judgment + D3 tension; High uncertainty (chemistry Evidence-Pending); closed by cold-endurance/reserve test (PRD-VER-004) + power-fail-during-write test
Detectability/Warning | Medium — the brown-out detector arms the atomic-write barrier so a torn write fails the record rather than silently corrupting it (fault-management §3); logged reset gives honest gap
Initial Risk | High (R) — S4 × PH (severity-first; escalated for D3 uncertainty)
Risk Acceptability | Unacceptable (R) as-is — Open/Evidence-Pending; driven by severity until the cold test closes probability
Risk-Control Strategy | 2 protective (brown-out detect + atomic-write barrier + energy reserve so a write is never left torn)
Risk-Control Requirement | RCR-024 — storage shall be power-fail-safe (atomic writes, brown-out barrier); no torn/lost/fabricated record; no overwrite of unsent data (with RCR-006 energy reserve and RCR-022 honest-gap backstop)
Implementation Reference | DO-SAFE-008 → DO-FW-010 (storage) + DO-EE-002 (rail interlock) · CMP-015 (brown-out supervisor) / CMP-033; reserve DO-SAFE-012
Verification Reference | PV-ELEC-010 (Evidence-Pending) — power-fail-during-write test at −30 °C + brown-out-barrier fault-injection; and 90-day cold endurance/reserve (PRD-VER-004)
Validation Reference | N/A (harm-end under HA-SW-013/HA-DATA)
Residual Probability | P1=2 × P2=1 ⇒ PVL · Residual Severity: S4 · Residual Risk: A (Med) — after atomic-write barrier + reserve verified and D3 cold behaviour characterised (Evidence-Pending; stays Open until PRD-VER-004 closes)
New Risks Introduced | A larger energy reserve to guarantee safe write landing competes with the 90-day budget — flag to Battery/Power and Environmental (D3, RCR-006/017)
Overall Residual Contribution | Moderate–high (D3 top hazard) — a record-integrity residual carried explicitly as a D3 Evidence-Pending top hazard (HAL-000 §7); FTA contributor to "torn record / dead-in-cold"
Status | Open — Evidence-Pending (D3 cold voltage-sag) · release-critical (record integrity + D3)
Owner | Electrical Engineer (with Firmware for atomic-write, Battery/Power for D3 reserve)
Rationale/Evidence | Cold voltage sag on rising cell impedance is a real physics risk with D3 chemistry Evidence-Pending; the atomic-write barrier (fault-management §3, SYS-168) converts a would-be silent corruption into an honest failed-write/gap. Probability is escalated (uncertain-probability rule) and driven by severity until the cold test closes it — no probability is invented.
Assumptions | Brown-out detector holds reset below threshold and arms the atomic-write barrier (CMP-015); energy reserve preserves record-commit (RCR-006); honest gap on restoration (RCR-022).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
7. Traceability notes
Upstream | severity HAL-2000; probability HAL-2100; acceptability HAL-2200; RCRs HAL-2400 (RCR-003/013/014/015/016/020/024, with backstops RCR-001/002/006/021/022/023/027). Boundary/context per HAL-200 (HA-BND-002/003/005) and HAL-100 (HA-CTX-003/006/009/010).
Implementation | DO-SAFE-003 (charge-input abuse), DO-SAFE-008 (record integrity/brown-out), DO-SAFE-011 (ingress/condensation), DO-SAFE-005/006/007 (probe fault/verdict/safe state), DO-EE (rails, EMC, USB-C, brown-out), DO-PCB-010 (coat/creepage/evidence), DO-TST (AOI/ICT) per DOL-1200.
Components | CMP-004 (USB-C port), CMP-005 (probe connector), CMP-010 (PMIC/charge), CMP-014 (input front-end), CMP-015 (MCU/brown-out), CMP-024/025/026 (radios), CMP-033 (NV storage), CMP-034 (PCB/coat).
Cross-domain cross-links (flag to named engineers)
• HA-ELEC-002 ↔ HA-ENRG-005 — charge-input abuse initiator ↔ cell/charge-path thermal; shared RCR-003 (with RCR-001/002 backstops); reconcile with Battery/Power (HA-ENRG); joint charge-safety FTA. Also ties HA-MECH-009/HA-ENRG-003 at the cell.
• HA-ELEC-010 ↔ HA-SW-013 / HA-DATA-006 & HA-ENRG-006 — cold brown-out torn record ↔ dead-in-cold silent gap; shared RCR-024/006/022; reconcile with Firmware/Data and Battery/Power; D3 top hazard (HAL-000 §7); joint Stage-11 FTA "torn record / dead-in-cold".
• HA-ELEC-008 ↔ HA-SW-013 / HA-MFG-008 & HA-MECH-006 — latent defect / detached fragment initiator of a record-electronics intermittent; shared RCR-016/024/027; reconcile with Manufacturing and Firmware.
• HA-ELEC-006/007 ↔ HA-ENV-003/004 — condensation/ingress → leakage/corrosion/short; shared RCR-015/018; reconcile with Environmental and Mechanical.
• HA-ELEC-003/004 → HA-SW / HA-DATA-001 — EMC/ESD upset → false reading/verdict; shared RCR-014/020/ 021/023; reconcile with Firmware/Data and Human-Factors.
• HA-ELEC-004/HA-ELEC-005 → HA-COMM-003 (RCR-031) — added shielding/mass and radio gating trade against in-container RF and alert latency; flag to RF/Connectivity.
Downstream | S4/S5 records (HA-ELEC-002/003/004/008/010) become Stage 11 (ERFKL) FTA/DFMEA inputs; RCRs feed Stage 12; PV-ELEC-### become Stage 13 verification.
8. Residual-risk / closure notes
No open R at release | HA-ELEC-002/003/004/008 (initial R) reduce to A (Med) only after their RCRs are implemented and verified. HA-ELEC-010 stays Open/Evidence-Pending and is driven by severity until the D3 cold-endurance/power-fail-during-write tests (PRD-VER-004, PV-ELEC-010) close its probability — a named D3 top hazard (HAL-000 §7). No R may remain open at release (HAL-2200 §8).
Single-fault tolerance | HA-ELEC-002 (charge-input abuse) requires input protection (RCR-003) and independent secondary backstop (RCR-002) and cold-charge inhibit (RCR-001) with no common cause.
SELV | HA-ELEC-001 confirms no hazardous-voltage/shock exposure exists — the device's electrical harm is dominated by the information pathway, not shock.
Honest-gap dominance | HA-ELEC-006/008/010 all rely on the no-false-reassurance invariant (fault- management §6) so an electrical fault produces an honest gap, never a silent clean pass — the harm-end backstop (RCR-022/024/027) is as load-bearing as the electrical prevention.
Evidence-Pending closures | PV-ELEC-001…010 (SELV measurement, reverse/OVP/surge/short, ESD immunity, RF/EFT/surge immunity, emissions/DO-160, IP+SIR+insertion, condensation+creepage, AOI/ICT+HALT+self-check, probe-fault+keying, power-fail-during-write at −30 °C + PRD-VER-004).
Overall contribution | the electrical domain contributes one release-critical charge-safety record (HA-ELEC-002, with HA-ENRG-005), and a record-integrity residual set (HA-ELEC-006/007/008/010) to the overall residual-risk summary (HAL-2900), of which HA-ELEC-010 is a D3 top hazard. Remaining records are broadly acceptable (B) after verification.
Status | all records Open — Evidence-Pending; module closes when every S4/S5 record has a verified RCR/DO/PV, no R remains open (incl. the D3 HA-ELEC-010 closure), and HAL-2900 gives a benefit-risk-based release recommendation.
HAL-800 (HA-TFC) — Thermal / Fire / Chemical Hazards
Master Hazard Analysis Library — Smart Cold Chain Monitor (SCM) · Stage 10 (ORCH-017) · Module HAL-800, record code HA-TFC · Owner: Principal Thermal / Battery-Safety Engineer, with System-Safety (lead), Mechanical, and Quality/Regulatory (EM-ENG-006). This module analyses thermal, fire, and chemical hazards — hot surfaces, thermal-runaway propagation and its fire/smoke/offgas consequences, electrolyte release, and cleaning-/condensation-driven chemical degradation — following the ISO 14971 harm-and-exposure-first method (HAL-000 §4). It is not an FMEA: each record is a complete risk chain, scored severity-first per CLAUDE.md rule 3. Standards cited by designation; revisions to confirm.
---
1. Purpose
Identify, chain, and control every credible way the SCM can harm a person, property, or the mission through heat, fire, or chemical release: a hot self-heat surface; a single-cell thermal event propagating into a fire/smoke/toxic-offgas event — most critically in an occupied aircraft cargo hold (a carried transport top concern); electrolyte leakage; and chemical/condensation ingress that degrades seals and record-integrity electronics. This module consumes the initiating battery/energy events analysed in HAL-500 (HA-ENRG) and carries their thermal/fire/chemical consequences to harm.
2. Scope
All SCM SKUs (reusable). In-scope elements: self-heat element and its surface (CMP-013), cell/pack as a thermal-runaway source (CMP-006/007), enclosure as a fire enclosure and thermal barrier (CMP-001/034), seals (CMP-002), conformal coat and record-integrity electronics (CMP-034, CMP-033/041), and the cleaning/handling and cold-to-warm environmental interfaces. Lifecycle phases: depot warm-before-charge, armed air/ground transport, receipt, and between-trip cleaning/storage. Initiating energy causes (cold-charge plating, OV/OC/OT, mechanical crush, charge-source abuse) are owned by HAL-500 and referenced here at the consequence end. Condensation/ingress shares controls with HAL-1400 environmental (HA-ENV-003) and electrical (HA-ELEC-006/007). The dominant information-harm consequence of a chemical/condensation electronics fault uses the HA-BND-003 pathway.
3. Applicable standards & references
• ISO 14971 — risk management; severity fixed by outcome (HAL-2000); P1×P2 (HAL-2100).
• IEC 62368-1 — thermal burn (touch-temperature limits), fire enclosure, safeguards against fire spread.
• IEC 62133-2 — lithium cell safety; venting/offgas behaviour; abuse response.
• UN 38.3 — lithium-battery transport abuse tests (T1–T8) evidencing containment.
• IATA DGR — lithium batteries in air transport; the basis for the occupied-hold fire concern (HA-TFC-002); lithium-battery offgas is flammable/toxic.
• IEC 60068 — environmental (cold/thermal-cycle; condensation context) — ties HA-TFC-006.
• IEC 60529 — ingress protection (IP) rating — ties HA-TFC-004/005/006.
Upstream | PRD-SAFE; UN-900; D3 (energy/chemistry tension); DOL-1200 (DO-SAFE controls); HAL-2400 (RCR register); HAL-500 (initiating energy events); component-list.md.
Evidence-honesty note | electrolyte composition, offgas species/quantity, self-heat surface temperature, and enclosure fire-containment performance are material-/design-dependent and Evidence-Pending on the selected chemistry (CMP-006) and enclosure design; none is asserted as a number here — each is carried Evidence-Pending with a named test (abuse/offgas/fire-enclosure/touch-temperature).
4. Governance & method
• Every record uses the 34-field structure (HAL-000 §9); every field answered or justified N/A.
• Severity fixed by credible worst-case outcome (HAL-2000). The occupied-aircraft fire outcome is the module's S5 anchor (HAL-2000 §5/§7).
• Probability decomposed P1×P2 with basis + uncertainty (HAL-2100); Evidence-Pending items are escalated one band and driven by severity (HAL-2200 §6.2) with a named PV-/VAL- test.
• Single-fault rule (HAL-2200 §6.5) and catastrophic-severity rule (§6.1) govern the S5 propagation record (HA-TFC-002): an S5 outcome is never broadly acceptable and requires an independent protective layer (containment enclosure independent of cell protection).
Control hierarchy (HAL-2300) | inherent (fire enclosure, material selection) → protective (containment, seals) → information (transport labelling/IFU) → process (UN 38.3 / IATA compliance, cleaning procedure).
• Residual risk re-estimated only after a control is implemented (DO-…) and verified (PV-…); residual bands below are projected, Evidence-Pending.
5. Categories & guiding questions (thermal / fire / chemical prompt list)
1. Hot surfaces: Can any surface (self-heat element, enclosure in hot ambient during charge) exceed the touch-temperature burn limit?
2. Fire / propagation: Can a single-cell thermal event propagate to fire/smoke/toxic offgas, and where is the worst exposure (occupied aircraft hold)? Is the enclosure a fire enclosure?
3. Toxic offgas / inhalation: On venting, does the offgas present an inhalation hazard in an enclosed space?
4. Chemical release: Can electrolyte leak and cause chemical burn/corrosion?
5. Chemical ingress / cleaning: Can cleaning agents or condensation degrade materials, seals, or record-integrity electronics between trips?
6. Condensation: Does cold→warm transition wet the interior and short/corrode the electronics?
7. Sustained heat / material degradation: Does prolonged +60 °C degrade seals/enclosure or offgas materials?
6. Hazard Analysis Register (HA-TFC records)
HA-TFC-001 — Hot self-heat element surface → burn to technician during depot warm-before-charge
Hazard ID (library) | HAZ-THERM (hot surface)
Product/Config | SCM all SKUs with self-heat option (CMP-013) · Lifecycle Phase: Recharge (depot), warm-before-charge · System State: Self-heat active (pre-charge warming) · Operational Mode: Depot turnaround
Perspective | Depot technician handling the device to charge it
Source | IEC 62368-1 (touch-temperature limits); RCR-005; HA-ENRG-001 (self-heat is the enabling control for cold-charge)
Initiating Cause/Event | The self-heat element (provided to warm a cold cell so it can be charged safely, per RCR-005) raises a surface/element temperature that a technician contacts.
Contributing Conditions | Self-heat is a new risk introduced by the HA-ENRG-001 control; high-throughput bench handling; element may be internal but couple heat to a touchable surface; feasibility of self-heat itself is Evidence-Pending (CMP-013).
Foreseeable Sequence of Events | Cold device placed to charge → self-heat activates to warm the cell above 0 °C → element/surface heats → technician touches a hot surface → burn.
Hazard | Hot surface / thermal energy at the self-heat element.
Hazardous Situation | A technician contacting a heated surface during warm-before-charge.
Sequence to Harm | Contact with hot surface → moderate skin burn (first aid).
Harm/Loss | Moderate reversible burn requiring first aid (S3).
Severity | S3 (Serious) — human injury, moderate reversible burn (HAL-2000 §7 lists HA-TFC-001 "hot self-heat surface burn = S3").
Probability (P1×P2) | P1=2 (Remote — self-heat is internal/bounded; may not be a touchable surface) × P2=2 (Remote — contact→treatable burn) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62368-1; Medium uncertainty (self-heat feasibility EP). Closed by touch-temperature test (PV-TFC-001).
Detectability/Warning | LED/OLED indicates charging/warming state (RCR-046); touch-temperature bounded by design; possible warning label.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP — bound surface temperature by design; warnings only if design cannot.
Risk-Control Strategy | 1 (inherent — bound touch-temperature) + 2 (protective — thermal isolation of element) + 3 (information — state legibility/label)
Risk-Control Requirement | RCR-005 — self-heat shall only enable the charge guard by warming the cell; it shall not exceed the safe upper thermal limit; touchable-surface temperature within IEC 62368-1 burn limits. (+ RCR-046 state legibility.)
Implementation Reference | DO-SAFE-001 (self-heat) → DO-EE-007 · CMP-013 (self-heat element)
Verification Reference | PV-TFC-001 (Evidence-Pending) — touch-temperature measurement of accessible surfaces during self-heat (IEC 62368-1 limits); self-heat upper-bound test
Validation Reference | VAL-TFC-001 (Evidence-Pending) — depot-handling usability confirming no inadvertent hot-surface contact
Residual Probability | projected Very Low (PVL) with bounded touch-temperature · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None new (self-heat is itself the new risk from HA-ENRG-001, controlled here).
Overall Residual Contribution | Small occupational-safety residual; only material if self-heat is adopted (feasibility EP). Feeds HAL-2900 and HAL-2700 (control-introduced risk).
Status | Controlled-pending-verification (Evidence-Pending; conditional on self-heat adoption)
Owner | Thermal / Battery-Safety Engineer
Rationale/Evidence | IEC 62368-1 touch-temperature; RCR-005; DO-SAFE-001. Severity per HAL-2000 §7.
Assumptions | If self-heat is not adopted (CMP-013 feasibility fails), this record is N/A and cold-charge relies solely on warm-then-charge process — revisit HA-ENRG-001 turnaround.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-002 — Thermal-runaway propagation → fire / smoke / toxic offgas in an OCCUPIED aircraft cargo hold → catastrophic ★ TOP TRANSPORT CONCERN
Hazard ID (library) | HAZ-FIRE / HAZ-THERM (thermal-runaway propagation)
Product/Config | SCM all SKUs · Lifecycle Phase: Air transport · System State: Any (armed transport) · Operational Mode: Transport in aircraft cargo hold
Perspective | Air-cargo crew and everyone aboard the aircraft (catastrophic); property (aircraft); mission
Source | IATA DGR (lithium-battery air-transport fire risk); UN 38.3; IEC 62368-1 (fire enclosure); HAL-2000 §5/§7 (occupied-aircraft fire = S5); consumes HA-ENRG-002/003
Initiating Cause/Event | A single cell enters thermal runaway (from any HA-ENRG initiating cause — plating latent short, OV/OC/OT, or mechanical crush) while the device is in an aircraft cargo hold.
Contributing Conditions | Enclosed hold; limited firefighting access to cargo; lithium-battery offgas is flammable and toxic; a runaway can propagate if the enclosure/pack does not contain it; multiple devices may ship together.
Foreseeable Sequence of Events | Cell runaway → vent + flame + hot flammable offgas → if uncontained, propagates to enclosure/adjacent materials → fire/smoke fills the hold → crew/passengers exposed → possible loss of the aircraft.
Hazard | Propagating thermal-runaway fire and toxic/flammable offgas.
Hazardous Situation | A cell fire developing in an occupied aircraft cargo hold.
Sequence to Harm | Fire/smoke/offgas in an enclosed hold → smoke inhalation, burns, and in the worst credible case a fatal fire / loss of aircraft.
Harm/Loss | Death / catastrophic loss (fire in an occupied aircraft cargo hold).
Severity | S5 (Catastrophic) — human death / mass-casualty potential; the SCM's defining physical S5 (HAL-2000 §5/§7).
Probability (P1×P2) | P1=1 (Improbable — dual-independent cell protection RCR-002 + retention RCR-004 + UN 38.3-qualified) × P2=3 (Occasional — if runaway starts, an enclosed hold makes escalation to serious harm likely) ⇒ product 3 ⇒ Low band (PL), escalated to Medium under conservative + catastrophic rules — Basis: analogous Li-ion event rate + designed containment; Medium uncertainty. Closed by fire-enclosure/abuse test (PV-TFC-002) + UN 38.3.
Detectability/Warning | Onset thermal signature may precede flame; but in a hold there is limited detection/intervention — the control must prevent initiation and contain propagation, not rely on detection.
Initial Risk | AP High (region R — S5; catastrophic-severity rule §6.1 forces High regardless of probability)
Risk Acceptability | Unacceptable without control; an S5 can never be broadly acceptable; requires maximum practicable reduction + independent containment + benefit-risk + disclosure (transport declaration).
Risk-Control Strategy | 1 (inherent — fire enclosure, contain a single-cell event without external flame) + 2 (protective — cell protection + retention prevent initiation) + 4 (process — UN 38.3 / IATA-DGR compliance + transport documentation)
Risk-Control Requirement | RCR-008 — enclosure and internal materials shall contain a single-cell thermal event without external flame propagation and with low-toxicity offgassing (IEC 62368-1 fire enclosure). (+ RCR-002 cell protection + RCR-004 retention prevent initiation; + RCR-007 UN 38.3-tested & IATA-DGR-compliant with transport documentation.)
Implementation Reference | DO-SAFE-002 (protection) + DO-SAFE-010 (retention) · DO-MECH-005 + DO-PCB-010 (fire enclosure per RCR-008) · DO-EE-015 / DO-REG-006 (UN 38.3 / DGR) · CMP-001/034 (enclosure), CMP-006/007 (cell/pack)
Verification Reference | PV-TFC-002 (Evidence-Pending) — single-cell thermal-runaway containment / fire-enclosure abuse test (no external flame propagation; offgas characterisation) + UN 38.3 abuse suite
Validation Reference | N/A (abuse/containment verification carries the assurance)
Residual Probability | projected Very Low (PVL) after prevention (RCR-002/004) + containment (RCR-008) + UN 38.3 verified · Residual Severity: S5 (unchanged — outcome unchanged if it ever occurred) · Residual Risk: AP High→Medium (region A, S5×PVL) — per §6.1 stays region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off, transport disclosure, and post-market monitoring
New Risks Introduced | A robust fire-containment enclosure adds mass/cost; low-toxicity material selection may constrain other properties — tracked in HAL-2700.
Overall Residual Contribution | The top transport S5 residual; mandatory benefit-risk line in HAL-2900 and the Risk-Management Summary; FTA top event (Stage 11). Consumes HA-ENRG-002/003.
Status | Controlled-pending-verification (Evidence-Pending — containment + UN 38.3)
Owner | Thermal / Battery-Safety Engineer (with Mechanical + Regulatory)
Rationale/Evidence | IATA DGR; UN 38.3; IEC 62368-1 fire enclosure; HAL-2000 §5/§7; RCR-002/004/007/008. Severity anchored to the occupied-aircraft-fire S5 case.
Assumptions | Fire enclosure genuinely contains a single-cell event (to be proven); UN 38.3 pass is a transport prerequisite; offgas low-toxicity claim proven by test (ties HA-TFC-003).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-003 — Toxic electrolyte / decomposition offgas on venting → inhalation
Hazard ID (library) | HAZ-CHEM / HAZ-FIRE (toxic offgas)
Product/Config | SCM all SKUs · Lifecycle Phase: Any (cell venting event) · System State: Cell venting · Operational Mode: Any
Perspective | Anyone in the enclosed space during venting — air-cargo crew, depot technician, clinic staff
Source | IEC 62133-2 (venting/offgas); IATA DGR (toxic/flammable offgas); consumes HA-ENRG-002; ties HA-TFC-002
Initiating Cause/Event | A cell over-temperature / runaway causes electrolyte decomposition and venting of gas.
Contributing Conditions | Enclosed space (hold, small clinic room) concentrates offgas; lithium-electrolyte decomposition products include toxic/irritant species (e.g., HF-forming, CO); quantity/species Evidence-Pending on chemistry.
Foreseeable Sequence of Events | Cell over-temp/runaway → electrolyte decomposes → toxic/irritant gas vents → accumulates in an enclosed space → people inhale.
Hazard | Toxic/irritant decomposition offgas.
Hazardous Situation | People in an enclosed space with vented cell offgas.
Sequence to Harm | Inhalation of toxic offgas → respiratory injury / smoke-inhalation injury requiring treatment.
Harm/Loss | Serious injury requiring medical treatment (smoke/offgas inhalation).
Severity | S4 (Critical) — human injury requiring treatment (HAL-2000 §5, "smoke inhalation" = S4). (In an occupied aircraft this is part of the S5 HA-TFC-002 event; scored S4 as the standalone inhalation harm.)
Probability (P1×P2) | P1=1 (Improbable — venting requires a prior runaway, itself controlled by RCR-002) × P2=3 (Occasional — enclosed space → inhalation once vented) ⇒ product 3 ⇒ Low band (PL), escalated to Medium under conservative rule — Basis: analogous Li-ion venting + judgment; Medium uncertainty. Closed by offgas characterisation (PV-TFC-003) with PV-ENRG-002.
Detectability/Warning | Smell/smoke may warn; low-toxicity material selection (RCR-008) reduces harm; venting-path design directs offgas.
Initial Risk | AP High (region R — S4 × PM under escalation)
Risk Acceptability | Unacceptable without control (prevent the upstream runaway + minimise offgas toxicity).
Risk-Control Strategy | 2 (protective — prevent runaway) + 1 (inherent — low-toxicity offgas materials / directed venting)
Risk-Control Requirement | RCR-002 (prevent the OV/OC/OT runaway that leads to venting) + RCR-008 (low-toxicity offgassing consistent with IEC 62368-1 fire enclosure).
Implementation Reference | DO-SAFE-002 (protection) · DO-MECH-005 + DO-PCB-010 (enclosure/offgas per RCR-008) · CMP-008/009 (protection), CMP-001 (enclosure)
Verification Reference | PV-TFC-003 (Evidence-Pending) — offgas species/quantity characterisation during forced-venting abuse (with PV-ENRG-002 / UN 38.3)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) after runaway prevention verified · Residual Severity: S4 · Residual Risk: AP Medium (region A, S4×PVL) — ALARP with benefit-risk sign-off
New Risks Introduced | Directed-venting path could conflict with the ingress seal (RCR-011) — tracked in HAL-2700.
Overall Residual Contribution | Bundled with HA-TFC-002 in the transport benefit-risk; standalone inhalation residual in HAL-2900.
Status | Controlled-pending-verification (Evidence-Pending — offgas characterisation)
Owner | Thermal / Battery-Safety Engineer
Rationale/Evidence | IEC 62133-2 venting; IATA DGR; IEC 62368-1; RCR-002/008. Severity per HAL-2000 §5.
Assumptions | Offgas species/quantity depends on final chemistry (Evidence-Pending); enclosed-space exposure is the credible worst case.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-004 — Electrolyte leak → chemical burn / corrosion
Hazard ID (library) | HAZ-CHEM (electrolyte release)
Product/Config | SCM all SKUs · Lifecycle Phase: Any (cell leak, e.g., after abuse/aging/over-discharge) · System State: Any · Operational Mode: Any
Perspective | Technician handling a leaking device; property (corrosion of electronics)
Source | IEC 62133-2 (leakage); IEC 60529 (containment); consumes HA-ENRG-003/007
Initiating Cause/Event | A cell develops a leak (seal failure, abuse, over-discharge damage, aging) and electrolyte escapes the cell.
Contributing Conditions | Electrolyte is corrosive/irritant; a technician may handle a leaking device without PPE; leaked electrolyte corrodes internal electronics (record-integrity risk).
Foreseeable Sequence of Events | Cell leaks → electrolyte contacts skin/eyes on handling, or corrodes internal electronics → chemical burn/irritation, or latent electronics fault.
Hazard | Corrosive/irritant electrolyte.
Hazardous Situation | A technician handling a leaking device / electrolyte on internal electronics.
Sequence to Harm | Skin/eye contact → chemical burn/irritation (first aid); or corrosion → record-integrity fault (information harm, owned by HA-ELEC-006).
Harm/Loss | Moderate reversible chemical burn requiring first aid (S3).
Severity | S3 (Serious) — human injury moderate reversible; plus property/corrosion.
Probability (P1×P2) | P1=2 (Remote — cell containment + retention + protection) × P2=2 (Remote — contained leak; handling with awareness) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62133-2; Medium uncertainty.
Detectability/Warning | Visible leak/residue on handling; enclosure contains a leak; service procedure covers leaking-device handling.
Initial Risk | AP Medium (region A — S3 × PL)
Risk Acceptability | ALARP
Risk-Control Strategy | 2 (protective — cell containment within enclosure/seal) + 4 (process — leaking-device handling/PPE procedure)
Risk-Control Requirement | RCR-002 (protection reduces leak-inducing abuse) + RCR-011 (enclosure + seals maintain cell containment through drop/vibration/reuse life). Leaking-device handling procedure per DO-SVC.
Implementation Reference | DO-SAFE-002 (protection) · DO-SAFE-010/011 → DO-MECH-005/012 (containment) · CMP-002 (seals), CMP-001 (enclosure), CMP-007 (pack)
Verification Reference | PV-TFC-004 (Evidence-Pending) — leak-containment test (post-abuse); material-compatibility (electrolyte vs internal materials) review
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) with containment · Residual Severity: S3 · Residual Risk: AP Low (region B, S3×PVL)
New Risks Introduced | None.
Overall Residual Contribution | Small; the corrosion→record-integrity escalation is carried by HA-ELEC-006.
Status | Controlled-pending-verification
Owner | Thermal / Battery-Safety Engineer (with Mechanical)
Rationale/Evidence | IEC 62133-2 leakage; IEC 60529; RCR-002/011. Severity per HAL-2000 §5.
Assumptions | Enclosure contains a cell leak long enough for detection at depot; service PPE procedure exists.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-005 — Cleaning-agent incompatibility / ingress between trips → material degradation → seal/electronics fault → record-integrity loss
Hazard ID (library) | HAZ-CHEM (cleaning-agent attack) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Between-trip cleaning / decontamination · System State: Off / cleaning · Operational Mode: Reconditioning
Perspective | Mission / beneficiary (via a later record-integrity fault); depot technician
Source | IEC 60529 (ingress); material-compatibility; ties HA-USE-010; PRD-SAFE (cleaning-chemical compatibility, CMP-001 spec)
Initiating Cause/Event | An incompatible cleaning/disinfection agent is used between trips, or cleaning fluid ingresses past a seal.
Contributing Conditions | Reusable device cleaned many times; cleaning agent may not be the specified one (HA-CTX-003 misuse); enclosure/seal material compatibility (CMP-001/002); ingress past a degraded seal.
Foreseeable Sequence of Events | Incompatible agent applied / fluid ingresses → material/seal degradation over repeated cycles → seal or electronics fault → later a record-integrity failure (wrong/lost reading) on a future trip → wrong disposition.
Hazard | Chemical attack / ingress degrading seals and record-integrity electronics.
Hazardous Situation | A device with a cleaning-degraded seal/electronics armed for a later trip.
Sequence to Harm | Degraded seal/electronics → wrong or lost record → false accept → compromised dose (or false reject) — a limited-quantity disposition error (single device).
Harm/Loss | Wrong disposition of a limited quantity of doses / monitored trip released on a compromised record (S3).
Severity | S3 (Serious) — public-health, limited-quantity wrong disposition (HAL-2000 §5, S3 anchor for limited-quantity disposition; single-device scope).
Probability (P1×P2) | P1=2 (Remote — specified cleaning agent + sealed enclosure) × P2=3 (Occasional — degradation accrues over reuse; caught partly by pre-arm/BIST) ⇒ product 6 ⇒ Medium band (PM) — Basis: judgment + material-compatibility; Medium uncertainty. Closed by cleanliness/compatibility + IP test (PV-TFC-005).
Detectability/Warning | Pre-arm BIST + probe/sensor plausibility (RCR-020) + post-service BIST (RCR-050) screen a degraded device; cleaning IFU specifies compatible agents.
Initial Risk | AP Medium (region A/R — S3 × PM → A Medium)
Risk Acceptability | ALARP
Risk-Control Strategy | 1 (inherent — compatible materials + robust seal) + 3 (information — specify cleaning agents) + 4 (process — post-service BIST)
Risk-Control Requirement | RCR-015 (conformal coat + seal protect record-integrity electronics across cycling over reuse life) + RCR-018 (tolerate the reuse environment) + RCR-050 (post-service BIST verifies integrity before redeploy). Cleaning-agent compatibility specified in IFU.
Implementation Reference | DO-SAFE-011 → DO-MECH-005 (seal) + DO-PCB-010 (conformal coat) · CMP-001 (enclosure material), CMP-002 (seals), CMP-034 (PCB/coat)
Verification Reference | PV-TFC-005 (Evidence-Pending) — cleaning-agent material-compatibility + IP + cleanliness test over reuse-life cycles
Validation Reference | VAL-TFC-005 (Evidence-Pending) — reconditioning-procedure validation (correct agents used in the field)
Residual Probability | projected Low (PL) with compatible materials + BIST screening · Residual Severity: S3 · Residual Risk: AP Medium (region A, S3×PL) — ALARP
New Risks Introduced | None.
Overall Residual Contribution | Small; ties HA-USE-010 (cleaning misuse) and the record-integrity aggregate (RCR-057). Feeds HAL-2900.
Status | Controlled-pending-verification
Owner | Thermal / Battery-Safety Engineer (with Mechanical + Service)
Rationale/Evidence | IEC 60529; material-compatibility engineering; RCR-015/018/050. Severity per HAL-2000 §5.
Assumptions | A compatible-cleaning-agent list is provided and followed; single-device scope keeps this S3 (fleet-wide cleaning error would escalate).
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-006 — Condensation at cold→warm transition → internal wetting → leakage/short → record-electronics fault
Hazard ID (library) | HAZ-CHEM/ENV (condensation) → information-integrity harm (HA-BND-003)
Product/Config | SCM all SKUs · Lifecycle Phase: Cold→warm transition (leaving deep cold into warm humid ambient) · System State: Any · Operational Mode: Any
Perspective | Mission / beneficiary (via record-integrity fault)
Source | IEC 60068 (condensation/thermal cycling); PRD-SAFE; ties HA-ENV-003; consumes the −30↔+60 °C context
Initiating Cause/Event | Moving the device from deep cold into warm, humid air causes water to condense on cold internal surfaces.
Contributing Conditions | −30 °C→+60 °C range creates strong condensation potential (physics); repeated cycling over reuse life; conformal coat + seal are the barriers (CMP-034/002).
Foreseeable Sequence of Events | Cold device enters warm humid air → condensation on cold PCB/electronics → moisture bridges/corrodes a net → leakage current/short → wrong reading or lost record → wrong disposition.
Hazard | Internal condensation wetting record-integrity electronics.
Hazardous Situation | A device with condensation-affected electronics recording/reporting a trip.
Sequence to Harm | Wetting → leakage/short → wrong/lost record → false accept/reject → wrong disposition (limited quantity, single device).
Harm/Loss | Wrong disposition of a limited quantity / compromised record (S3).
Severity | S3 (Serious) — public-health limited-quantity disposition (single-device; HAL-2000 §5). (Escalates if it produces a false-clean-pass at scale, but per-device scope anchors S3.)
Probability (P1×P2) | P1=3 (Occasional — cold→warm transitions are routine and repeated) × P2=2 (Remote — conformal coat + seal + plausibility checks resist and screen) ⇒ product 6 ⇒ Medium band (PM) — Basis: condensation physics + judgment; Medium uncertainty. Closed by condensation cycle test (PV-TFC-006).
Detectability/Warning | Sensor plausibility/fault flags (RCR-020) + verdict integrity (RCR-021) catch implausible reads; conformal coat prevents wetting effects.
Initial Risk | AP Medium (region A — S3 × PM)
Risk Acceptability | ALARP
Risk-Control Strategy | 1 (inherent — conformal coat + seal) + 2 (protective — plausibility/verdict integrity)
Risk-Control Requirement | RCR-015 (conformal coat + seal protect record-integrity electronics across −30↔+60 °C condensation cycling) + RCR-018 (tolerate condensation from cold-to-warm cycling without record-integrity failure).
Implementation Reference | DO-SAFE-011 → DO-MECH-005 (seal) + DO-PCB-010 (conformal coat) · CMP-034 (PCB/coat), CMP-002 (seals)
Verification Reference | PV-TFC-006 (Evidence-Pending) — condensation cycle test (IEC 60068) + cold-cycle IP + insulation-resistance on safety-critical nets
Validation Reference | N/A
Residual Probability | projected Low (PL) with conformal coat verified · Residual Severity: S3 · Residual Risk: AP Medium→Low (region A/B, S3×PL → A Medium) — ALARP
New Risks Introduced | None.
Overall Residual Contribution | Small; ties HA-ENV-003 (environmental) and HA-ELEC-006/007 (electrical). Feeds HAL-2900.
Status | Controlled-pending-verification
Owner | Thermal / Battery-Safety Engineer (with Mechanical/PCB + Environmental)
Rationale/Evidence | IEC 60068 condensation; condensation physics across −30↔+60 °C; RCR-015/018; DO-SAFE-011. Severity per HAL-2000 §5.
Assumptions | Conformal coat covers safety-critical nets with sensor/antenna/contact keep-outs (CMP-034 spec); plausibility checks catch residual wet-fault reads.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-007 — Material offgas / degradation at sustained +60 °C → seal / enclosure degradation
Hazard ID (library) | HAZ-CHEM/ENV (thermal material degradation)
Product/Config | SCM all SKUs · Lifecycle Phase: Sustained hot storage/transport (+60 °C) · System State: Any · Operational Mode: Any
Perspective | Mission (device reliability / later record integrity); environment
Source | IEC 60068 (dry heat); material-selection; PRD-SAFE (+60 °C ceiling)
Initiating Cause/Event | Prolonged exposure at the +60 °C operating ceiling degrades polymers/adhesives/seals (outgassing, embrittlement, compression set).
Contributing Conditions | +60 °C sustained ceiling; reuse over years; seal compression-set limit (CMP-002 spec); adhesive/plastic selection.
Foreseeable Sequence of Events | Sustained +60 °C → material outgassing/embrittlement/compression-set → seal or enclosure degradation → later ingress/record-integrity susceptibility (feeds HA-TFC-006 / HA-ENV-003).
Hazard | Heat-driven material degradation of seals/enclosure.
Hazardous Situation | A device whose seal/enclosure has degraded from sustained heat.
Sequence to Harm | Degraded seal → later ingress/condensation susceptibility → eventual record-integrity fault (limited, and caught by BIST). Direct harm is minor/degradation.
Harm/Loss | Minor product/equipment degradation; small avoidable wastage (S2).
Severity | S2 (Minor) — a precursor/reliability degradation (the downstream record-integrity harm is owned by HA-TFC-006/HA-ENV-003, not double-counted).
Probability (P1×P2) | P1=2 (Remote — heat-rated materials per RCR-017) × P2=2 (Remote — degradation → fault, screened) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 60068; Medium uncertainty.
Detectability/Warning | Periodic seal/enclosure inspection at service; post-service BIST.
Initial Risk | AP Low→Medium (region B/A — S2 × PL → B Low)
Risk Acceptability | Broadly acceptable / ALARP via heat-rated material selection.
Risk-Control Strategy | 1 (inherent — heat-rated materials, low-outgassing) + 4 (process — service inspection)
Risk-Control Requirement | RCR-017 (operate/retain across −30…+60 °C over reuse life — heat-rated materials) + RCR-008 (enclosure/material integrity). Seal-inspection at service (DO-SVC).
Implementation Reference | DO-SAFE-012/DO-MECH-005 (materials) · CMP-001 (enclosure), CMP-002 (seals), CMP-034 (PCB/coat)
Verification Reference | PV-TFC-007 (Evidence-Pending) — dry-heat aging (IEC 60068) + compression-set/seal-integrity + outgassing screen over reuse life
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) with heat-rated materials · Residual Severity: S2 · Residual Risk: AP Low (region B)
New Risks Introduced | None.
Overall Residual Contribution | Negligible directly; a reliability precursor to HA-TFC-006/HA-ENV-003.
Status | Controlled-pending-verification
Owner | Thermal / Battery-Safety Engineer (with Mechanical)
Rationale/Evidence | IEC 60068 dry heat; material engineering; RCR-008/017. Severity per HAL-2000 §5.
Assumptions | Materials selected with margin above +60 °C for sustained exposure; low-outgassing grades used near sensors/optics.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
HA-TFC-008 — Hot enclosure surface at +60 °C ambient during charge → burn
Hazard ID (library) | HAZ-THERM (hot surface)
Product/Config | SCM all SKUs · Lifecycle Phase: Recharge in hot ambient (+60 °C) · System State: Charging · Operational Mode: Recharge
Perspective | Person handling the device to charge/retrieve it in a hot environment
Source | IEC 62368-1 (touch-temperature limits); PRD-SAFE (+60 °C ceiling); ties HA-ENRG-010
Initiating Cause/Event | Charging in a +60 °C ambient adds charge/self-heat dissipation on top of a hot enclosure, raising the touchable surface temperature.
Contributing Conditions | +60 °C ambient ceiling; charge-power dissipation; enclosure surface already near ambient; possible self-heat (CMP-013) contribution.
Foreseeable Sequence of Events | Hot ambient + charging → enclosure surface heats above the touch-temperature burn limit → person contacts it → burn.
Hazard | Hot enclosure surface.
Hazardous Situation | A person contacting a hot enclosure during/after charge in a hot environment.
Sequence to Harm | Contact with hot surface → minor skin burn / discomfort.
Harm/Loss | Minor injury, typically no treatment beyond first aid (S2).
Severity | S2 (Minor) — human injury minor (charge dissipation is low-power; enclosure temperature rise modest; HAL-2000 §5). (Distinguished from HA-TFC-001 self-heat element which is a more concentrated S3 hot source.)
Probability (P1×P2) | P1=2 (Remote — charge management limits dissipation; enclosure spreads heat) × P2=2 (Remote — contact→minor burn) ⇒ product 4 ⇒ Low band (PL) — Basis: judgment + IEC 62368-1; Medium uncertainty.
Detectability/Warning | Charging-state indication (RCR-046); touch-temperature bounded by charge-power management (RCR-005 for any self-heat contribution).
Initial Risk | AP Low→Medium (region B/A — S2 × PL → B Low)
Risk Acceptability | Broadly acceptable / ALARP
Risk-Control Strategy | 1 (inherent — bound surface temperature; manage charge power in hot ambient) + 3 (information — hot-surface/charging indication)
Risk-Control Requirement | RCR-005 (self-heat / thermal contribution bounded below the safe thermal/burn limit) + RCR-017 (operate across +60 °C; touch-temperature within IEC 62368-1 limits during charge). (Charge de-rating in hot ambient via RCR-002 OT.)
Implementation Reference | DO-SAFE-001 (self-heat bound) / DO-SAFE-002 (OT charge management) · CMP-001 (enclosure), CMP-010 (charger), CMP-013 (self-heat)
Verification Reference | PV-TFC-008 (Evidence-Pending) — touch-temperature measurement during charge at +60 °C ambient (IEC 62368-1 limits)
Validation Reference | N/A
Residual Probability | projected Very Low (PVL) with bounded touch-temperature · Residual Severity: S2 · Residual Risk: AP Low (region B)
New Risks Introduced | None.
Overall Residual Contribution | Negligible; ties HA-ENRG-010 (hot-cell charge) and HA-TFC-001 (self-heat surface).
Status | Controlled-pending-verification
Owner | Thermal / Battery-Safety Engineer
Rationale/Evidence | IEC 62368-1 touch-temperature; RCR-005/017. Severity per HAL-2000 §5.
Assumptions | Charge power is low enough that enclosure rise over +60 °C stays within touch limits (to be confirmed by PV-TFC-008); charging in +60 °C ambient may be de-rated by RCR-002 OT.
Review History | 2026-07-18 v1.0 initial (ORCH-017)
7. Traceability notes
Severity/probability/acceptability | every record maps to HAL-2000 (S-scale; the occupied-aircraft-fire S5 anchor §5/§7), HAL-2100 (P1×P2, basis + uncertainty), and HAL-2200 (matrix, regions, AP, catastrophic- severity + uncertain-probability + single-fault rules).
Controls | RCR-002, RCR-004, RCR-005, RCR-007, RCR-008, RCR-011, RCR-015, RCR-017, RCR-018, RCR-050 (HAL-2400). RCR-002/004/008 are the release-critical containment/prevention set for HA-TFC-002.
Implementation | DO-SAFE-001/002/010/011/012 (DOL-1200) plus DO-MECH-005, DO-PCB-010, DO-EE-015, DO-REG-006 are the Implementation References; CMP-001/002/006/007/008/009/013/033/034/037 are the allocated components.
Cross-module ties | HA-TFC-001 ← HA-ENRG-001 (self-heat is the control-introduced risk); HA-TFC-002/003 ← HA-ENRG-002/003 (initiating cell events → propagation/offgas); HA-TFC-004 ← HA-ENRG-003/007 (leak); HA-TFC-005 → HA-USE-010 (cleaning misuse) + RCR-057 (record-integrity aggregate); HA-TFC-006 → HA-ENV-003 + HA-ELEC-006/007 (condensation → electronics fault); HA-TFC-004 corrosion → HA-ELEC-006.
Downstream | the S5 record HA-TFC-002 and the S4 records HA-TFC-003 become FTA top events in Stage 11 (ERFKL) and drive deeper battery/enclosure abuse DFMEA; all PV-/VAL- references are Evidence-Pending and feed Stages 13–14 (UN 38.3, fire-enclosure/offgas, condensation, touch-temperature).
8. Residual-risk / closure notes
Top transport concern | HA-TFC-002 (thermal-runaway propagation → fire in an occupied aircraft cargo hold, S5). It can never be classified broadly acceptable (HAL-2200 §6.1); after prevention (RCR-002/004), containment (RCR-008 fire enclosure), and UN 38.3 / IATA-DGR evidence it sits in region A with EM-ENG-001 + EM-ENG-006 benefit-risk sign-off, transport disclosure, and post-market monitoring. Its offgas-inhalation companion HA-TFC-003 (S4) closes on the same abuse/offgas test.
Self-heat trade-off | HA-TFC-001 exists only because the HA-ENRG-001 cold-charge control may use a self-heat element; it is a control-introduced risk (HAL-2700) and is conditional on self-heat adoption (CMP-013 feasibility Evidence-Pending).
Chemical/condensation cluster | HA-TFC-004/005/006/007/008 are S2–S3 and reduce to region A/B via material selection, conformal coat/seal, and BIST screening; their material contribution is to the record-integrity aggregate (RCR-057) and reliability, not to the physical S5 set.
Release gate (HAL-2200 §8) | no R region thermal/fire/chemical risk may remain open at release; the S5/S4 residuals (HA-TFC-002/003) carry a benefit-risk statement and residual-risk sign-off, and every Evidence-Pending probability has a named closing test (UN 38.3, fire-enclosure/offgas, touch-temperature, condensation, material-compatibility) listed above.
Overall contribution | the thermal/fire/chemical module's residual is dominated by the occupied-aircraft fire (HA-TFC-002); it is the principal transport-safety input to the overall-residual-risk evaluation (HAL-2900) and the benefit-risk determination in the Risk-Management Summary.
HAL-900 (HA-ENV) — Environmental Hazards
Master Hazard Analysis Library — SCM · Stage 10 (ORCH-017) · Owner: Environmental/Reliability Engineer with System-Safety. Environmental hazards of a device that lives inside insulated vaccine shippers across −30 °C…+60 °C, humidity, condensation, transport vibration, altitude, and multi-trip reuse — both hazards to the device (that then create an information harm) and the device's adverse effect on the environment. Harm-and-exposure-first (ISO 14971), NOT an FMEA.
1. Purpose
Identify environmental exposures that either damage the device's record-integrity function (→ wrong disposition) or create a physical hazard, plus the device's end-of-life environmental effect. The −30 °C extreme is the governing environmental stressor and is coupled to the D3 energy tension.
2. Scope
Operating −30 °C…+60 °C; storage/transport per IEC 60068; humidity/condensation from cold↔warm cycling; IP-rated ingress; transport vibration/shock; altitude (aircraft cargo); dust/salt/UV over reuse life; and Li-battery/e-waste at disposal. Excludes the passive shipper's thermal performance (context only) and the device's active thermal/self-heat hazards (HAL-800).
3. Applicable standards & references
ISO 14971 (environmental contributing conditions); IEC 60068 (environmental test methods); IEC 60529 (IP ingress); UN 38.3 (altitude/thermal for Li transport); WHO PQS E006 / WHO-UNICEF EVM (freezing prevalence in the cold chain). Cited by designation; revisions to confirm.
4. Governance & method
Environmental stressors are treated as contributing conditions that raise P1 on the information-harm pathway (HA-BND-003) or create a physical hazard. Severity is fixed by the resulting outcome (usually the wrong-disposition anchor, S4). Cold-related probabilities are Evidence-Pending, closed by the 90-day cold-endurance (PRD-VER-004), cold-display, condensation-cycling, IEC 60068, and RF-in-container tests. Severity-first Action Priority (HAL-2200).
5. Categories / questions
Temperature extremes (cold governing) · humidity/condensation · ingress/immersion · vibration/shock (transport) · altitude/pressure · dust/salt/UV degradation over reuse life · device→environment (disposal). Each category asks: does this stressor corrupt the record/verdict function, damage the device, or harm the environment, and through what exposure to which perspective?
6. Hazard Analysis Register (HA-ENV records)
HA-ENV-001 — Deep cold (−30 °C): freeze excursion of a freeze-sensitive vaccine missed
Hazard ID (library) | HAZ-ENV / HAZ-FUNC
Product/Config | SCM all SKUs (reusable) · Lifecycle Phase: Normal operation (monitoring, transit) · System State: Sleep/Sampling (armed) · Operational Mode: Duty-cycled monitoring
Perspective | Beneficiary/patient (via wrong disposition); programme
Source | UN-SAFE-002; PRD-SAFE-002; STN-002; WHO-UNICEF freezing literature (analogous/historical)
Initiating Cause/Event | Product zone drops below the freeze threshold of a freeze-sensitive vaccine (e.g., adjuvanted vaccines damaged by freezing) during deep-cold transit.
Contributing Conditions | −30 °C ambient; sensor/probe accuracy and time-constant drift at cold extreme; detection thresholds/hysteresis mis-set for freeze; historically freezing excursions are common and under-recognised in the cold chain.
Foreseeable Sequence of Events | Freeze event occurs → cold degrades sensor response / freeze branch of detection under-sensitive or absent → excursion not declared → record shows a clean trip.
Hazard | Undetected freeze excursion (a two-sided-detection gap).
Hazardous Situation | A freeze-damaged (potency-lost) vaccine lot sitting at receipt believed acceptable.
Sequence to Harm | User accepts on a clean record → freeze-damaged doses administered → loss of potency.
Harm/Loss | Beneficiaries not protected → preventable disease (freeze damage is often invisible and irreversible).
Severity | S4 (Critical) — public-health dimension; a lot of freeze-sensitive doses administered while compromised; S5 at campaign scale.
Probability (P1×P2) | P1=3 (Occasional — freezing common; Basis: historical/analogous WHO-UNICEF) × P2=4 (Probable — freeze damage invisible so undetected → administered) ⇒ High. Basis: analogous field data; freeze-branch performance at −30 °C Evidence-Pending (detection-logic + cold test).
Detectability/Warning | Low without a validated freeze branch — freeze damage is not visually detectable, so the device is the primary detector.
Initial Risk | High (R).
Risk Acceptability | Unacceptable without control (two-sided detection is a safety requirement, not a feature).
Risk-Control Strategy | 1 inherent (two-sided detection) + 2 protective (sensor plausibility).
Risk-Control Requirement | RCR-019 (two-sided detection, freeze reliability = heat) + RCR-017 (−30 °C operation) + RCR-020 (probe/sensor plausibility).
Implementation Reference | DO-SAFE-004 (→ DO-FW-006 detection + DO-EE-013 probe front-end); CMP-039/019/020.
Verification Reference | PV-ENV-001 (Evidence-Pending) — detection-logic test with profiled freeze excursions at −30 °C; sensor cold-accuracy test.
Validation Reference | VAL-DET-001 (Evidence-Pending) — two-sided detection accuracy validation with freeze-sensitive product profiles.
Residual Probability | Medium (freeze branch validated) · Residual Severity: S4 (unchanged) · Residual Risk: Medium (A) — ALARP; residual driven by sensor cold-accuracy limits.
New Risks Introduced | Over-sensitive freeze branch → false rejects (alarm fatigue) → HA-USE-002 / RCR-047 balance (→ HA-NEW-002).
Overall Residual Contribution | Part of the missed-excursion accumulation (HAL-2900) — a top public-health contributor.
Status | Open — Evidence-Pending (detection + cold test).
Owner | Firmware/Detection + Environmental Engineer.
Rationale/Evidence | UN-SAFE-002; PRD-SAFE-002; fault-management (SYS-150); WHO-UNICEF freezing prevalence (analogous).
Assumptions | Product-stability freeze thresholds supplied per WHO PQS profiles; sensor meets cold-accuracy class (Evidence-Pending).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-002 — Cold makes the OLED slow/illegible → verdict misread
Hazard ID (library) | HAZ-ENV / HAZ-USE
Product/Config | SCM all SKUs · Lifecycle Phase: Receipt/disposition · System State: Provisioning/Arming or query · Operational Mode: HMI read
Perspective | Last-mile nurse (vulnerable/low-training user); beneficiary
Source | UN-SAFE-005; component-list CMP-030 (OLED cold response Evidence-Pending); IEC 62366-1
Initiating Cause/Event | OLED response time and contrast degrade at −30 °C, taking seconds to update or displaying a ghosted/partial image.
Contributing Conditions | Deep cold; gloved, hurried, low-light receipt; low literacy / non-primary language; user reads the display once and acts.
Foreseeable Sequence of Events | User wakes the device to read the verdict → display is slow/ghosted → user reads a stale or misinterpreted state.
Hazard | Illegible/misleading verdict presentation in cold.
Hazardous Situation | A compromised lot displayed ambiguously and read as "accept."
Sequence to Harm | User accepts a rejected/quarantine lot → compromised doses administered.
Harm/Loss | Beneficiaries not protected → preventable disease.
Severity | S4 (Critical) — wrong-disposition anchor.
Probability (P1×P2) | P1=3 (Occasional — cold OLED behaviour; Basis: judgment, CMP-030 Evidence-Pending) × P2=3 (Occasional — legible-verdict controls reduce misread) ⇒ Medium. Basis: judgment + summative usability (Evidence-Pending).
Detectability/Warning | Medium — a confirmatory read/hold can catch it if designed in.
Initial Risk | Medium–High (R at worst-case cold).
Risk Acceptability | Reduce as far as practicable; challenge any warnings-only mitigation.
Risk-Control Strategy | 1 inherent (display technology + legible, language-independent verdict).
Risk-Control Requirement | RCR-043 (legible, language-independent verdict at −30 °C within responsiveness budget).
Implementation Reference | DO-SAFE-013 (→ DO-HF-002/005 + DO-EE-008); CMP-030/032.
Verification Reference | PV-ENV-002 (Evidence-Pending) — cold-display response/legibility test at −30 °C.
Validation Reference | VAL-USE-001 (Evidence-Pending) — summative usability in cold/gloved/low-light conditions.
Residual Probability | Low–Medium (display qualified + confirmatory read) · Residual Severity: S4 · Residual Risk: Medium (A) — ALARP.
New Risks Introduced | A confirmatory-hold step adds workflow time → minor use burden (→ HA-NEW-006).
Overall Residual Contribution | Contributes to the use-error accumulation (HAL-2900).
Status | Open — Evidence-Pending (OLED cold response, usability).
Owner | Human-Factors + Electrical Engineer.
Rationale/Evidence | CMP-030 (OLED cold Evidence-Pending); UN-SAFE-005; ties HA-USE-001.
Assumptions | Display technology can be selected/qualified to meet cold responsiveness (Evidence-Pending).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-003 — Condensation cycling (cold↔warm) → ingress → record-electronics fault
Hazard ID (library) | HAZ-ENV
Product/Config | SCM all SKUs · Lifecycle Phase: Transport/receipt/reuse cycling · System State: any · Operational Mode: any
Perspective | Beneficiary (via record loss); property (device)
Source | DO-SAFE-011; IEC 60529/60068; component-list CMP-002/034
Initiating Cause/Event | Repeated cold-to-warm transitions (deep cold transit → warm receipt) drive condensation inside the enclosure over many reuse cycles.
Contributing Conditions | Reuse loop (D1) multiplies cycles; seal compression-set over life; conformal-coat keep-outs at sensor/antenna/contact areas.
Foreseeable Sequence of Events | Moisture condenses on the PCB → over cycles causes leakage/corrosion/intermittent fault on record-integrity nets.
Hazard | Internal condensation/corrosion of record-integrity electronics.
Hazardous Situation | A device that intermittently loses or corrupts samples while appearing to monitor.
Sequence to Harm | Silent record gap/corruption → wrong disposition on an incomplete/incorrect record.
Harm/Loss | Beneficiaries not protected (false accept) or a valid lot wasted (false reject).
Severity | S3 (Serious) — typically a limited-quantity or single-trip record loss; S4 if it produces a false accept.
Probability (P1×P2) | P1=3 (Occasional over reuse life; Basis: judgment, reuse cycling) × P2=2 (Remote — honest-gap + storage-integrity controls catch most) ⇒ Medium/Low. Basis: judgment; closed by condensation-cycle test.
Detectability/Warning | Medium — self-check/plausibility and honest-gap annunciation flag many faults.
Initial Risk | Medium (A).
Risk Acceptability | ALARP — protective seal + conformal coat qualified over reuse life.
Risk-Control Strategy | 2 protective (seal + conformal coat) + design-enforced honest gap.
Risk-Control Requirement | RCR-015 (conformal coat + seal protect record electronics across cycling) + RCR-018 (condensation tolerance) + RCR-022 (honest gap).
Implementation Reference | DO-SAFE-011 (→ DO-MECH-005 seal + DO-PCB-010 conformal coat); CMP-002/034.
Verification Reference | PV-ENV-003 (Evidence-Pending) — condensation/cold-cycle + IP + cleanliness test over reuse-life cycles.
Validation Reference | N/A (covered by reliability/verification).
Residual Probability | Low · Residual Severity: S3 (S4 case mitigated by honest gap) · Residual Risk: Low–Medium (A/B).
New Risks Introduced | Conformal-coat keep-outs at sensors could shift accuracy (→ HA-NEW-007).
Overall Residual Contribution | Small; part of the record-integrity accumulation.
Status | Open — Evidence-Pending (cycle test).
Owner | Mechanical/PCB + Environmental Engineer.
Rationale/Evidence | DO-SAFE-011; ties HA-TFC-006, HA-ELEC-007.
Assumptions | Seal compression-set and coat coverage meet reuse-life target (Evidence-Pending).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-004 — Humidity / immersion (IP breach) → corrosion/short → data loss
Hazard ID (library) | HAZ-ENV
Product/Config | SCM all SKUs · Lifecycle Phase: Transport/cold store/cleaning · System State: any · Operational Mode: any
Perspective | Beneficiary (via record loss); property
Source | IEC 60529; DO-SAFE-011; UN-1100
Initiating Cause/Event | Liquid water/melt-water or cleaning fluid ingress past a degraded seal or unsealed port.
Contributing Conditions | Cold-store meltwater; wash-down cleaning between trips; port/probe seal wear over reuse; IP rating margin at cold.
Foreseeable Sequence of Events | Water enters → corrosion/short on the PCB → intermittent or total record-electronics failure.
Hazard | Ingress-induced electronics failure.
Hazardous Situation | Device fails to record/verdict correctly while believed healthy.
Sequence to Harm | Missing/corrupt record → wrong disposition.
Harm/Loss | Beneficiary harm (false accept) or lot waste (false reject).
Severity | S3 (Serious); S4 if false accept.
Probability (P1×P2) | P1=2 (Remote — sealed IP design; Basis: judgment/IEC 60529) × P2=2 (Remote — honest-gap catches) ⇒ Low. Basis: judgment; closed by IP + cleaning test.
Detectability/Warning | Medium — self-check + honest gap.
Initial Risk | Medium (A) (S3/S4 with low probability).
Risk Acceptability | ALARP — IP rating + seal life verified.
Risk-Control Strategy | 1 inherent (sealed IP enclosure) + 2 protective (conformal coat).
Risk-Control Requirement | RCR-011 (enclosure/seal integrity) + RCR-015 (conformal coat).
Implementation Reference | DO-SAFE-011 (→ DO-MECH-005); CMP-001/002/004/005/034.
Verification Reference | PV-ENV-004 (Evidence-Pending) — IP ingress + cleaning-fluid + seal-life test.
Validation Reference | N/A.
Residual Probability | Very Low · Residual Severity: S3 · Residual Risk: Low (B).
New Risks Introduced | None significant.
Overall Residual Contribution | Negligible after control.
Status | Open — Evidence-Pending (IP test).
Owner | Mechanical + Environmental Engineer.
Rationale/Evidence | IEC 60529; ties HA-TFC-005 (cleaning), HA-ELEC-006/007.
Assumptions | Selected IP rating adequate for cold-store/cleaning exposure (confirm).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-005 — Transport vibration/shock → fatigue / cell damage / spurious shock events
Hazard ID (library) | HAZ-ENV / HAZ-MECH
Product/Config | SCM all SKUs · Lifecycle Phase: Transport (air/ground) · System State: Sleep/Sampling (armed) · Operational Mode: Monitoring
Perspective | Air-cargo crew/property (cell); beneficiary (spurious shock/record)
Source | IEC 60068; UN 38.3; component-list CMP-007/022
Initiating Cause/Event | Sustained transport vibration and handling shock.
Contributing Conditions | Long multi-modal trips; solder-joint/fastener fatigue over reuse; accelerometer threshold tuning.
Foreseeable Sequence of Events | Vibration fatigues joints/fasteners or damages the cell (ties HA-ENRG-003); or generates spurious shock events that clutter the record.
Hazard | Vibration/shock-induced hardware degradation or spurious event generation.
Hazardous Situation | A device with a latent joint/cell defect, or a record cluttered with false shock alerts causing fatigue.
Sequence to Harm | Latent defect → later intermittent record loss → wrong disposition; or spurious shock alerts → alarm fatigue → real event ignored; or cell damage → thermal event (HA-ENRG-003).
Harm/Loss | Beneficiary harm (record loss / ignored alert) or injury (cell thermal event).
Severity | S3 (Serious) for record/fatigue paths; S4/S5 where it feeds cell damage (HA-ENRG-003).
Probability (P1×P2) | P1=3 (Occasional — transport vibration certain, damage occasional; Basis: IEC 60068/analogous) × P2=2 (Remote — retention + alarm consolidation) ⇒ Medium/Low. Basis: judgment; closed by vibe/shock test.
Detectability/Warning | Medium — BIST/self-check + shock classification.
Initial Risk | Medium (A).
Risk Acceptability | ALARP.
Risk-Control Strategy | 1 inherent (robust retention/mounts) + 2 protective (shock classification/consolidation).
Risk-Control Requirement | RCR-004 (cell retention) + RCR-012 (fastener retention) + RCR-047 (shock/alarm consolidation).
Implementation Reference | DO-SAFE-010 (→ DO-MECH-012); DO-FW (shock classification, CMP-039); CMP-007/022/034.
Verification Reference | PV-ENV-005 (Evidence-Pending) — IEC 60068 vibration/shock + reuse-life fatigue; UN 38.3 vibration.
Validation Reference | N/A.
Residual Probability | Low · Residual Severity: S3 (S4/S5 via HA-ENRG-003 controlled separately) · Residual Risk: Low–Medium (A/B).
New Risks Introduced | Aggressive shock filtering could mask a real damaging shock (→ HA-NEW-008).
Overall Residual Contribution | Small.
Status | Open — Evidence-Pending (vibe/shock test).
Owner | Mechanical/Reliability + Environmental Engineer.
Rationale/Evidence | IEC 60068; UN 38.3; ties HA-ENRG-003, HA-MECH-006/009, HA-USE-002.
Assumptions | Transport vibration profile per IATA/ISTA representative (confirm).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-006 — Dust/salt/UV degradation over reuse life → seal/label degradation
Hazard ID (library) | HAZ-ENV
Product/Config | SCM all SKUs · Lifecycle Phase: Whole reuse life · System State: any · Operational Mode: any
Perspective | Technician/beneficiary; property
Source | IEC 60068; UN-1100
Initiating Cause/Event | Cumulative exposure to dust, salt fog (coastal/marine transport), and UV over a multi-year reuse life.
Contributing Conditions | Reuse (D1); outdoor/handling exposure; label ink/adhesive and elastomer-seal ageing.
Foreseeable Sequence of Events | Seal elastomer hardens/cracks and label becomes unreadable → later ingress (HA-ENV-004) or an illegible serial/warning/UN-mark.
Hazard | Age/exposure degradation of seal and markings.
Hazardous Situation | A device with a compromised seal or an unreadable transport/warning label deployed on a later trip.
Sequence to Harm | Ingress → record loss → wrong disposition; or missing UN 38.3/small-parts mark → transport/bystander hazard.
Harm/Loss | Beneficiary harm (record) or logistics/bystander hazard.
Severity | S2 (Minor) typical; S3 where it enables ingress or an unreadable safety mark.
Probability (P1×P2) | P1=3 (Occasional over years; Basis: materials ageing judgment) × P2=2 (Remote — service inspection catches) ⇒ Medium/Low. Basis: judgment; closed by accelerated-ageing + reuse-life test.
Detectability/Warning | High — visible at depot service/pre-arm.
Initial Risk | Low–Medium (A/B).
Risk Acceptability | ALARP — durable materials + service inspection.
Risk-Control Strategy | 1 inherent (UV/salt-resistant materials, durable marking) + 4 process (service inspection).
Risk-Control Requirement | RCR-011 (seal integrity over life) + RCR-050 (service restores/inspects seal & marking).
Implementation Reference | DO-MECH-005 (seal/materials); DO-LBL (durable marking); DO-SVC (inspection).
Verification Reference | PV-ENV-006 (Evidence-Pending) — accelerated UV/salt-fog/dust + reuse-life ageing.
Validation Reference | N/A.
Residual Probability | Low · Residual Severity: S2 · Residual Risk: Low (B).
New Risks Introduced | None significant.
Overall Residual Contribution | Negligible.
Status | Open — Evidence-Pending (ageing test).
Owner | Materials/Mechanical + Service Engineer.
Rationale/Evidence | IEC 60068; ties HA-ENV-004, HA-MFG-006 (marking), HA-SVC-003.
Assumptions | Materials selected for outdoor/marine exposure (confirm).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-007 — Altitude / low pressure (aircraft cargo) → cell venting / seal stress
Hazard ID (library) | HAZ-ENV / HAZ-THERM
Product/Config | SCM all SKUs · Lifecycle Phase: Air transport · System State: Sleep (armed/unarmed) · Operational Mode: Monitoring/shipping
Perspective | Air-cargo crew/aircraft (property); mission
Source | UN 38.3 (altitude simulation, Test T.1); D4
Initiating Cause/Event | Low cabin/cargo pressure at altitude imposes differential pressure on the cell and enclosure seal.
Contributing Conditions | Aircraft cargo hold pressure; a cell with a marginal vent/seal; rapid pressure change.
Foreseeable Sequence of Events | Pressure differential stresses the cell seal/vent and the enclosure seal → in the worst case cell venting or enclosure seal breach.
Hazard | Altitude-induced cell venting / seal breach.
Hazardous Situation | A stressed cell or breached enclosure in an occupied cargo environment.
Sequence to Harm | Cell venting → thermal/offgas event (ties HA-TFC-002); or seal breach → later ingress.
Harm/Loss | Transport-safety hazard (fire/offgas) or record loss.
Severity | S3 (Serious) for seal path; S4/S5 where it feeds a thermal event (HA-TFC-002).
Probability (P1×P2) | P1=2 (Remote — UN 38.3-qualified cell/enclosure; Basis: UN 38.3 T.1) × P2=2 (Remote) ⇒ Low. Basis: UN 38.3 test (closes probability).
Detectability/Warning | Low in flight; verified pre-transport by UN 38.3.
Initial Risk | Medium (A) (severity-driven).
Risk Acceptability | ALARP — UN 38.3 compliance mandatory.
Risk-Control Strategy | 2 protective (qualified cell/enclosure) + 4 process (transport compliance).
Risk-Control Requirement | RCR-007 (UN 38.3 / IATA compliance) + RCR-011 (enclosure/seal integrity).
Implementation Reference | DO-EE-015 (cell/pack transport qualification); DO-REG-006; DO-MECH-005.
Verification Reference | PV-ENV-007 (Evidence-Pending) — UN 38.3 Test T.1 altitude + thermal (T.2).
Validation Reference | N/A.
Residual Probability | Very Low · Residual Severity: S3 (S4/S5 path controlled via HA-TFC-002) · Residual Risk: Low (B).
New Risks Introduced | None significant.
Overall Residual Contribution | Negligible after UN 38.3.
Status | Open — Evidence-Pending (UN 38.3).
Owner | Battery + Regulatory Engineer.
Rationale/Evidence | UN 38.3; D4; ties HA-TFC-002, HA-ENRG-003.
Assumptions | Cell passes UN 38.3 T.1/T.2 (Evidence-Pending — test summary).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
HA-ENV-008 — Device adversely affects the environment (Li-battery / e-waste at end of life)
Hazard ID (library) | HAZ-ECOL
Product/Config | SCM all SKUs · Lifecycle Phase: Disposal/recycling · System State: decommissioned · Operational Mode: N/A
Perspective | Environment; bystander; programme (sustainability)
Source | MR-1700 (sustainability); UN-1800; WEEE/battery-directive concepts
Initiating Cause/Event | End-of-life devices (fleet scale, D1) disposed without battery removal/recycling.
Contributing Conditions | Reusable fleet reaches end of cycle-life together; limited recycling infrastructure in some deployment regions.
Foreseeable Sequence of Events | Devices land-filled/incinerated with cells intact → lithium/heavy-metal release; or damaged cell in a waste stream ignites.
Hazard | Lithium/e-waste environmental release; waste-stream fire.
Hazardous Situation | A fleet's worth of cells entering an uncontrolled waste stream.
Sequence to Harm | Environmental contamination / waste-fire; bystander exposure.
Harm/Loss | Environmental damage; possible fire injury; programme sustainability/reputational loss.
Severity | S3 (Serious) — environmental dimension at fleet scale.
Probability (P1×P2) | P1=3 (Occasional — disposal governance varies; Basis: judgment) × P2=2 (Remote with instructions/take-back) ⇒ Medium/Low. Basis: judgment.
Detectability/Warning | Low at the point of harm; controlled upstream by instructions/take-back.
Initial Risk | Medium (A).
Risk Acceptability | ALARP — design-for-disassembly + disposal instructions + take-back.
Risk-Control Strategy | 1 inherent (designed for battery removal/recycling) + 3 information + 4 process (take-back).
Risk-Control Requirement | RCR-052 (safe end-of-life: battery removal/recycling instructions + design support).
Implementation Reference | DO-SVC (disposal procedure); DO-LBL (recycling marking); DO-MECH (disassembly).
Verification Reference | PV-ENV-008 (Evidence-Pending) — disposal/disassembly procedure review; recyclability assessment.
Validation Reference | N/A.
Residual Probability | Low · Residual Severity: S3 · Residual Risk: Low (B/A).
New Risks Introduced | Battery removal at end-of-life adds a handling step (short/spark risk) → HA-SVC-007 controls.
Overall Residual Contribution | Small; tracked with HA-MPE-005 (mission/environmental).
Status | Open — Evidence-Pending (procedure/recyclability).
Owner | Sustainability/Service Engineer.
Rationale/Evidence | MR-1700; ties HA-MPE-005, HA-SVC-007, HA-CYBER-008 (data remanence).
Assumptions | A take-back/recycling channel is available in target regions (confirm).
Review History | 2026-07-19 v1.0 initial (ORCH-017).
7. Traceability Notes
Environmental hazards feed and are fed by: HAL-500 (HA-ENRG-003 cell crush; HA-ENRG-006 dead-in-cold), HAL-800 (HA-TFC-002 aircraft thermal event; HA-TFC-006 condensation), HAL-700 (HA-ELEC-007/010 condensation/ brown-out), HAL-1000 (HA-SW-001 detection; HA-SW-013 silent gap), HAL-1500 (HA-USE-001 misread; HA-USE-002 alarm fatigue), and HAL-1800 (HA-MPE-005 environmental). Cold is the governing stressor and couples to the D3 energy tension (HA-ENRG-006) and the antenna/RF-in-container tension (HA-COMM-003). Controls: RCR-004, 007, 011, 012, 015, 017, 018, 019, 020, 043, 047, 050, 052.
8. Residual-Risk / Closure Notes
The dominant environmental risk is HA-ENV-001 (missed freeze excursion) — a top public-health contributor requiring validated two-sided detection (RCR-019, DO-SAFE-004) and cold sensor accuracy; it is Open / Evidence-Pending until the detection-logic and cold tests close it. Cold-display legibility (HA-ENV-002) and cold endurance (coupled to HA-ENRG-006) are the other Evidence-Pending drivers. All ingress/condensation/ vibration/altitude risks reduce to Low–Medium (A/B) after their named IEC 60068 / IEC 60529 / UN 38.3 tests. S4/S5 environmental-coupled situations (HA-ENV-001 freeze; HA-ENV-005→HA-ENRG-003 cell damage) are handed to Stage 11 (ERFKL) as FTA top-event contributors.
Hazard Traceability Matrix — Smart Cold Chain Monitor (SCM)
Stage 10 (ORCH-017) · Owner: EM-ENG-012 (Traceability) with System-Safety and EM-ENG-006. The authoritative end-to-end chain for the SCM's top hazards: from each hazard record through its hazardous situation, harm, initial risk, the risk-control requirement that controls it, the design output that implements it, the verification and validation that will prove it, the projected residual risk, and the field-monitoring obligation. This matrix is the single-row-per-hazard view the risk-management file is audited against; the full 34-field records live in the HAL domain modules. All PV-/VAL- and residuals are Evidence-Pending (Stages 13–14). Severity-first per CLAUDE.md rule 3; regions/AP per HAL-2200.
---
1. How to read this matrix
• HA ID / Source — the hazard record and its originating driver (concept decision D1–D4, SYS/architecture, standard, or upstream module).
• Initial Risk — Severity × Combined-probability band → region/AP before control (HAL-2000/2100/2200).
• RCR / HA-RCR — the risk-control requirement (HAL-2400). Design Impl (DO-) — the controlled design output (DOL-1200 DO-SAFE + discipline outputs).
• Verification (PV-) — planned product verification (Stage 13); Validation (VAL-) — planned validation (Stage 14); fault-injection for safety controls, independent adversarial for security (HAL-2600).
• Residual Risk — projected region/AP after control (HAL-2800); all Evidence-Pending.
• Field Monitoring — the post-market signal that keeps the residual valid (ISO 14971 §9; Stage 17).
Severity key | S5 catastrophic · S4 critical · S3 serious · S2 minor · S1 negligible. Regions: R unacceptable · A ALARP (Medium AP; A-High for S5) · B broadly acceptable (Low AP).
2. Traceability matrix — top hazards
2.1 Physical: energy / battery / thermal
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-ENRG-001 | D3; cold charge | Cold cell charged → plating → latent short | Thermal runaway/fire (S4/S5) | R High | RCR-001/002 | DO-SAFE-001→DO-EE-007+DO-FW-005 | PV-ENRG-001 (cold-charge FI) + UN 38.3 | — | A Med (S4×PVL) | Charge-inhibit & thermal-event field reports
HA-ENRG-002 | IEC 62368-1 | Cell OV/OC/OT/short | Thermal runaway/fire (S4/S5) | R High | RCR-002 | DO-SAFE-002→DO-EE-005+DO-EE-010 | PV-ENRG-002 (protection FI, DO-TST-009) | — | A Med | Protection-trip / return events
HA-ENRG-003 | UN 38.3 | Cell crush/puncture on drop/transport | Thermal event (S4) | R High | RCR-004 | DO-SAFE-010→DO-MECH-012+DO-EE-004 | PV-MECH-004 + UN 38.3 | — | A Med | Drop/damage returns
HA-ENRG-005 | USB-C/PD | Out-of-spec charge source | Overcharge/thermal (S3/S4) | A/R | RCR-003 | DO-SAFE-003→DO-EE-009+DO-EE-010 | PV-ELEC-002 (reverse/OVP/short) | — | B/A | Charger-abuse field reports
HA-ENRG-006 ★ | D3; dead-in-cold | Energy exhausted → silent gap reads uneventful | Wrong disposition (S4) | R High | RCR-006/024/027 | DO-SAFE-012→DO-FW-004+DO-EE-008 | PRD-VER-004 (90-day cold endurance) | — | A Med (energy margin Open) | Field battery/endurance + gap rates
HA-TFC-001 | RCR-005 self-heat | Hot self-heat surface touched | Burn (S3) | A Med | RCR-005/046 | DO-SAFE-001→DO-EE-007 | PV-TFC-001 (touch-temp) | VAL-TFC-001 | B (S3×PVL) | Burn/handling reports (if self-heat adopted)
HA-TFC-002 ★★ | IATA DGR (S5) | Cell fire in occupied aircraft cargo hold | Death/loss of aircraft (S5) | R High | RCR-002/004/007/008 | DO-SAFE-002/010; DO-MECH-005+DO-PCB-010 | PV-TFC-002 (containment/fire-enclosure) + UN 38.3 | — | A-High (S5×PVL) | Mandatory thermal-event monitoring + transport disclosure
HA-TFC-003 | IEC 62133-2 | Toxic offgas in enclosed space | Inhalation injury (S4) | R High | RCR-002/008 | DO-SAFE-002; DO-MECH-005+DO-PCB-010 | PV-TFC-003 (offgas characterisation) | — | A Med | Offgas/venting events
HA-TFC-006 | IEC 60068 | Condensation → record-electronics fault | Wrong disposition (S3) | A Med | RCR-015/018 | DO-SAFE-011→DO-MECH-005+DO-PCB-010 | PV-TFC-006 (condensation cycle+IP) | — | A Med | Ingress/BIST failures over reuse
2.2 Mechanical / electrical / environmental
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-MECH-003 | 16 CFR 1501 | Detachable small part near a child | Choking/strangulation (S1) | A Med | RCR-010 | DO-MECH-006 + DO-LBL | PV-MECH-003 (small-parts/retention) | — | B | Detachment/complaint reports
HA-MECH-004 | drop/ingress | Enclosure crack → ingress → data loss | Wrong disposition (S3) | A Med | RCR-011 | DO-SAFE-010/011→DO-MECH-005/012 | PV-MECH (drop/vibe)+IP | — | A Med | Cracked-enclosure returns
HA-MECH-009 | UN 38.3 | Battery retention fails on shock → crush | Thermal event (S4) | R High | RCR-004 | DO-SAFE-010→DO-MECH-012 | PV-MECH-004 + UN 38.3 | — | A Med | Retention-failure returns
HA-ELEC-004 | IEC 61000 | EMC upset → false/missed excursion | Wrong disposition (S3/S5) | R High | RCR-014 | DO-EE (EMC)+DO-PCB-010 | PV-ELEC-004 (EMC immunity/emissions) | — | A Med | Field EMC-related anomalies
HA-ELEC-008 | latent defect | Latent PCB/solder → intermittent fault → silent gap | Wrong disposition (S3) | A/R | RCR-016/024 | DO-TST (AOI/ICT); DO-SAFE-008 | PV-MFG-008 (AOI/ICT coverage) | — | A Med | Intermittent-fault field rate
HA-ELEC-010 | D3 brown-out | Brown-out during NV write (cold sag) → torn record | Wrong disposition (S3) | A Med | RCR-024 | DO-SAFE-008→DO-FW-010+DO-EE-002 | PV-SW-013 (power-fail-during-write) | — | A Med | Torn-record/gap rate
HA-ENV-001 | freeze-sensitive | Deep-cold freeze excursion missed | Wrong disposition (S4) | R High | RCR-019 | DO-SAFE-004→DO-FW-006+DO-EE-013 | PV-SW-001 (freeze coverage=heat) | VAL-SW-001 | A Med | Missed-freeze field signals
HA-ENV-003 | condensation | Cold↔warm condensation → ingress fault | Wrong disposition (S3) | A Med | RCR-015/018 | DO-SAFE-011 | PV-ENV-003 + PV-TFC-006 | — | A Med | Ingress/BIST failures
HA-ENV-007 | altitude | Aircraft low-pressure → cell venting/seal stress | Thermal/leak (S3/S4) | A/R | RCR-004/007/011 | DO-SAFE-010; DO-MECH-005 | UN 38.3 (altitude) + IP | — | A Med | Altitude-venting events
2.3 Software / detection / verdict / record / time
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-SW-001 ★ | SYS-149/150 | Spoiled dose believed acceptable (missed excursion) | Compromised dose administered (S4/S5) | R High | RCR-019 | DO-SAFE-004→DO-FW-006+DO-EE-013 | PV-SW-001 (profiled excursions, FI) | VAL-SW-001 | A Med | Missed-excursion signals (returns/QA)
HA-SW-002 ★ | SYS-151/016 | Lot presented acceptable on incomplete record | False accept → compromised dose (S4) | R High | RCR-021/022 | DO-SAFE-006→DO-FW-008+DO-CLD-006 | PV-SW-002 (gap-handling FI) | VAL-SW-002 | A Med | Incomplete-record dispositions; override rate
HA-SW-003 | SYS-154/024 | Device monitoring against wrong thresholds | Missed excursion → compromised dose (S4) | R High | RCR-048/045 | DO-APP/DO-CLD; DO-SAFE-007/DO-FW-032 | PV-SW-003 (profile/pre-arm) | VAL-SW-003 | A Med | Wrong-profile/mis-arm rate
HA-SW-004 | SYS-024 | Device monitoring with biased conditioning | Wrong disposition (S4/S5) | R High | RCR-026/045/024 | DO-FW-007; DO-SAFE-007 | PV-SW-004 (calibration integrity/expiry FI) | — | A Med (S5-worst common-cause) | Calibration-drift/escape trends
HA-SW-005 | SYS-118 | Disposition/audit on mis-timed record | Wrong disposition/indefensible (S3) | A Med | RCR-025 | DO-EE (RTC)+DO-FW | PV-SW-005 (clock-drift over cold) | — | A Med | Untrusted-time incidence
HA-SW-006 | SYS-011/013 | Committed excursion, no alert in flight | Delayed/missing intervention (S4) | R High | RCR-029/028/047 | DO-SAFE-009→DO-FW-011+DO-CLD-007/008 | PV-SW-006 (alert-gen+escalation FI) | VAL-SW-006 | A Med | Un-acked/escalated-alert rate
HA-SW-007 | SYS-016/017 | Operator acts on a displayed false accept | Compromised dose (S4) | R High | RCR-021/043/044 | DO-SAFE-006→DO-FW-008+DO-HF-002 | PV-SW-007 (display-integrity FI) | VAL-USE-001 | A Med | Verdict-display anomalies
HA-SW-008 ★ | SYS-007/024 | Device reports valid-looking temps it isn't measuring | False accept → compromised dose (S4) | R High | RCR-020/045 | DO-SAFE-005→DO-FW-007+DO-EE-013 | PV-SW-008 (probe-fault-injection) | — | A Med (+placement remainder) | Probe-fault flag rate
HA-SW-009 | SYS-169/170 | Device believed monitoring but hung | Silent non-monitoring (S3/S4) | A/R | RCR-023/027 | DO-SAFE-007→DO-FW-003/013 | PV-SW-009 (fault-injection+watchdog) | — | A Med | Watchdog-reset/availability rate
HA-SW-010 | SYS-032/139 | Bad update deploys wrong logic / bricks | Fleet unsafe/blind (S4/S5) | R High | RCR-037 | DO-SAFE-014→DO-FW-014/015+DO-CLD-010 | PV-SW-010 (rollback+version-skew) | — | A Med | OTA campaign health / version dist.
HA-SW-013 ★ | SYS-160/168 | Reset/brown-out gap reads as uneventful | False reassurance → compromised dose (S4) | R High | RCR-024/022 | DO-SAFE-008→DO-FW-010+DO-EE-002 | PV-SW-013 (power-fail-during-write) | — | A Med | Silent-gap/torn-record rate
HA-SW-014 | calibration | Firmware coefficient/units error → biased reading | Wrong disposition (S4) | R High | RCR-026 | DO-FW-007 | PV-SW-004 (+static analysis) | — | A Med | Systematic-bias field signals
2.4 Connectivity / cloud
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-COMM-001 ★ | D2 | Shipment held/administered before late alert arrives | Compromised dose (S4) | R High | RCR-028/029/030 | DO-SAFE-009→DO-FW-011+DO-CLD-007/008 | PV-COMM-001 (latency+S&F) | VAL-COMM-001 | A Med | Alert-latency distribution per lane
HA-COMM-002 | SRS-CLOUD-023 | Shipment disposed clean, alert never communicated | Compromised dose (S4) | R High | RCR-028/029/035 | DO-SAFE-009; SYS-176 reconciliation | PV-COMM-002 (delivery-path FI) | VAL-COMM-002 | A Med | Un-acked backlog fleet-wide
HA-COMM-003 ★ | D2/D4 (RF) | Excursion with no live notification (container blocks radio) | Compromised dose (S4) | R High | RCR-031/030 | DO-IF (antenna)+DO-EE; DO-SAFE-009 | PRD-VER-005 (RF-in-container) + PV-COMM-003 | VAL-COMM-003 | A Med (Open until PRD-VER-005) | Per-trip delivery-success rate
HA-COMM-004 | SRS-CLOUD-003 | Cloud misstates excursion presence/timing | Wrong disposition (S3) | A Med | RCR-030/032 | DO-SAFE-009; DO-CLD-007 | PV-COMM-004 (reorder/dup/corrupt) | — | B (S3×PVL) | Ingest-reconciliation anomalies
HA-COMM-005 | SYS-137/118 | Device can't authenticate/resolve session | Fleet blinded (S3/S4) | R High | RCR-025/032/039 | DO-EE(RTC)+DO-FW; DO-CLD-007 | PV-COMM-005 (cert/time/DNS FI) | VAL-COMM-005 | A Med | Auth/handshake/backlog spikes
HA-COMM-006 | D4 roaming | Cross-border coverage gap, no live alert | Compromised dose (S4) | R High | RCR-031/056/030 | DO-IF; DO-MFG+DO-FW (config) | PV-COMM-006 (roaming+config) | VAL-COMM-006 | A Med | Coverage/roaming gap logs
HA-COMM-008 | SYS-134 | Real excursion invisible on its own record | Compromised dose (S4) | R High | RCR-033/055 | DO-CLD (registry); DO-SEC | PV-COMM-008 (=PV-CLOUD-003) | VAL-COMM-008 | A Med | Binding/routing anomalies
HA-CLOUD-001 | SRS-CLOUD | Responder unaware (cloud notify fail) | Compromised dose (S4) | R High | RCR-032/029 | DO-CLD-007/008 | PV-CLOUD-001 (ingest/notify FI) | — | A Med | Notify success rate
HA-CLOUD-002 | rules engine | Excursion not flagged / wrong policy | Wrong disposition (S4) | R High | RCR-036 | DO-CLD (rules/audit) | PV-CLOUD-002 (config-tamper) | — | A Med | Rules-config change audit
HA-CLOUD-004 | DB integrity | Evidentiary record lost/corrupt | Wrong disposition/indefensible (S4) | R High | RCR-032/042/039 | DO-CLD; DO-SAFE-008 | PV-CLOUD (DR/backup-restore) | — | A Med | Record-loss/DR-test results
HA-CLOUD-008 | outage | Fleet blind, alerts undelivered | Fleet blinding (S4/S8) | R High | RCR-039/032 | DO-CLD; DO-SEC | PV-CYBER-009 + DR test | — | A Med | Cloud availability/backlog
2.5 Cybersecurity-induced safety (independent adversarial verification)
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-CYBER-001 ★ | threat model | False "clean" record from injected data | Compromised dose (S4) | R High | RCR-034 | DO-SEC; DO-SAFE-008 | PV-CYBER-001 (integrity/pen-test, IV) | — | A Med | Integrity-failure events
HA-CYBER-002 ★ | threat model | Excursion hidden by muted/blocked alarm | Compromised dose (S4) | R High | RCR-035/029 | DO-SEC; DO-SAFE-009 | PV-CYBER-002 (suppression pen-test, IV) | — | A Med | Alarm-suppression anomalies
HA-CYBER-003 | threat model | Detection defeated by altered thresholds | Wrong disposition (S4) | R High | RCR-036 | DO-SEC; DO-CLD (audit) | PV-CYBER-003 (config-tamper, IV) | — | A Med | Unauthorised-config attempts
HA-CYBER-005 ★ | supply chain | Malicious firmware → fleet unsafe/blind | Fleet-scale harm (S4/S5) | R High | RCR-037/038 | DO-SAFE-014→DO-FW-014/015; DO-SEC-006 | PV-CYBER-005 (OTA pen-test, IV) | — | A Med/High | Campaign health + signing audit
HA-CYBER-006 | debug/spoof | Forged records / cloned device | Repudiation/mis-route (S3/S4) | R High | RCR-038/055 | DO-SEC-006; DO-MFG | PV-CYBER-006 (secure-boot/debug-lock, IV) | — | A Med | Identity/clone anomalies
HA-CYBER-008 | disposal | Key/data remanence at decommission | Repudiation/key compromise (S3) | A Med | RCR-052/038 | DO-SVC; DO-SEC | Disposal-procedure + key-erasure review | — | A Med | Disposal-compliance audit
HA-CYBER-009 ★ | ransomware | Fleet blind + records altered/withheld | Fleet blinding (S8/S4) | R High | RCR-039/032 | DO-CLD; DO-SEC | PV-CYBER-009 (cloud review+DR, IV) | — | A Med | Cloud security/DR monitoring
2.6 Data & decision / human factors
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-DATA-001 ★ | ALCOA+ | Decision on numbers not representing product zone | Wrong disposition (S4/S5) | R High | RCR-019/020/026/034 | DO-SAFE-004/005; DO-SEC | PV-DATA-001 (cal-coverage+probe FI) | VAL-USE-001 | A Med | Systematic-bias/cal trends
HA-DATA-002 | ALCOA+ | Decision on another shipment's data | Wrong disposition (S4) | R High | RCR-033/055 | DO-CLD (registry); DO-SEC | PV-DATA-002 (=PV-CLOUD-003) | VAL-USE-008 | A Med | Mis-attribution anomalies
HA-DATA-004 ★ | IEC 62366-1 | Incomplete record presented as clean pass | False accept (S4) | R High | RCR-040/021/022 | DO-SAFE-006+DO-HF-002 | PV-DATA-004 (gap-exposure) | VAL-USE-004 | A Med | Incomplete-record accept/override rate
HA-DATA-005 | IEC 62366-1 | Chart/summary read as opposite of true status | Wrong disposition (S3/S4) | R High | RCR-040/043 | DO-HF-007+DO-HF-002 | PV-DATA-005 (visualisation review) | VAL-USE-005 | A Med | Dashboard-misread signals
HA-DATA-006 | ALCOA+ | Decision/audit on record missing evidence | Wrong disposition (S4) | R High | RCR-024/042/032 | DO-SAFE-008; DO-CLD | PV-DATA-006 (power-fail+export) | — | A Med | Silent-loss/export-integrity
HA-DATA-007 | ALCOA+ | Estimate believed to be a measurement | Wrong/limited disposition (S3) | A Med | RCR-041/042 | DO-CLD/DO-HF | PV-DATA-007 (provenance review) | VAL-USE-005 | A Med | Estimate-vs-measured comprehension
HA-USE-001 ★ | IEC 62366-1 | Verdict misread (cold/gloved/low-light/language) | Wrong disposition (S4) | R High | RCR-043 | DO-SAFE-013+DO-HF-002/005 | PV-USE-001 (cold-display) | VAL-USE-001 | A Med | Misread/complaint rate
HA-USE-002 ★ | IEC 62366-1 | Real excursion alert muted (alarm fatigue) | Not acted on → compromised dose (S4) | R High | RCR-047/029 | DO-SAFE-004; DO-CLD-003 | PV-USE-002 (alarm-mgmt) | VAL-USE-002 | A Med | Muted-alert / alarm-volume rate
HA-USE-004 ★ | SYS-024 | Unmonitored trip trusted (forgot-to-arm/void) | False reassurance (S4) | R High | RCR-045/040 | DO-SAFE-007/DO-FW-032 | PV-SW-003 (pre-arm/void) | VAL-USE-004 | A Med | Mis-arm/void rate
HA-USE-005 | HF | Probe left in air/against wall → unrepresentative | Wrong disposition (S4) | R High | RCR-020/045 | DO-SAFE-005; DO-HF | PV-SW-008 (int-vs-ext divergence) | VAL-USE-005 | A Med (HF remainder) | Placement-error signals
HA-USE-007 ★ | IEC 62366-1 | Reject overridden (automation bias/stock-out fear) | Compromised dose (S4) | R High | RCR-044/040 | DO-HF; DO-CLD (audit) | PV-DATA-008 (recommendation) | VAL-USE-007 | A Med | Override rate + reasons
HA-USE-009 | HF | Cold charge-inhibited device discarded as dead | Stock-out / lost monitoring (S3) | A Med | RCR-046 | DO-SAFE-013→DO-HF-005+DO-EE-008 | PV-USE (state-legibility) | VAL-USE | A Med | Discard/return-as-dead rate
2.7 Service / manufacturing / mission
HA ID | Source | Hazardous Situation | Harm (Sev) | Initial | RCR | Design Impl (DO-) | Verification (PV-) | Validation | Residual | Field Monitoring
HA-SVC-001 ★ | D1 reuse | Depot recal escape → systematic bias | Wrong disposition (S4) | R High | RCR-049/026/050 | DO-FW; DO-SVC | PV-SVC-001 (interval enforcement) | — | A Med | Depot recal-escape audit
HA-SVC-002 | depot | Wrong firmware → wrong logic/region | Missed excursion/coverage loss (S4) | R High | RCR-037/056/050 | DO-SAFE-014; DO-MFG/DO-FW | PV-SVC (post-service BIST) | — | A Med | Firmware-version audit
HA-SVC-006 | depot | Debug/dev mode left enabled | Falsified record/suppressed alarm (S3) | R High | RCR-051/038 | DO-SEC; DO-SVC | PV-CYBER-006 (mode-lock, IV) | — | A Med | Unexited-service-mode detection
HA-SVC-007 | disposal | Improper Li-battery disposal | Fire/environmental (S2/S3) | A Med | RCR-052 | DO-SVC; DO-LBL; DO-SEC | Disposal-procedure review | — | A Med | Disposal-compliance
HA-MFG-001 ★ | EOL | Miscalibrated unit passes EOL | Systematic bias → wrong disposition (S4) | R High | RCR-053 | DO-TST; DO-MFG | PV-MFG-001 (EOL cal coverage) | — | A Med | EOL cal-escape trends
HA-MFG-002 | supply chain | Wrong/counterfeit cell chemistry | Cold-charge/thermal + endurance (S2/S4) | R High | RCR-054 | DO-MFG; DO-TST | PV-MFG-002 (incoming+AOI) | — | A Med | Incoming-inspection reject rate
HA-MFG-003 | assembly | Reversed/wrong-rated protection part | Protection defeated → thermal (S2/S4) | R High | RCR-054 | DO-MFG; DO-TST | PV-MFG-003 (poka-yoke+AOI) | — | A Med | AOI defect trends
HA-MFG-005 | provisioning | Wrong/duplicate crypto identity | Mis-route/repudiation (S3) | A/R | RCR-055 | DO-MFG (provisioning); DO-SEC | PV-MFG-005 (identity-uniqueness) | — | A Med | Duplicate-identity detection
HA-MFG-008 | assembly | Latent solder/PCB damage → intermittent | Silent monitoring gap (S3) | A/R | RCR-016 | DO-TST (AOI/ICT); DO-PCB | PV-MFG-008 (coverage) | — | A Med | Intermittent-failure field rate
HA-MPE-001 | mission | Valid product destroyed (wrong reject/lost record) | Vaccine-lot loss (S4) | R High | RCR-044/047/057 | DO-SAFE-004…008; DO-HF | VAL-USE-002 (false-reject targets) | VAL-USE-002 | A Med | Wastage/false-reject rate
HA-MPE-002 ★★ | mission | Mass administration of compromised doses | Population outbreak (S5) | R High | RCR-057 (aggregate) | all DO-SAFE record-integrity | (aggregate of the spine PVs) | (aggregate) | A-High (S5) | Population-scale outbreak surveillance
HA-MPE-003 ★ | mission | False-reject stock-out → unvaccinated population | Preventable disease (S4) | R High | RCR-047/044/046 | DO-SAFE-004; DO-CLD-003 | VAL-USE-002 | VAL-USE-002 | A Med (accepted trade) | Stock-out/wastage as first-class metric
HA-MPE-004 | GDP | Data-integrity violation at scale | Certification loss/hold (S4) | R High | RCR-042/057/025/033 | DO-CLD (API/audit); DO-SAFE-008 | PV-DATA-006 + traceability test | — | A Med | Data-integrity deviations
HA-MPE-005 | env | Lithium/e-waste at fleet scale | Environmental release (S2/S3) | A Med | RCR-052/007 | DO-SVC; DO-LBL; DO-REG-006 | Disposal + UN 38.3 | — | A Med | Take-back/recycling compliance
3. Coverage & closure notes
• ~60 top hazard rows span every domain module (HAL-500…1800); each traces HA → hazardous situation → harm → initial risk → RCR → DO- → PV-/VAL- → residual → field monitoring. The full ~120-record library (incl. S1–S2 records) is in the domain modules; this matrix carries the release-relevant set.
• Every top hazard has a controlling RCR (HAL-2400), a controlled design output (DOL-1200 + disciplines), and a named verification — no hazard is "controlled on paper only" (HA-CIR / HAL-2500).
• All residuals are projected / Evidence-Pending; two are carried Open (HA-COMM-003 RF-in-container → PRD-VER-005; HA-ENRG-006 energy margin → PRD-VER-004). The S5 rows (HA-TFC-002, HA-MPE-002) stay region-A-High by rule and carry mandatory monitoring.
• Security rows (HA-CYBER-) require independent adversarial (pen-test) verification (IV) before residual acceptance.
Handoff | the ★ (S4/S5) rows become Stage 11 (ERFKL) FTA top events; the RCR column feeds Stage 12; the PV-/VAL- columns feed Stages 13–14; the Field-Monitoring column feeds the Stage 17 post-market surveillance plan; the whole matrix is the traceability spine consumed by ORCH-018 / EM-ENG-012.
Preliminary Hazard Analysis (PHA) — Smart Cold Chain Monitor (SCM)
Stage 10 (ORCH-017) · Owner: System-Safety with EM-ENG-001 (Chief Systems) and EM-ENG-006 (Quality/ Regulatory). This is the concept-stage hazard artefact per the hazard-analysis standard: the early, top-down identification of the SCM's hazards from the concept and intended use alone, before detailed design, producing the preliminary safety characteristics, top preliminary hazards, and the preliminary risk-control requirements (RCRs) that seeded the design. It is retained as the front of the risk-management file and is now reconciled with the detailed analysis (the HAL domain modules + Detailed-Product-Hazard- Analysis.md); where a preliminary hazard was refined or renamed downstream, the mapping is shown. ISO 14971 harm-and-exposure-first; severity-first per CLAUDE.md rule 3. Standards cited by designation.
---
1. Purpose & method
The PHA answers, at concept lock: given what the SCM is and does, what can go wrong, to whom, and how badly? — using the ISO 14971 / ISO/TR 24971 safety-characteristic questions applied to the intended use (concept-and-decisions.md §1–§4), without waiting for a design to critique. Its output is a set of preliminary risk-control requirements (pre-RCRs) that became design inputs (PRD-SAFE, the SYS-safety set, DO-SAFE) and are now the baseline the detailed HAL refined into RCR-001…058. The PHA is deliberately severity-first and coarse: it establishes which harms dominate and what must be controlled by design, not the full record set.
2. Scope, product & lifecycle
Product (concept) | reusable, cellular-connected environmental data logger placed inside insulated vaccine shipping containers; logs temperature/humidity/shock, alerts on excursion, uploads on a cadence, supports OTA, charges over USB-C; −30 °C…+60 °C, 90-day endurance; BLE/Wi-Fi/LTE-M/GPS, rechargeable battery, OLED, cloud dashboard.
Class | information device — monitors and informs disposition; does not control temperature and is not the primary means of preserving the vaccine (concept §2). Public-health-significant, non-therapeutic.
Lifecycle phases considered | provisioning/arming → armed transport (in-container, in-transit, cross- border, air cargo) → receipt/disposition → depot recharge/recalibration/triage → OTA → reuse over life → decommission/disposal. Reusability (D1) and the −30 °C/90-day/charge tension (D3) are the two lifecycle amplifiers.
3. Perspectives (who/what can be harmed)
Broadened beyond the operator, per HAL-100/200:
• Beneficiary / patient — the ultimate harmed party via a wrong disposition (spoiled dose administered, or a needed dose destroyed). The dominant perspective.
• Disposition decision-maker / responder — nurse, QA officer, fleet responder acting on the SCM's output.
• Depot / service technician — handles, charges, recalibrates, services the device (battery/thermal/ electrical/chemical exposure).
• Air-cargo crew & everyone aboard an aircraft — the catastrophic (S5) transport perspective (Li-battery fire in an occupied hold).
• Bystanders incl. children — small-parts/handling exposure.
• Mission / programme / property / environment — vaccine-lot loss, GDP/data-integrity, e-waste, programme trust.
4. Safety-characteristic questions — summary of answers
Applying the ISO/TR 24971 question set to the SCM concept (the full working-through is in HAL-100/400):
# | Safety-characteristic question | SCM answer (concept-stage)
1 | Intended use / purpose? | Monitor + record + alert on temperature/humidity/shock for vaccine cold-chain disposition; information device, not a controller.
2 | Is it a measuring / decision-support function? | Yes — its measurements drive an accept/reject/quarantine decision; measurement accuracy + honest reporting are safety-critical.
3 | Intended & vulnerable users? | Depot techs and low-training, last-mile clinic users; often gloved, cold, low-light, multiple languages (HA-CTX-005).
4 | Energy sources / hazardous energy? | Rechargeable lithium battery (fire/thermal/chemical), USB-C charge input, RF emitters — the physical hazard set.
5 | Materials / substances? | Lithium electrolyte, enclosure polymers, cleaning agents (reuse).
6 | Operating environment? | −30 °C to +60 °C, condensation cycling, vibration/shock, inside metallised/insulated shippers (RF-hostile), air cargo (altitude, occupied-hold fire risk).
7 | Dependence on connectivity / external systems? | Yes, materially — alert delivery and second-opinion recompute depend on LTE-M/Wi-Fi/cloud; offline-first mitigates (D2).
8 | Software / firmware safety role? | Central — detection, verdict, plausibility, record, alerting all in firmware (Class-C-equivalent).
9 | Foreseeable misuse? | Wrong profile, probe left in air, override a reject, ignore alerts, skip cleaning/recal, discard a cold charge-inhibited device as dead.
10 | Reuse / service / disposal? | Reusable (D1) — recalibration, seal restore, firmware/config restore, and safe Li-battery disposal become safety-relevant.
5. Top preliminary hazards (concept-stage), with risk chain & derived pre-RCRs
The PHA identified the following dominant hazards from the concept alone. Each is shown with its risk chain, preliminary severity, the derived preliminary risk-control requirement, and the detailed HAL record(s) + RCR(s) it became.
PHA-01 — Missed / incorrect / delayed excursion detection → spoiled dose administered ★ dominant harm
Risk chain | excursion occurs → not detected / mis-detected / detected-too-late → record looks clean → disposition = accept → spoiled dose administered → beneficiary not protected → preventable disease (lot-scale S4, population-scale S5).
Perspective | beneficiary/patient. Preliminary severity: S4 (S5 worst).
Derived pre-RCR | two-sided (heat and freeze) detection with freeze reliability equal to heat, no undetected-excursion path; honest verdict that never shows a clean pass over a gap/fault.
Became | HA-SW-001/002/003/008, HA-DATA-001/004, HA-ENV-001, HA-COMM-001/002/003, HA-USE-001/002/007 · RCR-019/020/021/022/040 (+ RCR-028/029 delivery, RCR-043/044 use).
PHA-02 — Silent loss of monitoring (dead-in-cold / gap reads as uneventful) ★
Risk chain | battery exhausted or device reset in deep cold → monitoring stops → gap not logged / reads as an uneventful trip → false reassurance → accept → spoiled dose. Preliminary severity: S4.
Derived pre-RCR | energy reserve + honest low-battery/dead-in-cold indication; power-fail-safe storage so any loss becomes an honest gap, never a silent absence; 90-day cold endurance.
Became | HA-ENRG-006, HA-SW-013, HA-ELEC-010 · RCR-006/024/027 (+ PRD-VER-004 cold endurance).
PHA-03 — Cold-charge battery damage → thermal runaway / fire ★ (D3)
Risk chain | rechargeable Li cell charged below 0 °C (reusability + cold depot) → lithium plating → latent internal short → thermal runaway → fire. Preliminary severity: S4/S5.
Derived pre-RCR | prevent charging below 0 °C by an MCU-independent hardware interlock; independent cell protection + secondary backstop; single-fault-safe.
Became | HA-ENRG-001/002, HA-ELEC-002 · RCR-001/002 (D3 charge-safety trio).
PHA-04 — Lithium-battery fire in an occupied aircraft cargo hold ★★ (S5 anchor)
Risk chain | any cell thermal event while in air cargo → fire/smoke/toxic offgas in an enclosed occupied hold → fatal fire / loss of aircraft. Preliminary severity: S5 (catastrophic) — the defining physical S5.
Derived pre-RCR | prevent initiation (protection + retention) and contain a single-cell event without external flame propagation (fire enclosure); UN 38.3 / IATA-DGR compliance + transport disclosure.
Became | HA-TFC-002/003, HA-ENRG-003, HA-ENV-007 · RCR-002/004/007/008.
PHA-05 — Delayed / missing / undeliverable excursion alert (connectivity-bound)
Risk chain | excursion detected → no bearer (in-container / cross-border / RF-hostile) or delivery lost → responder unaware → no intervention → spoiled dose. Preliminary severity: S4.
Derived pre-RCR | commit every excursion locally independent of connectivity; store-and-forward, alert- first, ack-required, escalate-if-unacknowledged; on-receipt edge read as the offline compensating check; qualify the in-container RF link.
Became | HA-COMM-001/002/003, HA-CLOUD-001/005 · RCR-028/029/030/031/032 (+ PRD-VER-005 RF-in-container).
PHA-06 — Wrong / biased measurement (calibration, probe, placement) presented as valid
Risk chain | biased sensor/probe/calibration, or probe not in the product zone → plausible-but-wrong values → wrong verdict → spoiled dose (or good product destroyed). Preliminary severity: S4 (S5 fleet- wide calibration bias).
Derived pre-RCR | integrity-protected calibration with expiry gate (block arming on invalid calibration); probe/sensor plausibility (never read absent/open/wrong probe as valid); EOL calibration coverage + birth record; reuse recalibration interval.
Became | HA-DATA-001, HA-SW-004/008/014, HA-USE-005/006/011, HA-MFG-001, HA-SVC-001 · RCR-020/026/053/049.
PHA-07 — Falsified / suppressed data or malicious OTA (security → safety)
Risk chain | attacker injects false-clean data, mutes an alarm, alters thresholds, or pushes malicious firmware → excursion hidden / fleet blinded → spoiled dose(s). Preliminary severity: S4 (S5 fleet-scale).
Derived pre-RCR | capture-time device-anchored data authenticity; non-suppressible alarm path; authorised bounded audited config change; signed rollback-protected never-brick OTA; secure boot + debug-lock.
Became | HA-CYBER-001/002/003/005/006/009, HA-CLOUD-008 · RCR-034/035/036/037/038/039.
PHA-08 — Use error: misread verdict / ignored alert / mis-armed / override
Risk chain | legible-but-misread verdict, alarm fatigue, forgot-to-arm/void trip believed monitored, or a reject overridden under stock-out pressure → wrong disposition. Preliminary severity: S4.
Derived pre-RCR | legible language-independent verdict (icon+colour+word, gloved/cold/low-light); tuned alerting (no fatigue muting a real excursion); pre-arm readiness gate (unmistakable un-armed/void state); deliberate recorded override.
Became | HA-USE-001/002/004/007/008/009 · RCR-043/044/045/047.
PHA-09 — Record-integrity loss from environment/ingress/condensation (reuse)
Risk chain | condensation/ingress/cleaning over reuse → record-electronics fault → wrong/lost record → wrong disposition. Preliminary severity: S3 (S4 if it produces a false clean pass).
Derived pre-RCR | conformal coat + seal across −30↔+60 °C condensation cycling over reuse life; plausibility/ verdict screen residual wet-fault reads; post-service BIST.
Became | HA-TFC-005/006, HA-ENV-003/004, HA-ELEC-006/007, HA-MECH-004 · RCR-011/015/018/050.
PHA-10 — False reject → vaccine-lot loss → stock-out → unvaccinated population
Risk chain | the other disposition error — a good lot wrongly rejected/destroyed, or a serviceable device not deployed → stock-out → unvaccinated beneficiaries → preventable disease. Preliminary severity: S4.
Derived pre-RCR | validated false-reject/false-alarm targets (ALARP in both directions); audited informed override so a genuinely-good marginal lot is not blindly destroyed; charge-inhibited ≠ dead legibility.
Became | HA-MPE-001/003, HA-USE-009, and the HAL-2700 accepted-trade set · RCR-044/046/047.
PHA-11 — Physical handling hazards (edge/pinch/small-parts/hot-surface/RF)
Risk chain | accessible sharp edge, pinch at battery door, detachable small part (child choking), hot self- heat/charge surface, RF exposure → injury. Preliminary severity: S1–S3.
Derived pre-RCR | no accessible hazardous edge/pinch; small-parts retention/sizing/warning; touch- temperature bound; SELV; RF within limits.
Became | HA-MECH-001/002/003, HA-TFC-001/008, HA-ELEC-001, HA-ENRG-009 · RCR-005/009/010/013.
6. Preliminary risk-control requirements (pre-RCR baseline)
The PHA's pre-RCRs (summarised in §5) became the PRD-SAFE preliminary safety requirements and the SYS-safety set (SYS-149…162), which the detailed HAL then refined into the authoritative RCR-001…058 register (HAL-2400). The PHA's severity-first priority ordering — physical S5 (aircraft fire), then the undetected-excursion information pathway (S4/S5), then connectivity/security/use, then reliability and handling — set the design's control hierarchy: inherent/hardware controls for the physical hazards (charge-inhibit, protection, fire enclosure) and inherent-by-construction controls for the information hazards (no-undetected-excursion, honest verdict, honest gap), with information-for-safety used only where design-enforced.
7. Reconciliation & handoff
• Every PHA hazard maps forward to one or more detailed HAL records and RCRs (§5); no preliminary hazard was dropped, several were split (e.g., PHA-01 → the detection/verdict/probe/alert/use cluster) and refined.
• The two carried unknowns identified at concept — D3 (cold battery) and RF-in-container — were flagged in the PHA as top risks and remain the two carried-Open items through the detailed analysis (HAL-2800/2900).
Handoff | the PHA seeds the Detailed-Product-Hazard-Analysis.md (post-design view), the Hazard- Traceability-Matrix.md (chain), and the Risk-Management-Summary.md (release recommendation); the detailed analysis lives in the HAL-500…1800 domain modules with the analysis modules HAL-2000…2900.
Risk-Management Summary — Smart Cold Chain Monitor (SCM)
Stage 10 (ORCH-017) · Owner: EM-ENG-006 (Quality/Regulatory) + EM-ENG-001 (Chief Systems), with System-Safety. The ISO 14971 risk-management report for the SCM at the close of the hazard-analysis stage: it states the risk-management activities performed, the open/unacceptable risks, the residual-risk profile, the overall benefit-risk conclusion, the unresolved assumptions, the V&V status, and the overall release recommendation. It consolidates the HAL library (HAL-000…2900) and the connected artefacts (PHA, Detailed PHA, Traceability Matrix) into the single document a design review / release board reads. Severity-first per CLAUDE.md rule 3; regions/AP per HAL-2200. Standards cited by designation.
---
1. Scope of this summary
Product: the reusable, cellular-connected environmental data logger for vaccine cold-chain monitoring (concept-and-decisions.md), all SKUs, full lifecycle, multi-market (WHO PQS / EU GDP / US / air-transport, D4). This report covers the risk-management activities completed through Stage 10 (hazard identification, analysis, risk evaluation, risk-control definition, control-implementation review, verification planning, new-risk review, and individual + overall residual-risk evaluation). It does not report executed V&V evidence (Stages 13–14), which is why the conclusion is conditional.
2. Risk-management activities performed (per ISO 14971)
ISO 14971 activity | SCM artefact | Status
Risk-management plan / governance | HAL-000 (governance-plan) | Complete
Context / intended use / users / misuse | HAL-100 (HA-CTX), HAL-400 (HA-IUE) | Complete
System boundary / exposure models | HAL-200 (HA-BND) | Complete
Lifecycle / hazard taxonomy | HAL-300, HAL-1900 | Complete
Hazard identification & analysis (per domain) | HAL-500…1800 (~120 records, 34-field) | Complete
Severity / probability / acceptability framework | HAL-2000 / 2100 / 2200 | Complete
Risk-control requirements | HAL-2400 (RCR-001…058) | Complete
Control-implementation review | HAL-2500 (HA-CIR) | Complete
Risk-control verification plan | HAL-2600 (HA-RCV) | Complete (plan)
New risks introduced by controls | HAL-2700 (HA-NEW) | Complete
Individual residual-risk evaluation | HAL-2800 (HA-RES) | Complete (projected)
Overall residual + benefit-risk | HAL-2900 (HA-ORES) | Complete (conditional)
Preliminary / detailed / traceability | PHA, Detailed PHA, Traceability Matrix | Complete
Verification & validation execution | Stages 13–14 (PV-/VAL-) | Pending
Post-market surveillance | Stage 17 | Planned
3. Open / unacceptable risks at this stage
No individual risk remains in region R (unacceptable) as projected — every R-region hazard is reduced to region A (or, for S5, region A-High that can never be broadly acceptable) assuming its named verification passes. The material open items are therefore Evidence-Pending closures and two carried-Open unknowns, not uncontrolled risks:
# | Open item | Hazard(s) | Why open | Closes on | Owner
O-1 | RF-in-container link margin / antenna decision | HA-COMM-003 | In-container attenuation & integral-vs-external antenna not yet quantified; largest connectivity unknown | PRD-VER-005 | Connectivity + RF/Antenna + Regulatory
O-2 | D3 cold-battery: −30 °C / 90-day endurance + charge safety + fade | HA-ENRG-001/006, HA-NEW-001 | Endurance margin & cold-charge safety Evidence-Pending on datasheet + test | PRD-VER-004 + PV-ENRG-001 | Battery/Power
O-3 | S5 aircraft-hold fire containment | HA-TFC-002/003 | Containment + offgas + transport qualification not yet demonstrated | PV-TFC-002 + UN 38.3 | Thermal/Battery-Safety + Mech + Reg
O-4 | The two pathway choke points (honest verdict; supported decision) | HA-SW-002, HA-DATA-004, HA-USE-001/004/007 | Non-defeatability + human comprehension not yet verified/validated | PV-SW-002, PV-DATA-004, VAL-USE-004/001 | SW-Safety + Human-Factors
O-5 | Fleet common causes (calibration; OTA/cloud) | HA-SW-004, HA-MFG-001, HA-CYBER-005, HA-CLOUD-008 | EOL+depot calibration coverage & OTA/cloud rollout discipline not yet proven | PV-MFG-001 + PV-CYBER-005 + rollout policy | Metrology + Cloud + Security
O-6 | Security controls | HA-CYBER-001/002/003/005/006/009 | Require independent adversarial (pen-test) verification | PV-CYBER- (IV) | Product-Security
O-7 | WHO PQS accuracy target / US classification | benefit basis (HA-MPE-002) | Exact PQS category/spec revision & US class not yet confirmed (D4) | Regulatory stage confirmation | EM-ENG-006
None of O-1…O-7 is an accepted unacceptable risk; each is a condition on release (§7).
4. Residual-risk profile
Physical (energy/battery/thermal) | reduced to region A by independent, single-fault-safe, non-defeatable hardware controls (charge-inhibit, protection + secondary backstop, retention, fire enclosure). One S5 (aircraft-hold fire) stays region-A-High permanently by rule, with mandatory disclosure + thermal-event monitoring.
Information pathway (the dominant harm) | ~20 S4 region-A residuals that converge on one beneficiary through one pathway, contained by two independent choke points (honest verdict + supported human decision) and the honest-gap / no-false-reassurance invariant. Reduced to region A projected; not yet demonstrated.
Connectivity/cloud | reduced to region A by local-commit + store-and-forward + ack/escalation + offline edge read; the RF-in-container margin (O-1) is the carried-Open driver.
Security→safety | reduced to region A by device-anchored authenticity, non-suppressible alarm, secure OTA, secure boot — pending independent adversarial verification (O-6).
Use/human-factors | reduced to region A, usability-gated — the last-mile low-training/vulnerable user keeps these human-mediated and closable only by summative usability.
Accepted public-health trade | the conservative safety bias produces a real false-reject/wastage/stock-out residual (HA-MPE-003) — region A, balanced not eliminated, monitored as a first-class field metric.
Lower-severity (S1–S2) | region B (broadly acceptable) via inherent/protective controls.
5. Overall benefit-risk conclusion (from HAL-2900)
The SCM's public-health benefit — protecting vaccine potency and availability by honest excursion detection, timely alerting, and an auditable cold-chain record for disposition — outweighs its overall residual risk, conditional on the planned V&V demonstrating the projected residuals and on the seven open items (§3) closing acceptably. This is a positive but conditional determination, not unconditional, because V&V is not yet executed and three load-bearing unknowns (D3, RF-in-container, WHO PQS target) remain. The S5 aircraft-fire residual is accepted only with EM-ENG-001+006 benefit-risk sign-off, UN 38.3 / IATA-DGR disclosure, and mandatory post-market thermal-event monitoring.
6. Unresolved assumptions (carried forward, not invented)
D3 battery (−30 °C / 90-day / charge safety / fade) | a low-temperature-tolerant chemistry can meet the 90-day cold endurance with charge-inhibit-below-0 °C and (optionally) self-heat — assumed, to be proven by PRD-VER-004 + the battery datasheet. The single largest program assumption.
RF-in-container | LTE-M (or external antenna) can close (or the consequence is bounded by store-and-forward for) the in-container link — assumed, to be proven by PRD-VER-005; integral-vs-external antenna undecided.
WHO PQS accuracy/performance target | the device meets the recognised cold-chain-monitor accuracy target; the exact PQS category/spec revision and US classification are to be confirmed in the regulatory stage (D4).
Actionable-window usefulness | on lanes where the live alert is delayed, the on-receipt edge read + the finite re-ice/segregate window are still useful — lane-dependent, Evidence-Pending (HA-COMM-001).
Depot/service preservation | reuse recalibration, seal restore, firmware/config restore, and disposal are performed to procedure — controlled by post-service BIST + pre-arm gate, assumed followed in the field.
Self-heat feasibility (CMP-013) | if not adopted, cold charging relies on warm-then-charge process alone (strengthens O-2); conditional records (HA-TFC-001) become N-A.
7. Verification & validation status and release conditions
Status: V&V pending (Stages 13–14). Every risk control is Design-adequate / Evidence-Pending (HAL-2500); a complete, safety-driven verification plan exists (HAL-2600) with fault-injection for every safety control and independent adversarial verification for every security control. No PV-/VAL- has yet been executed.
Release conditions (all must be met for the conditional benefit-risk conclusion to become unconditional):
1. All safety-control PV- pass to their acceptance criteria, with fault-injection (HAL-2600).
2. The two pathway choke points are verified/validated: PV-SW-002 + PV-DATA-004 (honest verdict) and VAL-USE-004 + VAL-USE-001 (supported decision).
3. The S5 aircraft-fire containment is verified: PV-TFC-002 + UN 38.3 / IATA-DGR; disclosure + monitoring in place.
4. The two carried-Open unknowns close acceptably: PRD-VER-004 (D3 cold endurance) and PRD-VER-005 (RF-in-container); the antenna decision is fixed.
5. The fleet common causes are verified and monitored: PV-MFG-001 (calibration) + depot-recal control; staged/canary OTA + halt-on-regression; cloud DR-tested.
6. Security controls pass independent pen-test (PV-CYBER-001/002/003/005/006/009).
7. The WHO PQS accuracy target (and US classification) is confirmed and met (D4).
8. Post-market surveillance is established for both harm directions (missed-excursion and false-reject/ wastage) plus thermal-event and security monitoring (ISO 14971 §9 loop).
8. Overall release recommendation
Conditionally Ready — release pending V&V. The SCM's hazard analysis is complete, its risk controls are
defined, allocated, implemented in controlled design outputs, and reviewed for adequacy, new risk, and
residual risk; no individual residual remains unacceptable as projected, and the overall benefit-risk is
positive but conditional. Release is recommended only after the eight conditions in §7 are met — in
particular the two pathway-choke-point verifications, the S5 aircraft-fire + UN 38.3 qualification, the two
carried-Open unknowns (D3 cold endurance PRD-VER-004, RF-in-container PRD-VER-005), independent security
verification, and confirmation the device meets the WHO PQS accuracy target. Until then the device is not
released. Where a condition cannot be met, the corresponding claim is withheld or restricted (e.g., no
air-transport claim without UN 38.3; no real-time-alert claim on lanes where PRD-VER-005 shows the link
cannot close — defaulting to the auditable-record + on-receipt-verdict benefit).
9. Sign-off & traceability
• Acceptance authority (HAL-2200 §7): S4/S5 region-A residuals require EM-ENG-006 + EM-ENG-001 + benefit-risk; the S5 aircraft-fire residual requires EM-ENG-001 + EM-ENG-006 sign-off + documented benefit-risk + post-market monitoring. All sign-offs are pending V&V (this report is the pre-V&V baseline).
Feeds | Stage 11 (ERFKL — the S4/S5 hazardous situations as FTA top events + the common-cause list); Stage 12 (Risk-Control Requirements + re-baseline of PRD-SAFE and Product/Design); Stages 13–14 (the V&V that closes the conditions); Stage 17 (post-market surveillance carrying §7.8); ORCH-018 / EM-ENG-012 (the complete lifecycle traceability spine).
Living document | per ISO 14971 §9, this summary is updated as V&V evidence returns and post-market data accrues, to keep the benefit-risk conclusion valid across the reusable device's service life.

## 11 Failure & Risk Analyses
Failure & Risk Analyses
How each part could fail, how likely and serious that is, and what we do about it.
↩ Contents
In this section:
• Failure Mode & Risk Analyses (ERFKL) — Index — Smart Cold Chain Monitor (SCM)
• FAL-1000 — Design FMEA (DFMEA) · Mechanical / Enclosure / Seal / Thermal — Smart Cold Chain Monitor (SCM)
• FAL-1200 — Software FMEA (SWFMEA) — Smart Cold Chain Monitor (SCM)
• FAL-1300 — Hardware FMEA (HWFMEA) — Smart Cold Chain Monitor (SCM)
• FAL-1400 — System FMEA (SYSFMEA) — Smart Cold Chain Monitor (SCM)
• FAL-1500 — Interface FMEA (IFMEA) — Smart Cold Chain Monitor (SCM)
• FAL-1600 — Use-Error Analysis (link to HAL-1500) — Smart Cold Chain Monitor (SCM)
• FAL-1800 — Fault Tree Analysis (FTA) — Smart Cold Chain Monitor (SCM)
• FAL-1900 — Reliability Analysis — Smart Cold Chain Monitor (SCM)
• FAL-2000 / FAL-2100 — RCA & CAPA Readiness Note — Smart Cold Chain Monitor (SCM)
• FAL-2200 / FAL-2300 — Historical & AI Failure-Knowledge Note — Smart Cold Chain Monitor (SCM)
• FAL-2700 — Common-Cause Failure (CCF) Analysis — Smart Cold Chain Monitor (SCM)
• FMEA-to-Requirements-Hazard Matrix — Smart Cold Chain Monitor (SCM)
Failure Mode & Risk Analyses (ERFKL) — Index — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 11 — Engineering Risk & Failure Knowledge (ERFKL) · Orchestrator: ORCH-016 · Owner: Principal Reliability/Safety/Quality Engineer with the discipline engineers and Chief Systems Engineer (EM-ENG-001). This stage authors the SCM's failure/risk analyses — DFMEA, HW/SW/System/ Interface FMEA, Fault Tree Analysis, Reliability, Common-Cause, Use-Error, and the RCA/CAPA + Historical/AI readiness notes — using FA-<CODE>-NNN IDs and the 20-field register, scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3; AIAG-VDA; ISO 14971). Every safety-significant failure mode is reconciled to a HAL hazard (HA-…); the reconciliation spine is the FMEA-to-Requirements-Hazard-Matrix. Evidence honesty: fact / standard / Judgment / Assumption / Evidence-Pending are marked throughout; O and D carry a stated basis; the two program-top items (D3 −30 °C/90-day battery and RF-in-container) are carried Open.
Position in the chain: consumes Stage 07 (subsystem requirements), Stage 08 (architecture: CMP-/FUN-), Stage 09 (Design Outputs: DO-…), and critically Stage 10 (HAL: HA-…, RCR-…). Feeds Stage 12 (Risk-Control Requirements), Stage 13 (Verification, PV-), Stage 14 (Validation, VAL-), and ORCH-018 / EM-ENG-012 (end-to-end traceability). Reconciles both ways with the HAL (Stage 10): FMEA safety end-effects ↔ HAL hazardous situations; high-severity HAL situations → FTA top events.
---
1. Module files
FAL module | Code | File | Records | Scope
FAL-1000 | FA-DFMEA | FAL-1000-dfmea.md | 20 | Design FMEA — mechanical/enclosure/seal/thermal, retention, HMI-mech
FAL-1200 | FA-SWFMEA | FAL-1200-software-fmea.md | 35 | Software FMEA — detection, verdict, state machine, memory, timing, storage, OTA, recovery, alarm, security
FAL-1300 | FA-HWFMEA | FAL-1300-hardware-fmea.md | 34 | Hardware FMEA — power/battery/BMS/charge, sensors, PCB, connectors, antenna
FAL-1400 | FA-SYSFMEA | FAL-1400-system-fmea.md | 22 | System FMEA — subsystem/cloud/user/service interactions, mission effects
FAL-1500 | FA-IFMEA | FAL-1500-interface-fmea.md | 22 | Interface FMEA — connectors, LTE-M/Wi-Fi/BLE, API, BLE-provisioning, USB-C
FAL-1800 | FA-FTA | FAL-1800-fault-tree-analysis.md | 4 top + ~60 basic | Fault trees for the four HAL ★ top events; cut sets, SPF, CCF
FAL-1900 | FA-REL | FAL-1900-reliability-analysis.md | 12 | Reliability — MTBF/availability, Weibull/accelerated-life, cold endurance, battery aging, flash endurance, warranty
FAL-2700 | FA-CCF | FAL-2700-common-cause-analysis.md | 14 | Common-cause — firmware/clock/calibration/power/cloud/OTA shared dependencies
FAL-1600 | FA-USE | FAL-1600-use-error-link.md | 13 | Use-error, reconciled 1:1 with HAL-1500 (HA-USE-…)
FAL-2000/2100 | FA-RCA/CAPA | FAL-2000-2100-rca-capa-readiness.md | readiness note | RCA methods + CAPA flow + triggers, pre-wired to FMEA/FTA/HAL
FAL-2200/2300 | FA-HIST/AIKB | FAL-2200-2300-historical-ai-knowledge.md | pattern note | Recurring IoT/battery/cold-chain patterns + feedback loop
— | — | FMEA-to-Requirements-Hazard-Matrix.md | ~65 spine rows | FMEA/FTA ↔ HA ↔ Req/DO ↔ RCR ↔ V&V reconciliation
— | — | index.md | — | This file: module table + applicability register
Failure-record total: ~232 (133 FMEA + ~60 FTA basic events + 13 use-error + 14 CCF + 12 reliability), within the target rich set. The FMEA-to-Requirements-Hazard-Matrix carries the ~65 release-relevant safety rows.
2. Applicability Register (which modules are active, and why)
FAL module | Status | Rationale
FAL-1000 DFMEA | Active | Sealed IP enclosure, cold environment, reuse — mechanical/seal/thermal design risks are significant.
FAL-1200 SWFMEA | Active | Detection/verdict/record logic is the product's core value and its dominant safety surface (IEC 62304).
FAL-1300 HWFMEA | Active | Li-battery/charge (S5), sensors, RF — the highest-severity and most numerous modes.
FAL-1400 SYSFMEA | Active | Device↔cloud↔operator↔depot interactions produce the emergent spine failures.
FAL-1500 IFMEA | Active | RF-in-container, API, BLE-provisioning, USB-C — interface contracts are safety-critical.
FAL-1600 Use-Error | Active | Human decision loop (arm, read, alert, override) directly determines disposition; reconciled with HAL-1500.
FAL-1700 Hazard (touchpoint) | Active (pointer) | Authoritative hazard analysis is the HAL (Stage 10); FAL-1700 points there — not re-authored here.
FAL-1800 FTA | Active | Four HAL ★ (S4/S5) top events warrant deductive cut-set/SPF/CCF analysis (IEC 61025).
FAL-1900 Reliability | Active | D3 90-day cold endurance + reuse-life aging + flash endurance are top program risks.
FAL-2000/2100 RCA/CAPA | Active (readiness) | No field data yet (pre-verification); processes set up and pre-wired.
FAL-2200/2300 Hist/AI | Active (living) | Recurring class patterns applied now; SCM-specific records accrue from Stage 17.
FAL-2400 Traceability | Deferred to ORCH-018 | End-to-end traceability is authored at Stage 18 (EM-ENG-012); the ERFKL segment is this stage's matrix.
FAL-2700 CCF | Active | Fleet-scale shared dependencies (firmware/calibration/cloud/cold/OTA) are a first-order concern.
FAL-1100 PFMEA | Deferred to Stage 15 | Manufacturing/process FMEA belongs to Design Transfer (Stage 15); mfg causes are cross-referenced here (HA-MFG-).
FAL-2500 ETA | Not Applicable (confirm) | Event-tree adds little beyond the FTA for this product; the FTA + HAL sequences cover the consequence paths. Revisit if a quantitative barrier study is needed for the S5 fire event.
FAL-2600 Sneak Circuit | Not Applicable (confirm) | Low analog/relay complexity; the charge-inhibit interlock and protection paths are covered by HWFMEA + FTA. Revisit if the interlock/self-heat analog design grows.
3. Scoring convention (all modules)
Severity (S) 1–10, outcome-fixed | S5 (9–10) catastrophic (death/aircraft); S4 (7–8) critical (compromised dose administered / thermal event / fleet-blinding); S3 (5–6) serious (wrong disposition / lost record / burn); S2 (3–4) minor; S1 (1–2) negligible.
• Occurrence (O) / Detection (D) 1–10 with a stated basis per record (Judgment / handbook-analogue / Evidence-Pending until a named test). D=10 is silent/undetectable.
Action Priority (severity-dominant) | an unmitigated safety end-effect (S≥7) is High unless both O≤3 and D≤4 (then Medium); S3 loss-of-function follows S/O/D bands; S1–S2 nuisance is Low unless detection is poor. AP is the residual direction after current controls; recommended actions drive it to target.
4. Top items carried into Stage 12+
1. D3 — 90-day cold (−30 °C) energy endurance + reuse-life aging (FA-HWFMEA-001, FA-SWFMEA-021, FA-HWFMEA-012, FA-REL-007/009; HA-ENRG-006; FTA TOP-A) — Open, closed by PRD-VER-004.
2. RF-in-container alert egress (FA-HWFMEA-029, FA-IFMEA-005, FA-CCF-010; HA-COMM-003; FTA TOP-B) — Open, closed by PRD-VER-005.
3. Li-battery thermal runaway/fire (S5) — two-tier protection + HW/FW charge-inhibit diversity + containment (FA-HWFMEA-002/006/007/008, FA-DFMEA-008/009; HA-TFC-002; FTA TOP-C).
4. Undetected excursion / false accept — two-sided detection, probe integrity, false-accept-suppressing verdict, device+cloud diversity (FA-SWFMEA-001/008, FA-HWFMEA-017; HA-SW-001/002; FTA TOP-A; CCF-001).
5. No/late alert — store-and-forward-with-escalation, persistent queue, delivery ack (FA-SWFMEA-023/024, FA-IFMEA-011; HA-COMM-001/002; FTA TOP-B).
6. Corrupted/false record — power-fail-safe writes, tamper-evidence (IV), DR (FA-SWFMEA-018/031; HA-CYBER-001, HA-SW-013; FTA TOP-D).
7. Fleet-scale common cause — firmware/calibration/cloud/OTA/cold shared dependencies (FA-CCF-001/002/004/007/ 009; HA-MPE-002 S5).
5. Evidence & honesty statement
• Action Priority, not RPN, everywhere (org rule 3).
No fabricated numbers | no MTBF/failure-rate, capacity, cycle-count, RF attenuation, or datasheet figure is asserted; all such quantities are Evidence-Pending with a named closing test. O/D are pre-validation estimates with a stated basis.
FMEA↔HAL reconciled | every safety-significant mode traces to a HA-…; non-safety items with no direct hazard (accelerometer, GNSS, Wi-Fi offload, humidity) are recorded as acceptable, not gaps; no HAL ★ hazard lacks an FMEA record.
• Security-induced safety modes require independent adversarial verification (IV) before their controls are credited.
FAL-1000 — Design FMEA (DFMEA) · Mechanical / Enclosure / Seal / Thermal — Smart Cold Chain Monitor (SCM)
Module FAL-1000 (FA-DFMEA) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Principal Mechanical/Materials Engineer. Design FMEA for the SCM device's mechanical, enclosure, sealing, retention, HMI-mechanical and thermal-coupling design. Scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3; AIAG-VDA). Every safety-significant failure mode is reconciled to a HAL hazard (HA-…); where no hazard link exists it is flagged as a gap. Feeds Risk-Control Requirements (Stage 12), Verification (Stage 13), FTA (FAL-1800), CCF (FAL-2700), and ORCH-018/EM-ENG-012 traceability.
---
1. Purpose
Systematically identify how the SCM's mechanical/enclosure design can fail, the physics of each mechanism, the local→system→user/safety effect, and the design controls and recommended actions that reduce risk to acceptable levels — with full traceability from requirement/design-output through hazard to verification.
2. Scope
S1 enclosure (top/bottom housing, seals/gaskets, USB-C aperture, probe gland, mount/retention), S1/S6 HMI window and button mechanics, S1/S2 battery retention/isolation, S1/S3 PCB mounting, S1/S4 thermal-coupling and internal-sensor airpath, materials aging/UV/chemical compatibility, tolerance stack, and drop/vibration/altitude structural response. Excludes electronic-component design (FAL-1300 HWFMEA), firmware/software (FAL-1200), interfaces as data/energy paths (FAL-1500), and manufacturing-process defects (FAL-1100 PFMEA) except where a design decision drives a process risk. Cell electro-thermal failure modes are cross-referenced to FAL-1300.
3. Applicable Standards & References
ISO 14971 (risk), AIAG-VDA FMEA & IEC 60812 (method), IEC 60529 (IP ingress), IEC 60068-2 (environmental: cold, damp-heat, thermal cycling, vibration, shock), IEC 62368-1 (touch temperature / mechanical), UN 38.3 & IATA DGR (transport: crush, altitude, vibration), 16 CFR 1501 (small parts), IEC 62133-2 (cell mechanical abuse, cross-ref). Cited by designation; revision to confirm — no clause number asserted as fact.
4. Governance & Method — scoring key (shared across ERFKL)
• Severity (S) 1–10, outcome-fixed, mapped to HAL bands: S5 catastrophic (death / loss of aircraft) = 9–10; S4 critical (compromised dose administered / thermal event / fleet-blinding) = 7–8; S3 serious (wrong disposition / lost evidentiary record / burn) = 5–6; S2 minor = 3–4; S1 negligible = 1–2.
• Occurrence (O) 1–10 and Detection (D) 1–10 (10 = silent/undetectable). The basis for each O and D is stated in Notes (Judgment / handbook-analogue / Evidence-Pending until named test) per evidence policy.
• Action Priority (AP), severity-dominant (AIAG-VDA; CLAUDE.md rule 3): an unmitigated safety end-effect (S≥7) is High unless both O≤3 and D≤4 (then Medium); S3 loss-of-function follows S/O/D bands; S1–S2 nuisance is Low unless detection is poor. AP shown is the residual direction after current controls; Recommended Actions drive it to target. Occurrence/detection numbers are pre-validation estimates (Evidence-Pending) unless a released test exists.
5. Categories
Sealing/ingress · Enclosure structural · Battery mechanical/thermal isolation · Thermal-coupling/measurement · HMI-mechanical · Retention/fit/tolerance · Materials aging/chemical · Transport/abuse.
6. Risk & Failure Register (FA-DFMEA)
6.1 Sealing / ingress (S1 — CMP-001/002/004/005)
FA-DFMEA-001 — Main enclosure gasket loses seal under cold→warm thermal cycling
Description | Primary housing O-ring/gasket fails to maintain IP seal after repeated −30 °C↔+60 °C reuse cycles.
Category / Subsystem-Component | Sealing/ingress · S1 · CMP-002 main gasket (DO-MECH-005).
Cause → Mechanism → Failure Mode | Elastomer glass-transition/compression-set + differential CTE between housing halves → loss of gland compression / seal gap → loss of ingress seal.
Effect (local→system→user/safety) | Moisture path into cavity → condensation on PCB → intermittent sensor/record-electronics fault → wrong disposition (spoiled dose accepted or good dose rejected).
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — elastomer selected for full temp range and low compression-set; gland per DO-MECH-011; conformal coat (DO-PCB) as second barrier (DO-SAFE-011). Det — IP retest sampling; BIST/humidity plausibility flags ingress indirectly.
Recommended Action | Qualify seal material by compression-set after N reuse thermal cycles; add humidity-in-cavity trend as an in-service ingress indicator; define seal as a serviceable item (D1 reuse).
Hazard trace | HA-ENV-003, HA-TFC-006, HA-MECH-004 · Req/DO trace: HWR-038/040, PRD-MECH / DO-MECH-005, DO-MECH-011, DO-SAFE-011 · V&V: PV-TFC-006 (condensation+IP), PV-ENV-003, PV-MECH (IP).
Owner / Status / Evidence / Notes | Principal Mechanical / Open / Judgment. O-basis: compression-set analogue for reused elastomer seals across wide ΔT; Evidence-Pending cold-cycle seal test. D moderate: ingress is only indirectly observable until it causes a fault (silent-until-fault).
FA-DFMEA-002 — USB-C port aperture gasket ingress path
Description | Seal around the sealed USB-C receptacle aperture admits moisture after connector-mating cycles.
Category / Subsystem-Component | Sealing/ingress · S1 · CMP-004 / DO-MECH-008.
Cause → Mechanism → Failure Mode | Repeated depot mating stresses the aperture gasket; abrasion/set → local ingress at port.
Effect | Corrosion of USB-C pins / cavity moisture → charge-path fault or condensation → charge failure or record-electronics fault → device unavailable / wrong disposition.
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — sealed receptacle + aperture gasket (DO-MECH-008), USB-C insertion-cycle rating (CMP-004 baselined). Det — IP retest; charge-fault annunciation (F-CHG).
Recommended Action | Cycle-life the aperture seal to depot reuse count; consider recessed/keyed port to reduce mating abrasion; conformal-coat port-adjacent nets.
Hazard trace | HA-ENV-003, HA-TFC-006 · Req/DO trace: HWR-039 / DO-MECH-008, DO-EE-009 · V&V: PV-TFC-006, PV-ELEC-002.
Owner / Status / Evidence / Notes | Principal Mechanical / Open / Judgment. O/D-basis: connector-seal wear analogue; Evidence-Pending mating-cycle + IP test.
FA-DFMEA-003 — External-probe gland/strain-relief seal failure
Description | Probe-lead gland loses seal or grips lead insufficiently, admitting moisture or allowing lead pull-out.
Category / Subsystem-Component | Sealing/ingress + retention · S1/S4 · CMP-005/CMP-020 / DO-MECH-007.
Cause → Mechanism → Failure Mode | Cold-stiffened gland + repeated flex/tug → seal leak and/or loss of strain relief at the probe entry.
Effect | Ingress at probe channel → cavity moisture; or lead tension transmitted to solder joint → intermittent probe → false/missing product-zone reading → wrong disposition.
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — gland + strain-relief (DO-MECH-007), probe-presence/continuity detect (HWR-028). Det — probe-fault flag (DO-SAFE-005) blocks clean pass; continuity monitoring.
Recommended Action | Qualify gland grip retention after cold flex cycles; route lead with service loop; verify probe-fault flag latency covers intermittent (not just open) faults.
Hazard trace | HA-SW-008, HA-USE-005, HA-DATA-001, HA-ENV-003 · Req/DO trace: HWR-040/028, HWR-022, SYS-003 / DO-MECH-007, DO-EE-013, DO-SAFE-005 · V&V: PV-SW-008 (probe fault injection), PV-TFC-006.
Owner / Status / Evidence / Notes | Principal Mechanical + Sensing / Open / Judgment. S=7: probe feeds the product-zone verdict (safety). O/D Evidence-Pending; intermittent probe faults are the hard-to-detect case.
FA-DFMEA-004 — Housing joint/latch relaxation loses seal compression over reuse
Description | Snap/screw joint retention relaxes, reducing gasket compression below sealing threshold.
Category / Subsystem-Component | Sealing/structural · S1 · CMP-001 / DO-MECH-002/004.
Cause → Mechanism → Failure Mode | Creep/stress-relaxation of housing bosses + fastener loosening under vibration → reduced seal compression → ingress.
Effect | Same ingress chain → condensation → wrong disposition.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — captive fasteners / retention design, tolerance-stack (DO-MECH-018). Det — service inspection at depot recal; IP retest sampling.
Recommended Action | Torque/retention spec with locking feature; add depot re-torque/inspect step to service procedure (D1 reuse); creep analysis on boss material at +60 °C.
Hazard trace | HA-MECH-004, HA-ENV-003 · Req/DO trace: PRD-MECH / DO-MECH-002/004, DO-MECH-018 · V&V: PV-MECH (vibe + IP), PV-TFC-006.
Owner / Status / Evidence / Notes | Principal Mechanical / Open / Judgment. O-basis: fastener-relaxation handbook analogue; Evidence-Pending vibration+IP retest.
6.2 Enclosure structural / transport (S1 — CMP-001)
FA-DFMEA-005 — Enclosure crack on transit drop → ingress + data-loss path
Description | Housing fractures on a handling/transport drop, breaching the seal and exposing internals.
Category / Subsystem-Component | Structural · S1 · CMP-001 / DO-MECH-001/003, DO-MECH-017.
Cause → Mechanism → Failure Mode | Impact stress concentration at wall/boss + cold-embrittled polymer → enclosure crack → ingress and/or PCB/retention disturbance.
Effect | Ingress + shock to internals → intermittent fault / silent gap → wrong disposition; worst case disturbs battery retention (see FA-DFMEA-011).
S / O / D / AP | 6 / 4 / 5 / Medium (S3)
Current Controls | Prev — drop/structural analysis (DO-MECH-017), material toughness at −30 °C, wall/rib design. Det — cracked-enclosure visible at depot return; BIST anomalies.
Recommended Action | Select impact-modified polymer validated at −30 °C; drop-test to transport profile; design ribs to protect battery-retention zone; damage-evident inspection at depot intake.
Hazard trace | HA-MECH-004, HA-ENRG-003, HA-MECH-009 · Req/DO trace: PRD-MECH / DO-MECH-001/003/017, DO-SAFE-010 · V&V: PV-MECH (drop/vibe), UN 38.3.
Owner / Status / Evidence / Notes | Principal Mechanical / Open / Judgment. O/D-basis: cold-drop brittleness analogue; Evidence-Pending drop test at −30 °C.
FA-DFMEA-006 — Cold embrittlement of enclosure polymer
Description | Housing polymer becomes brittle at −30 °C, lowering impact and fatigue resistance.
Category / Subsystem-Component | Materials/structural · S1 · CMP-001 / DO-MECH-010.
Cause → Mechanism → Failure Mode | Sub-Tg embrittlement → reduced fracture toughness → brittle crack on impact/flex that would survive at room temperature.
Effect | As FA-DFMEA-005 (ingress / internal disturbance) but occurrence concentrated in deep-cold lanes.
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — material Tg and low-temp impact spec (DO-MECH-010). Det — none in-service until crack.
Recommended Action | Specify and verify low-temperature impact (e.g., notched-impact at −30 °C); avoid glass loadings that raise brittleness; include in HALT (FAL-1900).
Hazard trace | HA-MECH-004, HA-ENV-007 · Req/DO trace: DO-MECH-010, DO-MECH-017 · V&V: PV-MECH (cold drop), IEC 60068-2 cold.
Owner / Status / Evidence / Notes | Materials / Open / Judgment. Common-cause with FA-DFMEA-005 (see CCF).
FA-DFMEA-007 — Altitude (low-pressure) cavity/seal stress in aircraft cargo
Description | Pressure differential across the sealed enclosure at cruise altitude stresses seals and vents.
Category / Subsystem-Component | Structural/sealing · S1 · CMP-001/002 / DO-MECH-005.
Cause → Mechanism → Failure Mode | ΔP across a sealed IP cavity at low cabin/hold pressure → seal distortion / relief-path opening, and stress on any cell venting.
Effect | Post-flight ingress; combined with cell venting risk (FA-HWFMEA cross-ref) → thermal/leak risk.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — seal designed for ΔP; UN 38.3 altitude test on pack; pressure-relief consideration. Det — post-transport IP/BIST.
Recommended Action | Confirm seal integrity after altitude cycling; consider a controlled pressure-equalisation membrane that preserves IP; coordinate with battery venting design (CCF).
Hazard trace | HA-ENV-007, HA-TFC-002 · Req/DO trace: DO-MECH-005 / RCR-004/007/011 · V&V: UN 38.3 (altitude), PV-ENV, IP.
Owner / Status / Evidence / Notes | Principal Mechanical / Open / Judgment; Evidence-Pending altitude+IP test.
6.3 Battery mechanical / thermal isolation (S1/S2 — CMP-007 / DO-MECH-012)
FA-DFMEA-008 — Battery retention fails on shock/drop → cell crush
Description | Retention feature cannot restrain the pack under transport shock, allowing cell movement/crush.
Category / Subsystem-Component | Battery isolation · S1/S2 · CMP-007 / DO-MECH-012, DO-SAFE-010.
Cause → Mechanism → Failure Mode | Retention under-design / creep / fastener loss → loss of battery retention → cell impacts structure → mechanical abuse of cell.
Effect | Cell crush/puncture → internal short → thermal runaway/fire (device → occupied aircraft hold in worst case).
S / O / D / AP | 9 / 3 / 6 / High (S5)
Current Controls | Prev — battery retention/isolation design (DO-MECH-012), cushioning, UN 38.3 vibration/shock/crush. Det — BMS/secondary protection detects resulting electrical fault; visible damage at return.
Recommended Action | Design retention for worst-case transport shock with margin; add mechanical isolation barrier around cell (thermal + crush); verify by UN 38.3 + drop; single-fault review with FTA (top event c).
Hazard trace | HA-MECH-009, HA-ENRG-003, HA-TFC-002 (S5) · Req/DO trace: HWR-002 / DO-MECH-012, DO-SAFE-010 / RCR-004 · V&V: PV-MECH-004, UN 38.3.
Owner / Status / Evidence / Notes | Principal Mechanical + Power / Open / Judgment. S=9 (S5 aircraft-hold branch, HA-TFC-002). O low by design intent but Evidence-Pending; D moderate (mechanical failure only detected after the fact). FTA basic event.
FA-DFMEA-009 — Inadequate cell thermal isolation → single-cell event propagates
Description | Enclosure lacks a thermal barrier to contain a cell thermal event, allowing propagation/enclosure breach.
Category / Subsystem-Component | Battery isolation · S1/S2 · DO-MECH-005/012, DO-PCB-010.
Cause → Mechanism → Failure Mode | No fire-containment/vent-management design → uncontained thermal event → hot gas/flame exits enclosure.
Effect | Fire/toxic offgas in an enclosed shipper or aircraft hold → catastrophic (S5) / inhalation injury.
S / O / D / AP | 9 / 2 / 7 / High (S5)
Current Controls | Prev — flame-rated internals, vent-management/containment (DO-MECH-005, DO-PCB-010), secondary protection reduces initiation. Det — thermal-event field monitoring (mandatory per HA-TFC-002).
Recommended Action | Add fire-containment enclosure feature + defined vent path; characterise offgas (PV-TFC-003); flame-rating on materials near cell; single-point-of-failure review.
Hazard trace | HA-TFC-002 (S5), HA-TFC-003 · Req/DO trace: RCR-002/008 / DO-SAFE-002, DO-MECH-005, DO-PCB-010 · V&V: PV-TFC-002 (containment/fire-enclosure), PV-TFC-003 (offgas), UN 38.3.
Owner / Status / Evidence / Notes | Principal Mechanical + Safety / Open / Judgment. Occurrence low (initiation is upstream, FAL-1300) but severity fixes AP High. Detection of the design gap is by analysis, not in-service.
6.4 Thermal-coupling / measurement (S1/S4 — DO-MECH-015)
FA-DFMEA-010 — Internal-sensor airpath thermal lag/self-heat bias
Description | Enclosure thermal design gives the internal temp sensor excessive lag or self-heating error.
Category / Subsystem-Component | Thermal-coupling/measurement · S1/S4 · CMP-019 / DO-MECH-015.
Cause → Mechanism → Failure Mode | Poor airpath / sensor placed near heat-dissipating parts → measurement bias / lag vs true cavity temperature.
Effect | Slow/biased internal reading → short excursion missed or falsely declared → wrong disposition.
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — thermal-coupling/airpath design (DO-MECH-015), low-self-heat signal chain (HWR-027). Det — cross-check internal vs external probe divergence (DO-SAFE-005).
Recommended Action | Thermal-model + characterise sensor time constant; isolate sensor from PCB heat sources; define time-constant spec so excursion-detection logic (FAL-1200) is based on real dynamics.
Hazard trace | HA-SW-001, HA-DATA-001, HA-ENV-001 · Req/DO trace: HWR-021/027, SYS-001/003 / DO-MECH-015, DO-EE-014 · V&V: PV-SW-001, PV-DATA-001, thermal-chamber characterisation.
Owner / Status / Evidence / Notes | Mechanical/Thermal + Sensing / Open / Judgment. S=7: feeds the excursion verdict. Cross-ref FA-HWFMEA-019 (sensor) and FA-SWFMEA (detection dynamics).
FA-DFMEA-011 — Enclosure blocks/insulates probe placement → unrepresentative reading
Description | Mount/geometry forces internal sensor or probe into an unrepresentative micro-environment.
Category / Subsystem-Component | Thermal-coupling · S1/S4 · CMP-003/019/020 / DO-MECH-009/015.
Cause → Mechanism → Failure Mode | Retention geometry places sensor against an insulating wall / dead-air pocket → spatial measurement error vs product zone.
Effect | Device measures its own micro-climate, not the vaccine zone → missed excursion → compromised dose.
S / O / D / AP | 7 / 4 / 7 / High (S4)
Current Controls | Prev — placement guidance (DO-HF), airpath design. Det — int-vs-ext divergence flag; humidity/temp plausibility.
Recommended Action | Design mount to promote representative coupling; require external product-zone probe for the release decision; reconcile with HF placement (HA-USE-005).
Hazard trace | HA-USE-005, HA-DATA-001, HA-SW-001 · Req/DO trace: DO-MECH-009/015, DO-HF · V&V: PV-SW-008 (int-vs-ext divergence), VAL-USE-005.
Owner / Status / Evidence / Notes | Mechanical + HF / Open / Judgment. Overlaps use-error; D high because a plausible-but-wrong reading is hard to detect on the device alone.
6.5 HMI-mechanical (S1/S6 — DO-MECH-014)
FA-DFMEA-012 — OLED window/HMI mechanical failure hides verdict
Description | HMI window cracks/fogs or display mount fails, making the offline verdict unreadable.
Category / Subsystem-Component | HMI-mechanical · S1/S6 · CMP-030 / DO-MECH-014.
Cause → Mechanism → Failure Mode | Cold-cracked window / internal fog / connector unseat → verdict not legible at point of decision.
Effect | Operator cannot read accept/reject/quarantine → falls back to guess → wrong disposition; or reads stale/garbled state.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — window material/mount spec (DO-MECH-014), anti-fog/sealed display cavity. Det — self-test of display; cloud verdict is the authoritative fallback (device is not sole channel).
Recommended Action | Cold-fog and impact test the window; ensure verdict is always independently retrievable from cloud/app so display loss is not decision-blocking.
Hazard trace | HA-USE-001, HA-DATA-005 · Req/DO trace: HWR-046, SYS-015/041 / DO-MECH-014, DO-HF-002 · V&V: PV-USE-001 (cold display), VAL-USE-001.
Owner / Status / Evidence / Notes | Mechanical + HF / Open / Judgment. D low because loss is obvious to the operator and a cloud fallback exists.
FA-DFMEA-013 — Button fails to actuate with gloves / sticks in cold
Description | Glove-friendly button binds, sticks, or requires excessive force at −30 °C.
Category / Subsystem-Component | HMI-mechanical · S1/S6 · CMP-031 / DO-MECH-014.
Cause → Mechanism → Failure Mode | Cold-stiffened seal/dome + ingress + tolerance → no actuation or stuck-closed.
Effect | Operator cannot arm/query (no-actuation) → unmonitored trip if not caught; or stuck-closed → spurious wake/energy drain.
S / O / D / AP | 6 / 4 / 4 / Medium (S3)
Current Controls | Prev — gloved force/travel spec, sealed dome (HWR-047/050), debounce (FW). Det — arming confirmation gate (device shows armed/unarmed unmistakably, SYS-023).
Recommended Action | Cold-actuation test with gloves; stuck-button diagnostic in FW; ensure arming requires positive confirmation so a failed press cannot be mistaken for armed.
Hazard trace | HA-USE-004, HA-USE-009 · Req/DO trace: HWR-047/050, SYS-023 / DO-MECH-014, DO-SAFE-007 · V&V: PV-USE (state-legibility/actuation), PV-SW-003 (pre-arm).
Owner / Status / Evidence / Notes | Mechanical + HF / Open / Judgment. Ties to arming-state legibility; stuck-closed also a HW energy-drain contributor (FA-HWFMEA).
6.6 Retention / fit / tolerance (S1/S3 — DO-MECH-013/018)
FA-DFMEA-014 — PCB retention loosens → connector/solder stress → intermittent fault
Description | Board mounting relaxes under vibration/thermal cycling, stressing connectors and solder joints.
Category / Subsystem-Component | Retention · S1/S3 · CMP-034 / DO-MECH-013.
Cause → Mechanism → Failure Mode | Standoff creep / fastener loss + CTE cycling → board movement → flex-fatigue of joints → intermittent open.
Effect | Intermittent sensor/storage/radio fault → silent gap or false reading → wrong disposition.
S / O / D / AP | 6 / 3 / 7 / Medium (S3)
Current Controls | Prev — board-retention design (DO-MECH-013), tolerance-stack (DO-MECH-018), conformal coat. Det — BIST; intermittent faults hard to catch (D high).
Recommended Action | Vibration + thermal-cycle the assembly; strain-relief connectors; add watchdog/BIST coverage for intermittent joint faults (link FAL-1300 HA-ELEC-008).
Hazard trace | HA-ELEC-008, HA-MFG-008, HA-MECH-004 · Req/DO trace: DO-MECH-013/018, DO-PCB · V&V: PV-MECH (vibe/thermal-cycle), PV-MFG-008.
Owner / Status / Evidence / Notes | Mechanical + PCB / Open / Judgment. D=7: intermittent latent-joint faults are a classic hard-to-detect escape.
FA-DFMEA-015 — Tolerance stack defeats seal or connector engagement
Description | Accumulated tolerances open a seal gap or under-engage the probe/USB-C connector.
Category / Subsystem-Component | Fit/tolerance · S1 · DO-MECH-018.
Cause → Mechanism → Failure Mode | Worst-case stack of housing/gland/connector features → insufficient seal compression or connector engagement.
Effect | Ingress or partial-seat connector → intermittent probe/charge → wrong disposition / unavailability.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — tolerance-stack analysis (DO-MECH-018), GD&T on mating drawings. Det — EOL fit/IP check; partial-seat detection on probe (HWR-028).
Recommended Action | Close the tolerance loop with worst-case + statistical stack; add poka-yoke to guarantee full seat; verify at EOL.
Hazard trace | HA-ENV-003, HA-SW-008 · Req/DO trace: DO-MECH-018, HWR-028 / DO-EE-013 · V&V: PV-MECH (fit/IP), PV-MFG.
Owner / Status / Evidence / Notes | Mechanical / Open / Judgment; Evidence-Pending stack close-out.
6.7 Materials aging / chemical (S1 — DO-MECH-010)
FA-DFMEA-016 — Cleaning-chemical / UV degradation of housing & seals over reuse
Description | Repeated depot cleaning/disinfection and UV exposure degrade housing surface and seal elastomer.
Category / Subsystem-Component | Materials aging · S1 · CMP-001/002 / DO-MECH-010, DO-CFG-008.
Cause → Mechanism → Failure Mode | Chemical attack / UV chain-scission / swelling → embrittlement, crazing, seal degradation over the reuse life (D1).
Effect | Cracks/seal loss → ingress → wrong disposition; label/legend loss → mis-handling.
S / O / D / AP | 5 / 4 / 6 / Medium (S3)
Current Controls | Prev — chemical-compatible material selection (DO-MECH-010), approved cleaning-materials baseline (DO-CFG-008), labeling durability (DO-LBL). Det — depot visual inspection.
Recommended Action | Compatibility-test against the specified cleaning agents for N cycles; define a maximum reuse life / retirement criterion (feeds reliability FAL-1900 and service FAL-1600).
Hazard trace | HA-MECH-004, HA-SVC-007 · Req/DO trace: DO-MECH-010, DO-CFG-008 / RCR-011 · V&V: chemical-compatibility test, IEC 60068 damp-heat/UV.
Owner / Status / Evidence / Notes | Materials + Service / Open / Judgment. Reuse-life is a program open item.
6.8 Transport / abuse / small parts
FA-DFMEA-017 — Detachable small part (probe/cap) choking hazard near a child
Description | A detachable small component separates and is accessible to a child.
Category / Subsystem-Component | Small parts/abuse · S1 · CMP-005/020 / DO-MECH-006.
Cause → Mechanism → Failure Mode | Under-retained small part + foreseeable misuse → detachment → small part accessible.
Effect | Choking/aspiration risk (low likelihood in intended distribution context; S1 per HAL).
S / O / D / AP | 3 / 2 / 3 / Low (S1)
Current Controls | Prev — small-parts retention design + warning label (DO-MECH-006, DO-LBL). Det — retention/detachment test.
Recommended Action | Retention per 16 CFR 1501; label per RCR-010; verify by small-parts/retention test.
Hazard trace | HA-MECH-003 · Req/DO trace: RCR-010 / DO-MECH-006, DO-LBL · V&V: PV-MECH-003 (small-parts/retention).
Owner / Status / Evidence / Notes | Mechanical / Open / Standard (16 CFR 1501). Retained though low because it is a labeled/verified control obligation.
FA-DFMEA-018 — Vibration fatigue of housing/mount over multi-trip life
Description | Sustained transport vibration fatigues housing ribs, mount, or fastener bosses over reuse.
Category / Subsystem-Component | Structural fatigue · S1 · CMP-001/003 / DO-MECH-017.
Cause → Mechanism → Failure Mode | Resonant vibration + reuse cycles → fatigue crack / loosened mount.
Effect | Loss of retention/seal → ingress or internal disturbance → wrong disposition; ties to FA-DFMEA-008/014.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — structural/vibration analysis (DO-MECH-017), avoid transport-band resonances. Det — return inspection; BIST anomalies.
Recommended Action | Modal analysis + random-vibration durability to transport profile over reuse count; design out resonances; feed reuse-life to FAL-1900.
Hazard trace | HA-MECH-004, HA-MECH-009 · Req/DO trace: DO-MECH-017 · V&V: PV-MECH (random vibe), IEC 60068-2-6/-64.
Owner / Status / Evidence / Notes | Mechanical / Open / Judgment; Evidence-Pending durability test.
FA-DFMEA-019 — Condensation on internal electronics from cold→warm ingress-free path
Description | Even without a seal breach, cavity air moisture condenses on cold PCB when the device warms.
Category / Subsystem-Component | Thermal/materials · S1/S3 · DO-SAFE-011, DO-PCB (conformal).
Cause → Mechanism → Failure Mode | Trapped humid air + rapid warm-up → internal condensation → leakage currents / corrosion on unprotected nets.
Effect | Intermittent record-electronics fault → wrong disposition; long-term corrosion.
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — conformal coat (DO-PCB) as primary barrier, controlled internal free volume, low-moisture assembly. Det — humidity-in-cavity trend, BIST.
Recommended Action | Conformal-coat all sensitive nets with sensor/antenna keep-outs; consider desiccant/low free-volume; validate by condensation-cycle test.
Hazard trace | HA-TFC-006, HA-ENV-003, HA-ELEC-008 · Req/DO trace: DO-SAFE-011, DO-PCB / RCR-015/018 · V&V: PV-TFC-006 (condensation cycle).
Owner / Status / Evidence / Notes | Mechanical + PCB / Open / Judgment. Distinct from seal-breach ingress; conformal coat is the load-bearing control.
FA-DFMEA-020 — Self-heat element (if adopted) creates hot touch surface / uncontrolled heating
Description | Optional pre-charge cell self-heater raises an external surface above the touch-temperature limit, or heats without bound.
Category / Subsystem-Component | Thermal · S1/S2 · CMP-013 / DO-EE-007, DO-SAFE-001.
Cause → Mechanism → Failure Mode | Heater placement/thermal path + control fault → hot accessible surface or over-temperature.
Effect | Burn on handling (S3); or heater bypasses charge-inhibit → cold-charge/thermal (cross-ref FA-HWFMEA).
S / O / D / AP | 5 / 3 / 5 / Medium (S3) (design gated on self-heat adoption)
Current Controls | Prev — heater cannot bypass charge-inhibit or exceed safe upper limit (CMP-013 design constraint), touch-temp design (DO-SAFE-001). Det — cell-temp sense; thermal cutoff.
Recommended Action | If self-heat is adopted, independent over-temp cutoff + touch-temperature verification; otherwise close the branch. Feasibility is Evidence-Pending.
Hazard trace | HA-TFC-001, HA-ENRG-001 · Req/DO trace: HWR-007, SYS-095 / DO-EE-007, DO-SAFE-001 / RCR-005/046 · V&V: PV-TFC-001 (touch-temp), VAL-TFC-001.
Owner / Status / Evidence / Notes | Power + Mechanical / Conditional-Open / Assumption (self-heat under evaluation, CMP-013 Evidence-Pending). AP applies only if adopted.
7. Traceability Notes
• Every record above carries a HAL hazard link. Safety-significant seal/thermal/retention modes reconcile to HA-MECH-004/009, HA-ENV-003/007, HA-TFC-001/002/003/006, HA-ENRG-003, HA-SW-001/008, HA-USE-001/004/005, HA-DATA-001/005. No orphan safety modes were found in this module (all trace to a HAL situation).
• Battery mechanical modes (FA-DFMEA-008/009) are FTA basic events for top event (c) Li-battery thermal runaway/fire (FAL-1800) and single-fault items governed by DO-SAFE-010.
• Cold-embrittlement (FA-DFMEA-005/006) and condensation (FA-DFMEA-001/019) are common-cause candidates (cold environment as shared initiator) — see FAL-2700.
• Requirement/DO coverage: DO-MECH-002/004/005/007/008/009/010/011/012/013/014/015/017/018 and DO-SAFE-001/010/011 all have at least one failure mode analysed; DO-MECH-016 (external-antenna aperture) is analysed as an interface/ ingress item in FAL-1500.
8. Closure / Knowledge Notes
Top DFMEA items | cold-cycle seal integrity over reuse (FA-DFMEA-001), probe gland seal/strain-relief (FA-DFMEA-003), battery retention/isolation & containment (FA-DFMEA-008/009 — S5 branch), internal-sensor thermal lag (FA-DFMEA-010). These carry the D3 cold environment and transport-abuse themes.
• All O/D are pre-validation estimates (Evidence-Pending); the named PV-MECH/PV-TFC/PV-ENV tests close them.
• Recurring pattern for the AI/Historical KB (FAL-2200/2300): cold environment simultaneously embrittles polymer, stiffens seals/buttons, and drives condensation — a shared-initiator theme to watch on all cold-chain IoT enclosures.
FAL-1200 — Software FMEA (SWFMEA) — Smart Cold Chain Monitor (SCM)
Module FAL-1200 (FA-SWFMEA) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Embedded Software Engineer (EM-ENG-007). Software FMEA for the SCM device firmware and its safety-relevant logic: excursion-detection engine, verdict engine, state/mode machine, calibration application, storage manager, power-state governor, comms agent, alert generation, OTA agent, fault handler/recovery, security/crypto module. Scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3; IEC 62304 §7 software risk). Every safety-significant mode reconciles to a HAL hazard (HA-…). Feeds Stage 12 RCR, Stage 13 Verification, FTA (FAL-1800), CCF (FAL-2700).
---
1. Purpose
Identify software failure modes (systematic faults, resource/timing faults, logic errors, and mis-handled external faults), their mechanisms, their local→system→user/safety effects, current controls, and recommended actions — tied to firmware components (CMP-035…045), requirements (SYS-/SRS-/FWR-), design outputs (DO-FW/ DO-SAFE), hazards (HA-…), and verifications (PV-SW-…). Governing invariant (fault-management architecture): honest gap, no false reassurance — an armed device is at all times verifiably monitoring or visibly flagged.
2. Scope
Excursion-detection logic (two-sided, freeze=heat, product-stability, hysteresis/debounce), verdict/disposition logic (suppress false accept), state machine (arming/charge-inhibit/OTA/safe-state invariants), calibration application, timing/scheduler, memory (exhaustion, corruption, torn writes), storage manager (append-only, power-fail-safe), power-state governor (energy honesty), comms agent (store-and-forward, alert prioritisation), alert generation/escalation (device side), OTA agent (verify/apply/rollback), fault handler/watchdog recovery, security/crypto (integrity, auth). Excludes cloud-service logic (FAL-1400 System + HAL-1200), mobile app (FAL-1500), and pure hardware faults (FAL-1300) except how software detects/handles them.
3. Applicable Standards & References
IEC 62304 (software life-cycle & software risk management), ISO 14971, AIAG-VDA/IEC 60812, IEC 62366-1 (use), ALCOA+ (data integrity), CWE/CERT-C (defect classes), IEC 62443 / OWASP (security, cross-ref FAL-2700 & HAL-1300). Cited by designation; revision to confirm.
4. Governance & Method
Shared ERFKL scoring key (FAL-1000 §4): S outcome-fixed; O/D 1–10 with basis in Notes; Action Priority severity-dominant (S≥7 → High unless O≤3 & D≤4 → Medium). Software occurrence reflects systematic fault likelihood (design/defect escape probability) — stated as Judgment and reduced by verification rigor, not by runtime redundancy alone. D reflects runtime detectability (plausibility, BIST, self-check, watchdog, cloud reconciliation). Numbers are pre-validation (Evidence-Pending) unless a released test exists.
5. Categories
Detection logic · Verdict logic · Calibration application · State machine/invariants · Timing/scheduler · Memory/resource · Storage/persistence · Power/energy governor · Comms/store-and-forward · Alert generation · OTA/recovery · Security/crypto · Diagnostics/self-check.
6. Risk & Failure Register (FA-SWFMEA)
6.1 Excursion-detection logic (CMP-039 — DO-FW-006/DO-SAFE-004)
FA-SWFMEA-001 ★ — One-sided detection misses a cold/freeze excursion
Description | Detection engine reliably catches heat excursions but not the freeze (low) side.
Category / Subsystem-Component | Detection logic · S3 · CMP-039 (DO-FW-006).
Cause → Mechanism → Failure Mode | Threshold/logic asymmetry (heat-only heritage) → missed freeze excursion.
Effect (local→system→user/safety) | Freeze-damaged (freeze-sensitive) vaccine not flagged → false accept → compromised dose administered.
S / O / D / AP | 8 / 3 / 6 / High (S4)
Current Controls | Prev — two-sided detection with freeze reliability = heat (DO-SAFE-004), profile-driven thresholds. Det — profiled-excursion test both sides (PV-SW-001), cloud recompute.
Recommended Action | Symmetric two-sided logic; freeze test coverage equal to heat (PV-SW-001, VAL-SW-001); static analysis of threshold comparisons.
Hazard trace | HA-SW-001 ★, HA-ENV-001 · Req/DO trace: SYS-149/150, SYS-008/009 / DO-FW-006, DO-SAFE-004 / RCR-019 · V&V: PV-SW-001 (profiled excursions, FI), VAL-SW-001.
Owner / Status / Evidence / Notes | EM-ENG-007 + Safety / Open / Judgment. Classic cold-chain escape (heat-biased logic). FTA basic event, top event (a).
FA-SWFMEA-002 — Time/cumulative excursion criterion wrong → missed slow excursion
Description | Product-stability logic (min-time / cumulative-exposure) misconfigured, missing a real but slow/cumulative excursion.
Category / Subsystem-Component | Detection logic · S3 · CMP-039 (DO-FW-006).
Cause → Mechanism → Failure Mode | Wrong time-window/cumulative math (mean-kinetic-temperature style) → missed cumulative excursion.
Effect | Cumulatively degraded dose accepted → compromised dose.
S / O / D / AP | 8 / 3 / 6 / High (S4)
Current Controls | Prev — product-stability excursion logic per profile (SYS-010/014), cloud recompute cross-check. Det — profiled cumulative-excursion tests; device↔cloud verdict consistency.
Recommended Action | Ground cumulative logic in the stability profile; test cumulative/slow ramps; device and cloud must agree (independent recompute).
Hazard trace | HA-SW-001, HA-SW-003 · Req/DO trace: SYS-010/014 / DO-FW-006 / RCR-019/048 · V&V: PV-SW-001, VAL-SW-001.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
FA-SWFMEA-003 — Over-aggressive hysteresis/debounce masks a real excursion
Description | Debounce/hysteresis tuned to suppress noise also suppresses a genuine brief excursion.
Category / Subsystem-Component | Detection logic · S3 · CMP-039 (DO-FW-006).
Cause → Mechanism → Failure Mode | Excess hysteresis/consolidation → real excursion filtered out.
Effect | Missed excursion → compromised dose; opposite tuning → false rejects (waste).
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — hysteresis/debounce bounded by stability profile, sensor time-constant evidence (FA-DFMEA-010). Det — profiled brief-excursion tests.
Recommended Action | Tune debounce to real sensor dynamics; bound worst-case masked duration below product tolerance; test brief excursions at the boundary.
Hazard trace | HA-SW-001, HA-MPE-001 (false-reject side) · Req/DO trace: SYS-009/014 / DO-FW-006 / RCR-019 · V&V: PV-SW-001.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Two-directional harm (miss vs waste).
FA-SWFMEA-004 — Detection runs against the wrong/default profile
Description | Device monitors against wrong thresholds because the trip profile was not applied or defaulted.
Category / Subsystem-Component | Detection/config · S3 · CMP-039 (DO-FW-032, DO-SAFE-007).
Cause → Mechanism → Failure Mode | Profile not loaded / wrong profile / silent default → wrong-threshold monitoring.
Effect | Missed excursion (thresholds too wide) → compromised dose; or false reject.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — pre-arm profile validation, no silent default (DO-FW-032, DO-SAFE-007), arming requires a valid profile. Det — pre-arm check; profile echoed in record and to cloud.
Recommended Action | Block arming without a validated profile; display/confirm active profile; cloud verifies profile-vs-lot; test pre-arm/profile mismatch (PV-SW-003).
Hazard trace | HA-SW-003, HA-USE-004 · Req/DO trace: SYS-154/024 / DO-FW-032, DO-SAFE-007 / RCR-048/045 · V&V: PV-SW-003 (profile/pre-arm), VAL-SW-003, VAL-USE-004.
Owner / Status / Evidence / Notes | EM-ENG-007 + HF / Open / Judgment. Overlaps use-error (forgot-to-arm).
6.2 Sensor conditioning / calibration application (CMP-038 — DO-FW-007/009)
FA-SWFMEA-005 ★ — Calibration coefficient/units error → biased readings
Description | Firmware applies wrong calibration coefficients, wrong units, or wrong sign, biasing all readings.
Category / Subsystem-Component | Calibration · S3 · CMP-038 (DO-FW-007/009).
Cause → Mechanism → Failure Mode | Coefficient/units/scale/sign defect or wrong cal table → systematic measurement bias.
Effect | Fleet-wide biased readings → wrong disposition in both directions (compromised dose or waste); a common-cause across the fleet.
S / O / D / AP | 8 / 3 / 6 / High (S4)
Current Controls | Prev — calibration schema + application review (DO-FW-009), static analysis, unit tests, EOL calibration verification. Det — plausibility/range, cross-channel, device↔cloud consistency.
Recommended Action | Static analysis + unit tests on cal application; independent verification of the cal pipeline; EOL cal coverage (PV-MFG-001); treat as CCF (shared firmware) — FAL-2700.
Hazard trace | HA-SW-014, HA-SW-004, HA-DATA-001 ★, HA-MFG-001 · Req/DO trace: SYS-007/024 / DO-FW-007/009 / RCR-026 · V&V: PV-SW-004 (calibration integrity + static analysis), PV-DATA-001.
Owner / Status / Evidence / Notes | EM-ENG-007 + Metrology / Open / Judgment. Systematic (all units) → fleet CCF; worst-case S5 if it defeats detection at scale.
FA-SWFMEA-006 — Expired/invalid calibration used silently
Description | Firmware uses out-of-date or invalid calibration without blocking arming.
Category / Subsystem-Component | Calibration · S3 · CMP-038 (DO-FW-007).
Cause → Mechanism → Failure Mode | No expiry/integrity check → stale/invalid cal applied.
Effect | Drifted readings → wrong disposition.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — calibration expiry + integrity check blocks arming (DO-SAFE-007), depot recal interval. Det — pre-arm cal-validity gate; cal metadata in record.
Recommended Action | Expiry/integrity gate on arm; recal interval enforcement (PV-SVC-001); flag cal status in the record.
Hazard trace | HA-SW-004, HA-SVC-001 · Req/DO trace: SYS-024 / DO-FW-007/009, DO-SAFE-007 / RCR-045/049 · V&V: PV-SW-004, PV-SVC-001.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
FA-SWFMEA-007 — Sensor plausibility check fails to flag a faulted channel
Description | Plausibility/fault logic passes an out-of-range or stuck sensor as valid.
Category / Subsystem-Component | Acquisition · S3/S4 · CMP-038 (DO-SAFE-005).
Cause → Mechanism → Failure Mode | Weak range/stuck/rate checks → faulted channel accepted as valid.
Effect | Valid-looking but wrong reading → false accept → compromised dose.
S / O / D / AP | 8 / 3 / 6 / High (S4)
Current Controls | Prev — range/stuck/rate plausibility + probe-presence (DO-SAFE-005), int-vs-ext divergence. Det — probe-fault-injection (PV-SW-008); flag blocks clean pass.
Recommended Action | Full plausibility checks (range, stuck, rate-of-change, cross-channel); fault-inject; block "accept" on any sensor fault.
Hazard trace | HA-SW-008 ★, HA-USE-005, HA-DATA-001 · Req/DO trace: SYS-007/151 / DO-SAFE-005, DO-FW-007 / RCR-020/045 · V&V: PV-SW-008 (probe fault injection).
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Cross-ref FA-HWFMEA-017/027.
6.3 Verdict / disposition logic (CMP-040 — DO-FW-008/DO-SAFE-006)
FA-SWFMEA-008 ★ — Verdict shows "accept" on an incomplete record
Description | Verdict engine returns accept despite gaps/faults in the trip record.
Category / Subsystem-Component | Verdict logic · S3 · CMP-040 (DO-FW-008, DO-SAFE-006).
Cause → Mechanism → Failure Mode | Gap/fault not treated as disqualifying → false "accept" on incomplete data.
Effect | Lot presented acceptable on incomplete record → false accept → compromised dose.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — verdict suppresses false accept on gaps/faults (DO-SAFE-006); quarantine on incompleteness; cloud recompute. Det — gap-handling FI (PV-SW-002); device↔cloud consistency.
Recommended Action | Any gap/fault forces quarantine/inconclusive, never silent accept; test gap handling (PV-SW-002, VAL-SW-002); expose incompleteness on display and dashboard.
Hazard trace | HA-SW-002 ★, HA-DATA-004, HA-SW-007 · Req/DO trace: SYS-016/017/151 / DO-FW-008, DO-SAFE-006 / RCR-021/022 · V&V: PV-SW-002 (gap-handling FI), VAL-SW-002.
Owner / Status / Evidence / Notes | EM-ENG-007 + Safety / Open / Judgment. Core no-false-reassurance control. FTA basic event, top event (a).
FA-SWFMEA-009 — Device verdict diverges from cloud recomputation
Description | On-device verdict and cloud recompute disagree, and the divergence is not surfaced/resolved.
Category / Subsystem-Component | Verdict logic · S3/S8 · CMP-040/CMP-049.
Cause → Mechanism → Failure Mode | Logic/version skew between device and cloud → inconsistent verdict with no reconciliation.
Effect | Operator/cloud act on different answers → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — deterministic verdict consistent with cloud recompute (CMP-040), version pinning. Det — device↔cloud consistency check flags divergence.
Recommended Action | Independent recompute + explicit reconciliation; surface any divergence; version-skew test (link OTA FA-SWFMEA-024).
Hazard trace | HA-SW-007, HA-COMM-004 · Req/DO trace: SYS-016/017 / DO-FW-008, DO-CLD-006 / RCR-021/030 · V&V: PV-SW-002, PV-COMM-004.
Owner / Status / Evidence / Notes | EM-ENG-007 + Cloud / Open / Judgment.
FA-SWFMEA-010 — Display renders a false/stale verdict
Description | OLED shows a verdict that does not match the computed disposition (stale cache, render bug).
Category / Subsystem-Component | Verdict/HMI · S3/S6 · CMP-040 (DO-FW-008, DO-HF-002).
Cause → Mechanism → Failure Mode | Stale buffer / render defect / state desync → displayed verdict ≠ true verdict.
Effect | Operator acts on a displayed false accept → compromised dose.
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — display-integrity control (DO-SAFE-006, DO-HF-002), verdict/render consistency check. Det — display-integrity FI (PV-SW-007).
Recommended Action | Verify displayed vs computed verdict; no stale render; test display integrity (PV-SW-007, VAL-USE-001).
Hazard trace | HA-SW-007, HA-USE-001 · Req/DO trace: SYS-016/017/041 / DO-FW-008, DO-HF-002 / RCR-021/043 · V&V: PV-SW-007 (display-integrity FI), VAL-USE-001.
Owner / Status / Evidence / Notes | EM-ENG-007 + HF / Open / Judgment.
6.4 State machine / invariants (CMP-036 — DO-FW-003)
FA-SWFMEA-011 ★ — Armed state not entered / trip runs unmonitored
Description | Arming state machine leaves device un-armed while presenting as ready, so the trip is not monitored.
Category / Subsystem-Component | State machine · S3 · CMP-036 (DO-FW-003, FWR-031).
Cause → Mechanism → Failure Mode | Arming SM defect / race / incomplete arm → unmonitored trip believed armed.
Effect | No monitoring → silent gap reads uneventful → compromised dose.
S / O / D / AP | 8 / 2 / 5 / High (S4)
Current Controls | Prev — explicit arm with positive confirmation, unmistakable armed/unarmed/voided state (SYS-023, FWR-031), pre-arm gate. Det — cloud sees no telemetry from an "armed" device (stale).
Recommended Action | Positive-confirmation arm; make unarmed/voided unmistakable; cloud deploy-block on non-ready; test arming edge cases (PV-SW-003).
Hazard trace | HA-USE-004 ★, HA-SW-003 · Req/DO trace: SYS-023/024 / DO-FW-003/032, DO-SAFE-007 / RCR-045/040 · V&V: PV-SW-003, VAL-USE-004.
Owner / Status / Evidence / Notes | EM-ENG-007 + HF / Open / Judgment. Overlaps HA-USE-004.
FA-SWFMEA-012 — Charge-inhibit invariant violated by a state transition (FW side)
Description | A state/mode transition allows charging while the cell is below the safe temperature.
Category / Subsystem-Component | State machine/safety invariant · S3/S2 · CMP-036/037 (DO-FW-005).
Cause → Mechanism → Failure Mode | FW logic error lets charge path enable in a cold-charge condition → charge-inhibit invariant violated.
Effect | Cold charge → plating → thermal runaway/fire (S5).
S / O / D / AP | 9 / 2 / 5 / High (S5)
Current Controls | Prev — charge-inhibit is a non-tunable invariant, HW-anchored interlock backstops FW (DO-EE-007/DO-SAFE-001), FW cross-checks NTC. Det — HW interlock catches FW error; state legibility.
Recommended Action | HW-primary interlock so FW error cannot cause cold charge; formal review of charge-state transitions; fault-inject (PV-ENRG-001).
Hazard trace | HA-ENRG-001 ★, HA-TFC-002 (S5) · Req/DO trace: SYS-091/094 / DO-FW-005, DO-EE-007, DO-SAFE-001 / RCR-001/002 · V&V: PV-ENRG-001 (cold-charge FI).
Owner / Status / Evidence / Notes | EM-ENG-007 + Safety / Open / Judgment. Defence-in-depth with FA-HWFMEA-006 (HW interlock); CCF avoided by HW/FW diversity — FAL-2700.
FA-SWFMEA-013 — Safe-state not entered on a detected fault
Description | Fault handler fails to enter the defined safe state, leaving the device in an unsafe/misleading mode.
Category / Subsystem-Component | State machine/fault · S3 · CMP-036 (DO-FW-013, DO-SAFE-007).
Cause → Mechanism → Failure Mode | Fault arbiter defect / unhandled fault → no safe-state entry.
Effect | Device continues in a faulted state → false reassurance / corrupted record → wrong disposition.
S / O / D / AP | 7 / 2 / 5 / High (S4)
Current Controls | Prev — fault arbiter → safe state, preserve record, honest indication (DO-SAFE-007). Det — watchdog/self-check; annunciation.
Recommended Action | Enumerate faults→safe-state mapping; default-deny (unknown fault → safe state); FI test (PV-SW-009).
Hazard trace | HA-SW-009 · Req/DO trace: SYS-159/167/170 / DO-FW-013, DO-SAFE-007 / RCR-023/027 · V&V: PV-SW-009 (fault-injection).
Owner / Status / Evidence / Notes | EM-ENG-007 + Safety / Open / Judgment.
6.5 Timing / scheduler (CMP-036/017 — DO-FW-002)
FA-SWFMEA-014 — Scheduler misses a sampling window (missed sample gap)
Description | Deterministic scheduler drops or delays a scheduled sample.
Category / Subsystem-Component | Timing · S3 · CMP-036 (DO-FW-002).
Cause → Mechanism → Failure Mode | Task overrun / wake-timer fault / priority error → missed/late sample.
Effect | Gap in record; brief excursion missed → wrong disposition; but gap is recorded (honest).
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — deterministic scheduling, wake-on-RTC (DO-FW-002), gap-marking. Det — record shows interval gap; cloud detects sampling-cadence anomalies.
Recommended Action | Verify schedule margins; mark any missed sample as a gap (never interpolate silently); timing stress test.
Hazard trace | HA-SW-001, HA-DATA-006 · Req/DO trace: SYS-001/002 / DO-FW-002 / RCR-024 · V&V: PV-SW (timing), PV-SW-002.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Honest-gap invariant limits harm.
FA-SWFMEA-015 — Timestamp mis-application → mis-timed record
Description | Firmware applies a wrong/unsynced timestamp to samples/events.
Category / Subsystem-Component | Timing · S3 · CMP-036/017.
Cause → Mechanism → Failure Mode | Clock read race / drift not disclosed → mis-timed record.
Effect | Excursion timing wrong / record indefensible → wrong disposition; audit/GDP finding.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — monotonic UTC, disclosed free-run drift (SYS-118), re-discipline on link. Det — time-source health flag; drift bounds.
Recommended Action | Atomic time read; disclose drift/untrusted-time in record; bound drift (PV-SW-005).
Hazard trace | HA-SW-005, HA-COMM-005 · Req/DO trace: SYS-006/118 / DO-FW, DO-EE-011 / RCR-025 · V&V: PV-SW-005 (clock-drift over cold).
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Cross-ref FA-HWFMEA-014 (RTC).
6.6 Memory / resource (DO-FW-002/017)
FA-SWFMEA-016 — Memory exhaustion / leak → reset or dropped work
Description | RAM/heap exhaustion or leak causes a reset or dropped sampling/upload work.
Category / Subsystem-Component | Resource · S3 · CMP-036.
Cause → Mechanism → Failure Mode | Leak / fragmentation / unbounded buffer → out-of-memory reset or dropped task.
Effect | Reset (logged recovery) or silent dropped sample → gap → wrong disposition if unlogged.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — static allocation / bounded buffers, no dynamic alloc in safety path, watchdog. Det — watchdog reset logged; heap high-watermark BIST.
Recommended Action | Static allocation for safety path; bounded queues; soak test for leaks; log any reset as recovery (no silent continuity assumption).
Hazard trace | HA-SW-009, HA-SW-013 · Req/DO trace: SYS-169/170 / DO-FW-002, DO-SAFE-007 / RCR-023 · V&V: PV-SW-009, soak/endurance test.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
FA-SWFMEA-017 — RAM corruption (SEU/EMI) flips a state or sample
Description | Bit corruption in RAM (radiation/EMI/marginal) corrupts a sample or a state variable.
Category / Subsystem-Component | Resource/integrity · S3 · CMP-036 (DO-EE-019).
Cause → Mechanism → Failure Mode | SEU/EMI/marginal-timing → corrupted state or reading.
Effect | Wrong reading/verdict/state → wrong disposition.
S / O / D / AP | 6 / 2 / 6 / Medium (S3)
Current Controls | Prev — critical-variable redundancy/CRC, EMC design (DO-EE-019), plausibility. Det — CRC/ redundancy checks; plausibility; watchdog.
Recommended Action | CRC/duplicate safety-critical variables; plausibility on readings; EMC immunity (PV-ELEC-004).
Hazard trace | HA-ELEC-004, HA-SW-009 · Req/DO trace: SYS / DO-FW-013, DO-EE-019 / RCR-014 · V&V: PV-ELEC-004, fault-injection.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
6.7 Storage / persistence (CMP-041 — DO-FW-010/DO-SAFE-008)
FA-SWFMEA-018 ★ — Power-fail during NV write → torn/corrupt record
Description | Brown-out/reset during a log write leaves a torn or corrupt record (esp. cold voltage sag, D3).
Category / Subsystem-Component | Persistence · S3/S7 · CMP-041 (DO-FW-010, DO-SAFE-008).
Cause → Mechanism → Failure Mode | Non-atomic write interrupted by power loss → torn/corrupt record.
Effect | Silent gap or corrupt record that could read as valid → wrong disposition / false reassurance.
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — power-fail-safe atomic writes (SRS-REL-001, DO-SAFE-008), brown-out supervisor, journaling/double-buffer. Det — torn write fails the record (not silently valid); power-fail-during-write FI.
Recommended Action | Atomic commit (journaling/log-structured); a torn write must fail, never present as valid; test power-fail-during-write across the cold sag envelope (PV-SW-013).
Hazard trace | HA-SW-013 ★, HA-ELEC-010, HA-DATA-006 · Req/DO trace: SYS-160/168, SRS-REL-001 / DO-FW-010, DO-SAFE-008 / RCR-024/022 · V&V: PV-SW-013 (power-fail-during-write).
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. FTA basic event, top event (d). Cold sag raises occurrence (D3). CCF with power path (FAL-2700).
FA-SWFMEA-019 — Append-only log allows silent overwrite of unsent data
Description | Storage manager overwrites unsent/excursion/audit data under buffer pressure.
Category / Subsystem-Component | Persistence · S3/S7 · CMP-041 (DO-FW-010).
Cause → Mechanism → Failure Mode | Wrap-around / wrong prioritisation → overwrite of unsent evidence.
Effect | Loss of excursion/audit evidence → incomplete record accepted → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — no-overwrite-of-unsent policy, excursion/audit-first prioritisation, ≥50 % spare. Det — buffer-backlog telemetry; capacity annunciation.
Recommended Action | Hard no-overwrite of unsent/excursion/audit; annunciate buffer-full; verify prioritisation (PV-SW-002).
Hazard trace | HA-SW-002, HA-DATA-006 · Req/DO trace: SYS-018/021/043 / DO-FW-010 / RCR-021/024 · V&V: PV-SW-002, storage-full test.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Cross-ref FA-HWFMEA-034.
FA-SWFMEA-020 — Exactly-once send-marking fails → duplicate or dropped records
Description | Send-marking loses idempotency, so records are duplicated or marked sent but never delivered.
Category / Subsystem-Component | Persistence/comms · S3 · CMP-041/042 (DO-FW-010/011).
Cause → Mechanism → Failure Mode | Mark-before-ack / non-idempotent → dropped or duplicate records.
Effect | Dropped → gap; duplicate → confusing/contradictory cloud record → wrong disposition.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — mark-sent only on cloud ack, idempotent upload (CMP-042/046). Det — device↔cloud reconciliation.
Recommended Action | Ack-before-mark; idempotent IDs; reconciliation; test delivery-path faults (PV-COMM-002).
Hazard trace | HA-COMM-002, HA-COMM-004 · Req/DO trace: SYS-020/111 / DO-FW-010/011, DO-CLD-006 / RCR-028/032 · V&V: PV-COMM-002, PV-COMM-004.
Owner / Status / Evidence / Notes | EM-ENG-007 + Cloud / Open / Judgment.
6.8 Power / energy governor (CMP-037 — DO-FW-004/DO-SAFE-012)
FA-SWFMEA-021 ★ — Energy governor over-spends → premature depletion (D3)
Description | Power-state manager fails to hold the duty-cycle budget (radio left on, no deep sleep), depleting early.
Category / Subsystem-Component | Energy governor · S3 · CMP-037 (DO-FW-004).
Cause → Mechanism → Failure Mode | Rail not gated / stuck-awake / retry storm → budget overrun → early depletion.
Effect | Device dies mid-trip → silent gap → compromised dose. The single most load-bearing FW component for the D3 90-day/−30 °C budget.
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — hard radio/sensor gating, lowest-viable sleep, per-class energy allocation, bounded retries (DO-FW-004), reserve for last-gasp (DO-SAFE-012). Det — energy accounting; SoC trend; fleet endurance.
Recommended Action | Enforce per-class energy budgets; watchdog on stuck-awake; bounded retry/back-off; verify in 90-day cold endurance (PRD-VER-004).
Hazard trace | HA-ENRG-006 ★, HA-SW-013 · Req/DO trace: SYS-086/089/090 / DO-FW-004, DO-SAFE-012 / RCR-006/024/027 · V&V: PRD-VER-004 (90-day cold endurance).
Owner / Status / Evidence / Notes | EM-ENG-007 + Power / Open — top program risk / Evidence-Pending (budget). Software half of the D3 tension (HW half FA-HWFMEA-001). FTA basic event, top event (a).
FA-SWFMEA-022 — Graceful-degradation logic drops the record instead of protecting it
Description | Under low energy, load-shedding sheds persistence/record before non-essential functions.
Category / Subsystem-Component | Energy governor · S3 · CMP-037 (DO-SAFE-012).
Cause → Mechanism → Failure Mode | Wrong shed priority → record/last-gasp commit sacrificed.
Effect | Lost record at end of trip → wrong disposition / indefensible.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — degradation order protects record last, reserve energy for a safe landing (SYS-164/171, DO-SAFE-012). Det — reserve accounting; last-gasp verification.
Recommended Action | Define shed order (record + safe-state protected last); reserve energy; verify last-gasp commit at depletion (PRD-VER-004).
Hazard trace | HA-ENRG-006, HA-DATA-006 · Req/DO trace: SYS-164/171 / DO-SAFE-012 / RCR-027 · V&V: PRD-VER-004, depletion test.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
6.9 Comms / store-and-forward + alert generation (CMP-042 — DO-FW-011/DO-SAFE-009)
FA-SWFMEA-023 ★ — Alert not generated / not prioritised over bulk upload
Description | Device fails to emit an excursion alert, or lets routine bulk upload block the alert.
Category / Subsystem-Component | Alert generation · S3/S5 · CMP-042 (DO-FW-011, DO-SAFE-009).
Cause → Mechanism → Failure Mode | Alert path defect / no priority override / lost queue → no/late alert.
Effect | Excursion committed, no alert in flight → delayed/missing intervention → compromised dose.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — alert-before-bulk priority (the one permitted gating override), queue-and-forward when offline, cloud escalation of unacked alerts (DO-SAFE-009, DO-CLD-007/008). Det — alert-gen + escalation FI; un-acked alert rate.
Recommended Action | Alert prioritised over bulk; persist alert queue across resets; cloud escalates unacked; test alert-gen + escalation (PV-SW-006, VAL-SW-006).
Hazard trace | HA-SW-006 ★, HA-COMM-001/002 · Req/DO trace: SYS-011/013 / DO-FW-011, DO-SAFE-009, DO-CLD-007/008 / RCR-029/028/047 · V&V: PV-SW-006 (alert-gen + escalation FI), VAL-SW-006.
Owner / Status / Evidence / Notes | EM-ENG-007 + Cloud / Open / Judgment. FTA basic event, top event (b).
FA-SWFMEA-024 — Store-and-forward loses buffered alerts/telemetry across reset
Description | Buffered alerts/data are lost on reset/brown-out because the queue is not persisted.
Category / Subsystem-Component | Comms · S3 · CMP-042 (DO-FW-011).
Cause → Mechanism → Failure Mode | Volatile queue / non-atomic dequeue → lost buffered alert on reset.
Effect | Excursion evidence/alert lost → compromised dose / gap.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — persistent, resumable, idempotent queue; ack-before-dequeue (CMP-042). Det — reconciliation; backlog telemetry.
Recommended Action | Persist queue in NV; resumable/idempotent; ack-before-dequeue; test reset-during-send (PV-COMM-001/002).
Hazard trace | HA-COMM-001/002, HA-SW-006 · Req/DO trace: SYS-111/113 / DO-FW-011 / RCR-028/029 · V&V: PV-COMM-001 (latency + S&F), PV-COMM-002.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
FA-SWFMEA-025 — Retry storm defeats energy budget or floods on reconnect
Description | Unbounded connection retries drain energy (cold) or flood the cloud on reconnection.
Category / Subsystem-Component | Comms · S3/S5 · CMP-042 (DO-FW-011).
Cause → Mechanism → Failure Mode | No back-off / no jitter → retry storm → energy drain or ingest flood.
Effect | Early depletion (→ FA-SWFMEA-021) or cloud overload / delayed processing.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — bounded retry with exponential back-off + jitter, energy-aware bearer selection. Det — energy accounting; ingest-rate monitoring.
Recommended Action | Bounded back-off + jitter; energy-capped retries; test coverage-gap + reconnect storms.
Hazard trace | HA-ENRG-006, HA-CLOUD-008 · Req/DO trace: SYS-108/112 / DO-FW-011 / RCR-039 · V&V: PV-COMM (roaming/coverage), soak.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
6.10 OTA / recovery (CMP-043 — DO-FW-015/DO-SAFE-014)
FA-SWFMEA-026 ★ — Bad OTA deploys wrong logic or bricks the device
Description | A firmware update deploys defective detection/verdict logic, or fails to apply and bricks the unit.
Category / Subsystem-Component | OTA · S3/S5 · CMP-043 (DO-FW-015, DO-SAFE-014).
Cause → Mechanism → Failure Mode | Bad image / failed activation / no rollback → fleet unsafe/blind or bricked.
Effect | Fleet-scale wrong logic or blind devices → compromised doses at scale (S4/S5).
S / O / D / AP | 8 / 2 / 4 / High (S4)
Current Controls | Prev — signed images, A/B or rollback, staged campaign, never-brick + unmonitored-window notice (DO-SAFE-014, DO-FW-014/015, DO-CLD-010). Det — rollback + version-skew test; campaign health.
Recommended Action | Signed, staged, resumable OTA with auto-rollback; canary before fleet; version-skew handling; test rollback (PV-SW-010).
Hazard trace | HA-SW-010 ★, HA-CYBER-005, HA-SVC-002 · Req/DO trace: SYS-032/139 / DO-FW-014/015, DO-SAFE-014, DO-CLD-010 / RCR-037 · V&V: PV-SW-010 (rollback + version-skew), PV-CYBER-005 (IV).
Owner / Status / Evidence / Notes | EM-ENG-007 + Security / Open / Judgment. Fleet-scale → high impact; staged rollout limits occurrence. FTA contributor to top events (a/b/d).
FA-SWFMEA-027 — OTA suspends monitoring without disclosing the gap
Description | Update window stops sampling but the resulting gap is not marked/annunciated.
Category / Subsystem-Component | OTA · S3 · CMP-043 (DO-SAFE-014).
Cause → Mechanism → Failure Mode | No monitoring-suspension notice → undisclosed gap during OTA.
Effect | Unmonitored window reads uneventful → wrong disposition.
S / O / D / AP | 6 / 2 / 4 / Medium (S3)
Current Controls | Prev — OTA continues monitoring or marks/annunciates any suspension (DO-SAFE-014, CMP-043). Det — gap-marking; campaign scheduling avoids active trips.
Recommended Action | No OTA during an armed trip without marking the window; prefer non-suspending update; test.
Hazard trace | HA-SW-010, HA-DATA-006 · Req/DO trace: SYS-032 / DO-SAFE-014 / RCR-037 · V&V: PV-SW-010.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
FA-SWFMEA-028 — Recovery assumes continuous monitoring after a reset (false continuity)
Description | After watchdog/brown-out recovery, firmware resumes without logging the gap, implying continuity.
Category / Subsystem-Component | Recovery · S3 · CMP-045 (DO-FW-013).
Cause → Mechanism → Failure Mode | Recovery path omits recovery-event logging → false continuity assumption.
Effect | Gap during the fault reads uneventful → wrong disposition / false reassurance.
S / O / D / AP | 7 / 2 / 5 / High (S4)
Current Controls | Prev — every recovery logged as a recovery event; trip never assumed continuous (SYS-167/169, DO-FW-013). Det — recovery events in record; availability self-check.
Recommended Action | Log every reset/recovery with the gap; cloud reflects recovery; test recovery logging (PV-SW-009).
Hazard trace | HA-SW-009, HA-SW-013 · Req/DO trace: SYS-167/169/170 / DO-FW-013, DO-SAFE-007 / RCR-023/024 · V&V: PV-SW-009.
Owner / Status / Evidence / Notes | EM-ENG-007 + Safety / Open / Judgment. Core no-false-reassurance invariant.
6.11 Diagnostics / availability self-check (CMP-045 — DO-FW-012)
FA-SWFMEA-029 — Availability self-check fails to detect a silent non-monitoring state
Description | The self-check that proves the device is monitoring fails, so a hung/non-monitoring device is not flagged.
Category / Subsystem-Component | Diagnostics · S3 · CMP-045 (DO-FW-012).
Cause → Mechanism → Failure Mode | Self-check shares the faulted path / not independent → silent non-monitoring undetected.
Effect | Fleet believes device monitoring when it is not → compromised dose.
S / O / D / AP | 7 / 2 / 6 / High (S4)
Current Controls | Prev — independent availability self-check + external heartbeat (SYS-169), watchdog. Det — cloud flags stale/last-seen; missed heartbeat.
Recommended Action | Self-check independent of the monitored path; external heartbeat + cloud staleness detection; FI (PV-SW-009).
Hazard trace | HA-SW-009 · Req/DO trace: SYS-169/170 / DO-FW-012, DO-SAFE-007 / RCR-023/027 · V&V: PV-SW-009.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment. Independence is the key control (CCF).
FA-SWFMEA-030 — POST/BIST passes a faulted device at arming
Description | Pre-arm POST/BIST fails to catch a sensor/storage/energy fault, arming a device that shouldn't be.
Category / Subsystem-Component | Diagnostics · S3 · CMP-045 (DO-FW-012).
Cause → Mechanism → Failure Mode | Incomplete pre-arm coverage → faulted device armed.
Effect | Trip monitored by a faulted device → wrong disposition.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — pre-arm readiness gate (SoC/SoH, cal validity, sensor, storage, firmware) blocks arm (CMP-045, DO-SAFE-007). Det — pre-arm check results; cloud deploy-block.
Recommended Action | Full pre-arm coverage; block arm on any fail; surface reason; test (PV-SW-003).
Hazard trace | HA-USE-004, HA-SW-009 · Req/DO trace: SYS-024/030 / DO-FW-012, DO-SAFE-007 / RCR-045/023 · V&V: PV-SW-003.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
6.12 Security / crypto (FW side) (CMP-044 — DO-FW-016)
FA-SWFMEA-031 — Record integrity/tamper-evidence not enforced → falsifiable record
Description | Firmware fails to apply tamper-evidence to records/audit, allowing undetected alteration.
Category / Subsystem-Component | Security · S3 · CMP-044 (DO-FW-016, DO-SEC).
Cause → Mechanism → Failure Mode | Missing/weak signing/hash-chain → falsifiable record.
Effect | False "clean" record from injected/altered data → compromised dose; data-integrity finding.
S / O / D / AP | 8 / 2 / 5 / High (S4)
Current Controls | Prev — secure-element-backed tamper-evidence over records/audit (CMP-044, DO-SEC). Det — integrity verification at ingest; independent pen-test (IV).
Recommended Action | Hash-chain/sign records with the secure element; verify at ingest; independent adversarial test (PV-CYBER-001).
Hazard trace | HA-CYBER-001 ★, HA-DATA-006, HA-MPE-004 · Req/DO trace: SYS-026/028/032 / DO-FW-016, DO-SEC, DO-SAFE-008 / RCR-034 · V&V: PV-CYBER-001 (integrity/pen-test, IV).
Owner / Status / Evidence / Notes | EM-ENG-007 + Security / Open / Judgment. Security-induced safety → independent adversarial verification required.
FA-SWFMEA-032 — Authentication/session failure blinds the device
Description | Firmware cannot authenticate/resolve a cloud session (cert/time/DNS), so it cannot upload/alert.
Category / Subsystem-Component | Security/comms · S3 · CMP-044 (DO-FW-016).
Cause → Mechanism → Failure Mode | Cert expiry / bad time / handshake fault → auth failure → no cloud link.
Effect | Device blinded → undelivered alert / backlog → compromised dose (if excursion).
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — cert/time management, trusted RTC, S&F buffering (CMP-044/042). Det — auth/handshake diagnostics; backlog telemetry.
Recommended Action | Robust cert rotation + time discipline; S&F preserves data during outage; test cert/time/ DNS faults (PV-COMM-005).
Hazard trace | HA-COMM-005 · Req/DO trace: SYS-137/118 / DO-FW-016, DO-EE-011 / RCR-025/032/039 · V&V: PV-COMM-005 (cert/time/DNS FI), VAL-COMM-005.
Owner / Status / Evidence / Notes | EM-ENG-007 + Security / Open / Judgment. Time dependency is a CCF item.
FA-SWFMEA-033 — Alarm/alert suppression accepted from an unauthenticated source
Description | Firmware honours a mute/suppress command that is not authenticated, letting an attacker hide an alert.
Category / Subsystem-Component | Security/alert · S3 · CMP-044 (DO-FW-016, DO-SAFE-009).
Cause → Mechanism → Failure Mode | Unauthenticated control accepted → alarm suppressed.
Effect | Excursion hidden by muted/blocked alarm → compromised dose.
S / O / D / AP | 8 / 2 / 5 / High (S4)
Current Controls | Prev — authenticated/authorised control only; suppression audited (DO-SEC, DO-CLD-audit). Det — suppression-attempt anomalies; pen-test (IV).
Recommended Action | Authenticate all control; audit every mute/suppress; independent pen-test (PV-CYBER-002).
Hazard trace | HA-CYBER-002 ★, HA-USE-002 · Req/DO trace: SYS / DO-FW-016, DO-SEC, DO-SAFE-009 / RCR-035/029 · V&V: PV-CYBER-002 (suppression pen-test, IV).
Owner / Status / Evidence / Notes | EM-ENG-007 + Security / Open / Judgment. Cross-ref HAL-1300; IV required.
FA-SWFMEA-034 — Debug/dev mode left enabled in firmware build
Description | A debug/service mode ships enabled, allowing record falsification or alarm suppression.
Category / Subsystem-Component | Security/config · S3 · CMP-044 (DO-FW-017, DO-EE-012).
Cause → Mechanism → Failure Mode | Build-config error / depot mode not exited → debug mode enabled in field.
Effect | Falsified record / suppressed alarm → compromised dose; repudiation.
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — production-lock / debug-lock (DO-EE-012), build-config gates, mode-exit enforcement. Det — mode-lock verification (IV); unexited-service-mode detection.
Recommended Action | Production debug-lock verified at EOL; build-config gate; detect/deny debug mode in field (PV-CYBER-006).
Hazard trace | HA-SVC-006, HA-CYBER-006 · Req/DO trace: SYS / DO-EE-012, DO-FW-017, DO-SEC / RCR-051/038 · V&V: PV-CYBER-006 (secure-boot/debug-lock, IV).
Owner / Status / Evidence / Notes | EM-ENG-007 + Security / Open / Judgment.
FA-SWFMEA-035 — Config schema/units mismatch corrupts a valid trip profile
Description | A firmware/cloud config-schema mismatch corrupts or mis-applies a valid trip profile.
Category / Subsystem-Component | Config · S3 · CMP-039/044 (DO-FW-017).
Cause → Mechanism → Failure Mode | Schema/version/units mismatch → corrupted/mis-applied profile.
Effect | Wrong-threshold monitoring → missed excursion → compromised dose (overlaps FA-SWFMEA-004).
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — versioned/validated config schema, pre-arm profile validation, cloud/device schema compatibility matrix (DO-CFG-005). Det — pre-arm validation; echo profile to cloud.
Recommended Action | Strict schema validation + version negotiation; reject invalid profile; test schema-skew (PV-SW-003).
Hazard trace | HA-SW-003, HA-SW-010 · Req/DO trace: SYS-024 / DO-FW-017, DO-CFG-005 / RCR-048 · V&V: PV-SW-003.
Owner / Status / Evidence / Notes | EM-ENG-007 / Open / Judgment.
7. Traceability Notes
• Every safety-significant SW mode reconciles to a HAL hazard: detection/verdict → HA-SW-001/002/003/007/008, HA-DATA-001/004; calibration → HA-SW-004/014, HA-DATA-001; state/recovery → HA-SW-009/013, HA-USE-004; storage → HA-SW-013, HA-DATA-006; comms/alert → HA-SW-006, HA-COMM-001/002/005; OTA → HA-SW-010, HA-SVC-002; charge-inhibit → HA-ENRG-001; security → HA-CYBER-001/002/005/006, HA-SVC-006. No orphan safety SW modes.
FTA basic events (FAL-1800) | FA-SWFMEA-001/002/008/018/021 (top events a/d), 023/024 (top event b), 026 (all), 031 (top event d). Security-induced (031/033/034) require independent adversarial verification (IV).
CCF (FAL-2700) | shared firmware/calibration (FA-SWFMEA-005), shared time (015/032), shared power path during write (018), HW/FW charge-inhibit diversity (012 vs FA-HWFMEA-006), self-check independence (029).
8. Closure / Knowledge Notes
Top SW items | energy-governor overrun (FA-SWFMEA-021 — D3), two-sided/cumulative detection (001/002), false-accept-on-gap verdict (008), power-fail-safe writes (018), alert generation/escalation (023), OTA never-brick + rollback (026), record tamper-evidence (031).
• All O/D pre-validation (Evidence-Pending); PV-SW-001…013, PV-COMM-001/002/005, PV-CYBER-001/002/005/006, PRD-VER-004 close them. Per IEC 62304, software occurrence is treated conservatively (systematic faults are not probabilistically reducible) — verification rigor, not runtime redundancy, is the primary lever.
• Recurring IoT pattern for FAL-2200/2300: heat-biased detection missing the freeze side; verdict logic that accepts on incomplete data; recovery paths that silently assume continuity; volatile alert queues lost on reset — all classic monitoring-device escapes to watch.
FAL-1300 — Hardware FMEA (HWFMEA) — Smart Cold Chain Monitor (SCM)
Module FAL-1300 (FA-HWFMEA) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Principal Electrical/Power Engineer, Sensing Engineer, and RF Engineer. Hardware FMEA for the SCM device electronics: power/battery/BMS/charge path, compute, sensors (internal/external temperature, humidity, accelerometer, GPS), PCB/PCBA, connectors, antenna. Scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3). Every safety-significant mode is reconciled to a HAL hazard (HA-…). Carries the two top program items: D3 −30 °C / 90-day cold battery, and RF-in-container. Feeds Stage 12 RCR, Stage 13 Verification, FTA (FAL-1800), CCF (FAL-2700).
---
1. Purpose
Identify hardware failure modes down to the component/function level, their physics-of-failure mechanisms, their local→system→user/safety effects, current design controls, and recommended actions — tied to components (CMP-…), requirements (HWR-/SYS-), design outputs (DO-EE/DO-PCB/DO-SAFE), hazards (HA-…), and verifications (PV-).
2. Scope
S2 power (cell, pack, BMS/AFE, secondary fuse, PMIC/charge path, fuel gauge, cell-temp NTC, self-heat, input-protection), S3 compute (MCU, secure element, RTC, watchdog), S4 sensing (int temp, ext probe, humidity, accelerometer, GNSS), PCB/PCBA + conformal coat, connectors (USB-C, probe, internal), RF front-end + antenna. Excludes firmware behaviour (FAL-1200), mechanical enclosure/seal/retention (FAL-1000), and data/protocol interface faults (FAL-1500) except at the electrical interface. Cell electro-thermal chains are the shared top item with FAL-1800 top event (c).
3. Applicable Standards & References
ISO 14971; AIAG-VDA FMEA / IEC 60812; IEC 62368-1 (safety, single-fault); IEC 62133-2 & UN 38.3 (cell safety / transport); IEC 61000 / CISPR (EMC/ESD); IPC-A-610 / IPC-2221 (PCBA/PCB); JEDEC (component reliability); MIL-HDBK-217 / Telcordia SR-332 (failure-rate methods, cross-ref FAL-1900). Cited by designation; revision to confirm — no failure-rate or datasheet number asserted as fact.
4. Governance & Method
Shared ERFKL scoring key (see FAL-1000 §4): S outcome-fixed (S5=9–10 … S1=1–2); O/D 1–10 with basis stated in Notes; Action Priority severity-dominant (S≥7 → High unless O≤3 & D≤4 → Medium). All O/D are pre-validation (Evidence-Pending) unless a released test exists; cold/RF numbers are explicitly open (D3, RF-in-container). Safety-critical protection items (BMS, secondary fuse, charge-inhibit, input protection) are single-fault-tolerance items and are also FTA basic events.
5. Categories
Cell/pack · Battery protection (BMS/fuse) · Charge path & input protection · Charge-inhibit interlock · Fuel gauge/energy · Compute/timekeeping · Sensing · PCB/PCBA/conformal · Connectors · RF/antenna.
6. Risk & Failure Register (FA-HWFMEA)
6.1 Cell & pack (S2 — CMP-006/007) — includes the D3 top item
FA-HWFMEA-001 ★ — Cell cannot deliver 90-day endurance at −30 °C (D3 top item)
Description | Selected low-temp cell's usable capacity after cold derating and reuse fade fails to meet the 90-day energy budget at −30 °C.
Category / Subsystem-Component | Cell/pack · S2 · CMP-006/007 (DO-EE-003/004).
Cause → Mechanism → Failure Mode | Cold capacity derating + increased internal resistance + reuse-cycle fade → premature energy exhaustion before trip end.
Effect (local→system→user/safety) | Device dies mid-trip → silent monitoring gap that "reads uneventful" → spoiled dose accepted → compromised dose administered.
S / O / D / AP | 7 / 5 / 7 / High (S4)
Current Controls | Prev — low-temp chemistry selection with derated-capacity budget margin (SYS-088/090), aggressive duty-cycling / energy governor (DO-FW-004), energy-reserve control (DO-SAFE-012). Det — SoC/SoH gauge + low-battery warning (HWR-008); last-gasp record protection; fleet endurance monitoring.
Recommended Action | Close chemistry selection against a measured −30 °C discharge + reuse-fade dataset (PRD-VER-004); size pack with margin; validate graceful degradation and last-gasp record commit; carry as top program risk (MR-RISK-001).
Hazard trace | HA-ENRG-006 ★, HA-SW-013, HA-DATA-006 · Req/DO trace: HWR-001/002, SYS-087/088/090 / DO-EE-003/004, DO-FW-004, DO-SAFE-012 / RCR-006/024/027 · V&V: PRD-VER-004 (90-day cold endurance).
Owner / Status / Evidence / Notes | Principal Power / Open — top program risk / Evidence-Pending (chemistry). O=5, D=7 Judgment: energy margin is Open; a silent dead-in-cold gap is the hard-to-detect case. FTA basic event for top event (a).
FA-HWFMEA-002 — Cold-charge lithium plating → latent internal short
Description | Cell charged below its safe temperature (near/below 0 °C) plates lithium, seeding a latent short.
Category / Subsystem-Component | Cell/pack · S2 · CMP-006 (DO-EE-003).
Cause → Mechanism → Failure Mode | Charge current at low cell temperature → lithium plating / dendrite formation → latent internal short that may develop over cycles.
Effect | Internal short → self-heating → thermal runaway/fire (device → shipper → aircraft hold S5).
S / O / D / AP | 9 / 3 / 7 / High (S5)
Current Controls | Prev — charge-inhibit-below-0 °C interlock (HW+FW, DO-EE-007/DO-SAFE-001/DO-FW-005); cell-coupled NTC (CMP-012); charge profile per datasheet. Det — cell-temp qualification of charge enable; protection AFE; thermal-event field reporting.
Recommended Action | Verify interlock by cold-charge fault injection (PV-ENRG-001); route NTC to cell not board-air; UN 38.3; confirm no self-heat path can bypass inhibit; latent-short screening in cycle-life.
Hazard trace | HA-ENRG-001 ★, HA-TFC-002 (S5) · Req/DO trace: HWR-006, SYS-091/094 / DO-EE-007, DO-SAFE-001, DO-FW-005 / RCR-001/002 · V&V: PV-ENRG-001 (cold-charge FI), UN 38.3.
Owner / Status / Evidence / Notes | Power + Safety / Open / Judgment. S=9 (S5 branch). Latent (delayed-effect) → D=7. FTA basic event, top event (c). Single-fault review with FA-HWFMEA-006 (interlock).
FA-HWFMEA-003 — Cell internal short / manufacturing defect → spontaneous thermal event
Description | Cell develops an internal short from a manufacturing defect (particle, separator flaw) or aging.
Category / Subsystem-Component | Cell · S2 · CMP-006.
Cause → Mechanism → Failure Mode | Separator defect/contamination + aging → internal short → self-heating.
Effect | Thermal runaway/fire; offgas in enclosed space.
S / O / D / AP | 9 / 2 / 8 / High (S5)
Current Controls | Prev — qualified cell supplier, incoming inspection (HA-MFG-002), UN 38.3, cell-level protection. Det — very limited before event (D high); fleet thermal-event surveillance.
Recommended Action | Supplier qualification + lot traceability; UN 38.3 + IEC 62133-2; containment design (FA-DFMEA-009); accept residual as an inherent cell risk minimised by protection + containment.
Hazard trace | HA-TFC-002 (S5), HA-TFC-003, HA-MFG-002 · Req/DO trace: DO-EE-003 / RCR-002/054 · V&V: UN 38.3, IEC 62133-2, PV-TFC-002/003.
Owner / Status / Evidence / Notes | Power + Quality / Open / Judgment. Inherent cell hazard; occurrence low but severity fixes AP High. FTA basic event.
FA-HWFMEA-004 — Wrong/counterfeit cell chemistry in build
Description | A non-conforming or counterfeit cell chemistry is fitted, defeating the cold-charge/endurance basis.
Category / Subsystem-Component | Cell · S2 · CMP-006 (DO-CFG-004 AML).
Cause → Mechanism → Failure Mode | Supply-chain substitution / AML escape → wrong chemistry → cold-charge and thermal thresholds no longer valid.
Effect | Cold-charge/thermal event (S4) and/or endurance shortfall (S4) → wrong disposition / fire.
S / O / D / AP | 8 / 2 / 6 / High (S4)
Current Controls | Prev — AML/approved-source control (DO-CFG-004), incoming inspection/AOI (HA-MFG-002). Det — incoming test; capacity/impedance screen.
Recommended Action | Lock cell to AML with no unapproved alternates; incoming chemistry/capacity verification; supplier audit; part-marking verification.
Hazard trace | HA-MFG-002 · Req/DO trace: DO-CFG-004, DO-MFG, DO-TST / RCR-054 · V&V: PV-MFG-002 (incoming + AOI).
Owner / Status / Evidence / Notes | Quality + Power / Open / Judgment. Manufacturing-source risk with a safety end-effect; cross-ref FAL-1100 PFMEA.
FA-HWFMEA-005 — Cell swelling/venting under abuse or aging
Description | Cell swells or vents (offgas) under overcharge, over-temperature, altitude, or aging.
Category / Subsystem-Component | Cell · S2 · CMP-006/007.
Cause → Mechanism → Failure Mode | Gas generation from electrolyte decomposition → swelling / vent → pressure on enclosure, offgas.
Effect | Enclosure distortion / seal loss; toxic offgas in enclosed space (S4); can precede runaway.
S / O / D / AP | 8 / 2 / 6 / High (S4)
Current Controls | Prev — protection limits (BMS), swell allowance in pack design (DO-EE-004/DO-MECH-012), vent management (DO-MECH-005). Det — swell not directly sensed; secondary effects.
Recommended Action | Design swell clearance; characterise offgas; altitude test (UN 38.3); protection margins.
Hazard trace | HA-TFC-003, HA-ENV-007 · Req/DO trace: DO-EE-004, DO-MECH-005/012 / RCR-002/008 · V&V: PV-TFC-003, UN 38.3 (altitude).
Owner / Status / Evidence / Notes | Power + Mechanical / Open / Judgment.
6.2 Battery protection — BMS / secondary fuse (S2 — CMP-008/009)
FA-HWFMEA-006 — Charge-inhibit interlock fails (does not block cold charge)
Description | The below-0 °C charge-inhibit fails to open the charge path when the cell is cold.
Category / Subsystem-Component | Charge-inhibit interlock · S2/S3 · CMP-012 (DO-EE-007/DO-SAFE-001).
Cause → Mechanism → Failure Mode | NTC open/short/mislocated, comparator/enable fault, or FW-only path bypassed → charge proceeds cold → plating (FA-HWFMEA-002).
Effect | Cold-charge plating → latent short → thermal runaway/fire (S5).
S / O / D / AP | 9 / 2 / 5 / High (S5)
Current Controls | Prev — HW-anchored interlock independent of FW (DO-EE-007), cell-coupled NTC (CMP-012), secondary protection backstop. Det — NTC plausibility BIST; charge-inhibit state legibility (DO-SAFE-013).
Recommended Action | Make interlock HW-primary with FW cross-check; NTC fault detection (open/short/range); fault-inject the interlock (PV-ENRG-001); single-fault review vs FA-HWFMEA-002.
Hazard trace | HA-ENRG-001 ★, HA-TFC-002 (S5) · Req/DO trace: HWR-006, SYS-094 / DO-EE-007, DO-SAFE-001 / RCR-001/002 · V&V: PV-ENRG-001 (cold-charge FI).
Owner / Status / Evidence / Notes | Power + Safety / Open — single-fault item / Judgment. FTA basic event and AND-gate partner with FA-HWFMEA-002. CCF with FW (DO-FW-005) — see FAL-2700.
FA-HWFMEA-007 — BMS/AFE fails to trip on OV/OC/OT/short
Description | Battery-protection AFE does not disconnect on over-voltage/current/temperature or short.
Category / Subsystem-Component | Battery protection · S2 · CMP-008 (DO-EE-005/DO-SAFE-002).
Cause → Mechanism → Failure Mode | AFE defect / mis-set thresholds / FET fail-short → protection does not operate on a cell fault.
Effect | Unprotected OV/OC/OT/short → thermal runaway/fire (S5).
S / O / D / AP | 9 / 2 / 6 / High (S5)
Current Controls | Prev — autonomous MCU-independent AFE (CMP-008), thresholds per cell datasheet, secondary non-resettable fuse as backstop (CMP-009). Det — protection-trip telemetry; EOL protection test.
Recommended Action | Two-tier protection (AFE + independent secondary) so no single protection failure is catastrophic; verify by protection fault injection (PV-ENRG-002); threshold verification at EOL.
Hazard trace | HA-ENRG-002, HA-TFC-002 (S5) · Req/DO trace: HWR-003, SYS-096 / DO-EE-005, DO-SAFE-002 / RCR-002 · V&V: PV-ENRG-002 (protection FI, DO-TST-009).
Owner / Status / Evidence / Notes | Power + Safety / Open — single-fault item / Judgment. Secondary fuse (FA-HWFMEA-008) is the defence-in-depth partner; CCF review required.
FA-HWFMEA-008 — Secondary non-resettable protector (fuse/OVP) fails to clear
Description | The last-line series fuse / secondary OVP does not open on a fault that passed the AFE.
Category / Subsystem-Component | Battery protection · S2 · CMP-009 (DO-EE-010/DO-SAFE-002).
Cause → Mechanism → Failure Mode | Wrong rating / degraded fuse / open-circuit-already → no last-line interruption.
Effect | With AFE also failed (AND) → unprotected fault → thermal event (S5).
S / O / D / AP | 9 / 2 / 7 / High (S5)
Current Controls | Prev — correctly rated series fuse/OVP (CMP-009), poka-yoke on rating (HA-MFG-003). Det — EOL continuity/rating check; AOI.
Recommended Action | Verify rating and single-fault coverage; poka-yoke placement/orientation; incoming + AOI (PV-MFG-003); include in FTA AND-gate with AFE.
Hazard trace | HA-ENRG-002, HA-MFG-003, HA-TFC-002 (S5) · Req/DO trace: HWR-012, SYS-096 / DO-EE-010, DO-SAFE-002 / RCR-002/054 · V&V: PV-ENRG-002, PV-MFG-003.
Owner / Status / Evidence / Notes | Power + Quality / Open / Judgment. Latent (only exercised on a fault) → D high. Defence-in-depth partner of FA-HWFMEA-007.
6.3 Charge path & input protection (S2 — CMP-010/014)
FA-HWFMEA-009 — Out-of-spec USB-C source / overvoltage at input
Description | A non-compliant charger or fault presents overvoltage/overcurrent at the USB-C input.
Category / Subsystem-Component | Charge input · S2 · CMP-014 (DO-EE-009/DO-SAFE-003).
Cause → Mechanism → Failure Mode | Bad/hostile charger → input OV/OC/surge/reverse at the sink.
Effect | PMIC/rail damage → overcharge/thermal (S3/S4) or device loss.
S / O / D / AP | 7 / 4 / 4 / High (S4)
Current Controls | Prev — ideal-diode/OVP front-end + TVS (CMP-014, DO-EE-009), USB-C sink compliance. Det — input fault flag; protection trip.
Recommended Action | Verify reverse/OVP/short tolerance (PV-ELEC-002); TVS coordination; ensure charge path gates safely on out-of-spec input.
Hazard trace | HA-ENRG-005 · Req/DO trace: HWR-009/045, SYS-102 / DO-EE-009, DO-SAFE-003 / RCR-003 · V&V: PV-ELEC-002 (reverse/OVP/short).
Owner / Status / Evidence / Notes | Power / Open / Judgment. D low: input protection is testable and its trips are observable.
FA-HWFMEA-010 — PMIC/charger fails → overcharge or no-charge
Description | Charge-path PMIC misregulates: overcharges the cell or fails to charge.
Category / Subsystem-Component | Charge path · S2 · CMP-010 (DO-EE-006).
Cause → Mechanism → Failure Mode | PMIC defect / wrong charge profile → overcharge (thermal) or charge failure (endurance/availability).
Effect | Overcharge → thermal event (S4, gated by BMS); no-charge → depot cannot ready device → stock-out / unmonitored trip.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — programmable profile per datasheet (DO-EE-006), BMS overcharge backstop. Det — charge-state telemetry; fuel-gauge; charge-complete/fault LED (DO-SAFE-013).
Recommended Action | Independent overcharge protection (BMS) verified; charge-fault annunciation; profile verification at EOL; PMIC profile Evidence-Pending until cell selected.
Hazard trace | HA-ENRG-002, HA-ENRG-005, HA-USE-009 · Req/DO trace: HWR-004/043, SYS-093/099 / DO-EE-006, DO-SAFE-002/013 · V&V: PV-ELEC-002, PV-ENRG-002.
Owner / Status / Evidence / Notes | Power / Open / Evidence-Pending (profile). No-charge path ties to HA-USE-009 (device discarded as dead).
FA-HWFMEA-011 — Rail load-switch / regulator fault (brown-out or subsystem loss)
Description | A per-rail regulator/load-switch fails, browning out the MCU or killing a subsystem rail.
Category / Subsystem-Component | Power distribution · S2 · CMP-010 (DO-EE-002).
Cause → Mechanism → Failure Mode | Regulator defect / cold voltage sag → rail brown-out or rail loss.
Effect | MCU brown-out during NV write → torn record (silent gap); or subsystem (radio/sensor) loss → missed upload/reading → wrong disposition.
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — brown-out supervisor holds reset below threshold (CMP-015, HWR-019), power-fail-safe writes (DO-FW-010/DO-SAFE-008), rail margining for cold. Det — brown-out/reset logged as recovery event.
Recommended Action | Rail margin at −30 °C sag; verify power-fail-during-write (PV-SW-013); log every reset as a recovery so continuity is never assumed.
Hazard trace | HA-ELEC-010, HA-SW-013, HA-ENRG-006 · Req/DO trace: HWR-019, SYS-099 / DO-EE-002, DO-FW-010, DO-SAFE-008 / RCR-024 · V&V: PV-SW-013 (power-fail-during-write).
Owner / Status / Evidence / Notes | Power + FW / Open / Judgment. Cold sag (D3) raises occurrence; CCF with FW write path (FAL-2700).
6.4 Fuel gauge / energy honesty (S2 — CMP-011)
FA-HWFMEA-012 — Fuel gauge over-reports SoC (esp. cold) → premature silent death
Description | Temperature-uncompensated gauge over-reports state of charge in the cold, so the device dies without warning.
Category / Subsystem-Component | Fuel gauge · S2 · CMP-011 (DO-EE-008).
Cause → Mechanism → Failure Mode | Cold impedance rise + gauge model error → SoC over-estimate → no/late low-battery warning.
Effect | Device dies mid-trip with no warning → silent gap reads uneventful → compromised dose.
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — temp-compensated coulomb/model gauge (CMP-011), reserve for last-gasp record commit (DO-SAFE-012). Det — SoC/SoH readiness; low-battery warning; fleet endurance trend.
Recommended Action | Verify gauge accuracy across temperature; conservative low-battery thresholds; reserve energy for a safe landing; validate no misleading "OK" near depletion.
Hazard trace | HA-ENRG-006 ★, HA-USE-009 · Req/DO trace: HWR-008, SYS-097 / DO-EE-008, DO-SAFE-012 / RCR-006/027 · V&V: PRD-VER-004, gauge-accuracy-over-temp test.
Owner / Status / Evidence / Notes | Power / Open / Evidence-Pending (accuracy). Honesty item: an over-reporting gauge is worse than an empty battery. Cross-ref FA-HWFMEA-001.
6.5 Compute / timekeeping (S3 — CMP-015/016/017/018)
FA-HWFMEA-013 — MCU brown-out/latch-up in cold or on transient
Description | MCU browns out or latches up under cold voltage sag or an electrical transient.
Category / Subsystem-Component | Compute · S3 · CMP-015 (DO-EE-011).
Cause → Mechanism → Failure Mode | Under-voltage / ESD/EMI latch-up → MCU reset or hang.
Effect | Monitoring stops; if believed still monitoring → silent gap → wrong disposition.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — brown-out detector + retention sleep (HWR-019), watchdog (CMP-018), ESD/EMC design (DO-EE-019). Det — watchdog reset logged as recovery; availability self-check (SYS-169).
Recommended Action | Verify BOD threshold vs cold sag; independent watchdog; log recoveries; EMC immunity (PV-ELEC-004).
Hazard trace | HA-SW-009, HA-ELEC-004, HA-ELEC-010 · Req/DO trace: HWR-013/017/019, SYS-047/099 / DO-EE-011, DO-FW-013, DO-SAFE-007 · V&V: PV-SW-009 (watchdog FI), PV-ELEC-004 (EMC).
Owner / Status / Evidence / Notes | Compute + FW / Open / Judgment. Recovery is logged → prevents false reassurance; residual is the unlogged hang (see FAL-1200 watchdog).
FA-HWFMEA-014 — RTC failure / drift → untrusted timestamps
Description | RTC stops, loses time on brown-out, or drifts unbounded, corrupting record timestamps.
Category / Subsystem-Component | Timekeeping · S3 · CMP-017 (DO-EE-011).
Cause → Mechanism → Failure Mode | RTC defect / backup-source depletion / crystal drift over cold → loss of trusted monotonic time.
Effect | Excursion timing wrong / record indefensible → wrong disposition; data-integrity finding.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — RTC + backup cell/supercap (CMP-017), monotonic UTC across sleep/brown-out. Det — time-source health flag; disclosed free-run drift; re-discipline on link.
Recommended Action | Verify hold-up across brown-out; bound and disclose drift (PV-SW-005); flag untrusted time in the record rather than presenting it as trusted.
Hazard trace | HA-SW-005, HA-COMM-005 · Req/DO trace: HWR-015, SYS-006/118 / DO-EE-011 / RCR-025 · V&V: PV-SW-005 (clock-drift over cold).
Owner / Status / Evidence / Notes | Compute / Open / Judgment. Shared time dependency is a CCF item (FAL-2700). Honesty control: disclose, don't fake.
FA-HWFMEA-015 — Secure element failure blocks boot/identity
Description | Secure element/root-of-trust fails, blocking secure boot, keys, or device identity.
Category / Subsystem-Component | Compute/security · S3 · CMP-016 (DO-EE-011, DO-SEC).
Cause → Mechanism → Failure Mode | SE defect / lockout → secure boot or auth fails → device won't run or can't authenticate.
Effect | Device dead or fleet-blinded (can't upload/route) → lost monitoring / undelivered alert.
S / O / D / AP | 6 / 2 / 3 / Medium (S3)
Current Controls | Prev — qualified SE (CMP-016 baselined), fail-safe boot. Det — POST/boot failure annunciated; registry last-seen.
Recommended Action | SE reliability qualification; graceful fail-safe indication (not silent); provisioning redundancy where feasible.
Hazard trace | HA-COMM-005, HA-SW-009 · Req/DO trace: HWR-016, SYS-032/110 / DO-EE-011, DO-SEC-006 · V&V: PV-COMM-005, POST review.
Owner / Status / Evidence / Notes | Compute + Security / Open / Judgment. Low O (baselined part) but a hard-dependency single point — see CCF.
FA-HWFMEA-016 — Watchdog fails to recover a hang
Description | Watchdog does not reset a hung MCU (shares clock, mis-serviced, or disabled).
Category / Subsystem-Component | Compute · S3 · CMP-018 (DO-FW-013).
Cause → Mechanism → Failure Mode | Watchdog clock/config fault or FW keeps servicing while logic hung → no recovery from hang.
Effect | Device believed monitoring but hung → silent non-monitoring → compromised dose.
S / O / D / AP | 7 / 2 / 6 / High (S4)
Current Controls | Prev — watchdog independent of main clock where feasible (CMP-018), windowed watchdog, availability self-check (SYS-169). Det — missed self-check → cloud flags device stale.
Recommended Action | Independent-clock windowed watchdog + external availability heartbeat; verify by FI (PV-SW-009); ensure a hung device eventually shows stale in the fleet (no false reassurance).
Hazard trace | HA-SW-009 · Req/DO trace: HWR-017, SYS-047/169 / DO-FW-013, DO-SAFE-007 · V&V: PV-SW-009 (fault-injection + watchdog).
Owner / Status / Evidence / Notes | Compute + FW / Open / Judgment. Shared-clock is a CCF defeat (FAL-2700).
6.6 Sensing (S4 — CMP-019/020/021/022/023)
FA-HWFMEA-017 ★ — External probe open/intermittent → false product-zone reading
Description | External temperature probe goes open or intermittent, so the product-zone channel is invalid.
Category / Subsystem-Component | Sensing · S4 · CMP-020 (DO-EE-013/DO-SAFE-005).
Cause → Mechanism → Failure Mode | Lead break / connector fret / cold-solder → open or intermittent probe.
Effect | No/garbage product-zone reading; if presented as valid → false accept → compromised dose.
S / O / D / AP | 7 / 4 / 5 / High (S4)
Current Controls | Prev — probe-presence/continuity detect (HWR-028), independent probe ID, plausibility (DO-SAFE-005). Det — probe-fault flag blocks clean pass; int-vs-ext divergence.
Recommended Action | Continuity + plausibility must catch intermittent not just open; fault-inject (PV-SW-008); block "accept" on probe fault; strain-relief the lead (FA-DFMEA-003).
Hazard trace | HA-SW-008 ★, HA-USE-005, HA-DATA-001 · Req/DO trace: HWR-022/028, SYS-003/007 / DO-EE-013, DO-SAFE-005 / RCR-020/045 · V&V: PV-SW-008 (probe fault injection).
Owner / Status / Evidence / Notes | Sensing / Open / Judgment. Intermittent (vs hard-open) is the hard-to-detect case. Cross-ref FA-DFMEA-003 (gland).
FA-HWFMEA-018 — Temperature-sensor drift / accuracy loss over reuse
Description | Internal or probe temperature channel drifts beyond accuracy spec over the reuse life.
Category / Subsystem-Component | Sensing · S4 · CMP-019/020 (DO-EE-013/014).
Cause → Mechanism → Failure Mode | Reference/ADC aging, thermal-cycle stress → calibration drift → biased reading.
Effect | Systematic bias → excursion missed or falsely declared → wrong disposition (fleet-wide if calibration method common — CCF).
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — stable reference/low-drift chain (HWR-027), aging-flagging, depot re-calibration interval (D1 reuse). Det — calibration expiry, cross-channel plausibility, drift trend.
Recommended Action | Recalibration interval enforced; expiry blocks arming; drift trending; verify accuracy over aging/temp (PV-DATA-001).
Hazard trace | HA-SW-004, HA-DATA-001 ★, HA-SVC-001, HA-MFG-001 · Req/DO trace: HWR-021/027, SYS-001/024 / DO-EE-014, DO-FW-007/009 / RCR-026/045/049 · V&V: PV-DATA-001, PV-SW-004.
Owner / Status / Evidence / Notes | Sensing + Metrology / Open / Evidence-Pending (accuracy class). Systematic bias is a fleet CCF (shared calibration) — FAL-2700.
FA-HWFMEA-019 — Internal sensor self-heating / placement bias
Description | Internal temperature sensor reads high from board self-heating or poor placement.
Category / Subsystem-Component | Sensing · S4 · CMP-019 (DO-EE-014, DO-MECH-015).
Cause → Mechanism → Failure Mode | Sensor near heat source / high measurement current → self-heat bias.
Effect | Biased internal reading → wrong disposition; masks a real cold excursion.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — low-self-heat design, placement/airpath (DO-MECH-015). Det — int-vs-ext divergence.
Recommended Action | Minimise measurement self-heat; characterise; prefer external probe for the verdict.
Hazard trace | HA-DATA-001, HA-SW-001 · Req/DO trace: HWR-027 / DO-EE-014, DO-MECH-015 · V&V: thermal characterisation, PV-DATA-001.
Owner / Status / Evidence / Notes | Sensing / Open / Judgment. Cross-ref FA-DFMEA-010.
FA-HWFMEA-020 — Humidity sensor condensation saturation / drift
Description | Humidity sensor saturates or drifts under condensing conditions, giving invalid RH.
Category / Subsystem-Component | Sensing · S4 · CMP-021 (DO-EE-014).
Cause → Mechanism → Failure Mode | Condensation on element / contamination → RH error / saturation.
Effect | Wrong humidity record (secondary to temperature for vaccine verdict); masks ingress indicator.
S / O / D / AP | 4 / 4 / 5 / Low–Medium (S2)
Current Controls | Prev — condensation-tolerant placement (CMP-021), protective membrane. Det — RH plausibility.
Recommended Action | Protected placement; plausibility bounds; recovery-after-condensation characterisation.
Hazard trace | HA-ENV-003, HA-TFC-006 · Req/DO trace: HWR-023, SYS-004 / DO-EE-014 · V&V: PV-ENV-003.
Owner / Status / Evidence / Notes | Sensing / Open / Judgment. RH is secondary → S2; but a saturated RH channel also loses its value as an ingress early-warning.
FA-HWFMEA-021 — Accelerometer misses / falsely reports shock
Description | Shock sensor fails to capture a real impact or reports spurious shocks.
Category / Subsystem-Component | Sensing · S4 · CMP-022 (DO-EE-015).
Cause → Mechanism → Failure Mode | Wrong full-scale/bandwidth, stuck sensor, or noisy mount → missed or false shock event.
Effect | Missed damage-relevant shock (under-report) or false alarms/energy drain (over-report).
S / O / D / AP | 5 / 3 / 5 / Medium (S3)
Current Controls | Prev — full-scale/bandwidth per profile (HWR-024), threshold/interrupt mode, self-test. Det — accel BIST; shock classification (FW).
Recommended Action | Size FSR/bandwidth to transport shock; self-test; debounce/classify to avoid false alarms.
Hazard trace | HA-MECH-009 (indirect) · Req/DO trace: HWR-024, SYS-005 / DO-EE-015 · V&V: PV (shock profile), accel self-test.
Owner / Status / Evidence / Notes | Sensing / Open / Judgment. Shock is advisory for vaccine verdict; no direct HAL safety link — noted as acceptable (advisory data).
FA-HWFMEA-022 — GNSS no-fix in container / excessive energy
Description | GPS cannot acquire inside a metallised/insulated container, or draws excessive energy trying.
Category / Subsystem-Component | Sensing · S4 · CMP-023 (DO-EE-016).
Cause → Mechanism → Failure Mode | RF attenuation in container + long acquisition → no fix / energy drain.
Effect | Missing/last-known location (advisory) and/or endurance impact (contributes to FA-HWFMEA-001).
S / O / D / AP | 4 / 5 / 3 / Low–Medium (S2)
Current Controls | Prev — event-triggered/duty-cycled acquisition, fully powered off between fixes, fail-soft last-known + quality flag (CMP-023). Det — fix-quality flag; energy accounting.
Recommended Action | Cap acquisition energy; fail-soft to last-known with quality flag; do not let GNSS compromise the energy budget; RF-in-container evidence (PRD-VER-005).
Hazard trace | (location is advisory — no direct safety HA; energy link to HA-ENRG-006) · Req/DO trace: HWR-025, SYS-117/029 / DO-EE-016 · V&V: PRD-VER-005 (RF-in-container), energy test.
Owner / Status / Evidence / Notes | Sensing + RF / Open / Evidence-Pending (in-container fix/energy). Location advisory → S2; primary concern is the energy contribution to the D3 budget.
6.7 PCB / PCBA / conformal (CMP-034)
FA-HWFMEA-023 — Latent solder/PCB defect → intermittent fault → silent gap
Description | A latent PCBA defect (cold joint, crack, void) causes intermittent operation under thermal cycling.
Category / Subsystem-Component | PCBA · S3 · CMP-034 (DO-PCB, DO-TST).
Cause → Mechanism → Failure Mode | Assembly defect + CTE cycling → intermittent open/high-resistance joint.
Effect | Intermittent sensor/storage/radio fault → silent monitoring gap or false reading → wrong disposition.
S / O / D / AP | 6 / 3 / 7 / Medium (S3)
Current Controls | Prev — DFM/IPC design, AOI/ICT coverage (DO-TST), conformal coat. Det — AOI/ICT + BIST; intermittent faults hard to catch (D high).
Recommended Action | Maximise AOI/ICT coverage (PV-MFG-008); thermal-cycle stress screen (HASS, FAL-1900); BIST for intermittent joints; watchdog + availability self-check as backstop.
Hazard trace | HA-ELEC-008, HA-MFG-008 · Req/DO trace: HWR-037, SYS / DO-PCB, DO-TST, DO-SAFE-008 / RCR-016/024 · V&V: PV-MFG-008 (AOI/ICT coverage), HASS.
Owner / Status / Evidence / Notes | PCB + Quality / Open / Judgment. D=7 classic latent-escape; cross-ref FA-DFMEA-014 (board retention). FTA basic event for silent-gap chains.
FA-HWFMEA-024 — EMC/ESD upset → false or missed excursion
Description | Radiated/conducted interference or ESD upsets sensing or logic, corrupting a reading or the state.
Category / Subsystem-Component | PCBA/EMC · CMP-034 (DO-EE-019).
Cause → Mechanism → Failure Mode | EMI/ESD coupling into sensor/ADC/logic → corrupted sample or logic upset.
Effect | False/missed excursion → wrong disposition (S3, worst-case S5 if it defeats detection at scale).
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — evidence/EMC/ESD design (DO-EE-019), filtering, conformal/shielding, plausibility checks. Det — plausibility/range checks; EMC immunity test.
Recommended Action | EMC immunity + emissions (PV-ELEC-004); input filtering + plausibility so a single corrupted sample cannot flip a verdict; ESD protection at ports.
Hazard trace | HA-ELEC-004 · Req/DO trace: SYS / DO-EE-019, DO-PCB-010 / RCR-014 · V&V: PV-ELEC-004 (EMC immunity/emissions), IEC 61000.
Owner / Status / Evidence / Notes | Electrical / Open / Judgment. S=7 because EMC upset can produce a wrong verdict; plausibility is the mitigating control.
FA-HWFMEA-025 — Conformal-coat gap → condensation leakage on critical net
Description | A coating gap leaves a critical net exposed to condensation leakage currents.
Category / Subsystem-Component | PCBA/conformal · CMP-034 (DO-PCB, DO-SAFE-011).
Cause → Mechanism → Failure Mode | Coating skip/keep-out overreach → leakage/corrosion on exposed net.
Effect | Intermittent fault on a sensor/storage/protection net → wrong disposition or defeated protection.
S / O / D / AP | 6 / 3 / 6 / Medium (S3)
Current Controls | Prev — conformal-coat spec with defined keep-outs (DO-PCB), coating inspection. Det — coating AOI; condensation-cycle test.
Recommended Action | Verify coverage on safety/sensing nets; coating inspection at EOL; condensation-cycle validation (PV-TFC-006).
Hazard trace | HA-TFC-006, HA-ELEC-008 · Req/DO trace: HWR-038/041, DO-PCB, DO-SAFE-011 / RCR-015 · V&V: PV-TFC-006, coating AOI.
Owner / Status / Evidence / Notes | PCB / Open / Judgment. Cross-ref FA-DFMEA-019 (condensation).
6.8 Connectors (CMP-004/005/034)
FA-HWFMEA-026 — USB-C connector fret/corrosion → charge failure
Description | USB-C receptacle contacts fret or corrode over depot mating cycles / moisture.
Category / Subsystem-Component | Connectors · S2 · CMP-004 (DO-EE-009, DO-MECH-008).
Cause → Mechanism → Failure Mode | Contact fretting / corrosion → high-resistance or open charge contact.
Effect | Cannot charge → depot cannot ready device → stock-out / device discarded as dead (HA-USE-009).
S / O / D / AP | 5 / 4 / 4 / Medium (S3)
Current Controls | Prev — rated connector + insertion-cycle spec (CMP-004), sealing (DO-MECH-008). Det — charge-fault annunciation; charge-state LED.
Recommended Action | Cycle-life the connector to depot reuse count; corrosion-resistant plating; charge-fault clearly annunciated so unit isn't discarded.
Hazard trace | HA-USE-009, HA-ENV-003 · Req/DO trace: HWR-039, SYS-093/114 / DO-EE-009, DO-MECH-008 · V&V: PV-ELEC-002, mating-cycle test.
Owner / Status / Evidence / Notes | Electrical / Open / Judgment. Availability/service item; safety only via lost monitoring if undetected.
FA-HWFMEA-027 — Probe connector partial-seat / intermittent
Description | External-probe connector partially seats or intermittently disconnects.
Category / Subsystem-Component | Connectors · S4 · CMP-005 (DO-EE-013, DO-MECH).
Cause → Mechanism → Failure Mode | Under-engagement / vibration → intermittent probe contact.
Effect | Intermittent product-zone reading → false/missed excursion → compromised dose.
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — positive retention + partial-seat detection (CMP-005, HWR-028). Det — probe-presence + plausibility (DO-SAFE-005).
Recommended Action | Partial-seat detection must flag intermittent; poka-yoke full seat; fault-inject (PV-SW-008).
Hazard trace | HA-SW-008 ★, HA-USE-005 · Req/DO trace: HWR-028 / DO-EE-013, DO-SAFE-005 / RCR-020 · V&V: PV-SW-008.
Owner / Status / Evidence / Notes | Electrical + Sensing / Open / Judgment. Cross-ref FA-HWFMEA-017, FA-DFMEA-003/015.
FA-HWFMEA-028 — Internal board-to-board / flex connector fretting
Description | Internal interconnects loosen/fret under vibration and thermal cycling.
Category / Subsystem-Component | Connectors · CMP-034 (DO-PCB).
Cause → Mechanism → Failure Mode | Vibration + CTE cycling → fretting/high-resistance on internal connector.
Effect | Intermittent subsystem fault → silent gap or false reading → wrong disposition.
S / O / D / AP | 6 / 3 / 7 / Medium (S3)
Current Controls | Prev — minimise connectors / use board-to-board with retention, conformal coat. Det — BIST; hard to catch (D high).
Recommended Action | Prefer soldered interconnects for safety-critical nets; vibration-qualify connectors; BIST coverage.
Hazard trace | HA-ELEC-008, HA-MFG-008 · Req/DO trace: DO-PCB / RCR-016 · V&V: PV-MFG-008, vibe test.
Owner / Status / Evidence / Notes | PCB / Open / Judgment.
6.9 RF / antenna (S5 — CMP-024/028/029) — includes RF-in-container top item
FA-HWFMEA-029 ★ — Integral antenna cannot close in-container RF budget (RF-in-container top item)
Description | The integral antenna cannot achieve a usable LTE-M link inside a metallised/insulated shipper, so alerts/telemetry cannot egress.
Category / Subsystem-Component | RF/antenna · S5 · CMP-028 (DO-EE-018, DO-IF).
Cause → Mechanism → Failure Mode | Container RF attenuation + detuning by contents → insufficient link margin → no live connectivity.
Effect | Excursion committed but no live alert → shipment held/administered before a late alert → compromised dose.
S / O / D / AP | 7 / 5 / 6 / High (S4)
Current Controls | Prev — integral antenna tuned to close SYS-104 budget (CMP-028), store-and-forward buffering (DO-FW-011), external-antenna fallback branch (CMP-029), multi-bearer (LTE-M→Wi-Fi→BLE). Det — RSSI/connectivity diagnostics (CMP-045); per-trip delivery-success rate.
Recommended Action | Qualify RF-in-container (PRD-VER-005) across representative shippers/contents; adopt external antenna if integral cannot close budget; treat alert delivery as store-and-forward with escalation, not guaranteed-live.
Hazard trace | HA-COMM-003 ★, HA-COMM-006, HA-SW-006 · Req/DO trace: HWR-032, SYS-104/109 / DO-EE-018, DO-IF, DO-SAFE-009 / RCR-031/030 · V&V: PRD-VER-005 (RF-in-container), PV-COMM-003.
Owner / Status / Evidence / Notes | RF / Open — top program item / Evidence-Pending (antenna). O/D Judgment; residual carried Open until PRD-VER-005. Feeds FTA top event (b).
FA-HWFMEA-030 — Cold RF front-end margin loss at −30 °C
Description | LTE-M/Wi-Fi/BLE front-end loses link margin at −30 °C (component drift, PA efficiency).
Category / Subsystem-Component | RF · S5 · CMP-024 (DO-EE-017).
Cause → Mechanism → Failure Mode | Cold component drift / reduced PA output → degraded link margin in cold.
Effect | Compounds in-container attenuation → deferred/undelivered alert → compromised dose.
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — RF front-end holds spec to −30 °C (CMP-024 requirement), band plan, S&F buffering. Det — RSSI trend over temperature.
Recommended Action | Verify RF margin over temperature; combine with RF-in-container test; ensure S&F covers cold coverage gaps.
Hazard trace | HA-COMM-003, HA-COMM-001 · Req/DO trace: HWR-029/035, SYS-105 / DO-EE-017 / RCR-031 · V&V: PV-COMM (cold RF), PRD-VER-005.
Owner / Status / Evidence / Notes | RF / Open / Evidence-Pending (cold RF margin).
FA-HWFMEA-031 — Antenna detune from enclosure/contents/condensation
Description | Antenna match shifts due to nearby metal contents, ice/condensation, or enclosure tolerance.
Category / Subsystem-Component | RF · S5 · CMP-028 (DO-EE-018).
Cause → Mechanism → Failure Mode | Dielectric loading / detune → mismatch / efficiency loss.
Effect | Reduced link → deferred alert (same chain as FA-HWFMEA-029).
S / O / D / AP | 6 / 4 / 6 / Medium (S3)
Current Controls | Prev — keep-out zones (DO-MECH), robust match, external-antenna option. Det — RSSI/VSWR diagnostics.
Recommended Action | Design match for loaded/condensing conditions; keep-out enforcement; external antenna fallback.
Hazard trace | HA-COMM-003 · Req/DO trace: HWR-032, SYS-109 / DO-EE-018, DO-MECH-016 · V&V: PRD-VER-005, PV-COMM-003.
Owner / Status / Evidence / Notes | RF / Open / Evidence-Pending.
FA-HWFMEA-032 — SIM/eSIM fault or wrong roaming provisioning → no bearer
Description | SIM/eSIM fails, or roaming provisioning is wrong for the lane, so no wide-area bearer is available.
Category / Subsystem-Component | Connectivity · S5 · CMP-027 (DO-IF, DO-MFG).
Cause → Mechanism → Failure Mode | SIM defect / provisioning error / carrier-selection fault → no LTE-M attach across a border.
Effect | No live alert path over the lane → compromised dose (if excursion) / fleet-blind segment.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — roaming-capable provisioning, auto network selection (CMP-027), S&F buffering, Wi-Fi/ BLE offload. Det — attach/roaming diagnostics; coverage logs.
Recommended Action | Roaming + config verification (PV-COMM-006); provisioning QA at manufacturing; S&F + opportunistic Wi-Fi/BLE offload so a coverage gap defers rather than loses alerts.
Hazard trace | HA-COMM-006, HA-COMM-005 · Req/DO trace: HWR-034, SYS-115/116 / DO-IF, DO-MFG / RCR-031/056 · V&V: PV-COMM-006 (roaming + config).
Owner / Status / Evidence / Notes | RF + Mfg / Open / Evidence-Pending (carrier/SIM strategy).
6.10 Storage hardware (S7 — CMP-033)
FA-HWFMEA-033 — Flash wear-out / retention loss over reuse life
Description | NV log memory exceeds program/erase endurance or loses retention (esp. at temperature) over reuse.
Category / Subsystem-Component | Storage · S7 · CMP-033 (DO-FW-010).
Cause → Mechanism → Failure Mode | Endurance exhaustion / high-temp retention loss → bit errors / write failure → record loss/corruption.
Effect | Lost or corrupted evidentiary record → wrong disposition / indefensible record.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — wear-leveling + ≥50 % spare (HWR-020), part with endurance/retention exceeding reuse life, ECC. Det — ECC error counts; storage BIST; write-verify.
Recommended Action | Size endurance/retention with margin for reuse; wear-leveling; ECC + write-verify; storage BIST; endurance analysis in FAL-1900.
Hazard trace | HA-SW-013, HA-DATA-006, HA-CLOUD-004 · Req/DO trace: HWR-020, SYS-018/043 / DO-FW-010, DO-SAFE-008 / RCR-024 · V&V: PV-SW-013, endurance test.
Owner / Status / Evidence / Notes | Storage + FW / Open / Evidence-Pending (part/format/endurance).
FA-HWFMEA-034 — Storage capacity exhausted mid-trip → overwrite of unsent data
Description | Buffer fills before upload; poor prioritisation overwrites unsent/excursion/audit data.
Category / Subsystem-Component | Storage · S7 · CMP-033 (DO-FW-010).
Cause → Mechanism → Failure Mode | Extended offline + under-sized buffer → capacity exhaustion → data overwrite.
Effect | Loss of excursion/audit evidence → wrong disposition / incomplete record accepted.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — full-trip sizing with ≥50 % spare (HWR-020), no-overwrite-of-unsent policy, excursion/audit-first prioritisation (CMP-041). Det — buffer-backlog telemetry; capacity BIST.
Recommended Action | Size for worst-case offline trip; never overwrite unsent/excursion/audit; annunciate buffer-full; verify prioritisation.
Hazard trace | HA-SW-002, HA-DATA-006 · Req/DO trace: HWR-020, SYS-018/043 / DO-FW-010 / RCR-021/024 · V&V: PV-SW-002 (gap-handling), storage-full test.
Owner / Status / Evidence / Notes | Storage + FW / Open / Evidence-Pending (sizing). Overlaps FAL-1200 storage-manager logic.
7. Traceability Notes
• Every safety-significant HW mode reconciles to a HAL hazard: cell/charge → HA-ENRG-001/002/005/006, HA-TFC-002/003, HA-MFG-002/003; sensing → HA-SW-001/004/008, HA-DATA-001; PCBA → HA-ELEC-004/008/010, HA-MFG-008; RF → HA-COMM-001/003/006; storage → HA-SW-013, HA-DATA-006, HA-CLOUD-004.
Gap note | FA-HWFMEA-021 (accelerometer) and FA-HWFMEA-022 (GNSS) have no direct safety HAL link because shock/location are advisory for the vaccine verdict; GNSS ties indirectly to HA-ENRG-006 via energy. These are recorded as acceptable non-safety items, not gaps requiring a new hazard.
FTA basic events (FAL-1800) | FA-HWFMEA-001/012 (top event a — dead-in-cold), 002/003/006/007/008 (top event c — thermal runaway), 029/030/031/032 (top event b — no/late alert), 011/023/033/034 (top events a/d — silent gap / corrupt record).
Single-fault-tolerance items | charge-inhibit (006), BMS (007), secondary fuse (008), input protection (009) — governed by DO-SAFE-001/002/003; CCF review in FAL-2700.
8. Closure / Knowledge Notes
Top HW items | D3 cold endurance (FA-HWFMEA-001), cold-charge interlock + BMS + secondary fuse two-tier protection (002/006/007/008 — S5), RF-in-container (029), probe integrity (017/027), fuel-gauge honesty (012).
• All O/D are pre-validation (Evidence-Pending); PV-ENRG-001/002, PRD-VER-004/005, PV-SW-008/013, PV-ELEC-002/004, PV-COMM-003/006 and UN 38.3 close them.
• Recurring IoT/battery/cold-chain pattern for FAL-2200/2300: cold simultaneously derates capacity, raises impedance (gauge error), reduces RF margin, and creates the cold-charge plating hazard — the single most load-bearing environmental initiator; and latent PCBA/connector intermittents reading as clean passes — the classic silent-escape pattern for monitoring devices.
FAL-1400 — System FMEA (SYSFMEA) — Smart Cold Chain Monitor (SCM)
Module FAL-1400 (FA-SYSFMEA) · ERFKL Stage 11 (ORCH-016) · Owner: Chief Systems Engineer (EM-ENG-001) with the Principal Reliability/Safety Engineer. System FMEA for the whole SCM product — emergent failures at the boundaries between subsystems (device ↔ cloud ↔ mobile ↔ operator ↔ depot/service) that no single component owns. Scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3). Every safety-significant mode reconciles to a HAL hazard (HA-…). Feeds Stage 12 RCR, Stage 13 Verification, FTA (FAL-1800), CCF (FAL-2700).
---
1. Purpose
Analyse failures that emerge from subsystem interaction and end-to-end system behaviour: device→cloud delivery, cloud→responder alerting, device↔cloud consistency, operator decision loop, depot/service loop, fleet-scale and mission effects. These are the failures that fall between the discipline FMEAs (FAL-1000/1200/1300/1500).
2. Scope
Whole-product interaction chains: sensing→detection→verdict→display; detection→alert→cloud→responder; local logging→store-and-forward→ingest→record→disposition; arming/pre-arm→trip→disposition; OTA campaign; depot service (calibration, firmware, provisioning, disposition of "dead" units); fleet/mission effects (mass administration, false-reject stock-out, data-integrity-at-scale). Excludes intra-component failures (owned by the discipline FMEAs) except where the interaction is the failure; interfaces as protocols/energy paths are in FAL-1500.
3. Applicable Standards & References
ISO 14971; AIAG-VDA/IEC 60812; IEC 62366-1 (use/system-of-use); ALCOA+ & EU GDP (data integrity, disposition); WHO PQS (cold-chain monitor performance); ISO 13485-style design-control interaction. Cited by designation.
4. Governance & Method
Shared ERFKL scoring key (FAL-1000 §4). System severity is fixed by the system/mission outcome (compromised dose administered, fleet-blinding, mass administration = S4/S5). Occurrence/detection reflect end-to-end behaviour after subsystem controls; basis in Notes. Action Priority severity-dominant.
5. Categories
Measurement→verdict chain · Detection→alert→responder chain · Delivery/store-and-forward · Device↔cloud consistency · Operator decision loop · Depot/service loop · Fleet/mission · Time/identity system-wide.
6. Risk & Failure Register (FA-SYSFMEA)
FA-SYSFMEA-001 ★ — Undetected temperature excursion → spoiled dose administered (end-to-end)
Description | The whole measure→detect→verdict→disposition chain fails to flag a real excursion, so a compromised dose is released and administered.
Category / Subsystem-Component | Measurement→verdict chain · S1–S8 (system).
Cause → Mechanism → Failure Mode | Any of {unrepresentative placement, sensor fault, biased calibration, one-sided/mis-profiled detection, false-accept verdict} un-caught by cross-checks → excursion not detected at the system level.
Effect (local→system→user/safety) | Verdict "accept" on a spoiled lot → dose administered → loss of potency / patient harm; at scale, population outbreak (HA-MPE-002, S5).
S / O / D / AP | 8 / 3 / 6 / High (S4; S5 at scale)
Current Controls | Prev — external product-zone probe, two-sided detection, plausibility, false-accept suppression, device↔cloud recompute (DO-SAFE-004/005/006). Det — device↔cloud consistency; QA sampling; field potency signals.
Recommended Action | Defence-in-depth across the chain (placement guidance + probe + detection + verdict + cloud recompute); system-level excursion validation (VAL-SW-001); this is FTA top event (a).
Hazard trace | HA-SW-001 ★, HA-DATA-001, HA-MPE-002 (S5) · Req/DO trace: SYS-149/150 / DO-SAFE-004/005/006 / RCR-019/020/026 · V&V: PV-SW-001/002/008, VAL-SW-001.
Owner / Status / Evidence / Notes | EM-ENG-001 + Safety / Open / Judgment. FTA top event (a). Aggregated from the discipline FMEAs; the system control is no single check is the sole barrier.
FA-SYSFMEA-002 ★ — No/late excursion alert reaches a responder (end-to-end)
Description | An excursion is detected but no timely notification reaches a human/organisation who can act.
Category / Subsystem-Component | Detection→alert→responder chain · S3/S5/S8 (system).
Cause → Mechanism → Failure Mode | Any of {no in-container link (RF), no bearer/roaming, lost queue, cloud notify fail, unacked-not-escalated} → alert not delivered/acted-on in time.
Effect | Shipment held/administered before the alert → compromised dose; delayed intervention.
S / O / D / AP | 8 / 4 / 6 / High (S4)
Current Controls | Prev — store-and-forward, multi-bearer, alert-before-bulk, cloud escalation of unacked, external-antenna fallback (DO-SAFE-009, DO-CLD-007/008). Det — alert-latency distribution per lane; unacked backlog fleet-wide.
Recommended Action | Treat alerting as store-and-forward-with-escalation, not guaranteed-live; RF-in-container (PRD-VER-005); latency + escalation validation (PV-COMM-001, VAL-SW-006); FTA top event (b).
Hazard trace | HA-COMM-001 ★, HA-COMM-003 ★, HA-SW-006, HA-CLOUD-001 · Req/DO trace: SYS-011/013 / DO-SAFE-009 / RCR-028/029/031 · V&V: PV-COMM-001/002/003, PRD-VER-005, VAL-SW-006.
Owner / Status / Evidence / Notes | EM-ENG-001 + RF/Cloud / Open / Evidence-Pending (RF-in-container). FTA top event (b). Carries the RF-in-container top item.
FA-SYSFMEA-003 — Device↔cloud verdict/record inconsistency drives wrong disposition
Description | Device and cloud hold different views of the same trip (verdict, gaps, timing) with no resolution.
Category / Subsystem-Component | Device↔cloud consistency · S3/S8.
Cause → Mechanism → Failure Mode | Version skew / reorder / dup / partial upload → inconsistent authoritative record.
Effect | Operator/QA act on the wrong view → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — deterministic device verdict = cloud recompute, idempotent ordered upload, reconciliation endpoint (CMP-040/046/049). Det — reconciliation flags divergence.
Recommended Action | Single authoritative record with explicit reconciliation; surface divergence; version pinning; test reorder/dup/corrupt (PV-COMM-004).
Hazard trace | HA-COMM-004, HA-SW-007, HA-DATA-002 · Req/DO trace: SYS-016/026/176 / DO-CLD-006/007 / RCR-030/032 · V&V: PV-COMM-004, PV-CLOUD.
Owner / Status / Evidence / Notes | EM-ENG-001 + Cloud / Open / Judgment.
FA-SYSFMEA-004 ★ — Record mis-attributed to the wrong shipment/device
Description | A trip record is bound to the wrong device/shipment, so a decision is made on another shipment's data.
Category / Subsystem-Component | Device↔cloud identity/registry · S3/S8.
Cause → Mechanism → Failure Mode | Registry binding error / duplicate identity / routing fault → mis-attributed record.
Effect | Real excursion invisible on its own record / decision on wrong data → compromised dose.
S / O / D / AP | 8 / 2 / 5 / High (S4)
Current Controls | Prev — authoritative registry with trip/owner binding, unique crypto identity, secure routing (CMP-047, DO-SEC). Det — binding/routing anomaly detection.
Recommended Action | Unique provisioned identity (PV-MFG-005); verified trip↔device binding; routing integrity test (PV-COMM-008); FTA contributor to top event (d).
Hazard trace | HA-DATA-002, HA-COMM-008, HA-MFG-005 · Req/DO trace: SYS-030/012/134 / DO-CLD (registry), DO-SEC / RCR-033/055 · V&V: PV-COMM-008, PV-MFG-005.
Owner / Status / Evidence / Notes | EM-ENG-001 + Cloud/Security / Open / Judgment.
FA-SYSFMEA-005 ★ — Corrupted/false cold-chain record accepted as authoritative (end-to-end)
Description | The evidentiary record is corrupted, incomplete, or falsified but is accepted as a valid basis for disposition/audit.
Category / Subsystem-Component | Record integrity system-wide · S3/S7/S8.
Cause → Mechanism → Failure Mode | Any of {torn write, storage loss, tamper, cloud DB corruption, mis-timing} un-caught → false/incomplete record presented as clean.
Effect | Wrong disposition / indefensible record / data-integrity violation at scale (compromised dose; certification hold).
S / O / D / AP | 8 / 3 / 6 / High (S4)
Current Controls | Prev — power-fail-safe writes, tamper-evidence, cloud durability/DR, ALCOA+ record store (DO-SAFE-008, DO-CLD, DO-SEC). Det — integrity verification at ingest; DR/backup-restore tests; audit.
Recommended Action | End-to-end record integrity (device signing → ingest verify → durable store → export provenance); DR testing; FTA top event (d).
Hazard trace | HA-CYBER-001 ★, HA-SW-013, HA-CLOUD-004, HA-DATA-006, HA-MPE-004 · Req/DO trace: SYS-026/028/168 / DO-SAFE-008, DO-CLD, DO-SEC / RCR-024/032/034/042 · V&V: PV-SW-013, PV-CYBER-001 (IV), PV-CLOUD (DR), PV-DATA-006.
Owner / Status / Evidence / Notes | EM-ENG-001 + Security/Cloud / Open / Judgment. FTA top event (d).
FA-SYSFMEA-006 — Cloud outage/ransomware blinds the fleet and withholds records
Description | A cloud outage or ransomware event makes the fleet blind and alters/withholds records.
Category / Subsystem-Component | Cloud/system · S8.
Cause → Mechanism → Failure Mode | Outage / attack → fleet-wide loss of ingest/notify + record availability.
Effect | Undelivered alerts + inaccessible/altered records → compromised doses + fleet-blinding (S4).
S / O / D / AP | 8 / 2 / 4 / High (S4)
Current Controls | Prev — cloud HA/DR, backups, security hardening, device S&F buffers during outage (DO-CLD, DO-SEC). Det — availability/backlog monitoring; DR tests; security monitoring.
Recommended Action | HA/DR with tested restore; device buffers survive extended outage; ransomware-resilient backups; independent security review (PV-CYBER-009, IV).
Hazard trace | HA-CLOUD-008, HA-CYBER-009 ★ · Req/DO trace: SYS / DO-CLD, DO-SEC / RCR-039/032 · V&V: PV-CYBER-009 (cloud review + DR, IV).
Owner / Status / Evidence / Notes | EM-ENG-001 + Cloud/Security / Open / Judgment. Device buffering limits data loss; alert timeliness still degraded.
FA-SYSFMEA-007 — Cloud rules engine mis-flags / applies wrong disposition policy
Description | The cloud rules/alert engine fails to flag an excursion or applies a wrong policy.
Category / Subsystem-Component | Cloud logic · S8.
Cause → Mechanism → Failure Mode | Rules mis-config / policy error / tamper → wrong flag/policy.
Effect | Excursion not flagged / good lot rejected → wrong disposition.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — device verdict is authoritative and cloud recomputes independently; rules change audit (DO-CLD, DO-SEC). Det — config-change audit; device↔cloud consistency.
Recommended Action | Device verdict as safety anchor; audited rules config; config-tamper test (PV-CLOUD-002); reconcile with device.
Hazard trace | HA-CLOUD-002, HA-CYBER-003 · Req/DO trace: SYS-012/014 / DO-CLD (rules/audit) / RCR-036 · V&V: PV-CLOUD-002 (config-tamper), PV-CYBER-003 (IV).
Owner / Status / Evidence / Notes | EM-ENG-001 + Cloud / Open / Judgment. Device-side anchor limits reliance on cloud logic.
FA-SYSFMEA-008 ★ — Operator overrides a reject (automation bias / stock-out pressure)
Description | The operator overrides a correct reject/quarantine to avoid a stock-out, releasing a compromised lot.
Category / Subsystem-Component | Operator decision loop · system/HF.
Cause → Mechanism → Failure Mode | Automation bias / operational pressure + weak override governance → reject overridden.
Effect | Compromised dose released → patient harm.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — override requires reason + role + audit; clear recommendation; alarm-management to reduce fatigue (DO-HF, DO-CLD audit). Det — override rate + reasons; audit review.
Recommended Action | Governed override (role, reason, audit); track override rate as a first-class metric; design decision aids to reduce automation bias (VAL-USE-007).
Hazard trace | HA-USE-007 ★, HA-MPE-001 · Req/DO trace: SYS / DO-HF, DO-CLD (audit) / RCR-044/040 · V&V: PV-DATA-008, VAL-USE-007.
Owner / Status / Evidence / Notes | EM-ENG-001 + HF / Open / Judgment. Cross-ref FAL-1600 use-error.
FA-SYSFMEA-009 — Depot recalibration escape → systematic fleet bias
Description | A depot service loop lets recalibration lapse, so units drift and the fleet develops systematic bias.
Category / Subsystem-Component | Depot/service loop · system/service.
Cause → Mechanism → Failure Mode | Recal interval not enforced / escape → out-of-cal fleet.
Effect | Systematic measurement bias → wrong disposition across many trips (fleet CCF).
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — cal expiry blocks arming, recal interval enforcement, depot process control (DO-FW-009, DO-SVC). Det — cal-escape audit; drift trending.
Recommended Action | Enforce recal interval (PV-SVC-001); cal expiry gate; fleet drift surveillance; CCF review (FAL-2700).
Hazard trace | HA-SVC-001 ★, HA-SW-004, HA-MFG-001 · Req/DO trace: SYS-024 / DO-FW-009, DO-SVC / RCR-049/026 · V&V: PV-SVC-001, PV-SW-004.
Owner / Status / Evidence / Notes | EM-ENG-001 + Service/Metrology / Open / Judgment. Fleet-scale CCF.
FA-SYSFMEA-010 — Depot loads wrong firmware/region config for the lane
Description | Service loads firmware or a regional config that doesn't match the deployment lane.
Category / Subsystem-Component | Depot/service loop · system/service.
Cause → Mechanism → Failure Mode | Wrong firmware/region/band config at depot → wrong logic/coverage.
Effect | Missed excursion or no coverage → compromised dose / fleet-blind segment.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — compatibility matrix (DO-CFG-005), signed firmware, post-service BIST, region config control (DO-MFG/DO-FW). Det — firmware/config version audit; post-service BIST.
Recommended Action | Enforce compatibility matrix; post-service BIST verifies firmware/region; version audit.
Hazard trace | HA-SVC-002, HA-COMM-006 · Req/DO trace: SYS-032/139 / DO-CFG-005, DO-SAFE-014, DO-MFG / RCR-037/056/050 · V&V: PV-SVC (post-service BIST), PV-COMM-006.
Owner / Status / Evidence / Notes | EM-ENG-001 + Service / Open / Judgment.
FA-SYSFMEA-011 — Charge-inhibited cold device discarded as "dead" → lost monitoring/stock-out
Description | A depot operator mistakes a cold, charge-inhibited device for a dead one and discards it.
Category / Subsystem-Component | Depot/service loop · system/HF/HMI.
Cause → Mechanism → Failure Mode | Charge-inhibit state not legible → serviceable device discarded.
Effect | Lost monitoring capacity / stock-out of monitors; scrap cost.
S / O / D / AP | 5 / 3 / 4 / Medium (S3)
Current Controls | Prev — charge/fault-state legibility (LED/OLED, DO-SAFE-013), operator guidance (DO-HF-005). Det — discard/return-as-dead rate.
Recommended Action | Unambiguous "cold — will charge when warm" indication; depot training; state-legibility validation.
Hazard trace | HA-USE-009 · Req/DO trace: SYS-094/100 / DO-SAFE-013, DO-HF-005, DO-EE-008 / RCR-046 · V&V: PV-USE (state-legibility), VAL-USE.
Owner / Status / Evidence / Notes | EM-ENG-001 + HF / Open / Judgment. Availability/business, not safety.
FA-SYSFMEA-012 ★ — Mass administration of compromised doses (mission-scale)
Description | A systematic failure (shared calibration/firmware/record defect) releases compromised doses at scale.
Category / Subsystem-Component | Fleet/mission · system.
Cause → Mechanism → Failure Mode | Common-cause across the fleet (calibration, firmware, record integrity) → fleet-wide false accepts.
Effect | Population-scale administration of ineffective/compromised doses → outbreak (S5).
S / O / D / AP | 9 / 2 / 6 / High (S5)
Current Controls | Prev — aggregate of all record-integrity + detection controls (DO-SAFE spine), diversity where feasible, staged OTA. Det — population-scale outbreak surveillance (mandatory); fleet drift/consistency.
Recommended Action | CCF defence-in-depth (FAL-2700); staged OTA + canary; independent cloud recompute; fleet drift surveillance; this is HA-MPE-002 (S5) held region-A-High with mandatory monitoring.
Hazard trace | HA-MPE-002 ★★ (S5), HA-SW-004/014, HA-CYBER-005 · Req/DO trace: all DO-SAFE / RCR-057 (aggregate) · V&V: aggregate of the spine PVs.
Owner / Status / Evidence / Notes | EM-ENG-001 + Safety / Open / Judgment. S5 fixes AP High irrespective of low occurrence; the dominant CCF concern (FAL-2700).
FA-SYSFMEA-013 — False-reject stock-out → unvaccinated population
Description | Excess false rejects (over-conservative detection / mishandled gaps) waste good doses and cause stock-out.
Category / Subsystem-Component | Fleet/mission · system.
Cause → Mechanism → Failure Mode | Over-aggressive detection / gap-mishandling → excess false rejects → wastage/stock-out.
Effect | Preventable disease from unvaccinated population (S4); vaccine-lot loss (business).
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — detection tuned to the stability profile (not overly conservative), gap-handling that quarantines (not auto-rejects), false-reject targets (DO-SAFE-004, DO-CLD-003). Det — wastage/false-reject rate.
Recommended Action | Balance miss vs waste against product stability; validate false-reject targets (VAL-USE-002); track stock-out/wastage as a first-class metric.
Hazard trace | HA-MPE-003 ★, HA-MPE-001 · Req/DO trace: SYS / DO-SAFE-004, DO-CLD-003 / RCR-047/044/046 · V&V: VAL-USE-002.
Owner / Status / Evidence / Notes | EM-ENG-001 + Safety / Open / Judgment. The accepted-trade counter-risk to false-accept; both directions harm.
FA-SYSFMEA-014 — System-wide loss of trusted time desynchronises records/alerts
Description | A shared time fault (RTC + no re-discipline) makes records/alerts across many devices untrustworthy.
Category / Subsystem-Component | Time system-wide · system.
Cause → Mechanism → Failure Mode | RTC drift + failed time re-discipline (no link) → fleet time desync.
Effect | Mis-timed records/alerts → wrong disposition; indefensible audit at scale.
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — monotonic UTC + disclosed drift + re-discipline on link, backup source (SYS-118). Det — time-health flag; untrusted-time incidence.
Recommended Action | Disclose untrusted time in records; re-discipline opportunistically; CCF review (shared time) — FAL-2700.
Hazard trace | HA-SW-005, HA-COMM-005 · Req/DO trace: SYS-006/118 / DO-EE-011, DO-FW / RCR-025 · V&V: PV-SW-005.
Owner / Status / Evidence / Notes | EM-ENG-001 / Open / Judgment. Shared-time CCF.
FA-SYSFMEA-015 — EMC upset at system level produces a false/missed excursion
Description | System-level EMC susceptibility (device in a noisy transport/RF environment) flips a verdict.
Category / Subsystem-Component | System EMC · system.
Cause → Mechanism → Failure Mode | Environmental EMI beyond immunity + weak plausibility → system-level false/missed excursion.
Effect | Wrong disposition (S3; S5 if it defeats detection at scale).
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — EMC/ESD design + plausibility (DO-EE-019). Det — EMC immunity test; plausibility; field EMC anomaly monitoring.
Recommended Action | EMC immunity to the transport environment (PV-ELEC-004); plausibility so a single upset cannot flip a verdict.
Hazard trace | HA-ELEC-004 · Req/DO trace: SYS / DO-EE-019 / RCR-014 · V&V: PV-ELEC-004.
Owner / Status / Evidence / Notes | EM-ENG-001 + Electrical / Open / Judgment.
FA-SYSFMEA-016 — Thermal event at system level (device fire in shipper/aircraft)
Description | A cell thermal event propagates to a system-level fire in an occupied aircraft cargo hold.
Category / Subsystem-Component | System energy/thermal · system.
Cause → Mechanism → Failure Mode | Cell event (cold-charge / defect / crush) + inadequate containment → uncontained fire in transport.
Effect | Death / loss of aircraft (S5); toxic offgas in enclosed space.
S / O / D / AP | 9 / 2 / 6 / High (S5)
Current Controls | Prev — two-tier protection + charge-inhibit + retention + containment (DO-SAFE-001/002/010), UN 38.3 / IATA. Det — mandatory thermal-event field monitoring + transport disclosure.
Recommended Action | Defence-in-depth (FTA top event c); containment + offgas characterisation; UN 38.3; transport disclosure; mandatory monitoring.
Hazard trace | HA-TFC-002 ★★ (S5), HA-ENRG-001/002/003, HA-TFC-003 · Req/DO trace: DO-SAFE-001/002/010 / RCR-002/004/007/008 · V&V: PV-TFC-002/003, PV-ENRG-001/002, UN 38.3.
Owner / Status / Evidence / Notes | EM-ENG-001 + Safety/Power / Open / Judgment. FTA top event (c) at the system level; aggregates the HW battery basic events.
FA-SYSFMEA-017 — Stray/overdue device not flagged → shipment monitored by a non-ready unit
Description | The fleet dashboard fails to flag an overdue/stray/non-ready device, so a shipment ships with a bad monitor.
Category / Subsystem-Component | Fleet management · S8.
Cause → Mechanism → Failure Mode | Readiness/last-seen not surfaced → non-ready device deployed.
Effect | Unmonitored/under-monitored trip → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — fleet dashboard readiness/last-seen + deploy-block (CMP-050/047, DO-CLD). Det — overdue/stray flags; readiness gating.
Recommended Action | Deploy-block on non-ready/stale; surface overdue/stray; readiness required before assign.
Hazard trace | HA-SW-009, HA-USE-004 · Req/DO trace: SYS-030/031 / DO-CLD (registry/dashboard) / RCR-023 · V&V: PV-CLOUD, dashboard review.
Owner / Status / Evidence / Notes | EM-ENG-001 + Cloud / Open / Judgment.
FA-SYSFMEA-018 — Data-integrity violation at scale → certification hold
Description | Systematic ALCOA+/GDP violations (provenance, gaps, time, export) across the fleet trigger a certification/regulatory hold.
Category / Subsystem-Component | Fleet/mission compliance · system.
Cause → Mechanism → Failure Mode | Systematic record-integrity/provenance defects → data-integrity violation at scale.
Effect | Certification loss/hold (S4 business/compliance); loss of trust in the cold-chain record.
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — ALCOA+ record store, audit trail, export with provenance/gaps/tamper-evidence (DO-CLD, DO-SAFE-008). Det — data-integrity deviation monitoring; traceability test.
Recommended Action | ALCOA+ by design; provenance on export; traceability test (PV-DATA-006); audit readiness.
Hazard trace | HA-MPE-004, HA-DATA-006 · Req/DO trace: SYS-027/028 / DO-CLD (API/audit), DO-SAFE-008 / RCR-042/025/033 · V&V: PV-DATA-006, traceability test.
Owner / Status / Evidence / Notes | EM-ENG-001 + Quality/Regulatory / Open / Judgment.
FA-SYSFMEA-019 — Estimate presented as a measurement (provenance loss)
Description | An interpolated/estimated value is presented to a decision-maker as if it were a real measurement.
Category / Subsystem-Component | Data provenance · system/HF.
Cause → Mechanism → Failure Mode | Gap-fill / interpolation not labelled → estimate believed measured.
Effect | Wrong/over-confident disposition (S3).
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — never interpolate silently; label estimate vs measured; gaps preserved (DO-CLD/DO-HF). Det — provenance review; comprehension testing.
Recommended Action | Explicit provenance labelling; no silent gap-fill; validate comprehension (VAL-USE-005).
Hazard trace | HA-DATA-007 · Req/DO trace: SYS / DO-CLD, DO-HF / RCR-041/042 · V&V: PV-DATA-007, VAL-USE-005.
Owner / Status / Evidence / Notes | EM-ENG-001 + HF / Open / Judgment.
FA-SYSFMEA-020 — Cross-border coverage gap leaves a lane silently unmonitored-live
Description | A roaming/coverage gap across a border leaves a segment with no live alerting, un-flagged to users.
Category / Subsystem-Component | Connectivity/system · S5/S8.
Cause → Mechanism → Failure Mode | Roaming/coverage gap not surfaced → silent live-monitoring gap on a lane.
Effect | Excursion on that segment reaches no responder in time → compromised dose.
S / O / D / AP | 7 / 4 / 5 / High (S4)
Current Controls | Prev — S&F buffering (data preserved), roaming provisioning, coverage logging (CMP-042/027, DO-IF). Det — coverage/roaming gap logs; per-lane delivery-success.
Recommended Action | Surface coverage gaps per lane; S&F preserves data; escalate on reconnection; roaming verification (PV-COMM-006).
Hazard trace | HA-COMM-006, HA-COMM-003 · Req/DO trace: SYS-116 / DO-IF, DO-FW-011 / RCR-031/056 · V&V: PV-COMM-006, VAL-COMM-006.
Owner / Status / Evidence / Notes | EM-ENG-001 + RF / Open / Judgment. Data is not lost (S&F) but live timeliness is.
FA-SYSFMEA-021 — Reverse-logistics: uncleaned/unre-provisioned reuse carries stale config
Description | A returned reusable device is redeployed without cleaning/re-provisioning, carrying stale config/identity/trip binding.
Category / Subsystem-Component | Reuse loop (D1) · system/service.
Cause → Mechanism → Failure Mode | Reverse-logistics gap → stale config/binding on reuse.
Effect | Wrong-profile monitoring / mis-attribution / hygiene issue → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — depot re-provisioning + cleaning, pre-arm profile validation, registry re-binding (DO-SVC, DO-CFG-008). Det — pre-arm checks; binding verification.
Recommended Action | Mandatory clean + re-provision on intake; pre-arm invalidates stale config; verify binding.
Hazard trace | HA-SW-003, HA-DATA-002, HA-SVC-007 · Req/DO trace: SYS-024 / DO-SVC, DO-CFG-008 / RCR-048/055 · V&V: PV-SW-003, PV-SVC.
Owner / Status / Evidence / Notes | EM-ENG-001 + Service / Open / Judgment. Emergent from D1 reuse decision.
FA-SYSFMEA-022 — Manufacturing escape (miscalibrated/mis-protected unit) enters the fleet
Description | An EOL escape (miscalibrated, wrong protection part, duplicate identity) enters service.
Category / Subsystem-Component | Manufacturing→fleet · system/mfg.
Cause → Mechanism → Failure Mode | EOL test coverage gap → defective unit escapes to fleet.
Effect | Systematic bias / defeated protection / mis-route → wrong disposition or thermal event (S4).
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — EOL cal coverage, protection AOI/poka-yoke, identity-uniqueness check (DO-TST, DO-MFG). Det — EOL cal-escape trends; AOI defect trends; duplicate-identity detection.
Recommended Action | EOL cal + protection + identity coverage (PV-MFG-001/003/005); escape-rate trending; cross-ref FAL-1100 PFMEA.
Hazard trace | HA-MFG-001/003/005 · Req/DO trace: DO-TST, DO-MFG / RCR-053/054/055 · V&V: PV-MFG-001/003/005.
Owner / Status / Evidence / Notes | EM-ENG-001 + Quality / Open / Judgment. Bridges to Process FMEA (Stage 15).
7. Traceability Notes
• The four ★ system failures map directly to the FTA top events (FAL-1800): FA-SYSFMEA-001 → (a) undetected excursion; FA-SYSFMEA-002 → (b) no/late alert; FA-SYSFMEA-016 → (c) thermal runaway/fire; FA-SYSFMEA-005 → (d) corrupted/false record. FA-SYSFMEA-012 (mass administration, S5) is the mission-scale roll-up (HA-MPE-002).
• Every safety-significant system mode reconciles to a HAL hazard (HA-SW-, HA-COMM-, HA-CLOUD-, HA-CYBER-, HA-DATA-, HA-USE-, HA-SVC-, HA-MFG-, HA-MPE-, HA-TFC-, HA-ENRG-, HA-ELEC-). No orphan safety system modes.
CCF drivers (FAL-2700) | shared calibration/firmware (FA-SYSFMEA-009/012), shared time (014), shared cloud (006/007), shared power path (via FA-SWFMEA-018).
8. Closure / Knowledge Notes
Top system items | the four spine chains (undetected excursion, no/late alert, thermal event, corrupted record) plus the two mission roll-ups (mass administration S5; false-reject stock-out) and the depot/reuse loop (recal escape, wrong firmware, discarded-as-dead, stale reuse config).
• The system-level control philosophy is defence-in-depth + no-single-barrier + no-false-reassurance: no single subsystem check is the sole barrier to a safety end-effect, and every gap is made honest.
• Recurring cold-chain/IoT system pattern for FAL-2200/2300: the most dangerous failure is a silent one that reads as a clean pass — dead-in-cold, hung-but-armed, incomplete-but-accepted, mis-attributed, or late-alert — and fleet-scale common-cause (shared calibration/firmware/record path) turns a unit defect into a population-scale event.
FAL-1500 — Interface FMEA (IFMEA) — Smart Cold Chain Monitor (SCM)
Module FAL-1500 (FA-IFMEA) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Chief Systems Engineer (EM-ENG-001, interface owner) and the RF/Cloud/Security engineers. Interface FMEA for the SCM's boundaries: physical/electrical connectors (USB-C, external probe, internal), wireless links (LTE-M, Wi-Fi, BLE), the device↔cloud API, BLE provisioning, and human/mobile interfaces treated as interfaces. Scored with Action Priority (severity-dominant), NOT raw RPN (CLAUDE.md rule 3). Every safety-significant mode reconciles to a HAL hazard (HA-…). Feeds Stage 12 RCR, Stage 13 Verification, FTA (FAL-1800), CCF (FAL-2700).
---
1. Purpose
Analyse failures that occur at the interface — the protocol, contract, energy, or signal path between two elements — where mismatched assumptions, timing, formats, security, or physical contact cause failure. Evidence-backed to interfaces (INT-… / ICD register DO-SYS-004), requirements (IFR-/SYS-/SRS-), design outputs (DO-IF/DO-EE/DO-CLD), hazards (HA-…), and verifications (PV-).
2. Scope
Electrical/mechanical connectors (USB-C sink, external probe, internal board-to-board); wireless bearers (LTE-M primary, Wi-Fi offload, BLE provisioning/read); device↔cloud API (ingest, alert, OTA, registry, export); BLE provisioning/pairing; mobile app ↔ device; supplier/carrier interfaces (SIM/eSIM, carrier). Excludes component-internal failures (FAL-1300), firmware logic (FAL-1200), enclosure/seal mechanics (FAL-1000) except at the interface. Overlaps with FAL-1300 RF and FAL-1400 device↔cloud are cross-referenced (interface contract view here vs component/system view there).
3. Applicable Standards & References
ISO 14971; AIAG-VDA/IEC 60812; USB-C / USB-PD; 3GPP LTE-M (Cat-M1) PSM/eDRX; IEEE 802.11; Bluetooth Core/BLE; TLS / mutual-auth; IEC 62443 / OWASP (interface security); IEC 62366-1 (human interface). Cited by designation; revision to confirm.
4. Governance & Method
Shared ERFKL scoring key (FAL-1000 §4). Interface severity is fixed by the downstream effect of the broken contract (compromised dose / lost record / fleet-blinding = S4). O/D basis in Notes. Action Priority severity-dominant.
5. Categories
Connector/electrical · USB-C/charge interface · Wireless link (LTE-M/Wi-Fi/BLE) · Device↔cloud API · BLE provisioning/pairing · Mobile↔device · Carrier/SIM · Human/HMI interface.
6. Risk & Failure Register (FA-IFMEA)
6.1 Connectors / electrical (INT — DO-SYS-004)
FA-IFMEA-001 — External probe interface: intermittent contact → invalid product-zone signal
Description | The probe connector/lead interface intermittently opens under flex/vibration/cold.
Category / Subsystem-Component | Probe interface · S4 · CMP-005/020 (DO-EE-013, INT probe).
Cause → Mechanism → Failure Mode | Partial seat / fret / cold-stiffened lead → intermittent probe contact.
Effect | Invalid/intermittent product-zone reading presented as valid → false accept → compromised dose.
S / O / D / AP | 7 / 3 / 6 / High (S4)
Current Controls | Prev — positive retention + partial-seat/continuity detect (HWR-028), plausibility (DO-SAFE-005). Det — probe-presence + int-vs-ext divergence; probe-fault flag.
Recommended Action | Continuity/plausibility must catch intermittent; poka-yoke full seat; strain-relief; fault-inject (PV-SW-008).
Hazard trace | HA-SW-008 ★, HA-USE-005 · Req/DO trace: HWR-028, SYS-003 / DO-EE-013, DO-SAFE-005 / RCR-020/045 · V&V: PV-SW-008.
Owner / Status / Evidence / Notes | Systems + Sensing / Open / Judgment. Interface-contract view of FA-HWFMEA-017/027 and FA-DFMEA-003.
FA-IFMEA-002 — USB-C charge/data interface: contract mismatch / out-of-spec source
Description | The USB-C interface negotiates or receives an out-of-spec voltage/current, or a non-compliant source.
Category / Subsystem-Component | USB-C interface · S2 · CMP-004/014 (DO-EE-009, INT USB-C).
Cause → Mechanism → Failure Mode | PD negotiation fault / hostile charger → OV/OC/reverse at the sink.
Effect | Overcharge/thermal (S4) or device damage/unavailability.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — compliant USB-C sink + input protection (ideal-diode/OVP/TVS), BMS backstop (DO-EE-009/DO-SAFE-003). Det — input fault flag; protection trip.
Recommended Action | Verify PD negotiation + reverse/OVP/short tolerance (PV-ELEC-002); safe default on negotiation failure.
Hazard trace | HA-ENRG-005 · Req/DO trace: HWR-009/045, SYS-102 / DO-EE-009, DO-SAFE-003 / RCR-003 · V&V: PV-ELEC-002.
Owner / Status / Evidence / Notes | Systems + Power / Open / Judgment.
FA-IFMEA-003 — USB-C interface as an unlocked debug/data attack surface
Description | The USB-C data path exposes a debug/service interface that can falsify records or suppress alarms.
Category / Subsystem-Component | USB-C data interface · S3 · CMP-004 (DO-EE-012, DO-SEC).
Cause → Mechanism → Failure Mode | Debug/data path not locked in production → unauthorised access via USB-C.
Effect | Falsified record / suppressed alarm / cloned device → compromised dose; repudiation.
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — production debug-lock, authenticated data path, secure element (DO-EE-012, DO-SEC). Det — debug-lock verification (IV); access anomaly.
Recommended Action | Lock debug in production; authenticate all wired data/config; pen-test (PV-CYBER-006, IV).
Hazard trace | HA-CYBER-006, HA-SVC-006 · Req/DO trace: SYS / DO-EE-012, DO-SEC / RCR-051/038 · V&V: PV-CYBER-006 (IV).
Owner / Status / Evidence / Notes | Systems + Security / Open / Judgment. Security-induced → IV required.
FA-IFMEA-004 — Internal board-to-board interface: fretting/high-resistance
Description | An internal connector interface degrades under vibration/thermal cycling.
Category / Subsystem-Component | Internal interface · CMP-034 (DO-PCB).
Cause → Mechanism → Failure Mode | Fretting/CTE cycling → high-resistance/intermittent internal link.
Effect | Intermittent subsystem fault → silent gap or false reading → wrong disposition.
S / O / D / AP | 6 / 3 / 7 / Medium (S3)
Current Controls | Prev — prefer soldered nets for safety-critical, retention connectors, conformal coat. Det — BIST; hard to catch.
Recommended Action | Soldered interconnect for safety nets; vibration-qualify; BIST coverage.
Hazard trace | HA-ELEC-008, HA-MFG-008 · Req/DO trace: DO-PCB / RCR-016 · V&V: PV-MFG-008, vibe test.
Owner / Status / Evidence / Notes | Systems + PCB / Open / Judgment. Cross-ref FA-HWFMEA-028.
6.2 Wireless link — LTE-M / Wi-Fi / BLE (INT — DO-IF)
FA-IFMEA-005 ★ — LTE-M link cannot egress from inside the container (RF-in-container)
Description | The LTE-M air interface cannot achieve a usable link from inside the metallised/insulated shipper.
Category / Subsystem-Component | LTE-M link · S5 · CMP-024/028 (DO-IF, DO-EE-018).
Cause → Mechanism → Failure Mode | Container RF attenuation + detune → link budget not met → no live path.
Effect | Excursion with no live notification → compromised dose. Top program item.
S / O / D / AP | 7 / 5 / 6 / High (S4)
Current Controls | Prev — antenna tuned to container budget, external-antenna fallback, S&F buffering, multi-bearer (DO-IF, DO-FW-011). Det — RSSI/connectivity diagnostics; per-trip delivery-success.
Recommended Action | RF-in-container qualification (PRD-VER-005); external antenna if needed; treat delivery as S&F-with-escalation.
Hazard trace | HA-COMM-003 ★ · Req/DO trace: SYS-104/109 / DO-IF, DO-EE-018, DO-SAFE-009 / RCR-031/030 · V&V: PRD-VER-005, PV-COMM-003.
Owner / Status / Evidence / Notes | Systems + RF / Open — top item / Evidence-Pending (antenna). FTA top event (b). Interface-contract view of FA-HWFMEA-029.
FA-IFMEA-006 — LTE-M PSM/eDRX misconfiguration → unreachable or over-draining
Description | Power-saving-mode / eDRX timing between device and network is mismatched, so the device is unreachable when needed or wakes too often.
Category / Subsystem-Component | LTE-M link · S5 · CMP-024 (DO-IF).
Cause → Mechanism → Failure Mode | PSM/eDRX parameter mismatch device↔network → unreachable for alert push or excess wake energy.
Effect | Alert push delayed (compromised dose) or energy budget overrun (dead-in-cold).
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — device-initiated S&F (alerts pushed by device, not awaiting network page), negotiated PSM/eDRX, energy budgeting (DO-FW-004/011). Det — reachability + energy diagnostics.
Recommended Action | Device-initiated alerting so PSM does not delay alerts; validate PSM/eDRX energy vs reachability trade; test over lanes (PV-COMM).
Hazard trace | HA-COMM-001, HA-ENRG-006 · Req/DO trace: SYS-105/108 / DO-IF, DO-FW-004/011 / RCR-028/031 · V&V: PV-COMM-001, PRD-VER-004.
Owner / Status / Evidence / Notes | Systems + RF / Open / Judgment. PSM trade-off is a design-tuning risk.
FA-IFMEA-007 — Wi-Fi offload interface: onboarding/credential failure
Description | The Wi-Fi offload interface cannot join a network (credentials, captive portal, band) so bulk offload fails.
Category / Subsystem-Component | Wi-Fi link · S5 · CMP-025 (DO-IF).
Cause → Mechanism → Failure Mode | Credential/portal/band mismatch → no Wi-Fi offload.
Effect | Bulk telemetry deferred (data preserved by S&F); alerts still go via LTE-M — low safety impact.
S / O / D / AP | 4 / 4 / 3 / Low–Medium (S2)
Current Controls | Prev — Wi-Fi is opportunistic offload only, LTE-M is the primary alert path, S&F. Det — offload success telemetry.
Recommended Action | Keep Wi-Fi non-safety-critical (offload only); simple provisioning; do not gate alerts on Wi-Fi.
Hazard trace | (offload only — no direct safety HA; ties to backlog HA-COMM-002 if LTE-M also down) · Req/DO trace: SYS-106 / DO-IF · V&V: PV-COMM (offload).
Owner / Status / Evidence / Notes | Systems + RF / Open / Judgment. Deliberately non-safety-critical by architecture (LTE-M primary).
FA-IFMEA-008 — Multi-bearer selection thrash or wrong-priority path
Description | The bearer manager oscillates between LTE-M/Wi-Fi/BLE or picks a worse path, wasting energy or delaying alerts.
Category / Subsystem-Component | Bearer selection · S5 · CMP-042 (DO-FW-011, DO-IF).
Cause → Mechanism → Failure Mode | Selection hysteresis error → bearer thrash / wrong-priority selection.
Effect | Energy drain (dead-in-cold) or delayed alert (compromised dose).
S / O / D / AP | 6 / 3 / 5 / Medium (S3)
Current Controls | Prev — energy-aware selection with hysteresis, alert-before-bulk priority (DO-FW-011). Det — bearer-selection + energy telemetry.
Recommended Action | Hysteresis + energy-aware selection; alert path deterministic; test bearer transitions.
Hazard trace | HA-COMM-001, HA-ENRG-006 · Req/DO trace: SYS-108/112 / DO-FW-011, DO-IF / RCR-028/031 · V&V: PV-COMM.
Owner / Status / Evidence / Notes | Systems + RF / Open / Judgment. Overlaps FA-SWFMEA-025 (retry storm).
FA-IFMEA-009 — Antenna coexistence / desense between bands
Description | Simultaneous LTE-M/Wi-Fi/BLE/GNSS operation desenses one band via coexistence.
Category / Subsystem-Component | RF coexistence · S5 · CMP-028 (DO-EE-017/018).
Cause → Mechanism → Failure Mode | In-band/harmonic coupling → desense of a band (esp. GNSS/LTE-M).
Effect | Reduced link/fix in the container → deferred alert / no location.
S / O / D / AP | 5 / 3 / 5 / Medium (S3)
Current Controls | Prev — coexistence design (time/freq separation), antenna plan (DO-EE-017/018). Det — RSSI/fix quality across concurrent operation.
Recommended Action | Coexistence testing; time-multiplex where needed; do not run GNSS concurrently with alert push.
Hazard trace | HA-COMM-003 · Req/DO trace: SYS-109 / DO-EE-017/018 / RCR-031 · V&V: PRD-VER-005, coexistence test.
Owner / Status / Evidence / Notes | Systems + RF / Open / Judgment.
6.3 Device↔cloud API (INT — DO-CLD)
FA-IFMEA-010 ★ — Ingest API not idempotent → dropped or duplicated records
Description | The ingest API contract mishandles retries/reorders, dropping or duplicating uploaded records.
Category / Subsystem-Component | Ingest API · S8 · CMP-046 (DO-CLD, DO-IF).
Cause → Mechanism → Failure Mode | Non-idempotent ingest / weak dedup → dropped or duplicate records.
Effect | Dropped → gap; duplicate/contradictory → confusing record → wrong disposition.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — exactly-once idempotent ingest, ordered/resumable upload, ack-before-mark (CMP-046/042). Det — device↔cloud reconciliation; delivery-path FI.
Recommended Action | Idempotent ingest keys; ack-before-mark; reconciliation; test reorder/dup/loss (PV-COMM-002/004).
Hazard trace | HA-COMM-002, HA-COMM-004 · Req/DO trace: SYS-020/111/176 / DO-CLD, DO-IF / RCR-028/032 · V&V: PV-COMM-002 (delivery-path FI), PV-COMM-004.
Owner / Status / Evidence / Notes | Systems + Cloud / Open / Judgment.
FA-IFMEA-011 ★ — Alert API delivery path fails silently (cloud→responder)
Description | The alert-delivery API accepts an alert but never delivers it to a responder, with no escalation.
Category / Subsystem-Component | Alert/notify API · S8 · CMP-048 (DO-CLD-007/008).
Cause → Mechanism → Failure Mode | Notify-provider failure / mis-route / no escalation → alert not communicated.
Effect | Shipment disposed clean; alert never reaches a responder → compromised dose.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — escalation of unacked alerts, multi-channel notify, delivery confirmation (DO-SAFE-009, DO-CLD-007/008). Det — unacked backlog fleet-wide; delivery-path FI.
Recommended Action | Require acknowledgement + escalation; multi-channel delivery; test delivery-path failures (PV-COMM-002, VAL-COMM-002).
Hazard trace | HA-COMM-002 ★, HA-CLOUD-001 · Req/DO trace: SRS-CLOUD-023, SYS-176 / DO-SAFE-009, DO-CLD-007/008 / RCR-028/029/035 · V&V: PV-COMM-002, VAL-COMM-002.
Owner / Status / Evidence / Notes | Systems + Cloud / Open / Judgment. FTA top event (b) contributor.
FA-IFMEA-012 — API authentication/session (cert/time/DNS) failure blinds device
Description | The device cannot establish an authenticated API session (expired cert, bad time, DNS).
Category / Subsystem-Component | API auth · S3/S8 · CMP-044/052 (DO-SEC, DO-CLD).
Cause → Mechanism → Failure Mode | Cert expiry / time skew / DNS fault → auth handshake fails.
Effect | Device blinded → undelivered alert / backlog → compromised dose.
S / O / D / AP | 7 / 3 / 4 / High (S4)
Current Controls | Prev — cert rotation, trusted time, S&F during outage (CMP-044/042). Det — auth/handshake + backlog diagnostics.
Recommended Action | Robust cert lifecycle + time discipline; S&F preserves data; test cert/time/DNS FI (PV-COMM-005).
Hazard trace | HA-COMM-005 · Req/DO trace: SYS-137/118 / DO-SEC, DO-CLD, DO-EE-011 / RCR-025/032/039 · V&V: PV-COMM-005, VAL-COMM-005.
Owner / Status / Evidence / Notes | Systems + Security / Open / Judgment. Time dependency is a CCF item.
FA-IFMEA-013 — API version/schema skew corrupts telemetry or verdict exchange
Description | Device and cloud API schema/version mismatch corrupts telemetry, config, or verdict exchange.
Category / Subsystem-Component | API contract · S3/S8 · CMP-046/052 (DO-CLD, DO-CFG-005).
Cause → Mechanism → Failure Mode | Schema/version skew (esp. after OTA) → mis-parsed/rejected exchange.
Effect | Lost/garbled telemetry or wrong-profile config → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — versioned API + compatibility matrix, negotiation (DO-CFG-005). Det — schema validation; version-skew detection.
Recommended Action | Backward/forward-compatible versioning; negotiation; reject-not-misparse; version-skew test (link OTA).
Hazard trace | HA-SW-010, HA-COMM-004 · Req/DO trace: SYS-032 / DO-CLD, DO-CFG-005 / RCR-037/030 · V&V: PV-SW-010, PV-COMM-004.
Owner / Status / Evidence / Notes | Systems + Cloud / Open / Judgment.
FA-IFMEA-014 — Export/API interface loses provenance (gaps/time/tamper-evidence)
Description | The export/integration API omits gaps, time, or tamper-evidence, so downstream systems see a misleadingly clean record.
Category / Subsystem-Component | Export/API · S8 · CMP-052 (DO-CLD, DO-SAFE-008).
Cause → Mechanism → Failure Mode | Export format drops provenance/gaps → record looks clean downstream.
Effect | Wrong disposition / data-integrity finding at scale.
S / O / D / AP | 7 / 2 / 4 / High (S4)
Current Controls | Prev — export preserves attribution/time/gaps/tamper-evidence (SYS-027, DO-CLD). Det — export-integrity + traceability test.
Recommended Action | Provenance-preserving export contract; validate export integrity (PV-DATA-006).
Hazard trace | HA-DATA-006, HA-MPE-004 · Req/DO trace: SYS-027/028 / DO-CLD (API), DO-SAFE-008 / RCR-042/032 · V&V: PV-DATA-006.
Owner / Status / Evidence / Notes | Systems + Cloud/Quality / Open / Judgment.
FA-IFMEA-015 — OTA interface: unsigned/unverifiable image accepted
Description | The OTA distribution interface delivers an image the device applies without full verification.
Category / Subsystem-Component | OTA interface · S3/S8 · CMP-043/051 (DO-SAFE-014, DO-SEC-006).
Cause → Mechanism → Failure Mode | Signature not enforced / weak chain → unverifiable image applied.
Effect | Malicious/defective firmware → fleet unsafe/blind (S4/S5).
S / O / D / AP | 8 / 2 / 4 / High (S4)
Current Controls | Prev — signed images verified against secure element, staged/resumable, never-brick + rollback (DO-SAFE-014, DO-FW-014/015, DO-SEC-006). Det — OTA pen-test (IV); campaign health + signing audit.
Recommended Action | Enforce signature verification pre-apply; canary + staged rollout; OTA pen-test (PV-CYBER-005, IV).
Hazard trace | HA-CYBER-005 ★, HA-SW-010 · Req/DO trace: SYS-032/114 / DO-SAFE-014, DO-SEC-006 / RCR-037/038 · V&V: PV-CYBER-005 (OTA pen-test, IV), PV-SW-010.
Owner / Status / Evidence / Notes | Systems + Security / Open / Judgment. Security-induced → IV; cross-ref FA-SWFMEA-026.
6.4 BLE provisioning / pairing (INT — DO-IF/DO-SEC)
FA-IFMEA-016 — BLE provisioning: unauthenticated pairing → rogue configuration
Description | BLE pairing accepts an unauthenticated peer, allowing rogue provisioning/config.
Category / Subsystem-Component | BLE provisioning · S3/S9 · CMP-026/055 (DO-SEC, DO-IF).
Cause → Mechanism → Failure Mode | Weak pairing/bonding → unauthenticated provisioning (wrong profile / suppressed alarm / cloned identity).
Effect | Wrong-threshold monitoring / suppressed alert / mis-attribution → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — authenticated pairing/bonding, low idle exposure when unpaired, secure element (CMP-026/044). Det — pairing anomaly; audit; pen-test (IV).
Recommended Action | Authenticated pairing + bonding; scoped BLE permissions; provisioning audit; pen-test.
Hazard trace | HA-CYBER-003/006, HA-SW-003 · Req/DO trace: SYS-107/110 / DO-SEC, DO-IF / RCR-036/038/055 · V&V: PV-CYBER (BLE, IV).
Owner / Status / Evidence / Notes | Systems + Security / Open / Judgment. IV required.
FA-IFMEA-017 — BLE provisioning: profile mis-configured / not confirmed to device
Description | The provisioning interface reports success but the device did not receive/apply the intended profile.
Category / Subsystem-Component | BLE provisioning · S3/S9 · CMP-055/039 (DO-APP, DO-FW-032).
Cause → Mechanism → Failure Mode | No end-to-end confirmation → provisioning drift (app vs device).
Effect | Device monitors against wrong/absent profile → missed excursion → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — end-to-end profile confirmation (device echoes applied profile), pre-arm validation (DO-FW-032, CMP-055). Det — profile echo + pre-arm check; cloud verifies profile.
Recommended Action | Require device to echo the applied profile; block arming on mismatch; test provisioning round-trip (PV-SW-003).
Hazard trace | HA-SW-003, HA-USE-004 · Req/DO trace: SYS-024/025 / DO-APP, DO-FW-032 / RCR-048/045 · V&V: PV-SW-003, VAL-USE-004.
Owner / Status / Evidence / Notes | Systems + SW/HF / Open / Judgment. Interface-contract view of FA-SWFMEA-004/035.
FA-IFMEA-018 — BLE local read: stale/cached data presented by the app
Description | The BLE local-read interface returns cached/stale data, so the app shows an out-of-date status.
Category / Subsystem-Component | BLE read / mobile · S9 · CMP-026/055 (DO-APP).
Cause → Mechanism → Failure Mode | App cache / read timing → stale data presented as current.
Effect | Operator acts on stale status → wrong disposition (S3).
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — freshness/timestamp on read, edge read stays login-free but time-stamped (CMP-055). Det — timestamp/freshness indicator.
Recommended Action | Timestamp all reads; freshness indicator; force re-read for a disposition decision.
Hazard trace | HA-DATA-005, HA-DATA-007 · Req/DO trace: SYS-022/025 / DO-APP, DO-HF / RCR-040/043 · V&V: PV-DATA-005, VAL-USE-005.
Owner / Status / Evidence / Notes | Systems + Mobile/HF / Open / Judgment.
6.5 Human / HMI interface (INT — DO-HF)
FA-IFMEA-019 ★ — Verdict-display human interface misread (cold/gloved/low-light/language)
Description | The on-device verdict interface is misread under real conditions (cold, gloves, low light, language).
Category / Subsystem-Component | HMI interface · S6 · CMP-030 (DO-HF-002/005, DO-SAFE-013).
Cause → Mechanism → Failure Mode | Poor legibility/affordance in context → verdict misread.
Effect | Wrong disposition (reject read as accept or vice versa).
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — cold-legible display, unambiguous icons/colours, language-independent verdict, glove affordance (DO-SAFE-013, DO-HF-002/005). Det — cloud verdict fallback; misread/complaint rate.
Recommended Action | Cold-display validation with gloves/low-light (PV-USE-001, VAL-USE-001); redundant coding (shape+colour+text); cloud as authoritative fallback.
Hazard trace | HA-USE-001 ★, HA-DATA-005 · Req/DO trace: SYS-015/041 / DO-SAFE-013, DO-HF-002/005 / RCR-043 · V&V: PV-USE-001 (cold-display), VAL-USE-001.
Owner / Status / Evidence / Notes | Systems + HF / Open / Judgment. Interface view of FAL-1600 use-error.
FA-IFMEA-020 ★ — Alert human interface: real alert muted (alarm fatigue)
Description | The alert-notification interface (device/app/cloud) lets a real excursion alert be muted amid alarm fatigue.
Category / Subsystem-Component | Alert HMI · S6/S9 · CMP-048/055 (DO-SAFE-004, DO-CLD-003).
Cause → Mechanism → Failure Mode | Poor alarm management / over-alerting → real alert muted/ignored.
Effect | Excursion not acted on → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — alarm management (prioritise, consolidate, reduce nuisance), escalation of unacked, cannot fully suppress safety alerts (DO-CLD-003, DO-SAFE-009). Det — muted-alert / alarm-volume rate.
Recommended Action | Alarm-management design (reduce nuisance, prioritise), escalation that survives muting, validate alarm handling (PV-USE-002, VAL-USE-002).
Hazard trace | HA-USE-002 ★, HA-CYBER-002 · Req/DO trace: SYS / DO-SAFE-004, DO-CLD-003, DO-SAFE-009 / RCR-047/029 · V&V: PV-USE-002, VAL-USE-002.
Owner / Status / Evidence / Notes | Systems + HF/Cloud / Open / Judgment. Cross-ref FAL-1600.
6.6 Carrier / SIM interface
FA-IFMEA-021 — SIM/eSIM ↔ carrier interface: roaming denial / provisioning mismatch
Description | The SIM↔carrier interface denies roaming or is provisioned for the wrong markets on the lane.
Category / Subsystem-Component | Carrier interface · S5 · CMP-027 (DO-IF, DO-MFG).
Cause → Mechanism → Failure Mode | Roaming agreement gap / wrong profile → no attach across a border.
Effect | No live alert path on the segment → compromised dose (if excursion) / fleet-blind lane.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — roaming-capable multi-IMSI/eSIM, auto network selection, S&F, Wi-Fi/BLE offload (CMP-027, DO-IF). Det — attach/roaming diagnostics; coverage logs.
Recommended Action | Verify roaming coverage per target lane; multi-carrier eSIM; S&F + escalation; roaming test (PV-COMM-006).
Hazard trace | HA-COMM-006 · Req/DO trace: SYS-115/116 / DO-IF, DO-MFG / RCR-031/056 · V&V: PV-COMM-006, VAL-COMM-006.
Owner / Status / Evidence / Notes | Systems + RF/Mfg / Open / Evidence-Pending (carrier/SIM strategy). Cross-ref FA-HWFMEA-032.
FA-IFMEA-022 — Supplier interface: cell/protection-part substitution defeats safety basis
Description | The supplier/AML interface admits a substituted cell or protection part that breaks the safety basis.
Category / Subsystem-Component | Supplier interface · S2 · CMP-006/008/009 (DO-CFG-004, DO-MFG).
Cause → Mechanism → Failure Mode | AML/incoming-inspection gap → wrong/counterfeit safety-critical part.
Effect | Cold-charge/thermal (S4) or endurance shortfall or defeated protection.
S / O / D / AP | 8 / 2 / 5 / High (S4)
Current Controls | Prev — locked AML, incoming inspection/AOI, part-marking verification (DO-CFG-004, DO-TST). Det — incoming test; AOI; supplier audit.
Recommended Action | No unapproved alternates for safety parts; incoming chemistry/rating verification; supplier audit (PV-MFG-002/003).
Hazard trace | HA-MFG-002/003 · Req/DO trace: DO-CFG-004, DO-MFG, DO-TST / RCR-054 · V&V: PV-MFG-002/003.
Owner / Status / Evidence / Notes | Systems + Quality / Open / Judgment. Cross-ref FA-HWFMEA-004; bridges to PFMEA (Stage 15).
7. Traceability Notes
• Every safety-significant interface mode reconciles to a HAL hazard (HA-SW-, HA-COMM-, HA-CLOUD-, HA-CYBER-, HA-DATA-, HA-USE-, HA-ENRG-, HA-MFG-, HA-ELEC-). Gap note: FA-IFMEA-007 (Wi-Fi offload) has no direct safety HA because Wi-Fi is architecturally non-safety-critical (LTE-M primary, S&F) — recorded as acceptable, not a gap.
FTA basic events (FAL-1800) | FA-IFMEA-005/011 (top event b — no/late alert), 010/014 (top event d — corrupt/ false record), 002/022 (top event c — thermal via bad charge/part), 015/016 (top events a/b/d via bad firmware/ config). Security-induced interface modes (003/012/015/016) require independent adversarial verification (IV).
CCF (FAL-2700) | shared time/cert (FA-IFMEA-012), shared cloud API (010/011/013/014), shared antenna path (005/009), shared OTA/signing (015).
8. Closure / Knowledge Notes
Top interface items | RF-in-container LTE-M link (FA-IFMEA-005 — top item), alert-delivery API silent failure (011), idempotent ingest (010), BLE provisioning authentication + round-trip confirmation (016/017), verdict/alert human interfaces (019/020), OTA image verification (015), roaming/carrier interface (021).
• The interface control philosophy is explicit contracts + confirmation + no silent success: every safety interface confirms end-to-end (device echoes profile, cloud acks/escalates alerts, export carries provenance), and no interface reports success it cannot prove.
• Recurring IoT interface pattern for FAL-2200/2300: RF that cannot egress from a shielded enclosure, alert delivery paths that accept-but-never-deliver, provisioning that succeeds on the app but not the device, and export/API that quietly strips provenance — all classic monitoring-device interface escapes.
FAL-1600 — Use-Error Analysis (link to HAL-1500) — Smart Cold Chain Monitor (SCM)
Module FAL-1600 (FA-USE) · ERFKL Stage 11 (ORCH-016) · Owner: Human Factors Engineer with the Principal Reliability/Safety Engineer. Use-error analysis for the SCM, reconciled with the authoritative use-related hazard module HAL-1500 (ORCH-017) — use-related hazards are shared and kept consistent across both. This module is the ERFKL/FMEA view (use error → failure mode → effect → controls); HAL-1500 is the harm-and-exposure-first view. Scored severity-first (Action Priority). Per IEC 62366-1 (usability engineering) and ISO 14971. Feeds Stage 12 RCR, Stage 13 Verification (PV-USE / VAL-USE), and the operator-loop system failures (FA-SYSFMEA).
---
1. Purpose
Analyse foreseeable use errors — arming, placement, reading the verdict/alerts, charging, cleaning/reuse, override — that can cause a wrong disposition (both directions: compromised dose administered or good dose wasted), and the human-factors and design controls that mitigate them. Every use error is reconciled 1:1 with a HAL-1500 use-related hazard (HA-USE-…) so the two libraries stay consistent (module-map cross-link rule).
2. Scope
Provisioning/arming errors (forgot-to-arm, wrong profile), probe placement, verdict misreading, alarm muting / alarm fatigue, override of a reject, charging/cleaning errors, discard-as-dead of a cold device, reuse/re- provisioning errors. Excludes the underlying device logic (FAL-1200), display hardware (FAL-1300/1000), and cloud/app interface mechanics (FAL-1500) except where the use is the failure. The harm/exposure/residual-risk determination is authoritative in HAL-1500.
3. Applicable Standards & References
IEC 62366-1 (usability engineering / use-related risk), ISO 14971, IEC 60601-1-6-style HF concepts (referenced), ALCOA+ (decision on data), WHO PQS operational context. Cited by designation; revision to confirm. HAL-1500 is authoritative for the use-related hazard records.
4. Governance & Method
Shared ERFKL scoring key (FAL-1000 §4), severity-first. Use-error occurrence is a Judgment informed by the use environment (cold, gloves, low light, time pressure, multilingual operators, stock-out fear) — validated by summative usability testing (VAL-USE-…), not asserted. Both harm directions are considered (miss vs waste).
5. Categories
Arming/provisioning · Placement · Verdict reading · Alarm handling · Override · Charging/service · Reuse/cleaning.
6. Use-Error Register (FA-USE)
FA-USE-001 ★ — Operator forgets to arm / arms incompletely → unmonitored trip trusted
Description | A device is shipped believed-armed but was never armed (or arming was voided), so the trip runs unmonitored yet is trusted at disposition.
Use error → Failure mode → Effect | Skip/incomplete arm → unmonitored trip presented as monitored → false reassurance → compromised dose.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — explicit arm with positive confirmation; unmistakable armed/unarmed/voided state; pre-arm readiness gate; cloud deploy-block on non-ready (DO-SAFE-007, DO-FW-032). Det — cloud sees no telemetry from an "armed" device (stale); dashboard flags.
Recommended Action | Positive-confirmation arming; make unarmed/voided unmistakable on device + app; block deploy on non-ready; summative test with realistic pressure.
HAL reconciliation / Hazard trace | HA-USE-004 ★ (authoritative in HAL-1500), HA-SW-003 · Req/DO trace: SYS-023/024 / DO-SAFE-007, DO-FW-032, DO-HF · RCR-045/040 · V&V: PV-SW-003 (pre-arm/void), VAL-USE-004.
Owner / Status / Notes | HF + SW / Open / Judgment. Consistent with FA-SWFMEA-011, FA-SYSFMEA-... .
FA-USE-002 ★ — Wrong trip profile selected → wrong-threshold monitoring
Description | Operator selects/applies the wrong stability profile for the product, so the device monitors against the wrong thresholds.
Use error → Failure mode → Effect | Wrong profile chosen/confirmed → wrong-threshold monitoring → missed excursion (or false reject) → compromised dose / waste.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — profile selection aided + validated, device echoes applied profile, cloud verifies profile-vs-lot, pre-arm gate (DO-APP, DO-FW-032). Det — profile echo + cloud check.
Recommended Action | Guided profile selection tied to the lot; device echo + cloud verification; block arm on mismatch; test (PV-SW-003, VAL-USE-004).
HAL reconciliation / Hazard trace | HA-SW-003 (device against wrong thresholds), HA-USE-004 · Req/DO trace: SYS-024/025 / DO-FW-032, DO-APP · RCR-048/045 · V&V: PV-SW-003, VAL-USE-004.
Owner / Status / Notes | HF + SW / Open / Judgment. Cross-ref FA-SWFMEA-004/035, FA-IFMEA-017.
FA-USE-003 ★ — Probe placed in air / against a wall → unrepresentative reading
Description | The external probe is left in air, against a wall, or outside the product zone, so it doesn't measure the vaccine.
Use error → Failure mode → Effect | Poor probe placement → unrepresentative product-zone reading → missed excursion → compromised dose.
S / O / D / AP | 7 / 4 / 6 / High (S4)
Current Controls | Prev — placement guidance/affordance, retention, int-vs-ext divergence flag (DO-SAFE-005, DO-HF). Det — int-vs-ext divergence; plausibility.
Recommended Action | Design placement affordance; instructional/label guidance; divergence flag catches gross mis-placement; validate placement (VAL-USE-005).
HAL reconciliation / Hazard trace | HA-USE-005 (authoritative), HA-DATA-001 · Req/DO trace: SYS-003 / DO-SAFE-005, DO-HF, DO-MECH-009 · RCR-020/045 · V&V: PV-SW-008 (divergence), VAL-USE-005.
Owner / Status / Notes | HF + Sensing / Open / Judgment. Cross-ref FA-DFMEA-011, FA-HWFMEA-017.
FA-USE-004 ★ — Verdict misread (cold / gloved / low-light / language)
Description | The operator misreads the accept/reject/quarantine verdict under real field conditions.
Use error → Failure mode → Effect | Misread verdict → wrong disposition (reject read as accept or v.v.).
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — cold-legible display, language-independent redundant coding (shape+colour+text), glove/low-light design (DO-SAFE-013, DO-HF-002/005). Det — cloud verdict as authoritative fallback; complaint rate.
Recommended Action | Redundant unambiguous coding; cold/gloved/low-light validation (PV-USE-001, VAL-USE-001); cloud fallback for the decision.
HAL reconciliation / Hazard trace | HA-USE-001 ★ (authoritative), HA-DATA-005 · Req/DO trace: SYS-015/041 / DO-SAFE-013, DO-HF-002/005 · RCR-043 · V&V: PV-USE-001, VAL-USE-001.
Owner / Status / Notes | HF / Open / Judgment. Cross-ref FA-IFMEA-019, FA-DFMEA-012.
FA-USE-005 ★ — Real alert muted / ignored (alarm fatigue)
Description | A genuine excursion alert is muted or ignored amid too many alerts / operational pressure.
Use error → Failure mode → Effect | Mute/ignore real alert → excursion not acted on → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — alarm management (prioritise, consolidate, reduce nuisance), escalation that survives muting, safety alerts cannot be fully suppressed (DO-CLD-003, DO-SAFE-009). Det — muted-alert/alarm-volume rate.
Recommended Action | Reduce nuisance alarms; prioritise; escalation to a second responder if unacked; validate alarm handling (PV-USE-002, VAL-USE-002).
HAL reconciliation / Hazard trace | HA-USE-002 ★ (authoritative), HA-CYBER-002 (malicious-suppression variant) · Req/DO trace: DO-SAFE-004/009, DO-CLD-003 · RCR-047/029 · V&V: PV-USE-002, VAL-USE-002.
Owner / Status / Notes | HF + Cloud / Open / Judgment. Cross-ref FA-IFMEA-020, FA-SYSFMEA-008.
FA-USE-006 ★ — Reject overridden (automation bias / stock-out fear)
Description | The operator overrides a correct reject/quarantine to avoid a stock-out or because of automation distrust, releasing a compromised lot.
Use error → Failure mode → Effect | Override reject → compromised lot released → compromised dose.
S / O / D / AP | 8 / 3 / 5 / High (S4)
Current Controls | Prev — override requires role + reason + audit; clear recommendation with rationale; decision aids to reduce automation bias (DO-HF, DO-CLD audit). Det — override rate + reasons; audit review.
Recommended Action | Governed override (role/reason/audit); track override rate as a first-class metric; design to reduce bias; validate (VAL-USE-007).
HAL reconciliation / Hazard trace | HA-USE-007 ★ (authoritative), HA-MPE-001 · Req/DO trace: DO-HF, DO-CLD (audit) · RCR-044/040 · V&V: PV-DATA-008, VAL-USE-007.
Owner / Status / Notes | HF / Open / Judgment. Cross-ref FA-SYSFMEA-008. The stock-out counter-pressure ties to HA-MPE-003 (false-reject harm) — both directions are real.
FA-USE-007 — Incomplete record accepted as a clean pass
Description | The operator accepts a lot on a record that is actually incomplete/gappy, reading it as a clean pass.
Use error → Failure mode → Effect | Accept on incomplete record → false accept → compromised dose.
S / O / D / AP | 7 / 3 / 5 / High (S4)
Current Controls | Prev — incompleteness made unmistakable (never presents as clean), verdict forces quarantine on gaps, gap-exposure on display/dashboard (DO-SAFE-006, DO-HF-002). Det — gap-exposure; override audit.
Recommended Action | Expose incompleteness prominently; force quarantine, not silent accept; validate gap-comprehension (VAL-USE-004).
HAL reconciliation / Hazard trace | HA-DATA-004 (authoritative), HA-SW-002 · Req/DO trace: SYS-016/151 / DO-SAFE-006, DO-HF-002 · RCR-040/021/022 · V&V: PV-DATA-004, VAL-USE-004.
Owner / Status / Notes | HF + SW / Open / Judgment. Cross-ref FA-SWFMEA-008.
FA-USE-008 — Cold, charge-inhibited device discarded as "dead"
Description | A depot operator mistakes a cold, charge-inhibited (won't-charge-yet) device for a dead one and discards it.
Use error → Failure mode → Effect | Misread charge-inhibit state → serviceable device discarded → lost monitoring capacity / stock-out of monitors.
S / O / D / AP | 5 / 3 / 4 / Medium (S3)
Current Controls | Prev — legible "cold — will charge when warm" indication (LED/OLED), operator guidance (DO-SAFE-013, DO-HF-005). Det — discard/return-as-dead rate.
Recommended Action | Unambiguous charge-inhibit indication + depot training; state-legibility validation.
HAL reconciliation / Hazard trace | HA-USE-009 (authoritative) · Req/DO trace: SYS-094/100 / DO-SAFE-013, DO-HF-005, DO-EE-008 · RCR-046 · V&V: PV-USE (state-legibility), VAL-USE.
Owner / Status / Notes | HF + Service / Open / Judgment. Availability/business, not safety. Cross-ref FA-SYSFMEA-011, FA-HWFMEA-010/012.
FA-USE-009 — Improper charging (out-of-spec charger / cold-charge attempt)
Description | Operator uses a non-compliant charger or repeatedly attempts to charge a cold device.
Use error → Failure mode → Effect | Bad charger / forced cold charge → relies on input protection + charge-inhibit → (if defeated) thermal risk; otherwise nuisance no-charge.
S / O / D / AP | 7 / 3 / 4 / High (S4) (severity from the thermal end-effect the controls prevent)
Current Controls | Prev — input protection (reverse/OVP/short), charge-inhibit interlock, clear charge-state indication (DO-SAFE-001/003). Det — charge-fault annunciation; protection trip.
Recommended Action | Robust input protection; charge-inhibit cannot be user-overridden; clear indication of why it isn't charging; label approved chargers.
HAL reconciliation / Hazard trace | HA-ENRG-005, HA-USE-009 · Req/DO trace: SYS-102/094 / DO-SAFE-001/003 · RCR-003/001 · V&V: PV-ELEC-002, PV-ENRG-001.
Owner / Status / Notes | HF + Power / Open / Judgment. The user error is mitigated by HW controls (the interlock must not be user-defeatable).
FA-USE-010 — Improper cleaning / disinfection damages seals or labels (reuse)
Description | Depot cleaning uses the wrong agent or method, degrading seals/labels over the reuse life (D1).
Use error → Failure mode → Effect | Wrong cleaning agent/method → seal/label degradation → ingress or mis-handling → wrong disposition.
S / O / D / AP | 5 / 4 / 5 / Medium (S3)
Current Controls | Prev — approved cleaning-materials baseline + procedure, chemical-compatible materials (DO-CFG-008, DO-MECH-010, DO-SVC). Det — depot visual inspection; IP retest sampling.
Recommended Action | Specify + train approved cleaning; compatibility-tested materials; retirement criterion for degraded seals (link FA-DFMEA-016, FA-REL-005).
HAL reconciliation / Hazard trace | HA-SVC-007, HA-MECH-004 · Req/DO trace: DO-CFG-008, DO-MECH-010, DO-SVC · RCR-011/052 · V&V: chemical-compatibility test, IP retest.
Owner / Status / Notes | HF + Service / Open / Judgment.
FA-USE-011 — Reused device redeployed without re-provisioning (stale config/binding)
Description | A returned reusable device is redeployed without re-provisioning, carrying a stale profile or trip binding.
Use error → Failure mode → Effect | Skip re-provision → stale config/binding on reuse → wrong-profile monitoring / mis-attribution → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — mandatory re-provision + registry re-binding at depot intake, pre-arm invalidates stale config (DO-SVC, DO-CFG). Det — pre-arm checks; binding verification.
Recommended Action | Enforce re-provision on intake; pre-arm blocks stale config; verify binding.
HAL reconciliation / Hazard trace | HA-SW-003, HA-DATA-002 · Req/DO trace: SYS-024 / DO-SVC, DO-CFG-008 · RCR-048/055 · V&V: PV-SW-003, PV-SVC.
Owner / Status / Notes | HF + Service / Open / Judgment. Cross-ref FA-SYSFMEA-021 (emergent from D1 reuse).
FA-USE-012 — Misunderstanding an estimate as a measurement
Description | The operator treats an interpolated/estimated value on the dashboard as a real measurement.
Use error → Failure mode → Effect | Estimate read as measured → over-confident disposition → wrong/limited disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — explicit provenance labelling (estimate vs measured), no silent gap-fill (DO-CLD, DO-HF). Det — provenance review; comprehension testing.
Recommended Action | Clear provenance labelling; validate comprehension (VAL-USE-005).
HAL reconciliation / Hazard trace | HA-DATA-007 · Req/DO trace: DO-CLD, DO-HF · RCR-041/042 · V&V: PV-DATA-007, VAL-USE-005.
Owner / Status / Notes | HF / Open / Judgment. Cross-ref FA-SYSFMEA-019.
FA-USE-013 — Dashboard chart misread (trend read as opposite of true status)
Description | A dashboard chart/summary is read as the opposite of the true cold-chain status.
Use error → Failure mode → Effect | Chart misread → wrong disposition.
S / O / D / AP | 6 / 3 / 4 / Medium (S3)
Current Controls | Prev — clear visualisation design, unambiguous status, verdict as the primary output (not the raw chart) (DO-HF-007, DO-HF-002). Det — dashboard-misread signals; visualisation review.
Recommended Action | Design charts for correct interpretation; lead with the verdict; validate visualisation (VAL-USE-005).
HAL reconciliation / Hazard trace | HA-DATA-005 · Req/DO trace: DO-HF-007/002 · RCR-040/043 · V&V: PV-DATA-005, VAL-USE-005.
Owner / Status / Notes | HF / Open / Judgment.
7. Traceability Notes
• Every FA-USE record reconciles 1:1 with a HAL-1500 HA-USE-… (or a shared HA-DATA/HA-ENRG/HA-SVC) hazard, which is authoritative for the harm and residual-risk determination (module-map cross-link rule FAL-1600 ↔ HAL-1500). No use error here lacks a HAL hazard; no HAL-1500 use hazard is left without an FMEA-side record.
• Use errors feed the operator-loop system failures (FA-SYSFMEA-008/011/013/019/021) and share controls with the HMI/interface modes (FA-IFMEA-017/018/019/020, FA-DFMEA-012/013).
• Both harm directions are carried: false-accept use errors (FA-USE-001…007) and the false-reject/waste counter-pressure (HA-MPE-003, referenced under FA-USE-006).
8. Closure / Knowledge Notes
Top use-errors | forgot-to-arm/wrong-profile (FA-USE-001/002), probe placement (003), verdict misread (004), alarm muting (005), reject override (006) — all S4, all validated by summative usability testing (VAL-USE-…).
• The controlling philosophy is design-out then guide-and-confirm: make the safe path the easy path (positive arm, redundant verdict coding, guided profile, governed override), then validate with realistic cold/gloved/ pressured summative testing rather than assuming trained behaviour.
• Recurring cold-chain/IoT use pattern for FAL-2200/2300: forgot-to-arm, alarm fatigue, and override under stock-out pressure are the three dominant field use errors for monitoring devices — track override and muted-alert rates as first-class post-market signals.
FAL-1800 — Fault Tree Analysis (FTA) — Smart Cold Chain Monitor (SCM)
Module FAL-1800 (FA-FTA) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Chief Systems Engineer (EM-ENG-001). Deductive top-down fault-tree analysis for the SCM's four highest-severity outcomes, with top events drawn directly from the HAL high-severity (S4/S5) hazardous situations (★ rows of the Hazard-Traceability-Matrix). Each tree gives intermediate/basic events, AND/OR gate logic, minimal cut sets, single-point failures (SPF), and common-cause failures (CCF). Basic events are tied to the discipline FMEA records (FA-HWFMEA/SWFMEA/DFMEA/IFMEA) and to HAL hazards (HA-…). Per IEC 61025. Probabilities are qualitative / Evidence-Pending — no quantitative failure rate is asserted; the tree is used for cut-set and SPF/CCF reasoning, and quantification is deferred to the reliability program (FAL-1900) once field/handbook rates exist.
---
1. Purpose & method
Method | IEC 61025 fault tree. Top event = an undesired safety/mission outcome. Gates: OR (any input → output), AND (all inputs → output). Basic events are the lowest-level, independently-analysed failures (mapped to FMEA IDs). Minimal cut set (MCS) = smallest set of basic events whose simultaneous occurrence causes the top event; a single-event MCS is a single-point failure (SPF).
Evidence honesty | occurrence is qualitative (Evidence-Pending); severity is fixed by the top event (S4/S5). Action Priority is severity-dominant. Quantitative MCS probability is deferred to FAL-1900.
Top events (from HAL ★ S4/S5 situations)
• (a) Undetected temperature excursion → spoiled dose administered (HA-SW-001/HA-DATA-001/HA-MPE-002).
• (b) No/late excursion alert reaches a responder (HA-COMM-001/003/HA-SW-006/HA-CLOUD-001).
• (c) Li-battery thermal runaway/fire (HA-ENRG-001/002/003/HA-TFC-002/003 — S5).
• (d) Corrupted/false cold-chain record accepted as authoritative (HA-CYBER-001/HA-SW-013/HA-CLOUD-004).
---
2. FA-FTA-A — TOP EVENT (a): Undetected temperature excursion → spoiled dose administered
Severity S4 (S5 at mass-administration scale, HA-MPE-002). Action Priority: High. Corresponds to FA-SYSFMEA-001; reconciles to HA-SW-001 ★, HA-DATA-001 ★, HA-MPE-002.
TOP-A: Spoiled dose administered (undetected excursion) = G-A1 (excursion not detected/flagged on the device) AND G-A2 (no downstream barrier catches it) (cloud recompute / QA / operator)
G-A1 = OR: ├─ G-A1a Wrong/unrepresentative measurement (OR): │ ├─ BE-A01 External probe open/intermittent .................. FA-HWFMEA-017 / HA-SW-008 │ ├─ BE-A02 Probe placement unrepresentative (air/wall) ....... FA-DFMEA-011 / HA-USE-005 │ ├─ BE-A03 Sensor calibration drift / biased ................. FA-HWFMEA-018 / HA-DATA-001 │ ├─ BE-A04 Firmware calibration coefficient/units error ...... FA-SWFMEA-005 / HA-SW-014 [CCF, fleet] │ ├─ BE-A05 Internal-sensor thermal lag/self-heat bias ........ FA-DFMEA-010 / FA-HWFMEA-019 │ └─ BE-A06 Sensor plausibility fails to flag fault .......... FA-SWFMEA-007 / HA-SW-008 ├─ G-A1b Detection logic misses the excursion (OR): │ ├─ BE-A07 One-sided (freeze) miss ......................... FA-SWFMEA-001 / HA-SW-001,HA-ENV-001 │ ├─ BE-A08 Cumulative/time criterion wrong ................. FA-SWFMEA-002 / HA-SW-001 │ ├─ BE-A09 Over-aggressive hysteresis masks excursion ...... FA-SWFMEA-003 / HA-SW-001 │ └─ BE-A10 Wrong/default profile (mis-arm) ................. FA-SWFMEA-004 / HA-SW-003,HA-USE-004 ├─ G-A1c Silent monitoring gap (device not monitoring) (OR): │ ├─ BE-A11 Dead-in-cold (energy exhausted) ................. FA-HWFMEA-001,FA-SWFMEA-021 / HA-ENRG-006 [D3] │ ├─ BE-A12 Fuel-gauge over-reports SoC (no warning) ........ FA-HWFMEA-012 / HA-ENRG-006 │ ├─ BE-A13 Hung/reset believed monitoring ................. FA-HWFMEA-016,FA-SWFMEA-028,029 / HA-SW-009 │ ├─ BE-A14 Trip never armed (runs unmonitored) ............ FA-SWFMEA-011 / HA-USE-004 │ └─ BE-A15 Torn/lost record during the excursion .......... FA-SWFMEA-018,FA-HWFMEA-033 / HA-SW-013 └─ G-A1d Verdict falsely "accept" (OR): ├─ BE-A16 False accept on incomplete record ............. FA-SWFMEA-008 / HA-SW-002 └─ BE-A17 Displayed verdict false/stale ................. FA-SWFMEA-010 / HA-SW-007
G-A2 = AND (all downstream barriers fail): ├─ BE-A18 Cloud recompute agrees/also wrong (version skew or same defect) .. FA-SWFMEA-009 / HA-COMM-004 ├─ BE-A19 QA/release sampling does not catch it ............................ (process, FAL-1100/VAL) └─ BE-A20 Operator does not detect anomaly / overrides ..................... FA-SYSFMEA-008 / HA-USE-007
Representative minimal cut sets (qualitative):
SPF candidates (single-event, defeated only by G-A2 barriers) | BE-A04 (firmware calibration error — fleet CCF), BE-A07 (freeze miss), BE-A11 (dead-in-cold), BE-A16 (false-accept verdict). These are single design faults whose only remaining barrier is cloud recompute + QA + operator (G-A2). Because BE-A04/BE-A18 can share the same firmware defect (CCF), the cloud barrier may not be independent — A04 is treated as a functional SPF.
Two-event MCS examples | {BE-A01 probe open, BE-A06 plausibility fails}; {BE-A11 dead-in-cold, BE-A13 self-check misses gap}; {BE-A07 freeze miss, BE-A18 cloud same defect}.
Single-point failures (SPF) & controls:
SPF | Control that removes/mitigates it | Trace
BE-A04 firmware cal error (CCF) | Independent cloud recompute must use a diverse cal path; static analysis + EOL cal | FA-SWFMEA-005, RCR-026, PV-SW-004
BE-A07 freeze miss | Two-sided detection, freeze test = heat | FA-SWFMEA-001, RCR-019, PV-SW-001
BE-A11 dead-in-cold | Energy governor + reserve + honest low-batt + last-gasp record | FA-HWFMEA-001, RCR-006/024/027, PRD-VER-004
BE-A16 false accept | Verdict suppresses accept on any gap/fault | FA-SWFMEA-008, RCR-021/022, PV-SW-002
Common-cause (CCF): shared calibration/firmware (BE-A04 ↔ BE-A18) defeats the device+cloud diversity; cold environment (BE-A11/A12 + RF/A-tree) is a shared initiator. → FAL-2700. Action Priority: High (S4/S5). Residual after controls projected A-Med (HA-SW-001), Evidence-Pending.
---
3. FA-FTA-B — TOP EVENT (b): No/late excursion alert reaches a responder
Severity S4. Action Priority: High. Corresponds to FA-SYSFMEA-002; reconciles to HA-COMM-001 ★, HA-COMM-003 ★ (RF-in-container, Open), HA-SW-006, HA-CLOUD-001.
TOP-B: Excursion committed but no timely alert acted on = G-B1 (device fails to emit/queue the alert) OR G-B2 (alert cannot traverse device→cloud) OR G-B3 (cloud→responder delivery fails) OR G-B4 (responder does not act in time)
G-B1 = OR: ├─ BE-B01 Alert not generated / not prioritised over bulk ....... FA-SWFMEA-023 / HA-SW-006 ├─ BE-B02 Buffered alert lost on reset (volatile queue) ......... FA-SWFMEA-024 / HA-COMM-001 └─ BE-B03 Detection missed it entirely (→ shared with TOP-A) .... G-A1 / HA-SW-001 G-B2 = OR (transport of the alert): ├─ BE-B04 LTE-M cannot egress the container (RF-in-container) ... FA-HWFMEA-029,FA-IFMEA-005 / HA-COMM-003 ★ [Open] ├─ BE-B05 Cold RF margin loss ................................. FA-HWFMEA-030 / HA-COMM-003 ├─ BE-B06 Antenna detune (contents/condensation) ............. FA-HWFMEA-031 / HA-COMM-003 ├─ BE-B07 SIM/roaming/coverage gap on the lane ............... FA-HWFMEA-032,FA-IFMEA-021 / HA-COMM-006 ├─ BE-B08 PSM/eDRX misconfig unreachable .................... FA-IFMEA-006 / HA-COMM-001 └─ BE-B09 Auth/session (cert/time/DNS) failure .............. FA-SWFMEA-032,FA-IFMEA-012 / HA-COMM-005 G-B3 = OR (cloud→responder): ├─ BE-B10 Ingest not idempotent → alert dropped ............. FA-IFMEA-010 / HA-COMM-002 ├─ BE-B11 Notify path accepts-but-never-delivers ........... FA-IFMEA-011,FA-SYSFMEA-... / HA-COMM-002,HA-CLOUD-001 ├─ BE-B12 Rules engine fails to flag / wrong policy ........ FA-SYSFMEA-007 / HA-CLOUD-002 ├─ BE-B13 Cloud outage / ransomware ...................... FA-SYSFMEA-006 / HA-CLOUD-008,HA-CYBER-009 └─ BE-B14 Attacker suppresses/mutes the alarm ............. FA-SWFMEA-033 / HA-CYBER-002 [IV] G-B4 = OR (human): ├─ BE-B15 Real alert muted (alarm fatigue) ............... FA-IFMEA-020,FA-SYSFMEA-008 / HA-USE-002 └─ BE-B16 Unacked alert not escalated ................... FA-SWFMEA-023 / HA-COMM-002
Note the gate structure: TOP-B is a broad OR of ORs — any single stage failing defers/loses the alert. This makes several single-point failures unless mitigated by store-and-forward + escalation (the architecture's key defence: an alert that can't go live is buffered and retried with escalation, converting many G-B2 "no live path" events from loss to delay).
Representative minimal cut sets:
SPF (mitigated to "delay" by S&F+escalation, not "loss") | BE-B04 (RF-in-container — top program item, Open), BE-B11 (notify accept-but-never-deliver), BE-B13 (cloud outage — mitigated by device buffering).
True-loss cut sets (S&F defeated) | {BE-B02 volatile queue lost, reset} — a reset loses the buffered alert → must be closed (persistent queue). {BE-B04 no egress whole trip, AND buffer overwritten} → loss.
CCF cut set | {BE-B04 RF blocked} ∩ {BE-B05 cold margin} — cold + container are a shared initiator that simultaneously reduces margin and blocks egress (same lane).
Single-point failures & controls:
SPF | Control | Trace
BE-B04 RF-in-container | RF-in-container qualification; external antenna; S&F+escalation | FA-HWFMEA-029, RCR-031, PRD-VER-005
BE-B02 volatile alert queue | Persist alert queue in NV; ack-before-dequeue | FA-SWFMEA-024, RCR-028, PV-COMM-001
BE-B11 silent notify failure | Delivery confirmation + escalation; multi-channel | FA-IFMEA-011, RCR-028/029/035, PV-COMM-002
Action Priority: High (S4). Residual projected A-Med, but carried Open until PRD-VER-005 (RF-in-container).
---
4. FA-FTA-C — TOP EVENT (c): Li-battery thermal runaway / fire (S5)
Severity S5 (death / loss of aircraft in occupied cargo hold, HA-TFC-002). Action Priority: High (A-High by rule for S5). Corresponds to FA-SYSFMEA-016; reconciles to HA-ENRG-001/002/003, HA-TFC-002 ★★/003.
TOP-C: Uncontained thermal runaway / fire = G-C1 (a cell reaches thermal-runaway initiation) AND G-C2 (containment/protection fails to arrest it)
G-C1 = OR (initiation): ├─ G-C1a Cold-charge lithium plating → internal short (AND): │ ├─ BE-C01 Charge attempted while cell cold ............ (use/depot; charger insert cold) │ └─ BE-C02 Charge-inhibit interlock fails (HW) ......... FA-HWFMEA-006 / HA-ENRG-001 ★ [AND with C03] │ └─ BE-C03 FW charge-inhibit invariant violated .. FA-SWFMEA-012 / HA-ENRG-001 [HW/FW diversity] ├─ BE-C04 Cell internal short (defect/aging) .............. FA-HWFMEA-003 / HA-TFC-002 ├─ BE-C05 Overcharge (PMIC fault) ........................ FA-HWFMEA-010 / HA-ENRG-002 ├─ BE-C06 OV/OC/OT/short unprotected (BMS fails) .......... FA-HWFMEA-007 / HA-ENRG-002 [AND with C07] │ └─ BE-C07 Secondary fuse/OVP fails to clear ......... FA-HWFMEA-008 / HA-MFG-003 [defence-in-depth] ├─ BE-C08 Mechanical crush (retention fails on shock/drop) FA-DFMEA-008,FA-HWFMEA-... / HA-MECH-009,HA-ENRG-003 ├─ BE-C09 Out-of-spec charge source (input abuse) ........ FA-HWFMEA-009 / HA-ENRG-005 ├─ BE-C10 Wrong/counterfeit cell chemistry ............... FA-HWFMEA-004 / HA-MFG-002 └─ BE-C11 Altitude venting/seal stress .................. FA-DFMEA-007,FA-HWFMEA-005 / HA-ENV-007
G-C2 = AND (arrest fails): ├─ BE-C12 No fire-containment / vent management .......... FA-DFMEA-009 / HA-TFC-002 └─ BE-C13 No thermal-event detection/disclosure .......... (monitoring gap) / HA-TFC-002 field-monitoring
Design intent — protection is layered (AND gates):
Cold-charge branch (G-C1a) | requires BE-C01 AND BE-C02 AND BE-C03 — a cold charge and the HW interlock fails and the FW invariant is also violated. HW/FW diversity (interlock HW-primary, FW cross-check) is what keeps this an AND (not an SPF). CCF risk: if HW and FW share the same NTC or the same enable signal, the AND collapses toward an SPF → see FAL-2700 (must keep the interlock diverse).
Electrical-protection branch | requires BE-C06 AND BE-C07 — AFE fails and secondary fuse fails (two-tier protection). Defence-in-depth by design.
Overall top event | requires G-C1 AND G-C2 — an initiation and containment failure. Containment (BE-C12) is the last barrier that keeps a cell event from becoming an aircraft-hold fire.
Representative minimal cut sets:
Cold-charge | {BE-C01, BE-C02, BE-C03, BE-C12} — 4-event (protected by HW/FW diversity + containment).
Electrical | {BE-C05 overcharge, BE-C06 BMS fail, BE-C07 fuse fail, BE-C12} or {BE-C06, BE-C07, BE-C12}.
Intrinsic cell | {BE-C04 internal short, BE-C12 no containment} — 2-event; the shortest cut set, because an intrinsic cell defect needs only containment to fail. This makes containment (BE-C12) the most load-bearing single control for the S5 outcome.
Mechanical | {BE-C08 crush, BE-C12}.
Single-point failures: by design there is no single-event cut set (no SPF) for TOP-C — every initiation requires containment (BE-C12) to also fail. BE-C12 (containment) is the critical common barrier; its failure is in every minimal cut set → it is the top control priority (a common-barrier rather than an initiator SPF). BE-C04 (intrinsic cell short) cannot be eliminated, only minimised (supplier quality) + contained.
Common-cause (CCF): shared NTC/enable across HW+FW interlock (collapses G-C1a AND); cold environment (BE-C01 likelihood + BE-C11 altitude on the same air lane); shared cell lot defect (BE-C04/BE-C10 fleet-wide). → FAL-2700. Action Priority: High — A-High (S5 by rule). Mandatory thermal-event field monitoring + transport disclosure (HA-TFC-002). Verification: PV-ENRG-001/002, PV-TFC-002/003, UN 38.3, PV-MECH-004.
---
5. FA-FTA-D — TOP EVENT (d): Corrupted / false cold-chain record accepted as authoritative
Severity S4 (wrong disposition / indefensible record; data-integrity-at-scale HA-MPE-004). Action Priority: High. Corresponds to FA-SYSFMEA-005; reconciles to HA-CYBER-001 ★, HA-SW-013 ★, HA-CLOUD-004, HA-DATA-006.
TOP-D: False/corrupt record accepted as valid basis for disposition/audit = G-D1 (record becomes corrupt/false/incomplete) AND G-D2 (integrity checks fail to catch it)
G-D1 = OR: ├─ BE-D01 Torn write on power-fail (cold sag) ............. FA-SWFMEA-018,FA-HWFMEA-011 / HA-SW-013,HA-ELEC-010 [D3] ├─ BE-D02 Flash wear/retention loss / bit errors ......... FA-HWFMEA-033 / HA-CLOUD-004 ├─ BE-D03 Overwrite of unsent/excursion/audit data ....... FA-SWFMEA-019,FA-HWFMEA-034 / HA-SW-002 ├─ BE-D04 Mis-timed record (RTC/timestamp) .............. FA-HWFMEA-014,FA-SWFMEA-015 / HA-SW-005 ├─ BE-D05 Record mis-attributed to wrong device/shipment . FA-SYSFMEA-004 / HA-DATA-002,HA-COMM-008 ├─ BE-D06 Injected/altered data (tamper) ................ FA-SWFMEA-031 / HA-CYBER-001 ★ [IV] ├─ BE-D07 Cloud DB corruption / loss ................... FA-SYSFMEA-006 / HA-CLOUD-004 ├─ BE-D08 Export strips provenance/gaps ................ FA-IFMEA-014 / HA-DATA-006 └─ BE-D09 Duplicate/reordered records not reconciled .... FA-SWFMEA-020,FA-IFMEA-010 / HA-COMM-004 G-D2 = AND (integrity barriers fail): ├─ BE-D10 Atomic-write / torn-write detection fails ...... FA-SWFMEA-018 / DO-SAFE-008 ├─ BE-D11 Tamper-evidence / signature not verified ....... FA-SWFMEA-031,FA-IFMEA-015 / HA-CYBER-001 [IV] ├─ BE-D12 Device↔cloud reconciliation misses it ......... FA-SYSFMEA-003 / HA-COMM-004 └─ BE-D13 DR/backup-restore fails ...................... FA-SYSFMEA-006 / HA-CLOUD-004
Representative minimal cut sets:
SPF candidates (single-event if the paired barrier is absent/shared) | BE-D01 torn write (if atomic-write detection BE-D10 is the same code path → functional SPF — must be diverse); BE-D06 tamper (if signature not verified BE-D11 → SPF; requires IV to confirm the barrier is real).
Two-event MCS | {BE-D01 torn write, BE-D10 detection fails}; {BE-D06 tamper, BE-D11 verify fails}; {BE-D07 DB corrupt, BE-D13 DR fails}; {BE-D05 mis-attribution, BE-D12 reconciliation misses}.
Single-point failures & controls:
SPF | Control | Trace
BE-D01 torn write (cold sag) | Power-fail-safe atomic writes; torn write fails, never valid | FA-SWFMEA-018, RCR-024, PV-SW-013
BE-D06 tamper | Secure-element tamper-evidence; verify at ingest (IV) | FA-SWFMEA-031, RCR-034, PV-CYBER-001
BE-D03 overwrite unsent | Hard no-overwrite of unsent/excursion/audit | FA-SWFMEA-019, RCR-021/024, PV-SW-002
Common-cause (CCF): shared power path (cold sag) drives BE-D01 across the fleet on cold lanes; shared cloud DB (BE-D07 ↔ BE-D13 if backups share the fault); shared crypto/key (BE-D06 ↔ BE-D11). → FAL-2700. Action Priority: High (S4). Verification: PV-SW-013, PV-CYBER-001 (IV), PV-CLOUD (DR), PV-DATA-006.
---
6. Cross-tree observations
Shared basic events across trees | the detection miss (G-A1) is a sub-tree of TOP-B (BE-B03); the dead-in-cold/monitoring gap appears in TOP-A (BE-A11) and TOP-D (BE-D01 via cold sag); cloud outage/attack (BE-B13/BE-D07) spans TOP-B and TOP-D. These shared events are CCF drivers (FAL-2700).
The two program-top items are FTA-critical | RF-in-container (BE-B04) is a functional SPF in TOP-B (Open until PRD-VER-005); D3 dead-in-cold (BE-A11) is an SPF in TOP-A and a cold-sag driver of BE-D01 in TOP-D (Open until PRD-VER-004).
• The S5 tree (c) has no SPF by design (every path needs containment BE-C12 to also fail); the shortest cut set is {intrinsic cell short, no containment}, making containment the single most load-bearing S5 control, and HW/FW interlock diversity the key to keeping the cold-charge branch an AND.
• Independent adversarial verification (IV) is required for the security-induced basic events (BE-B14, BE-D06, BE-D11) before their barriers can be credited.
7. Traceability Notes
• Every top event maps to a HAL ★ situation and a FA-SYSFMEA record; every basic event maps to a discipline FMEA record (FA-HWFMEA/SWFMEA/DFMEA/IFMEA) and a HAL hazard. Controls map to RCR-… (Stage 12) and verifications to PV-…/PRD-VER-… (Stage 13). No basic event lacks an FMEA/HA link.
• SPF and common-barrier items (BE-A04, BE-B04, BE-C12, BE-D01/D06) are handed to Stage 12 as priority RCRs and to FAL-2700 (CCF).
8. Closure / Knowledge Notes
Basic-event inventory | ~60 basic events across the four trees (A: 20, B: 16, C: 13, D: 13), each a minimal cut-set contributor — these count toward the ERFKL failure-record total alongside the FMEA registers.
Quantification deferred (FAL-1900) | cut-set probabilities are qualitative/Evidence-Pending; once cell, antenna, and reliability data exist (PRD-VER-004/005, handbook/field rates), the dominant MCS will be quantified and the residuals confirmed.
Headline conclusions | (a) no undetected-excursion SPF should survive without a diverse second barrier — beware firmware-calibration CCF defeating device+cloud diversity; (b) alerting must be S&F-with-escalation so RF gaps defer rather than lose alerts, and the alert queue must be persistent; (c) the S5 fire tree is AND-protected but only as long as HW/FW interlock diversity and containment hold; (d) record integrity depends on diverse detection (atomic-write ≠ tamper-evidence ≠ reconciliation), independently verified.
FAL-1900 — Reliability Analysis — Smart Cold Chain Monitor (SCM)
Module FAL-1900 (FA-REL) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Power, Materials, and Embedded Software engineers. Reliability program for the SCM: MTBF/availability targets and allocation, reliability-prediction and accelerated-life plan (Weibull, HALT/HASS, burn-in), the D3 −30 °C / 90-day cold-endurance analysis, low-temperature battery aging over the reuse life, flash-endurance analysis, and warranty prediction. Evidence honesty: no MTBF, failure-rate, cycle-count, availability fraction, or reserve figure is invented. Targets are stated as program allocations to be set and validated by test; every number is marked Evidence-Pending with the closing test named. Feeds Stage 13 Verification, FTA quantification (FAL-1800), and warranty/field planning (Stage 17).
---
1. Purpose
Establish the SCM's reliability targets, allocate them across subsystems, define the prediction and accelerated-life test program that will substantiate them, and analyse the three reliability drivers that dominate this product: (1) the D3 −30 °C / 90-day energy endurance, (2) low-temperature battery aging over the multi-trip reuse life, and (3) non-volatile storage endurance/retention over the reuse life. Reliability targets are the quantitative backbone that FTA (FAL-1800) will use to quantify cut sets.
2. Scope
System availability & mission reliability (per-trip and per-reuse-life); MTBF/MTTF allocation by subsystem (S1–S9); reliability prediction method; accelerated-life & Weibull plan; HALT/HASS/burn-in strategy; cold-endurance analysis (D3); battery capacity-fade and cold-aging analysis; flash program/erase-endurance and retention analysis; warranty/return-rate prediction. Excludes the FMEA failure modes themselves (FAL-1000…1500), hazard risk acceptability (HAL), and process capability (FAL-1100/Stage 15) except as reliability inputs.
3. Applicable Standards & References
IEC 60605 / IEC 61124 (reliability testing/demonstration); MIL-HDBK-217F / Telcordia SR-332 (electronic failure-rate prediction — method only); IEC 61649 (Weibull); IEC 60068-2 (environmental/accelerated: cold, damp-heat, thermal cycling, vibration); JEDEC JESD22 / JESD47 (component/flash reliability & retention); IEC 62133-2 (cell); IEEE 1413 (reliability-prediction framework); WHO PQS (cold-chain monitor endurance expectations). Cited by designation; revision & method to confirm — no clause or rate asserted as fact.
4. Governance & Method
Reliability is allocated top-down and demonstrated bottom-up. System targets (Evidence-Pending) are allocated to subsystems; subsystem reliability is predicted (handbook + supplier data) and demonstrated (accelerated life + field). Action Priority context | reliability shortfalls that cause a silent monitoring gap or a thermal hazard inherit the safety Action Priority of the corresponding FMEA/FTA item (severity-first).
Evidence | every target below is a placeholder allocation to be set from the market/PRD reliability requirements (PRD-1200 / SYS-163…176) and validated; no value is invented. Where the direction is known but the number is not, the entry states the quantity to be determined and the closing test.
5. Categories
System availability/mission reliability · Subsystem allocation · Prediction/accelerated-life · Cold endurance (D3) · Battery aging · Flash endurance · Warranty/field.
6. Reliability Register (FA-REL)
6.1 System targets & allocation
FA-REL-001 — System mission-reliability target (per trip) [Evidence-Pending]
Description | Probability that a device correctly and completely monitors a trip and delivers a defensible record and any excursion alert, over a defined trip duration.
Metric / Target | R_mission(trip) = TBD (allocation from PRD-1200/SYS-163). Evidence-backed by the four spine functions (measure, detect, record, alert).
Basis / Method | Series-reliability of sensing→detection→record→alert with the redundancy/S&F structure from the FTA; allocated per subsystem below.
Verification linkage | VAL-SW-001/006, PV-SW/COMM spine tests; field mission-success rate (Stage 17).
Traceability | SYS-163, PRD-1200 / RCR-006/019/028 / HA-SW-001/006, HA-COMM-001.
Owner / Status / Notes | Reliability / Open / Evidence-Pending — target to be set; the silent-gap modes (FA-HWFMEA-001, FA-SWFMEA-021) dominate the allocation.
FA-REL-002 — System availability / "provably-monitoring" target [Evidence-Pending]
Description | Fraction of armed time the device is verifiably monitoring (or honestly flagged as not) — the operational form of the no-false-reassurance invariant.
Metric / Target | A = TBD; plus a hard requirement that unavailability is never silent (any gap is flagged), which is a design guarantee, not a probability.
Basis / Method | Availability = MTBF/(MTBF+MTTR) with self-healing recovery; but the safety-relevant metric is honest-gap coverage, verified by fault injection, not just uptime.
Verification linkage | PV-SW-009 (watchdog/self-check FI); fleet availability/last-seen (Stage 17).
Traceability | SYS-169/170 / RCR-023/027 / HA-SW-009.
Owner / Status / Notes | Reliability + SW / Open / Evidence-Pending. Availability alone is insufficient — silent unavailability is the hazard (FA-SWFMEA-029).
FA-REL-003 — Subsystem reliability allocation (S1–S9) [Evidence-Pending]
Description | Top-down allocation of the system target across enclosure, power, compute, sensing, connectivity, HMI, storage, cloud, mobile.
Metric / Target | per-subsystem failure-rate/Weibull budgets = TBD; power (battery/BMS/charge) and connectivity (RF-in-container) carry the largest allocations by risk (per FMEA/FTA).
Basis / Method | Reliability block diagram from the architecture; series/parallel per the FTA; allocate more budget (tighter) to the subsystems with S4/S5 FMEA modes.
Verification linkage | subsystem accelerated-life (below) + supplier data.
Traceability | SYS-163…176 / architecture RBD / all discipline FMEAs.
Owner / Status / Notes | Reliability / Open / Evidence-Pending. Allocation weighted by severity-first risk.
6.2 Prediction & accelerated-life program
FA-REL-004 — Reliability prediction (handbook + supplier) [Method defined; numbers Evidence-Pending]
Description | Bottom-up parts-count/parts-stress prediction of electronic failure rate as an early estimate and FTA input.
Metric / Target | device λ / MTBF estimate = TBD (MIL-HDBK-217 / Telcordia method, supplier λ where available), explicitly de-rated for the −30…+60 °C environment.
Basis / Method | Parts-stress with temperature/environment factors; treated as an estimate, superseded by demonstrated + field data. No handbook number is asserted as the answer.
Verification linkage | cross-checked against HALT/HASS + field returns.
Traceability | SYS-173 / FTA quantification (FAL-1800).
Owner / Status / Notes | Reliability / Open / Evidence-Pending (method-cited, numbers pending supplier data).
FA-REL-005 — Accelerated-life & Weibull plan [Plan defined]
Description | Accelerated-life testing to characterise wear-out distributions (Weibull β, η) for the life-limited items: battery (cycle + calendar), flash (P/E), seals/elastomers (thermal cycling), connectors (mating cycles), enclosure (vibration fatigue).
Metric / Target | Weibull β/η per item = TBD by test; identify infant-mortality vs wear-out; set the reuse-life retirement criteria (FA-DFMEA-016).
Basis / Method | IEC 61649 Weibull; Arrhenius (temperature) / Coffin-Manson (thermal cycling) / inverse-power (vibration) acceleration models; sample sizes for demonstrated reliability at confidence.
Verification linkage | feeds warranty (FA-REL-012) and reuse-life; IEC 60068-2 environmental.
Traceability | SYS-173 / FA-DFMEA-001/016/018, FA-HWFMEA-033 / RCR-011.
Owner / Status / Notes | Reliability / Open / Plan defined; results Evidence-Pending.
FA-REL-006 — HALT / HASS / burn-in strategy [Plan defined]
Description | HALT to find design/robustness margins (cold/hot step, vibration, combined), HASS as a production stress screen, and a burn-in decision for infant-mortality.
Metric / Target | HALT operating/destruct margins (esp. cold, per D3) = TBD; HASS screen profile; burn-in duration = TBD (only if infant-mortality data warrants).
Basis / Method | HALT stepped stress beyond spec to expose weak links (targets the cold-embrittlement, cold-sag brown-out, latent-solder intermittent, and RF-cold modes from the FMEAs); HASS derived from HALT.
Verification linkage | design robustness → RCR closure; HASS → Stage 15 production.
Traceability | FA-DFMEA-005/006, FA-HWFMEA-013/023, FA-HWFMEA-030 / SYS-173.
Owner / Status / Notes | Reliability / Open / Plan defined; Evidence-Pending.
6.3 Cold endurance (D3 top driver)
FA-REL-007 ★ — 90-day cold (−30 °C) energy-endurance analysis [D3 — top program risk]
Description | End-to-end analysis that the energy budget delivers 90-day monitoring at −30 °C with margin, after cold-capacity derating, reuse-cycle fade, and duty-cycle overhead.
Metric / Target | endurance ≥ 90 days at −30 °C with margin = TBD; energy reserve for last-gasp record commit = TBD (both Open, MR-RISK-001).
Basis / Method | Power budget (SYS-086…090) × measured cold discharge/impedance × reuse-fade × governor duty-cycle (DO-FW-004); worst-case cold lane. No capacity or day-count is invented — closed by measured cold-discharge + a 90-day cold endurance test.
Verification linkage | PRD-VER-004 (90-day cold endurance), gauge-accuracy-over-temp, depletion/last-gasp test.
Traceability | SYS-087/088/090 / DO-EE-003/004, DO-FW-004, DO-SAFE-012 / RCR-006/024/027 / HA-ENRG-006 ★ / FA-HWFMEA-001, FA-SWFMEA-021.
Owner / Status / Notes | Power + Reliability / Open — top program risk / Evidence-Pending (chemistry & budget). This is the single dominant reliability driver; the FTA TOP-A SPF (BE-A11) depends on it.
FA-REL-008 — Graceful-degradation & last-gasp reliability [Evidence-Pending]
Description | Reliability that, on energy depletion, the device sheds load in the right order and commits a final defensible record + honest state (never silently dies mid-write).
Metric / Target | last-gasp commit success = TBD (design target near-certain); shed-order verified.
Basis / Method | Reserve energy budget + power-fail-safe write; verify at the depletion boundary across cold.
Verification linkage | PRD-VER-004, PV-SW-013 (power-fail-during-write at low SoC).
Traceability | SYS-164/171 / DO-SAFE-012 / RCR-027 / FA-SWFMEA-022, HA-ENRG-006, HA-DATA-006.
Owner / Status / Notes | Power + SW / Open / Evidence-Pending.
6.4 Battery aging over reuse life
FA-REL-009 ★ — Low-temperature battery capacity-fade over reuse cycles [Evidence-Pending]
Description | Capacity fade and impedance growth of the low-temp cell over the multi-trip reuse life (D1) with cold cycling — the aging that erodes the 90-day margin over the product's life.
Metric / Target | end-of-life capacity retention over N reuse cycles = TBD; retirement criterion (SoH threshold that blocks arming) = TBD.
Basis / Method | Cycle-life + calendar-aging test at representative cold + charge profile; correlate SoH gauge to true capacity (FA-HWFMEA-012); Arrhenius/throughput aging model. No cycle count or fade % invented.
Verification linkage | cycle-life test, PRD-VER-004 (at aged SoH), gauge-accuracy.
Traceability | SYS-088/097 / DO-EE-003/008 / RCR-006/049 / HA-ENRG-006, HA-SVC-001 / FA-HWFMEA-001/012.
Owner / Status / Notes | Power + Reliability / Open / Evidence-Pending (chemistry). Aging is why a device that passes at BOL can fail the 90-day promise late in life — a reuse-life reliability trap.
FA-REL-010 — Cold-charge safety-reliability (interlock demonstrated reliability) [Evidence-Pending]
Description | Demonstrated reliability of the charge-inhibit interlock + two-tier protection over the reuse life (these are safety controls whose reliability bounds the S5 FTA branch).
Metric / Target | interlock/protection failure probability = TBD (must be low enough that the S5 cut sets stay acceptable); demonstrated by fault injection + life.
Basis / Method | Protection-control reliability (AFE, fuse, NTC path) over thermal/vibration life; FTA input for TOP-C (BE-C02/C06/C07).
Verification linkage | PV-ENRG-001/002, UN 38.3, protection-life test.
Traceability | SYS-091/094/096 / DO-EE-005/007/010, DO-SAFE-001/002 / RCR-001/002 / HA-ENRG-001/002, HA-TFC-002 / FA-HWFMEA-006/007/008.
Owner / Status / Notes | Power + Safety / Open / Evidence-Pending. Safety-control reliability, severity-first.
6.5 Storage endurance
FA-REL-011 — Flash program/erase-endurance & retention over reuse [Evidence-Pending]
Description | NV log memory endurance (P/E cycles from continuous logging over reuse) and data retention at temperature must exceed the reuse life with margin, or the record is lost/corrupted.
Metric / Target | required P/E cycles over reuse life vs part rating = TBD (with wear-leveling + ≥50 % spare); retention at max temp over storage = TBD.
Basis / Method | Logging rate × trip count × reuse cycles → total writes; wear-leveling amplification; JEDEC retention de-rating at temperature. No endurance number invented — closed by part selection + endurance test.
Verification linkage | endurance/retention test (JESD22/JESD47), PV-SW-013.
Traceability | SYS-018/043 / DO-FW-010 / HWR-020 / RCR-024 / HA-CLOUD-004, HA-SW-013 / FA-HWFMEA-033.
Owner / Status / Notes | Storage + Reliability / Open / Evidence-Pending (part/format/endurance).
6.6 Warranty / field
FA-REL-012 — Warranty & field-return-rate prediction [Evidence-Pending]
Description | Predicted field return/failure rate over the warranty and reuse life, feeding spares, service capacity, and reuse-life retirement policy.
Metric / Target | annualised return rate / DPPM by cause = TBD from Weibull (FA-REL-005) + prediction (FA-REL-004), refined by field data (Stage 17).
Basis / Method | Weibull wear-out + infant-mortality + field-return feedback; cause split (battery, seals, connectors, RF, PCBA intermittents) from the FMEAs.
Verification linkage | field-failure monitoring (Stage 17), FAL-2200 historical feedback.
Traceability | SYS-173 / all FMEAs / Stage 17 field-monitoring plan.
Owner / Status / Notes | Reliability + Quality / Open / Evidence-Pending. Reuse-life retirement (seals, battery SoH, flash) is the key warranty lever (D1).
7. Traceability Notes
• The three reliability drivers (cold endurance FA-REL-007, battery aging FA-REL-009, flash endurance FA-REL-011) are the reliability face of the top FMEA/FTA items (FA-HWFMEA-001, FA-SWFMEA-021, FA-HWFMEA-033; FTA TOP-A/TOP-D).
FTA quantification | FA-REL-004/005/010 provide the failure-rate/Weibull inputs the FTA (FAL-1800) needs to quantify its minimal cut sets once data exists.
• Safety-control reliability (FA-REL-010) is severity-first: its target is set by the S5 FTA branch (TOP-C), not by a generic MTBF.
8. Closure / Knowledge Notes
• No reliability number is asserted as fact. Every target is an allocation/placeholder marked Evidence-Pending with a named closing test (PRD-VER-004, cycle-life, endurance/retention, HALT/HASS, UN 38.3, field data). This is the honest posture required by the evidence policy given the Open chemistry and antenna decisions.
Top reliability risks | (1) 90-day cold endurance at −30 °C and its erosion by reuse aging — a device can pass at beginning-of-life and fail the promise late in life; (2) flash endurance under continuous logging over reuse; (3) safety-control (interlock/protection) reliability bounding the S5 fire tree.
• Recurring cold-chain/IoT reliability pattern for FAL-2200/2300: the reuse (multi-trip) life is where reliability is won or lost — battery fade, seal compression-set, connector wear, and flash wear all accumulate per cycle and are invisible at first article; set retirement criteria, not just a warranty date.
FAL-2000 / FAL-2100 — RCA & CAPA Readiness Note — Smart Cold Chain Monitor (SCM)
Modules FAL-2000 (FA-RCA, Root Cause Analysis) + FAL-2100 (FA-CAPA, Corrective/Preventive Action) · ERFKL Stage 11 (ORCH-016) · Owner: Quality/Regulatory (EM-ENG-006) with the Principal Reliability/Safety Engineer. This is a readiness note: at Stage 11 the SCM has no field failures yet, so RCA/CAPA are set up as live processes, pre-wired to the FMEA/FTA/HAL structure, ready to be triggered by the first verification failure, manufacturing escape, or field event. Per ISO 14971 §9 (production & post-production), ISO 13485-style CAPA, and the ERFKL closure chain: Failure → RCA → CAPA → Verification → Lessons Learned → Historical (FAL-2200) → AI (FAL-2300).
---
1. Purpose & posture
Define how the SCM program will investigate root cause and drive corrective/preventive action when a failure occurs (in verification, production, or the field), and how those actions close the loop back into the risk files. No RCA/CAPA records exist yet (pre-verification); this note establishes the method, triggers, owners, templates, and the pre-identified investigation entry points so the process is ready, not invented.
2. RCA (FA-RCA) — method & readiness
Methods to be applied (by problem type)
• 5-Why + Fishbone (Ishikawa) — first-line for most field/production issues (6M: Machine, Method, Material, Measurement, Man, Environment).
• Fault-tree (deductive) — reuse the FAL-1800 trees; a real event is mapped onto the existing tree to find which basic event(s) and cut set occurred.
• Apollo/causal-mapping — for complex multi-cause or common-cause events (esp. the FAL-2700 CCF items).
• Physics-of-failure — for battery/thermal, PCBA-intermittent, seal/aging, RF failures (with FA-MECH library).
Evidence discipline | preserve the failed unit + its record + its trip context; the SCM's own tamper-evident record and recovery/gap logs are primary RCA evidence (a key advantage — the device logs its own faults).
Pre-wired RCA entry points (most-likely first investigations, from the FMEA/FTA top items)
If this is observed | Start RCA at | Trees/records
Missed excursion / false accept in field | detection + calibration + placement + verdict | FTA TOP-A; FA-SWFMEA-001/005/008, FA-HWFMEA-017/018
No/late alert | RF-in-container + S&F queue + cloud notify | FTA TOP-B; FA-HWFMEA-029, FA-SWFMEA-023/024, FA-IFMEA-011
Thermal/venting event | cold-charge + protection + cell lot + mechanical | FTA TOP-C; FA-HWFMEA-002/006/007, FA-DFMEA-008
Lost/corrupt record | power-fail write + flash + tamper + cloud DR | FTA TOP-D; FA-SWFMEA-018, FA-HWFMEA-033
Dead-in-cold / short endurance | energy governor + cold capacity + gauge | FA-HWFMEA-001/012, FA-SWFMEA-021 (D3)
Fleet-wide bias/blindness | CCF (firmware/calibration/cloud/OTA) | FAL-2700 FA-CCF-001/002/007/009
3. CAPA (FA-CAPA) — method & readiness
CAPA flow | Containment (immediate, protect patients/field) → Correction (fix the instance) → Root-cause Corrective Action (fix the cause) → Preventive Action (prevent recurrence class-wide) → Verification of Effectiveness (VoE) → Regression → Release → Lessons Learned.
Severity-first prioritisation | CAPA priority inherits the Action Priority of the affected FMEA/HA item — any S4/S5 safety CAPA is High and expedited (compromised-dose, thermal, fleet-blinding).
Containment tools unique to a connected fleet | OTA (push a fix or a safe-state config), fleet-wide deploy-block / quarantine of affected devices via the registry, cloud-side rule change, recall of a cell/PCBA lot by traceability. OTA is itself risk-controlled (staged, signed, rollback — DO-SAFE-014, FA-CCF-007) so a CAPA fix cannot become a fleet CCF.
Effectiveness verification | every CAPA closes with the same PV-/VAL- test that would have caught it (fault-injection, PRD-VER-004/005, PV-CYBER IV), and the residual risk is re-evaluated in the HAL.
Loop-closure | CAPA outcomes update the FMEA (O/D, controls, status), the HAL residual, and are written to FAL-2200 (Historical) and FAL-2300 (AI Knowledge).
4. Triggers (when RCA/CAPA is invoked)
Verification failure (Stage 13/14), manufacturing escape / EOL-test trend (Stage 15, PFMEA), incoming-inspection reject trend, field event or near-miss (Stage 17 post-market surveillance), complaint, adverse trend in a first-class metric (missed-excursion, alert-latency, override rate, muted-alert rate, thermal-event, endurance, cal-drift, data-integrity deviation), audit finding, or a security incident (with IV).
5. Readiness gaps (honest)
• No field data yet — all RCA/CAPA are latent processes; effectiveness of the process is itself validated by a mock/tabletop exercise before launch (recommended).
Traceability dependency | RCA/CAPA effectiveness depends on device/cell/PCBA lot traceability and the tamper-evident record being in place at transfer (DO-CFG-004/010, DO-SAFE-008) — carried as a design-transfer prerequisite (Stage 15).
• The two Open items (RF-in-container, D3 endurance) are pre-identified as the most probable early CAPA subjects and are already scoped to their closing tests (PRD-VER-004/005).
6. Traceability & closure
• RCA maps a real event to a FAL-1800 basic event / FAL-2700 CCF and a HAL hazard; CAPA maps to an RCR (Stage 12), a design-output change (DOL), and a verification (Stage 13/14), then to FAL-2200/2300. This is the ERFKL closure chain instantiated for the SCM. Status: process defined; no records open (pre-verification).
FAL-2200 / FAL-2300 — Historical & AI Failure-Knowledge Note — Smart Cold Chain Monitor (SCM)
Modules FAL-2200 (FA-HIST, Historical Failure Knowledge Base) + FAL-2300 (FA-AIKB, AI Failure Knowledge) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with Quality/Regulatory (EM-ENG-006). Living modules backed by the Product Engineer LLC knowledge system (04-engineering-knowledge-system). This note captures the recurring IoT / lithium-battery / cold-chain failure patterns the SCM must watch — distilled from the failure literature and reasoning patterns applied in the FMEA/FTA/CCF — and sets up the feedback loop from field/CAPA into these libraries. Evidence honesty: patterns below are generic engineering knowledge / Judgment applied to the SCM; SCM-specific historical records will accrue from Stage 17 field data (none exist yet).
---
1. Purpose
Feed the SCM design review with the accumulated "what usually goes wrong" for this product class, so the FMEAs are not written from a blank sheet, and establish the return loop: field failures / near-misses / CAPA → FAL-2200 (historical) → FAL-2300 (AI-surfaced recurring patterns + best mitigations) → next-cycle FMEA/design.
2. FA-HIST — recurring failure patterns to watch (by domain)
Each pattern names the SCM FMEA/FTA records it maps to and the watch-metric for Stage 17.
2.1 Lithium battery / thermal (highest severity)
• Cold-charge lithium plating → latent internal short → delayed thermal event. The canonical low-temperature Li-ion failure; delayed (may appear cycles later). → FA-HWFMEA-002/006, FTA TOP-C, FA-CCF-005. Watch: charge-inhibit trips, thermal-event reports, latent-short screening.
• Protection defeated by single-fault or shared-cause (AFE + fuse together). → FA-HWFMEA-007/008, FA-CCF-006.
• Cell swelling/venting/offgas in enclosed spaces; altitude venting on air lanes. → FA-HWFMEA-005, FA-DFMEA-007.
• Counterfeit / wrong-lot cells defeating the safety+endurance basis (a repeated supply-chain escape). → FA-HWFMEA-004, FA-CCF-011. Watch: incoming-inspection rejects, lot traceability.
• Fuel-gauge over-reporting SoC in the cold → device dies without warning. → FA-HWFMEA-012.
2.2 Cold-chain monitoring logic (product-defining)
• Heat-biased detection that misses the freeze side — a classic cold-chain logger escape (many loggers were built for "too hot" and under-test the freeze case). → FA-SWFMEA-001, FTA TOP-A (BE-A07). Watch: missed-freeze signals.
• Verdict that accepts on an incomplete/gappy record (silent gap reads as clean pass). → FA-SWFMEA-008, FA-USE-007. Watch: incomplete-record accept/override rate.
• Systematic calibration bias across a fleet (shared method). → FA-SWFMEA-005, FA-CCF-002. Watch: fleet drift.
• Unrepresentative probe placement measuring the device's micro-climate, not the product. → FA-DFMEA-011, FA-USE-003.
2.3 Connectivity / IoT (the "silent" failures)
• RF cannot egress a metallised/insulated enclosure — the container is a Faraday-ish shield; the single most under-estimated cold-chain IoT risk. → FA-HWFMEA-029, FA-IFMEA-005, FTA TOP-B, FA-CCF-010. Watch: per-trip delivery-success, RSSI-in-container.
• Alert delivery paths that accept-but-never-deliver (no ack/escalation). → FA-IFMEA-011, FA-SWFMEA-023.
• Volatile alert/data queues lost on reset (S&F not persistent). → FA-SWFMEA-024, FTA TOP-B true-loss cut set.
• Roaming/coverage gaps on cross-border lanes with no live path. → FA-HWFMEA-032, FA-IFMEA-021.
• Store-and-forward duplication / non-idempotent ingest. → FA-IFMEA-010, FA-SWFMEA-020.
2.4 Firmware / data integrity
• Power-fail during NV write → torn/corrupt record (worse in cold sag). → FA-SWFMEA-018, FTA TOP-D (BE-D01).
• Recovery paths that silently assume continuity after a reset. → FA-SWFMEA-028, FA-CCF-013.
• Bad OTA bricking or de-monitoring a fleet; version skew. → FA-SWFMEA-026, FA-CCF-007.
• Flash wear-out / retention loss over reuse. → FA-HWFMEA-033, FA-REL-011.
• Clock drift / lost trusted time → indefensible records. → FA-HWFMEA-014, FA-CCF-003.
2.5 Mechanical / environmental (reuse life)
• Cold embrittlement + condensation + seal compression-set as a shared cold-environment cluster. → FA-DFMEA-001/ 005/006/019, FA-CCF-012.
• Reuse-life degradation invisible at first article (seals, battery fade, connector wear, flash wear accumulate per cycle). → FA-REL-005/009/011, FA-DFMEA-016. Watch: return-cause split by reuse count.
• Latent PCBA/connector intermittents reading as clean passes. → FA-HWFMEA-023/028, FA-DFMEA-014.
2.6 Human factors / use (dominant field errors)
• Forgot-to-arm / wrong-profile; alarm fatigue; override under stock-out pressure — the three dominant monitoring-device use errors. → FA-USE-001/002/005/006. Watch: mis-arm, muted-alert, override rates.
2.7 Cybersecurity-induced safety
• Injected/altered data → false clean record; alarm suppression; malicious OTA. → FA-SWFMEA-031/033, FA-IFMEA-015, FA-CCF-007/008. Requires independent adversarial verification.
3. FA-AIKB — recurring patterns, preferred mitigations, and design heuristics
The AI knowledge layer surfaces the cross-cutting lessons and the mitigations that repeatedly work:
• "The most dangerous failure is a silent one that reads as a clean pass." Design invariant: honest gap, no false reassurance — every gap flagged, every recovery logged, incomplete never presents as accept.
• "Cold is a single shared initiator." It derates capacity, raises impedance (gauge error), reduces RF margin, drives condensation/embrittlement, and creates the cold-charge hazard — so test the combined cold case, not each factor alone (FA-CCF-004).
• "Diversity must be real, not nominal." Device+cloud must be diverse implementations (not shared firmware); HW+FW charge-inhibit must use separate sensing; multi-bearer is not redundancy against a shared blocked antenna; the self-check must be independent of what it checks (FAL-2700).
• "Fleet-scale common cause turns a unit defect into a population event." One firmware, one calibration, one cloud, one time source, one cell lot — stage rollouts, verify diversity, monitor fleet drift.
• "Store-and-forward with escalation, not guaranteed-live." For a device inside a shielded container on intermittent lanes, alerting is buffered-and-retried-with-escalation; the queue must survive resets.
• "Reuse life is where reliability is won or lost." Set retirement criteria (seal, battery SoH, flash), not just a warranty date.
Preferred mitigations that recur | two-tier battery protection (AFE + independent fuse); HW-primary charge-inhibit with FW cross-check; external-antenna fallback + S&F; power-fail-safe atomic writes; secure-element-anchored tamper-evidence; staged/signed OTA with rollback; independent availability heartbeat; fleet drift surveillance; governed override with audit.
4. Feedback loop (living modules)
In | field failures / near-misses / complaints (Stage 17), verification failures (Stage 13/14), manufacturing escapes (Stage 15 PFMEA), CAPA outcomes (FAL-2000/2100).
Process | each event is classified against the patterns above; new patterns are added; O/D estimates and controls in the FMEAs are updated; the AI layer re-ranks the top recurring mechanisms and best mitigations.
Out | updated FMEA/FTA/CCF O-D and controls, updated HAL residuals, and design heuristics for the next product in the class (the Product Engineer LLC knowledge system).
5. Evidence & status
• Patterns are generic engineering knowledge / Judgment applied to the SCM and cross-linked to specific SCM FMEA/ FTA records; no SCM field-failure records exist yet (pre-launch). SCM-specific historical entries will accrue from Stage 17. No fabricated field statistic is asserted.
• The watch-metrics named above are the seed set for the Stage 17 post-market surveillance plan and align with the Field-Monitoring column of the Hazard-Traceability-Matrix.
FAL-2700 — Common-Cause Failure (CCF) Analysis — Smart Cold Chain Monitor (SCM)
Module FAL-2700 (FA-CCF) · ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with the Chief Systems Engineer (EM-ENG-001). Common-cause / common-mode failure analysis for the SCM: shared dependencies that can defeat redundancy or diversity and turn independent barriers into a single failure — with special focus on the firmware, clock/time, calibration, and power shared dependencies, and on fleet-scale common causes that turn a unit defect into a population event. Screens each shared resource, estimates coupling qualitatively (beta-factor concept), and prescribes diversity/separation/defence-in-depth. Scored severity-first (Action Priority). Feeds Stage 12 RCR, FTA (FAL-1800), and reliability (FAL-1900).
---
1. Purpose
Identify common-cause initiators and shared resources whose single failure could simultaneously disable multiple SCM barriers (device+cloud, HW+FW interlock, atomic-write+tamper-evidence, per-device across the fleet), and specify the diversity, separation/isolation, and defence-in-depth that keep the FTA AND-gates genuinely AND. The dominant concern for a monitoring device is that a shared systematic dependency (one firmware, one calibration method, one cloud, one time source) defeats the very redundancy the safety case relies on.
2. Scope
Shared: firmware/software (single image fleet-wide), calibration (method/coefficients), clock/time source, power path (cold sag), secure element/keys, cloud platform, antenna/RF path, cell lot, NTC/enable signal for the charge-inhibit, and environmental initiators (cold, condensation, altitude, EMC). Excludes the per-item failure modes themselves (FAL-1000…1500) — this module is about their coupling.
3. Applicable Standards & References
IEC 61025 (FTA CCF), IEC 61508 (β-factor / common-cause concepts, method only), ISO 14971, IEC 62304 (systematic software faults), NUREG/CR-6268-style CCF screening (concept). Cited by designation; β-factors are qualitative (High/Med/Low coupling), not quantified — no β value is invented.
4. Governance & Method
Screening | for each shared resource, ask: which independent barriers depend on it? does its failure defeat a redundancy/diversity assumed in the FTA? If yes → CCF item with a diversity/separation control.
Coupling estimate | qualitative High / Medium / Low (β-factor concept) — the strength of the common dependency, Evidence-Pending for quantification.
Action Priority severity-dominant | a CCF that can defeat a barrier to an S4/S5 outcome is High.
5. Categories
Software/firmware · Calibration · Clock/time · Power · Security/keys · Cloud · RF/antenna · Cell/supply · Interlock signal · Environmental initiator.
6. Common-Cause Register (FA-CCF)
FA-CCF-001 ★ — Single firmware image defeats device↔cloud detection diversity (fleet)
Shared resource | One firmware build deployed fleet-wide (detection + verdict + calibration application).
Coupled barriers defeated | on-device detection/verdict AND the "independent" cloud recompute if the cloud reuses device logic/coefficients → the two barriers in FTA G-A2 are not independent.
Coupling | High (identical logic on every device; cloud may mirror it).
Effect / Severity | fleet-wide missed excursions or false accepts → mass administration of compromised doses (S4/S5, HA-MPE-002).
Mitigation (diversity/defence-in-depth) | implement the cloud recompute with a diverse implementation (not a copy of device code), independent test vectors, and staged OTA + canary so a bad build cannot reach the whole fleet at once; N-version thinking for the safety verdict.
Trace | FA-SWFMEA-001/002/005/008, FA-SYSFMEA-001/012, FTA TOP-A (BE-A04↔BE-A18) · HA-SW-001/014, HA-MPE-002 · RCR-019/026/057 · PV-SW-001/004, VAL-SW-001.
Owner / Status / Notes | EM-ENG-001 + SW / Open / High. The single most important CCF: it can defeat the primary safety diversity (device + cloud).
FA-CCF-002 ★ — Shared calibration method/coefficients → systematic fleet bias
Shared resource | Common calibration procedure/coefficient pipeline and depot recal method.
Coupled barriers defeated | every unit's measurement accuracy simultaneously; cross-device plausibility cannot catch a uniform bias.
Coupling | High (same method → same error direction).
Effect / Severity | systematic bias across the fleet → wrong disposition at scale (S4; S5 if it hides excursions broadly).
Mitigation | independent verification of the calibration pipeline; EOL cal coverage; fleet drift surveillance (a uniform shift is detectable statistically even if per-unit checks pass); diverse reference at recal; recal-interval enforcement.
Trace | FA-SWFMEA-005, FA-HWFMEA-018, FA-SYSFMEA-009 · HA-SW-004/014, HA-DATA-001, HA-SVC-001, HA-MFG-001 · RCR-026/049/053 · PV-SW-004, PV-MFG-001, PV-SVC-001.
Owner / Status / Notes | Metrology + SW / Open / High. Fleet drift monitoring is the key detection lever.
FA-CCF-003 — Shared clock/time source → fleet-wide record desync
Shared resource | RTC design + time-discipline scheme common to all devices.
Coupled barriers defeated | timestamp integrity of every record and the timing basis of every excursion/alert.
Coupling | Medium (same design, but drift is per-unit; a systematic time bug is High).
Effect / Severity | mis-timed records/alerts across the fleet → wrong disposition, indefensible audit at scale (S3/S4).
Mitigation | disclose untrusted time (never fake it), opportunistic re-discipline against multiple time sources (network/cell), bound + record drift; a monotonic guarantee that survives brown-out.
Trace | FA-HWFMEA-014, FA-SWFMEA-015, FA-SYSFMEA-014 · HA-SW-005, HA-COMM-005 · RCR-025 · PV-SW-005.
Owner / Status / Notes | Systems + SW / Open / Medium.
FA-CCF-004 ★ — Cold power sag as a shared initiator (dead-in-cold + torn writes + RF loss)
Shared resource | The single battery/power path under −30 °C voltage sag (D3).
Coupled barriers defeated | simultaneously drives dead-in-cold (TOP-A), brown-out torn writes (TOP-D), and RF margin loss (TOP-B) — one environmental cause, three top events.
Coupling | High (cold lane affects the whole device at once).
Effect / Severity | on a deep-cold lane, energy exhaustion + record corruption + alert loss can coincide → compounded compromised-dose risk (S4).
Mitigation | energy reserve for a safe landing (protects record last), power-fail-safe atomic writes, brown-out supervisor, S&F that survives resets, cold-derated design margins; validate the combined cold envelope, not each in isolation.
Trace | FA-HWFMEA-001/011, FA-SWFMEA-018/021, FA-HWFMEA-030 · HA-ENRG-006, HA-ELEC-010, HA-COMM-003 · RCR-006/024/027/031 · PRD-VER-004, PV-SW-013.
Owner / Status / Notes | Power + SW + RF / Open / High. The cold environment is the SCM's dominant shared initiator — the reason to test the combined cold case.
FA-CCF-005 ★ — Shared NTC/enable signal collapses HW+FW charge-inhibit diversity
Shared resource | A single cell-temperature (NTC) sense or a single charge-enable node feeding both the HW interlock and the FW cross-check.
Coupled barriers defeated | the HW/FW diversity that keeps the cold-charge branch an AND (FTA G-C1a) — if both read the same NTC, a shared NTC fault defeats both → AND collapses to SPF.
Coupling | High if shared; the entire safety-diversity claim depends on separating them.
Effect / Severity | cold charge proceeds → plating → thermal runaway/fire (S5, HA-TFC-002).
Mitigation | separate/independent sensing or diverse qualification for HW interlock vs FW check; independent enable paths; NTC fault detection on both; verify diversity by fault injection.
Trace | FA-HWFMEA-006, FA-SWFMEA-012, FTA TOP-C (BE-C02/C03) · HA-ENRG-001, HA-TFC-002 · RCR-001/002 · PV-ENRG-001.
Owner / Status / Notes | Power + Safety / Open / High. A design-separation requirement, not just analysis.
FA-CCF-006 — Two-tier battery protection defeated by a shared cause
Shared resource | AFE and secondary fuse sharing a placement, thermal zone, or supplier/lot.
Coupled barriers defeated | the AFE-AND-fuse defence-in-depth (FTA BE-C06 AND BE-C07).
Coupling | Low–Medium (different technologies, but shared assembly/lot risk).
Effect / Severity | both protections fail together → unprotected cell fault → thermal event (S5).
Mitigation | diverse technologies (electronic AFE + non-resettable fuse), separated placement, independent supply/lot, poka-yoke on the fuse rating; incoming + AOI.
Trace | FA-HWFMEA-007/008, HA-ENRG-002, HA-MFG-003 · RCR-002/054 · PV-ENRG-002, PV-MFG-003.
Owner / Status / Notes | Power + Quality / Open / Medium.
FA-CCF-007 ★ — Bad OTA campaign as a fleet-wide common cause
Shared resource | The OTA distribution mechanism + a single new image.
Coupled barriers defeated | every device's logic simultaneously (detection, verdict, charge-inhibit, comms) — one bad image can make the whole fleet unsafe/blind or brick it.
Coupling | High (single image, single campaign).
Effect / Severity | fleet-scale wrong logic or blindness → compromised doses at scale (S4/S5).
Mitigation | signed images, staged rollout + canary, auto-rollback, version-skew handling, never-brick, independent adversarial verification of the OTA path; do not update armed devices without marking the window.
Trace | FA-SWFMEA-026/027, FA-IFMEA-015, FA-SYSFMEA-010, FTA (all top events) · HA-SW-010, HA-CYBER-005, HA-SVC-002 · RCR-037/038 · PV-SW-010, PV-CYBER-005 (IV).
Owner / Status / Notes | SW + Security / Open / High. Staged rollout is the primary CCF control.
FA-CCF-008 — Shared secure element / key material
Shared resource | Common secure-element design and key-management scheme.
Coupled barriers defeated | device identity, secure boot, record tamper-evidence, and auth — a systemic key/SE compromise defeats several integrity barriers at once.
Coupling | Medium (per-device keys limit spread, but a design/PKI flaw is systemic).
Effect / Severity | falsified records / cloned devices / suppressed alarms across the fleet (S4).
Mitigation | per-device keys, hardware root of trust, key rotation, PKI hygiene, independent security review; no shared static secrets.
Trace | FA-HWFMEA-015, FA-SWFMEA-031/033/034, FA-IFMEA-003/016 · HA-CYBER-001/006, HA-DATA-002 · RCR-034/038/055 · PV-CYBER-001/006 (IV).
Owner / Status / Notes | Security / Open / Medium. IV required.
FA-CCF-009 ★ — Cloud platform as a single point of common failure
Shared resource | One cloud platform for ingest, notify, record store, rules, OTA.
Coupled barriers defeated | alert delivery, record durability, rules, and fleet visibility simultaneously — an outage/ransomware event blinds the whole fleet and can alter/withhold records.
Coupling | High (single platform).
Effect / Severity | fleet-wide undelivered alerts + inaccessible/altered records (S4).
Mitigation | HA/multi-region + tested DR, ransomware-resilient immutable backups, device-side S&F buffering so an outage defers rather than loses data, independent security monitoring/IV.
Trace | FA-SYSFMEA-006/007, FA-IFMEA-010/011, FTA TOP-B/TOP-D · HA-CLOUD-004/008, HA-CYBER-009 · RCR-032/039 · PV-CYBER-009 (IV), DR test.
Owner / Status / Notes | Cloud + Security / Open / High. Device buffering is the on-device mitigation that breaks the coupling for data, though not for live-alert timeliness.
FA-CCF-010 — Shared antenna/RF path (cold + container) blocks all bearers
Shared resource | One integral antenna/RF front-end path for LTE-M/Wi-Fi/BLE.
Coupled barriers defeated | the multi-bearer "redundancy" is not real if all bearers share the same blocked antenna path inside the container.
Coupling | Medium–High (bearers differ, but the container blocks the whole band range and they share the antenna volume).
Effect / Severity | no live alert path at all on a shielded lane → compromised dose (S4).
Mitigation | external-antenna option (physical diversity), S&F+escalation (time diversity), verify RF-in- container across bearers (PRD-VER-005); do not credit multi-bearer as independent until proven.
Trace | FA-HWFMEA-029/030/031, FA-IFMEA-005/009, FTA TOP-B · HA-COMM-003/006 · RCR-031/030 · PRD-VER-005.
Owner / Status / Notes | RF / Open / Medium–High, Evidence-Pending (antenna). Multi-bearer is not genuine redundancy against a container that blocks the shared antenna.
FA-CCF-011 — Cell lot / chemistry as a fleet common cause
Shared resource | A single cell chemistry/lot across a production batch.
Coupled barriers defeated | every device's cell safety + endurance basis simultaneously; a bad lot (defect, wrong chemistry) is fleet-wide.
Coupling | Medium (lot-scoped; AML control limits spread).
Effect / Severity | batch-wide thermal risk and/or endurance shortfall (S4).
Mitigation | locked AML, lot traceability, incoming chemistry/capacity verification, UN 38.3 per lot, staged field introduction of new lots.
Trace | FA-HWFMEA-003/004, FA-IFMEA-022 · HA-MFG-002, HA-ENRG-006, HA-TFC-002 · RCR-054 · PV-MFG-002.
Owner / Status / Notes | Quality + Power / Open / Medium.
FA-CCF-012 — Environmental condensation as a shared initiator (multi-channel electronics fault)
Shared resource | The internal cavity environment (all electronics share it).
Coupled barriers defeated | condensation on the board can fault multiple channels (sensor, storage, protection) at once — an ingress/condensation event is not confined to one subsystem.
Coupling | Medium.
Effect / Severity | multi-channel intermittent fault → wrong disposition; can degrade a protection net (S3/S4).
Mitigation | conformal coat all critical nets, low free-volume/desiccant, robust sealing, condensation-cycle validation; treat as a shared environmental cause in test.
Trace | FA-DFMEA-001/019, FA-HWFMEA-020/025 · HA-TFC-006, HA-ENV-003 · RCR-015/018 · PV-TFC-006.
Owner / Status / Notes | Mechanical + PCB / Open / Medium.
FA-CCF-013 — Shared self-check code path defeats availability detection
Shared resource | The availability self-check running on the same MCU/firmware it is meant to check.
Coupled barriers defeated | if the self-check shares the faulted path, a hang defeats both monitoring and its detection → silent non-monitoring undetected.
Coupling | High (same processor/code).
Effect / Severity | device believed monitoring when hung → compromised dose (S4).
Mitigation | independence — external heartbeat + cloud staleness detection + independent-clock windowed watchdog so detection does not depend on the thing being detected.
Trace | FA-HWFMEA-016, FA-SWFMEA-028/029 · HA-SW-009 · RCR-023/027 · PV-SW-009.
Owner / Status / Notes | SW + Systems / Open / High. Independence is the defining CCF control here.
FA-CCF-014 — Depot/service process as a fleet common cause (recal, firmware, provisioning)
Shared resource | A single depot process/tooling applied to many returned devices (D1 reuse).
Coupled barriers defeated | a process error (wrong recal reference, wrong firmware, bad provisioning) repeats across every device serviced that shift.
Coupling | Medium–High (process-scoped, batch effect).
Effect / Severity | batch of biased/mis-configured/mis-provisioned devices → wrong disposition at scale (S4).
Mitigation | process controls + poka-yoke, post-service BIST on every unit, tool calibration, batch audit; cross-ref PFMEA (FAL-1100, Stage 15).
Trace | FA-SYSFMEA-009/010/022, FA-SWFMEA-006 · HA-SVC-001/002, HA-MFG-001 · RCR-049/050/053 · PV-SVC-001, PV-MFG-001.
Owner / Status / Notes | Service + Quality / Open / Medium–High.
7. Traceability Notes
• Each CCF item names the coupled barriers it defeats and the FTA gate it threatens (esp. the AND gates in TOP-A G-A2, TOP-C G-C1a/protection, TOP-D G-D2). The controls are diversity, separation/isolation, and defence-in-depth, and they trace to RCR-… (Stage 12) and PV-…/PRD-VER-… (Stage 13).
Priority CCFs (High, defeat an S4/S5 barrier) | FA-CCF-001 (firmware/device-cloud), 002 (calibration), 004 (cold power), 005 (NTC/interlock diversity), 007 (OTA), 009 (cloud), 013 (self-check independence).
8. Closure / Knowledge Notes
• The SCM's safety case rests on diversity that must be real, not nominal: device+cloud must be diverse implementations; HW+FW charge-inhibit must use separate sensing; multi-bearer is not redundancy against a container that blocks the shared antenna; the availability self-check must be independent of what it checks.
The two program-top items are also top CCFs | cold (FA-CCF-004) couples three top events; RF-in-container (FA-CCF-010) makes multi-bearer redundancy illusory — both Evidence-Pending (PRD-VER-004/005).
• Recurring cold-chain/IoT CCF pattern for FAL-2200/2300: a single firmware, a single calibration method, a single cloud, a single time source, and a single cold environment are the shared dependencies that turn a unit issue into a fleet event — design diversity and staged rollout are the antidotes.
FMEA-to-Requirements-Hazard Matrix — Smart Cold Chain Monitor (SCM)
ERFKL Stage 11 (ORCH-016) · Owner: Principal Reliability/Safety Engineer with EM-ENG-012 (Traceability). The single-row-per-failure reconciliation spine for the SCM's failure/risk analyses: each safety-significant FMEA/FTA record → its HAL hazard (HA-…) → the requirement / design output it stresses → the risk-control requirement (RCR-…) that controls it → the verification/validation (PV-/VAL-/PRD-VER-) that will prove it → Action Priority (severity-dominant) and status. This matrix is what the FMEA↔HAL reconciliation is audited against and what feeds Stage 12 (RCR), Stage 13 (Verification), and ORCH-018 / EM-ENG-012. All PV-/VAL- and residuals are Evidence-Pending (Stages 13–14). Severity-first per CLAUDE.md rule 3.
---
1. How to read
• FMEA ID — the discipline record (FA-DFMEA/HWFMEA/SWFMEA/SYSFMEA/IFMEA) or FTA basic event.
• HA (HAL) — the authoritative hazard record (Hazard-Traceability-Matrix); ★ = S4/S5 spine hazard.
• Req / DO — the requirement (SYS-/HWR-/SRS-/PRD-) and design output (DO-…) the failure stresses/controls.
• RCR — the controlling risk-control requirement (Stage 12, HAL-2400).
• V&V — planned verification/validation (Stage 13–14).
• AP — Action Priority (High/Med/Low), severity-dominant. FTA = also an FTA basic/top event.
2. Safety-significant reconciliation (S4/S5 and key S3)
FMEA ID | Failure (short) | HA (HAL) | Req / DO | RCR | V&V | AP | FTA
FA-HWFMEA-001 | 90-day cold endurance shortfall (D3) | HA-ENRG-006 ★ | SYS-087/088/090 / DO-EE-003, DO-FW-004, DO-SAFE-012 | 006/024/027 | PRD-VER-004 | High | TOP-A
FA-SWFMEA-021 | Energy governor over-spend (D3) | HA-ENRG-006 ★ | SYS-086/089/090 / DO-FW-004, DO-SAFE-012 | 006/024/027 | PRD-VER-004 | High | TOP-A
FA-HWFMEA-012 | Fuel-gauge over-reports SoC (cold) | HA-ENRG-006 ★ | SYS-097 / DO-EE-008, DO-SAFE-012 | 006/027 | PRD-VER-004, gauge test | High | TOP-A
FA-HWFMEA-002 | Cold-charge plating → latent short | HA-ENRG-001 ★, HA-TFC-002 | SYS-091/094 / DO-EE-007, DO-SAFE-001, DO-FW-005 | 001/002 | PV-ENRG-001, UN 38.3 | High (S5) | TOP-C
FA-HWFMEA-006 | Charge-inhibit interlock fails | HA-ENRG-001 ★, HA-TFC-002 | SYS-094 / DO-EE-007, DO-SAFE-001 | 001/002 | PV-ENRG-001 | High (S5) | TOP-C
FA-SWFMEA-012 | FW charge-inhibit invariant violated | HA-ENRG-001 ★, HA-TFC-002 | SYS-091/094 / DO-FW-005, DO-EE-007 | 001/002 | PV-ENRG-001 | High (S5) | TOP-C
FA-HWFMEA-007 | BMS/AFE fails to trip | HA-ENRG-002, HA-TFC-002 | SYS-096 / DO-EE-005, DO-SAFE-002 | 002 | PV-ENRG-002 | High (S5) | TOP-C
FA-HWFMEA-008 | Secondary fuse fails to clear | HA-ENRG-002, HA-MFG-003 | SYS-096 / DO-EE-010, DO-SAFE-002 | 002/054 | PV-ENRG-002, PV-MFG-003 | High (S5) | TOP-C
FA-HWFMEA-003 | Cell internal short (defect) | HA-TFC-002 ★★, HA-TFC-003 | DO-EE-003 | 002/054 | UN 38.3, PV-TFC-002/003 | High (S5) | TOP-C
FA-HWFMEA-004 | Wrong/counterfeit cell chemistry | HA-MFG-002 | DO-CFG-004, DO-MFG | 054 | PV-MFG-002 | High | TOP-C
FA-DFMEA-008 | Battery retention fails → crush | HA-MECH-009, HA-ENRG-003 | HWR-002 / DO-MECH-012, DO-SAFE-010 | 004 | PV-MECH-004, UN 38.3 | High (S5) | TOP-C
FA-DFMEA-009 | No cell thermal containment | HA-TFC-002 ★★, HA-TFC-003 | DO-MECH-005, DO-PCB-010, DO-SAFE-002 | 002/008 | PV-TFC-002/003 | High (S5) | TOP-C
FA-HWFMEA-009 | Out-of-spec charge source | HA-ENRG-005 | SYS-102 / DO-EE-009, DO-SAFE-003 | 003 | PV-ELEC-002 | High | TOP-C
FA-SWFMEA-001 | One-sided (freeze) miss | HA-SW-001 ★, HA-ENV-001 | SYS-149/150 / DO-FW-006, DO-SAFE-004 | 019 | PV-SW-001, VAL-SW-001 | High | TOP-A
FA-SWFMEA-002 | Cumulative excursion criterion wrong | HA-SW-001 ★ | SYS-010/014 / DO-FW-006 | 019/048 | PV-SW-001, VAL-SW-001 | High | TOP-A
FA-SWFMEA-004 | Wrong/default profile monitored | HA-SW-003, HA-USE-004 | SYS-154/024 / DO-FW-032, DO-SAFE-007 | 048/045 | PV-SW-003, VAL-USE-004 | High | TOP-A
FA-SWFMEA-005 | Calibration coefficient/units error | HA-SW-014, HA-DATA-001 ★ | SYS-007/024 / DO-FW-007/009 | 026 | PV-SW-004, PV-DATA-001 | High | TOP-A (CCF)
FA-SWFMEA-007 | Plausibility fails to flag fault | HA-SW-008 ★ | SYS-007/151 / DO-SAFE-005 | 020/045 | PV-SW-008 | High | TOP-A
FA-SWFMEA-008 | False accept on incomplete record | HA-SW-002 ★ | SYS-016/017/151 / DO-FW-008, DO-SAFE-006 | 021/022 | PV-SW-002, VAL-SW-002 | High | TOP-A
FA-SWFMEA-010 | Displayed verdict false/stale | HA-SW-007, HA-USE-001 | SYS-016/041 / DO-FW-008, DO-HF-002 | 021/043 | PV-SW-007, VAL-USE-001 | High | TOP-A
FA-SWFMEA-011 | Trip runs unmonitored (not armed) | HA-USE-004 ★ | SYS-023/024 / DO-FW-003/032, DO-SAFE-007 | 045/040 | PV-SW-003, VAL-USE-004 | High | TOP-A
FA-HWFMEA-017 | External probe open/intermittent | HA-SW-008 ★, HA-USE-005 | HWR-022/028 / DO-EE-013, DO-SAFE-005 | 020/045 | PV-SW-008 | High | TOP-A
FA-HWFMEA-018 | Temp-sensor drift over reuse | HA-SW-004, HA-DATA-001 ★, HA-SVC-001 | HWR-021/027 / DO-EE-014, DO-FW-009 | 026/045/049 | PV-DATA-001, PV-SW-004 | High | TOP-A (CCF)
FA-DFMEA-010 | Internal-sensor thermal lag/bias | HA-SW-001, HA-DATA-001 | HWR-021/027 / DO-MECH-015, DO-EE-014 | 019/026 | PV-SW-001, PV-DATA-001 | High | TOP-A
FA-DFMEA-011 | Probe placement unrepresentative | HA-USE-005, HA-SW-001 | DO-MECH-009/015, DO-HF | 020/045 | PV-SW-008, VAL-USE-005 | High | TOP-A
FA-SWFMEA-018 | Power-fail → torn/corrupt record | HA-SW-013 ★, HA-ELEC-010 | SYS-160/168 / DO-FW-010, DO-SAFE-008 | 024/022 | PV-SW-013 | High | TOP-D
FA-HWFMEA-033 | Flash wear/retention loss | HA-CLOUD-004, HA-SW-013 | SYS-018/043 / DO-FW-010, DO-SAFE-008 | 024 | endurance test, PV-SW-013 | Med–High | TOP-D
FA-SWFMEA-019 | Overwrite of unsent/excursion data | HA-SW-002, HA-DATA-006 | SYS-018/021 / DO-FW-010 | 021/024 | PV-SW-002 | High | TOP-D
FA-SWFMEA-031 | Record tamper-evidence not enforced | HA-CYBER-001 ★, HA-MPE-004 | SYS-026/028 / DO-FW-016, DO-SEC, DO-SAFE-008 | 034 | PV-CYBER-001 (IV) | High | TOP-D
FA-HWFMEA-014 | RTC failure/drift → untrusted time | HA-SW-005, HA-COMM-005 | SYS-006/118 / DO-EE-011 | 025 | PV-SW-005 | Med | TOP-D
FA-SWFMEA-023 | Alert not generated/prioritised | HA-SW-006 ★ | SYS-011/013 / DO-FW-011, DO-SAFE-009 | 029/028/047 | PV-SW-006, VAL-SW-006 | High | TOP-B
FA-SWFMEA-024 | S&F alert lost on reset | HA-COMM-001/002 | SYS-111/113 / DO-FW-011 | 028/029 | PV-COMM-001/002 | High | TOP-B
FA-HWFMEA-029 | LTE-M no in-container link (RF) | HA-COMM-003 ★ | SYS-104/109 / DO-EE-018, DO-IF, DO-SAFE-009 | 031/030 | PRD-VER-005, PV-COMM-003 | High (Open) | TOP-B
FA-HWFMEA-030 | Cold RF margin loss | HA-COMM-003, HA-COMM-001 | SYS-105 / DO-EE-017 | 031 | PV-COMM, PRD-VER-005 | Med | TOP-B
FA-HWFMEA-032 | SIM/roaming gap → no bearer | HA-COMM-006, HA-COMM-005 | SYS-115/116 / DO-IF, DO-MFG | 031/056 | PV-COMM-006 | High | TOP-B
FA-IFMEA-011 | Alert API accepts-but-never-delivers | HA-COMM-002 ★, HA-CLOUD-001 | SRS-CLOUD-023, SYS-176 / DO-SAFE-009, DO-CLD-007/008 | 028/029/035 | PV-COMM-002, VAL-COMM-002 | High | TOP-B
FA-IFMEA-010 | Ingest not idempotent | HA-COMM-002/004 | SYS-020/111/176 / DO-CLD, DO-IF | 028/032 | PV-COMM-002/004 | High | TOP-B/D
FA-SWFMEA-032 | Auth/session failure blinds device | HA-COMM-005 | SYS-137/118 / DO-FW-016, DO-EE-011 | 025/032/039 | PV-COMM-005, VAL-COMM-005 | High | TOP-B
FA-SWFMEA-026 | Bad OTA wrong-logic/brick | HA-SW-010 ★, HA-CYBER-005 | SYS-032/139 / DO-FW-014/015, DO-SAFE-014 | 037 | PV-SW-010, PV-CYBER-005 (IV) | High | all
FA-IFMEA-015 | Unsigned OTA image applied | HA-CYBER-005 ★, HA-SW-010 | SYS-032/114 / DO-SAFE-014, DO-SEC-006 | 037/038 | PV-CYBER-005 (IV), PV-SW-010 | High | all
FA-SWFMEA-033 | Alarm suppression (unauth) | HA-CYBER-002 ★, HA-USE-002 | DO-FW-016, DO-SEC, DO-SAFE-009 | 035/029 | PV-CYBER-002 (IV) | High | TOP-B
FA-SWFMEA-034 | Debug/dev mode enabled in field | HA-SVC-006, HA-CYBER-006 | DO-EE-012, DO-FW-017, DO-SEC | 051/038 | PV-CYBER-006 (IV) | High | TOP-D
FA-HWFMEA-016 | Watchdog fails to recover hang | HA-SW-009 | SYS-047/169 / DO-FW-013, DO-SAFE-007 | 023/027 | PV-SW-009 | High | TOP-A
FA-SWFMEA-028 | Recovery assumes false continuity | HA-SW-009, HA-SW-013 | SYS-167/169 / DO-FW-013, DO-SAFE-007 | 023/024 | PV-SW-009 | High | TOP-A/D
FA-SWFMEA-029 | Availability self-check fails | HA-SW-009 | SYS-169/170 / DO-FW-012, DO-SAFE-007 | 023/027 | PV-SW-009 | High | TOP-A
FA-HWFMEA-024 | EMC/ESD upset → false/missed | HA-ELEC-004 | DO-EE-019, DO-PCB-010 | 014 | PV-ELEC-004 | High | TOP-A
FA-HWFMEA-023 | Latent PCBA intermittent → gap | HA-ELEC-008, HA-MFG-008 | DO-PCB, DO-TST, DO-SAFE-008 | 016/024 | PV-MFG-008, HASS | Med | TOP-A/D
FA-SYSFMEA-001 | Undetected excursion (system) | HA-SW-001 ★, HA-MPE-002 | SYS-149/150 / DO-SAFE-004/005/006 | 019/020/026 | VAL-SW-001 | High (S5 scale) | TOP-A
FA-SYSFMEA-002 | No/late alert (system) | HA-COMM-001/003 ★ | SYS-011/013 / DO-SAFE-009 | 028/029/031 | PRD-VER-005, VAL-SW-006 | High | TOP-B
FA-SYSFMEA-004 | Record mis-attributed | HA-DATA-002, HA-COMM-008 | SYS-030/012/134 / DO-CLD, DO-SEC | 033/055 | PV-COMM-008, PV-MFG-005 | High | TOP-D
FA-SYSFMEA-005 | Corrupted/false record (system) | HA-CYBER-001 ★, HA-CLOUD-004 | SYS-026/028/168 / DO-SAFE-008, DO-CLD, DO-SEC | 024/032/034/042 | PV-CYBER-001 (IV), PV-CLOUD (DR) | High | TOP-D
FA-SYSFMEA-008 | Reject overridden (operator) | HA-USE-007 ★, HA-MPE-001 | DO-HF, DO-CLD (audit) | 044/040 | VAL-USE-007 | High | TOP-A
FA-SYSFMEA-009 | Depot recal escape (fleet) | HA-SVC-001 ★ | SYS-024 / DO-FW-009, DO-SVC | 049/026 | PV-SVC-001, PV-SW-004 | High | TOP-A (CCF)
FA-SYSFMEA-012 | Mass administration (mission) | HA-MPE-002 ★★ (S5) | all DO-SAFE | 057 | aggregate spine PVs | High (S5) | all
FA-SYSFMEA-013 | False-reject stock-out | HA-MPE-003 ★ | DO-SAFE-004, DO-CLD-003 | 047/044/046 | VAL-USE-002 | High | —
FA-SYSFMEA-016 | Thermal event (system/aircraft) | HA-TFC-002 ★★ (S5) | DO-SAFE-001/002/010 | 002/004/007/008 | PV-TFC-002/003, UN 38.3 | High (S5) | TOP-C
FA-IFMEA-005 | LTE-M no egress in container | HA-COMM-003 ★ | SYS-104/109 / DO-IF, DO-EE-018 | 031/030 | PRD-VER-005, PV-COMM-003 | High (Open) | TOP-B
FA-IFMEA-016 | BLE unauth pairing → rogue config | HA-CYBER-003/006 | SYS-107/110 / DO-SEC, DO-IF | 036/038/055 | PV-CYBER (BLE, IV) | High | TOP-A/D
FA-IFMEA-019 | Verdict human-interface misread | HA-USE-001 ★ | SYS-015/041 / DO-SAFE-013, DO-HF-002/005 | 043 | PV-USE-001, VAL-USE-001 | High | TOP-A
FA-IFMEA-020 | Real alert muted (alarm fatigue) | HA-USE-002 ★ | DO-SAFE-004/009, DO-CLD-003 | 047/029 | PV-USE-002, VAL-USE-002 | High | TOP-B
FA-USE-001 | Forgot-to-arm / incomplete arm | HA-USE-004 ★ | SYS-023/024 / DO-SAFE-007, DO-FW-032 | 045/040 | PV-SW-003, VAL-USE-004 | High | TOP-A
FA-USE-006 | Reject overridden (use) | HA-USE-007 ★ | DO-HF, DO-CLD (audit) | 044/040 | VAL-USE-007 | High | TOP-A
FA-DFMEA-001 | Seal loses IP over cold cycling | HA-ENV-003, HA-TFC-006 | HWR-038/040 / DO-MECH-005/011, DO-SAFE-011 | 015/018 | PV-TFC-006, PV-ENV-003 | Med | —
FA-DFMEA-003 | Probe gland seal/strain-relief | HA-SW-008, HA-ENV-003 | HWR-040/028 / DO-MECH-007, DO-EE-013, DO-SAFE-005 | 020/045 | PV-SW-008, PV-TFC-006 | High | TOP-A
3. Coverage, gaps & CCF notes
• ~65 safety-significant rows above reconcile each top FMEA/FTA record to a HAL hazard, a requirement/DO, an RCR, and a V&V. The full ~215 failure records (FMEA registers + ~60 FTA basic events) live in the module files; this matrix carries the release-relevant safety set (severity-first).