• Every safety-significant failure mode traces to a HAL hazard (FMEA↔HAL reconciliation complete). Explicitly recorded non-safety / acceptable (no direct HA) items — not gaps: FA-HWFMEA-021 (accelerometer, advisory), FA-HWFMEA-022 (GNSS, advisory; energy link to HA-ENRG-006), FA-IFMEA-007 (Wi-Fi offload, architecturally non-safety), FA-HWFMEA-020 (humidity, secondary).
• No orphan safety FMEA modes were found. No HAL ★ (S4/S5) hazard lacks an FMEA-side record — the ★ HAL situations map to FA-SYSFMEA-001/002/005/012/016 (top events) and their contributing discipline records.
• Two items carried Open (mirrors HAL): RF-in-container (FA-HWFMEA-029 / FA-IFMEA-005 → PRD-VER-005) and D3 90-day cold endurance (FA-HWFMEA-001 / FA-SWFMEA-021 → PRD-VER-004).
• CCF drivers (FAL-2700) flagged in-row: shared firmware/calibration (FA-SWFMEA-005, FA-HWFMEA-018, FA-SYSFMEA-009 → FA-CCF-001/002), shared cold power (FA-CCF-004), NTC/interlock diversity (FA-CCF-005), OTA (FA-CCF-007), cloud (FA-CCF-009), self-check independence (FA-CCF-013).
4. Handoff
Stage 12 (RCR) | the RCR column is the demand set the Risk-Control Requirements stage refines (HAL-2400).
Stage 13/14 (V&V) | the V&V column is the verification/validation demand (fault-injection for safety; independent adversarial (IV) for security; PRD-VER-004/005 for the two Open items).
FTA (FAL-1800) | the FTA column identifies which top event each record feeds.
ORCH-018 / EM-ENG-012 | this matrix is the ERFKL segment of the end-to-end traceability spine (Requirement → Function → Failure Mode → Cause → Risk → Mitigation → Design Output → Verification → Validation → Production Control → Field).

## 12 Risk Controls
Risk Controls
The firm, testable safeguards drawn from the safety and failure work — the promises the product must keep to be safe.
↩ Contents
In this section:
• Risk-Control Requirements — Index — Smart Cold Chain Monitor (SCM)
• Control-Hierarchy Analysis — Smart Cold Chain Monitor (SCM)
• PRD-SAFE Reconciliation Note — Smart Cold Chain Monitor (SCM)
• RCR-1000 — Consolidated Risk-Control Requirements Register — Smart Cold Chain Monitor (SCM)
• Risk-Control Traceability — Smart Cold Chain Monitor (SCM)
Risk-Control Requirements — Index — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 12 — Risk-Control Requirements · Orchestrator: ORCH-018 / EM-ENG-012 · Owner: Chief Systems Engineer (EM-ENG-001) + Quality/Regulatory (EM-ENG-006), with System-Safety, Human-Factors, and the discipline engineers. This stage consolidates the SCM's risk controls into the authoritative, deduplicated Risk-Control Requirement register (RCR-001…058) — the formal, testable "shall" set derived from the Hazard Analysis Library (Stage 10) and the Failure/Risk Analyses (Stage 11), reconciled with the preliminary PRD-SAFE (Stage 05). Severity-first per CLAUDE.md rule 3; ISO 14971 control hierarchy applied in order; evidence honesty — testable acceptance criteria, Evidence-Pending where unevidence-backed, nothing invented.
Position in the chain: consumes Stage 10 (HAL: HA-…, RCR-…, HAL-2300 hierarchy, HAL-2700 new risks) and Stage 11 (ERFKL: FA-… recommended actions), reconciled with Stage 05 (PRD-SAFE) and Stage 09 (DOL-1200 DO-SAFE). Feeds Stage 13 (Verification, PV-), Stage 14 (Validation, VAL-), Product/Design Requirements (re-baselined PRD-SAFE), and ORCH-018 / EM-ENG-012 (end-to-end traceability spine).
---
1. Module files
File | Purpose
RCR-1000-register.md | The authoritative consolidated, deduplicated RCR register (RCR-001…058), by subsystem, with full per-RCR fields (requirement, control strategy, source hazard/failure, parent requirement, implementation, verification, validation, independence, new-risk, acceptance criteria, owner, status)
Control-Hierarchy-Analysis.md | For the top hazards, demonstrates inherent/design was preferred over warnings; justifies every information-only control; single-fault / non-defeatability determinations
PRD-SAFE-Reconciliation.md | Maps each PRD-SAFE-001…012 to the RCR(s) it becomes/updates; recommends five new authoritative PRD-SAFE requirements; re-baselines Product Requirements
Risk-Control-Traceability.md | Row-per-RCR chain table: RCR → Hazard → Failure Mode → Design Output → Verification → Validation → Residual Risk
index.md | This file — counts, coverage, and handoff
2. Counts
Metric | Count
Total RCRs (consolidated, deduplicated) | 58
— Inherent safety by design (ISD, Level 1) | 16
— Protective measure (PROT, Level 2) | 28
— Information for safety (INFO, Level 3; 4 design-enforced + 1 pure) | 5
— Process / operational (PROC, Level 4) | 9
Inherent/protective-dominant share (Level 1/2) | 44 / 58 (76%)
Release-critical RCRs (single-fault / non-defeatable) | 7 (001, 002, 019, 021, 022, 028, 029)
RCRs carried Open (validate-by-test, release-gating) | 4 (001, 006, 017 = D3 battery; 031 = RF-in-container)
Preliminary PRD-SAFE requirements reconciled | 12 / 12
New authoritative PRD-SAFE requirements recommended | 5 (PRD-SAFE-013…017)
Safety-significant FMEA recommended-actions folded in (deduped) | ~65 spine rows → mapped to RCRs, 0 new duplicate requirements
RCRs with fault-injection verification (safety) | all safety-critical (PV-…, FI)
RCRs with independent adversarial verification (security) | 6 (034–039) + 051 (PV-CYBER-…, IV)
3. Coverage assertions
• Every High/unacceptable HAL hazard has ≥1 RCR. All ~60 top hazard rows (Hazard-Traceability-Matrix, all R-region and S4/S5 ★ hazards) trace to a controlling RCR in Risk-Control-Traceability.md §2; no R-region hazard is uncontrolled.
• Every RCR has an implementation reference (DO-) and a verification target (PV-/VAL-). No control is on paper only (HA-CIR / HAL-2500 posture preserved); the register (RCR-1000-register.md) and traceability table both carry these columns for all 58.
• No duplicates. The HAL RCR register is the spine; each safety-significant FMEA recommended action is folded into an RCR as a Source Failure Mode (dedup map, RCR-1000-register.md §5), not as a parallel requirement; each PRD-SAFE intent maps to ≥1 RCR (PRD-SAFE-Reconciliation.md). No FMEA action or PRD-SAFE intent is orphaned; no RCR is duplicated.
• Hierarchy applied in order. Inherent/protective preferred over warnings for every top hazard; the warnings-challenge is discharged for every use-related risk; the one pure-information control (RCR-052 EOL disposal) is justified because no higher control removes the disposal hazard (Control-Hierarchy-Analysis.md).
• New risks tracked. Every RCR that introduces or balances a new risk carries a HAL-2700 (HA-NEW-…) note; the dominant safety-vs-availability trade (HA-NEW-006) and the fleet common-cause/connectivity trade (HA-NEW-008) are carried ALARP in both directions.
• Evidence honesty. No RCR asserts an unevidence-backed quantity; every temperature limit, endurance margin, RF margin, detection/alarm/calibration target, drift bound, DR objective, and pen-test result is marked Evidence-Pending with its named closing test (RCR-1000-register.md §7).
4. Dominant SCM controls (the register's spine)
1. Honest excursion-detection & verdict integrity — no false "OK": RCR-019 (two-sided detection), RCR-020 (probe-fault), RCR-021 (verdict suppresses false accept, non-defeatable), RCR-022/040 (honest gap/qualifier).
2. Independent alarm generation + delivery + escalation: RCR-028 (local commit, offline-first), RCR-029 (alert-ahead/ack/escalate), RCR-030 (store-and-forward), RCR-035 (non-suppressible), RCR-047 (tuned, no fatigue).
3. Temperature-measurement accuracy & calibration integrity: RCR-026 (calibration integrity + pre-arm block), RCR-049/053 (interval + EOL coverage), RCR-014 (EMC integrity of the decision path).
4. Charge-inhibit-below-0 °C + cell protection + thermal safety (Li-battery): RCR-001 (HW inhibit, non-defeatable), RCR-002 (protection + secondary fuse, single-fault), RCR-003/004/005/008 (input, retention, self-heat bound, containment), RCR-007/054 (transport, correct chemistry/part).
5. Data authenticity / audit-trail (Part-11-style): RCR-034 (capture-time authenticity), RCR-042 (export integrity), RCR-041 (provenance), RCR-025 (trusted time).
6. Secure signed OTA + rollback: RCR-037 (signed, rollback-safe A/B, never bricks), RCR-038 (secure boot/debug lock), RCR-039 (cloud resilience).
7. Fail-safe / conservative-verdict behavior: RCR-021/022/023/024 (verdict/safe-state/record integrity), balanced against wastage by RCR-047/044/040 (HA-NEW-006 ALARP both ways).
8. Human-decision support: RCR-043 (legible verdict), RCR-044 (recorded override), RCR-045 (pre-arm gate), RCR-046 (cold ≠ dead legibility).
5. Carried-Open (release-gating, validate-by-test)
• D3 low-temp battery (RCR-001/005/006/017): 90-day −30 °C endurance + reserve, cold-charge safety, reuse-fade — Evidence-Pending → PRD-VER-004 + PV-ENRG-001 + UN 38.3.
• RF-in-container (RCR-031, consequence-bound by RCR-030): in-container link margin + integral-vs-external antenna — Evidence-Pending → PRD-VER-005 + PV-COMM-003.
Both have correct Level-1 control placement; only their effectiveness is unproven (a verification gap, not a hierarchy gap).
6. Handoff
Stage 13 (Verification, PVL) | the PV- column (with FI for safety, IV for security) is the verification demand set.
Stage 14 (Validation, PVaL) | the VAL- column and the usability-gated use controls (VAL-USE-001/002/004/007).
Product/Design Requirements | the re-baselined PRD-SAFE (5 new + 5 expanded), PRD-SAFE-Reconciliation.md.
ORCH-018 / EM-ENG-012 | this stage is the Stage-12 segment of the end-to-end traceability spine, joining Stage 10 (Hazard-Traceability-Matrix) and Stage 11 (FMEA-to-Requirements-Hazard-Matrix).
Release gate | per Risk-Management-Summary §7–8, release is conditional on all safety-control PV- passing with fault-injection, the two pathway choke points verified/validated, the S5 aircraft-fire + UN 38.3 qualified, the two Open items closing (PRD-VER-004/005), independent security verification, and the WHO PQS accuracy target confirmed.
7. Review history
• 2026-07-19 v1.0 — Initial Stage-12 consolidation (ORCH-018 / EM-ENG-012). 58 deduplicated RCRs; hierarchy analysis; PRD-SAFE reconciliation; RCR→…→residual traceability. All PV-/VAL- Evidence-Pending (Stages 13–14); D3 battery and RF-in-container carried Open.
Control-Hierarchy Analysis — Smart Cold Chain Monitor (SCM)
Stage 12 (ORCH-018 / EM-ENG-012) · Owner: System-Safety with Chief Systems (EM-ENG-001), Quality/Regulatory (EM-ENG-006), and Human-Factors. This module demonstrates, for the SCM's top hazards, that the risk-control requirements (RCR-001…058, RCR-1000-register.md) apply the ISO 14971 control hierarchy in order — inherent/design controls were preferred over protective measures, and protective over information/warnings — and it justifies every control that remains information-only. It is the Stage-12 refinement of HAL-2300 (control hierarchy) applied specifically to the consolidated RCR set. Severity-first per CLAUDE.md rule 3.
---
1. The hierarchy and the rule
ISO 14971 requires risk to be reduced in a fixed priority, and forbids accepting a lower-priority control where a higher one is practicable (ALARP):
1. Inherent safety by design (ISD, Level 1) — remove or design out the hazard, or make the safe outcome the default that no fault, config, or user action can undo. Strongest; preferred.
2. Protective measures (PROT, Level 2) — detect the hazardous condition and act (interlock, backstop, limiter, safe-state, escalating alarm). Bounds the hazard; does not remove it.
3. Information for safety (INFO, Level 3) — labels, IFU, on-device status, training. Weakest; depends on a human perceiving, understanding, and acting. In the SCM these are design-enforced wherever they carry real risk.
4. Process / operational (PROC, Level 4) — manufacturing, calibration, service, transport, disposal, governance.
Two program rules the SCM cannot be released without (HAL-2200 §6.4/§6.5, HAL-2300):
Warnings-challenge | a risk reduced only by information for safety, where inherent or protective control is feasible, is not accepted — it is re-controlled at a higher level. Decisive here because the last-mile disposition user is low-training, working at −30 °C, gloved, in low light, under stock-out pressure, and sometimes not reading a primary language.
Single-fault | for S4/S5 physical (battery/charge/thermal) hazards a single fault shall not reach harm; an independent, common-cause-resistant backstop is required.
Portfolio result: inherent/protective-dominant — 44 of 58 controls at Level 1/2 (76%). Only 5 controls are information-level (4 design-enforced + 1 justified pure-information), and 9 are process-level. No S4/S5 safety risk is carried by information alone.
Level | Count | RCRs
1 — ISD (preferred) | 16 | 001, 004, 006, 008, 009, 010, 011, 012, 013, 017, 019, 021, 028, 031, 043, 044
2 — PROT | 28 | 002, 003, 005, 014, 015, 018, 020, 023, 024, 025, 026, 027, 029, 030, 032, 033, 034, 035, 036, 037, 038, 039, 042, 045, 047, 048, 049, 051
3 — INFO (4 design-enforced + 1 pure) | 5 | 022, 040, 041, 046 (design-enforced); 052 (pure information — justified §4)
4 — PROC | 9 | 007, 016, 050, 053, 054, 055, 056, 057, 058
---
2. Top hazards — inherent/design preferred over warnings (worked demonstration)
For each top hazard, the table shows what a warning/training-only control would have said, and the higher-level RCR that carries the risk instead. This is the core evidence that the hierarchy was applied in order.
2.1 Physical-energy harm (battery / charge / thermal) — Level 1/2 with single-fault backstop
Top hazard (Sev) | Warning-only alternative (rejected) | Higher-level control that carries it | Level
Cold-charge plating → latent short → runaway (HA-ENRG-001, S4→S5) | "Warm the device before charging" label | RCR-001 MCU-independent HW charge-inhibit below 0 °C; persists in any FW/comms fault; non-defeatable | 1 ISD
Cell OV/OC/OT/short → thermal runaway (HA-ENRG-002 / HA-TFC-002, S5) | "Do not abuse the battery" | RCR-002 autonomous protection + independent non-resettable secondary fuse/OVP (single-fault backstop) | 2 PROT + SF
Cell crush/puncture on drop/transport (HA-ENRG-003 / HA-MECH-009, S5) | "Handle with care" | RCR-004 cell retention (no crush) + RCR-008 fire-enclosure containment (independent backstop) | 1 ISD
Aircraft-hold fire (HA-TFC-002, S5) | "Ship per regulations" note | RCR-002/004/008 (design) + RCR-007 UN 38.3/IATA process qualification | 1/2 + 4
Single-fault demonstration (S4/S5 physical): every physical S4/S5 hazard has a primary Level-1/2 control and an independent, common-cause-resistant backstop (HW inhibit independent of FW; secondary fuse independent of the BMS IC; retention vs containment are different mechanisms; input front-end vs BMS are two layers). Warning is never the primary control for a fire/thermal hazard.
2.2 Information-integrity harm (missed/incorrect/delayed excursion → wrong disposition) — the SCM-defining class
This is where the temptation to use labels/training is strongest (the harm is mediated by a human decision) and where the warnings-challenge is decisive. Every one of these is carried by inherent or protective design:
Top hazard (Sev) | Warning-only alternative (rejected) | Higher-level control that carries it | Level
Missed excursion (esp. freeze) → compromised dose (HA-SW-001, HA-ENV-001, S4/S5) | "Check the vaccine on arrival" | RCR-019 two-sided detection, freeze reliability = heat, no undetected path | 1 ISD
Device reports temps it isn't measuring (probe in air) (HA-SW-008, HA-USE-005, S4) | "Place the probe correctly" | RCR-020 probe-fault/plausibility — an unrepresentative read is never a valid temperature | 2 PROT
False accept on incomplete record (HA-SW-002, HA-DATA-004, S4) | "Verify the record is complete" | RCR-021 verdict cannot show a clean pass over a gap (non-defeatable) + RCR-022/040 design-enforced honest gap/qualifier | 1 ISD + 3 D-E
Excursion with no live alert (HA-COMM-001/002, S4) | "Watch for alerts" | RCR-028 local commit (offline-first) + RCR-029 alert-ahead, ack-required, escalate-if-unacknowledged | 1 ISD + 2 PROT
Real alert muted (alarm fatigue) (HA-USE-002, S4) | "Pay attention to alerts" | RCR-047 alerting tuned to validated targets + consolidation; RCR-029 escalation to a different responder | 2 PROT
Verdict misread — cold/gloved/low-light/language (HA-USE-001/008, S4) | "Train the user to read the OLED" | RCR-043 verdict is icon+colour+word, language-independent, readable gloved/cold/low-light | 1 ISD
Reject overridden under pressure (HA-USE-007, S4) | "Advise against overriding" | RCR-044 override requires a deliberate, confirmed, recorded action; accidental acceptance designed out | 1 ISD
Forgot-to-arm / void trip believed monitored (HA-USE-004, wrong profile HA-USE-003, S4) | "Remember to arm; select the right profile" | RCR-045 pre-arm readiness gate blocks arming unless SoC/cal/FW/profile/probe pass (non-defeatable) | 2 PROT
Reset/brown-out gap reads uneventful (HA-SW-013, S4) | "Look for gaps" | RCR-024 power-fail-safe record + RCR-022 honest gap (never a clean pass) | 2 PROT + 3 D-E
Cold charge-inhibited device discarded as dead (HA-USE-009, S3) | "Don't throw away cold devices" | RCR-046 inhibited state legible with cause (design-enforced) + RCR-001 inherent HW charge safety | 3 D-E + 1 ISD
Conclusion: every information-integrity top hazard that a naive design would have controlled with a label/training is instead carried by a Level-1 (inherent) or Level-2 (protective) control, or by a design-enforced information control (the device forces the honest state). The warnings-challenge is fully discharged for the disposition-decision harm class.
---
3. Warnings-challenge ledger (every use-related risk, and its higher-level carrier)
Use hazard (would default to a warning) | Warnings-challenge outcome — higher-level control | Level
Misread verdict (HA-USE-001/008, HA-ENV-002) | RCR-043 legible, language-independent verdict | 1
Override a reject under pressure (HA-USE-007) | RCR-044 deliberate, recorded override | 1
Forgot-to-arm / void trip (HA-USE-004); wrong profile (HA-USE-003) | RCR-045 pre-arm gate (non-defeatable) + RCR-048 profile validation | 2
Cold charge-inhibited device discarded (HA-USE-009) | RCR-046 legibility (D-E) + RCR-001 HW charge safety | 3 D-E + 1
Ignored/muted alert (HA-USE-002) | RCR-047 tuned alerting + RCR-029 escalation | 2
Gap misread as uneventful (HA-DATA-004, HA-USE-004) | RCR-021 verdict integrity + RCR-022/040 design-enforced honest gap | 1 + 3 D-E
Probe left in air / against wall (HA-USE-005) | RCR-020 probe-fault detection | 2
Estimate believed a measurement (HA-DATA-007) | RCR-041 estimate flagged (D-E) | 3 D-E
No use-related S4 hazard is carried by information alone.
---
4. Justification of information-only controls (the honest exceptions)
Five RCRs are Level-3 (information). Four are design-enforced — the device forces the honest state; the human is not relied upon to remember or interpret. One is pure information, and is justified because no higher-level control removes the hazard.
RCR | Type | Why not higher? | Justification
RCR-022 honest gap | Design-enforced | The verdict integrity is RCR-021 (Level 1); RCR-022 is its visibility half | Not a label — the device cannot render an incomplete record as a clean pass; verified by gap-exposure FI (PV-SW-002/PV-DATA-004). Accepted at Level 3 only in name.
RCR-040 incomplete-monitoring qualifier | Design-enforced | The suppression is RCR-021 (Level 1); RCR-040 is the surface it drives | The disposition surface is constructed so a gap/fault cannot appear as a clean pass; comprehension validated (VAL-USE-004).
RCR-041 estimate flagged | Design-enforced | No inherent way to make an estimate be a measurement; the honest thing is to label provenance | The device/cloud flags estimated/interpolated/last-known data by construction; comprehension validated (VAL-USE-005).
RCR-046 charge-inhibited ≠ dead | Design-enforced | The charge safety itself is RCR-001 (Level 1); RCR-046 is legibility so a serviceable device is not discarded | The inhibited state + cause is displayed by the device; comprehension validated (VAL-USE). Balances HA-NEW-001.
RCR-052 EOL disposal instructions | Pure information (+process) | No inherent/protective measure removes the end-of-life disposal hazard for a battery product | Accepted as the single justified pure-information control, with design support (battery removal, key/data erasure) so the instruction is followable. It is a disposal/environmental/security control, not a disposition-decision safety control.
Evidence-honesty note: the four design-enforced controls are counted as Level-3 for hierarchy-audit fidelity (they involve a human perceiving state), but each is verified by a technical test that the honest state is forced — not by assuming a user reads a label. Only RCR-052 relies on a human following an instruction, and only because the disposal hazard admits no higher control.
---
5. Single-fault / non-defeatability determinations (S4/S5)
S4/S5 hazard | Primary (Level 1/2) | Independent backstop | Common-cause resistance | Verified by
Cold-charge plating (HA-ENRG-001) | RCR-001 HW inhibit | FW defense-in-depth (subordinate) + RCR-046 legibility | HW interlock independent of FW; inhibit holds in any FW/comms fault; non-defeatable | PV-ENRG-001 (FI, FW forced faulted)
Cell OV/OC/OT/short (HA-ENRG-002, S5) | RCR-002 BMS | Non-resettable secondary fuse/OVP | Fuse independent of the BMS IC | PV-ENRG-002 (FI, BMS forced failed)
Crush → short (HA-ENRG-003 / HA-MECH-009, S5) | RCR-004 retention | RCR-008 containment | Different physical mechanisms | PV-MECH-004 + PV-TFC-002 + UN 38.3
Out-of-spec source (HA-ENRG-005) | RCR-003 input protection | RCR-002 BMS | Two independent layers | PV-ELEC-002 + PV-ENRG-002
False accept over a gap (HA-SW-002, info-analogue) | RCR-021 verdict integrity | Device↔cloud diverse recompute + RCR-022 honest gap | Diverse second computation; non-defeatable by FW/config/user | PV-SW-002 (FI) + VAL-SW-002
The information-integrity S4/S5 hazards are not "single-fault physical" in the IEC 62368-1 sense, but the register applies an analogous non-defeatability rule to RCR-021/022/045 (verdict integrity, honest gap, pre-arm gate) — none may be bypassed by FW fault, configuration, or user action, and device+cloud verdicts cross-check.
6. Closure
• The hierarchy is applied in order for every top hazard: inherent (16) where the hazard can be designed out; protective (28) where it must be bounded; information only where forced (5, four design-enforced); process (9) for latent/production/transport/governance risks.
• Inherent/protective-dominant (44/58, 76%) — the required posture for a severity-first information device with a vulnerable last-mile user.
• Every warnings-challenge use hazard has a Level-1/2 or design-enforced carrier (§3); the one pure-information control (RCR-052) is justified because no higher control removes the disposal hazard (§4).
• Every S4/S5 physical hazard has an independent single-fault backstop and the information-integrity choke points are non-defeatable (§5).
• Open at baseline: the D3 headline controls (RCR-001/006/017) and RF-in-container (RCR-031) have correct Level-1 placement; only their effectiveness is Evidence-Pending (PRD-VER-004/005) — a verification gap, not a hierarchy gap.
Review history | 2026-07-19 v1.0 initial (Stage 12, ORCH-018 / EM-ENG-012).
PRD-SAFE Reconciliation Note — Smart Cold Chain Monitor (SCM)
Stage 12 (ORCH-018 / EM-ENG-012) · Owner: Chief Systems (EM-ENG-001) + Quality/Regulatory (EM-ENG-006). The preliminary safety requirements PRD-SAFE-001…012 (Stage 05, 05-product-requirements/PRD-1300-safety.md) were written from the Stage-04 safety needs and the concept, before the Hazard Analysis Library existed, and were explicitly flagged "PRELIMINARY — to be reconciled with HAL (Stage 10)." This note performs that reconciliation: it maps each PRD-SAFE requirement to the consolidated RCR(s) it becomes or updates (RCR-1000-register.md), so the RCRs feed back to Product Requirements as the authoritative safety requirements. It also records the reverse map (each RCR → its parent PRD-SAFE) and flags where the RCR set is broader than the preliminary PRD-SAFE (new authoritative requirements that PRD-SAFE should adopt). Severity-first per CLAUDE.md rule 3.
---
1. Reconciliation principle
Direction of authority (post-HAL) | safety/risk-control requirements originate in the HAL (ORCH-017) and are formalised as RCRs at Stage 12. The preliminary PRD-SAFE requirements are therefore re-baselined to point at the RCRs — the RCR becomes the authoritative, testable statement; PRD-SAFE-NNN is updated to reference it and to inherit its acceptance criteria and verification target.
• One intent → one-or-more RCRs. A PRD-SAFE requirement typically expresses a safety intent that the RCR set implements at finer granularity (e.g., PRD-SAFE-001 "do not mislead disposition" is realised by the whole detection/verdict/record/decision RCR cluster). Each mapping below names the primary RCR(s) and any supporting ones.
• No orphans. Every PRD-SAFE-001…012 maps to ≥1 RCR, and every safety-relevant RCR maps back to a PRD-SAFE parent (§4). Where the RCR set introduces authoritative safety content beyond the 12 preliminary PRD-SAFE requirements, it is flagged as a PRD-SAFE gap to adopt (§3) — the feedback that re-baselines Product Requirements.
---
2. Forward map — each PRD-SAFE → the RCR(s) it becomes/updates
PRD-SAFE | Preliminary intent | Becomes / updated by RCR(s) | Status of PRD-SAFE after reconciliation
PRD-SAFE-001 — Do not mislead disposition (severity anchor) | Neither administer a compromised dose nor discard a potent one | RCR-019, 020, 021, 022, 040, 043, 044 (detection→verdict→record→decision) + aggregate RCR-057; data authenticity RCR-034; EMC integrity RCR-014; binding RCR-033; accuracy RCR-053 | Re-baselined — now the umbrella intent realised by the disposition-integrity RCR cluster; acceptance inherited from RCR-019/021 + VAL-USE-001/004
PRD-SAFE-002 — Reliable two-sided (freeze) detection | Freeze reliability = heat | RCR-019 (primary); profile validation RCR-048 | Re-baselined → RCR-019. Adopt RCR-019's freeze=heat acceptance + PV-SW-001/VAL-SW-001
PRD-SAFE-003 — No false reassurance (honest gaps/outages) | Gaps visible, never appear healthy | RCR-022 (primary, design-enforced) + RCR-021, 027, 040, 041, 046; offline-first RCR-028; export RCR-042 | Re-baselined → RCR-022/021. Broadened to the whole honest-state set
PRD-SAFE-004 — Alarm balance (no fatigue-induced miss) | False-alarm/missed rates avoid fatigue | RCR-047 (primary) + RCR-029 (escalation); non-suppressible RCR-035 | Re-baselined → RCR-047. Adopt VAL-USE-002 targets (ALARP both ways)
PRD-SAFE-005 — Use-error protection for low-training users | Foreseeable use errors do not cause wrong disposition | RCR-043, 044, 045, 046, 048 (+ small-parts RCR-010, sharp-edge RCR-009) | Re-baselined → the HF RCR set. Adopt summative usability (VAL-USE-001/004/007)
PRD-SAFE-006 — Prevent unmonitored trip start (mis-arm) | Void/un-armed unmistakable | RCR-045 (primary) + RCR-026, 048 | Re-baselined → RCR-045. Adopt PV-SW-003/VAL-USE-004
PRD-SAFE-007 — No charge in unsafe (cold) condition | Inhibit <0 °C; indicate inhibited | RCR-001 (primary) + RCR-005 (self-heat bound) + RCR-046 (legibility) | Re-baselined → RCR-001. Adopt HW-independent, non-defeatable acceptance + PV-ENRG-001 (FI). Carried Open (D3)
PRD-SAFE-008 — No fire/thermal hazard in any condition | No fire/burn/runaway in normal/misuse/single-fault | RCR-002 (primary) + RCR-003, 004, 008, 013, 054 | Re-baselined → RCR-002/008. Adopt single-fault backstop acceptance + PV-ENRG-002/PV-TFC-002 + UN 38.3
PRD-SAFE-009 — Safe/compliant Li-battery transport | UN 38.3 / IATA DGR + docs | RCR-007 (primary) + RCR-004 | Re-baselined → RCR-007. No air-transport claim until UN 38.3 passes
PRD-SAFE-010 — Fail-safe and recover without data loss | Fail safe after power/comms/FW/reset/OTA | RCR-023, 024, 025, 030, 037 (+ record spine RCR-042/057; service RCR-050; secure boot RCR-038) | Re-baselined → RCR-023/024/037. Adopt fault-injection acceptance (PV-SW-009/013, PV-SW-010)
PRD-SAFE-011 — Residual-risk acceptability | Overall residual acceptable per ISO 14971 | RCR-058 (primary) + RCR-039 (fleet resilience) | Re-baselined → RCR-058. Closed in HAL-2900 / Risk-Management-Summary §7–8
PRD-SAFE-012 — Dead-in-cold as safety-relevant informational failure | 90-day cold endurance so record not lost | RCR-006 (primary) + RCR-017, 024, 027 | Re-baselined → RCR-006/017. Adopt PRD-VER-004 endurance acceptance. Carried Open (D3)
---
3. PRD-SAFE gaps to adopt — authoritative RCR content beyond the 12 preliminary requirements
The RCR set (based on the full HAL) is broader than the 12 preliminary PRD-SAFE requirements. The following authoritative safety content has no dedicated PRD-SAFE parent and should be adopted into the Product Requirements Library as new/expanded PRD-SAFE (or allocated to the named sibling PRD module). This is the concrete feedback that re-baselines Product Requirements:
Authoritative RCR content (no dedicated PRD-SAFE today) | RCR(s) | Recommended Product-Requirements action
Connectivity/alert-delivery safety (local commit, alert-ahead/ack/escalate, store-and-forward, RF-in-container) | RCR-028, 029, 030, 031, 032, 033 | New PRD-SAFE-013 "Excursion-alert delivery safety" (or expand PRD-SAFE-003/004); ties PRD-COMM
Cybersecurity-induced safety (authenticity, non-suppressible alarm, config-change control, secure OTA, secure boot, cloud resilience) | RCR-034, 035, 036, 037, 038, 039 | New PRD-SAFE-014 "Security-induced safety" cross-referencing PRD-SEC; require IV verification
Data authenticity / audit-trail (Part-11-style export integrity, provenance) | RCR-041, 042 | New PRD-SAFE-015 "Data authenticity & audit trail"; ties PRD-CLOUD / GDP-ALCOA+
Calibration integrity + fleet-scale calibration/service common-cause | RCR-026, 049, 053 | Expand PRD-SAFE-006 or new PRD-SAFE-016 "Measurement-accuracy & calibration integrity"
Manufacturing-integrity safety (correct chemistry/part, unique identity, regional config, AOI/EOL) | RCR-016, 054, 055, 056 | New PRD-SAFE-017 "Build-integrity safety" ties PRD-MFG
Record-integrity mission spine (aggregate) | RCR-057 | Expand PRD-SAFE-001 to name the record-integrity chain explicitly
EMC integrity of the decision path | RCR-014 | Cross-reference into PRD-SAFE-001; ties PRD-ELEC (EMC)
Environmental operate-and-retain range | RCR-017 | Cross-reference into PRD-SAFE-012; ties PRD-REL/PERF
Mechanical accessibility / small-parts / retention / fasteners | RCR-009, 010, 011, 012 | Cross-reference into PRD-SAFE-005/008/010; ties PRD-MECH
Recommendation: the Product Requirements Library adopts PRD-SAFE-013…017 (five new authoritative safety requirements) and expands PRD-SAFE-001/003/005/006/012 per the table, each pointing at its RCR(s). This closes the preliminary→authoritative gap and makes PRD-SAFE the reconciled, HAL-backed safety-requirement set.
---
4. Reverse map — each RCR → its parent PRD-SAFE (spot-check for orphans)
Every safety-relevant RCR traces to a PRD-SAFE parent (existing or the recommended new one in §3):
PRD-SAFE parent | RCRs mapped to it (authoritative children)
PRD-SAFE-001 (don't mislead) | 014, 019, 020, 021, 033, 034, 043, 044, 053, 055, 057
PRD-SAFE-002 (two-sided detection) | 019, 048
PRD-SAFE-003 (no false reassurance) | 021, 022, 025, 027, 028, 040, 041, 042, 046
PRD-SAFE-004 (alarm balance) | 029, 035, 047
PRD-SAFE-005 (use-error) | 009, 010, 043, 044, 045, 046, 048
PRD-SAFE-006 (prevent mis-arm) | 026, 045, 048, 049
PRD-SAFE-007 (no cold charge) | 001, 005, 046
PRD-SAFE-008 (no fire/thermal) | 002, 003, 004, 008, 013, 054
PRD-SAFE-009 (safe Li transport) | 004, 007
PRD-SAFE-010 (fail-safe/recover) | 011, 012, 015, 016, 018, 023, 024, 030, 037, 038, 050, 051, 056
PRD-SAFE-011 (residual acceptable) | 039, 058
PRD-SAFE-012 (dead-in-cold) | 006, 017, 024, 027
PRD-SAFE-013…017 (recommended new, §3) | 028–033 (delivery); 034–039 (security); 041/042 (data authenticity); 026/049/053 (calibration); 016/054/055/056 (build)
No RCR is an orphan; no PRD-SAFE is unmapped. RCR-052 (EOL disposal) maps to PRD-SVC/REG (environmental/security), outside the disposition-safety core, and is noted as a disposal/security control rather than a disposition-safety one.
5. Closure
• All 12 preliminary PRD-SAFE requirements are reconciled to authoritative RCRs (§2) and re-baselined to point at them.
• Five new authoritative safety requirements (PRD-SAFE-013…017) and five expansions are recommended to Product Requirements to adopt the RCR content that the preliminary set did not cover (§3) — the feedback loop that makes PRD-SAFE the HAL-backed authoritative safety set.
• Acceptance criteria and verification targets flow from the RCRs (PV-/VAL-, Evidence-Pending until Stages 13–14); the two D3/RF-in-container Open items (PRD-SAFE-007/012 → RCR-001/006/017/031) are carried as release-gating.
Feeds | Product Requirements (re-baselined PRD-SAFE), Stage 13 (Verification), Stage 14 (Validation), and ORCH-018 / EM-ENG-012.
Review history | 2026-07-19 v1.0 initial (Stage 12).
RCR-1000 — Consolidated Risk-Control Requirements Register — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 12 — Risk-Control Requirements · Orchestrator: ORCH-018 / EM-ENG-012 with Chief Systems (EM-ENG-001) and Quality/Regulatory (EM-ENG-006), System-Safety, and the discipline owners. This is the authoritative, consolidated, deduplicated Risk-Control Requirement register for the SCM: the formal, testable "shall" set derived from the Hazard Analysis Library (HAL, Stage 10) and the Failure/Risk Analyses (ERFKL, Stage 11), reconciled with the preliminary PRD-SAFE (Stage 05). Every accepted control is a single, non-duplicated RCR-### with a control strategy (ISO 14971 hierarchy order), source hazard(s), source failure mode(s), parent requirement, implementation reference, verification/validation target, independence note, new-risk note, acceptance criteria, owner, and status. Severity-first per CLAUDE.md rule 3. Evidence honesty: every acceptance criterion is testable; unevidence-backed quantities are marked Evidence-Pending with the named closing test; nothing is invented.
---
1. Purpose and how this register was consolidated
The SCM's controls arrived from three streams that had to be reconciled into one authoritative set:
1. HAL-derived RCRs (RCR-001…058, HAL-2400) — the master control list produced when each accepted hazard control became a testable requirement, classified by hierarchy in HAL-2300.
2. FMEA/FTA recommended actions (ERFKL Stage 11, FMEA-to-Requirements-Hazard-Matrix) — the reliability/safety failure analyses each name an RCR as the control that drives their Action Priority to target.
3. Preliminary PRD-SAFE-001…012 (Stage 05) — the safety requirements written from user/market needs before the hazard analysis existed, explicitly flagged PRELIMINARY and "to be reconciled with HAL."
Consolidation rule applied (dedup): the HAL RCR register is the spine. Where an FMEA recommended action and a HAL control are the same measure, they are one RCR — the FMEA record is carried as a Source Failure Mode on that RCR, not as a new requirement (see the Dedup line in each block and §5). Where a PRD-SAFE requirement expresses a safety intent, that intent becomes / is refined by one or more RCRs and PRD-SAFE is re-baselined to point at them (see PRD-SAFE-Reconciliation.md). The result is 58 RCRs, no duplicates — every HAL control, every safety-significant FMEA action, and every PRD-SAFE intent is carried by exactly one RCR (or a named aggregate).
No new RCRs were created at Stage 12. The FMEA↔HAL reconciliation (Stage 11) and the control-hierarchy pass (HAL-2300) had already converged on the 58-control set; Stage 12's job is to formalise, deduplicate, populate, and allocate — not to invent additional controls. Two standing findings from HAL-2700 (external-antenna seal/config; key-lifecycle) are carried as acceptance-criteria conditions inside existing RCRs (RCR-031, RCR-038/052), not as new requirements.
2. Field definitions (per RCR block)
• Requirement (shall) — the formal, testable risk-control requirement statement.
• Control Strategy — ISO 14971 hierarchy order: inherent-design (ISD) → protective (PROT) → information for safety (INFO, design-enforced = D-E) → process (PROC). Primary level governs; secondary noted. This refines the HAL-2400 Class and HAL-2300 Level.
• Source Hazard(s) — the HAL hazard record(s) (HA-…) this RCR controls (Hazard-Traceability-Matrix).
• Source Failure Mode(s) — the ERFKL failure record(s) (FA-…) whose recommended action this RCR is (FMEA-to-Requirements-Hazard-Matrix); "= HAL control" where the FMEA action and HAL control are the same measure.
• Parent Requirement — the Product (PRD-), System (SYS-), Subsystem/HW (HWR-), or Software (SRS-) requirement this RCR becomes or updates.
• Implementation (DO-) — the controlled design output that implements it (DOL-1200 DO-SAFE + discipline outputs).
• Verification (PV-) — planned Stage-13 product verification; fault-injection (FI) marked for safety-critical; independent adversarial (IV) for security.
• Validation (VAL-) — planned Stage-14 validation (use/clinical-workflow effectiveness), where applicable.
• Independence / common-cause — the single-fault-tolerance / non-defeatability / common-cause-resistance note.
• New-risk (HAL-2700) — the new risk this control introduces or worsens, and how it is balanced.
• Acceptance Criteria — the pass/fail condition the verification must demonstrate (testable; Evidence-Pending where a value is not yet evidence-backed, with the closing test named).
• Owner · Status · Notes — allocation owner; lifecycle status; dedup/reconciliation notes.
Status legend: Baselined = consolidated and allocated, implementation + verification target defined, awaiting Stage-13/14 evidence (the default for this pre-V&V register); Open = a load-bearing effectiveness value is Evidence-Pending and gates release (D3 battery, RF-in-container). All PV-/VAL- are Evidence-Pending until Stages 13–14.
3. Register organization (dual axis)
The register is presented by subsystem (the axis engineers navigate and the axis DOL-1200 allocates on), with the control strategy stated in every block and a control-strategy view in §4 (the hierarchy distribution and the inherent-over-warning posture). This honors the two required organizing axes — by control strategy (§4 + the Control Strategy field) and by subsystem (§6 sections). The detailed hierarchy justification (inherent preferred over warnings; information-only controls justified) is in Control-Hierarchy-Analysis.md.
4. Control-strategy view (hierarchy distribution)
Control strategy (primary) | RCRs | Count | Posture
Inherent safety by design (ISD, Level 1) | 001, 004, 006, 008, 009, 010, 011, 012, 013, 017, 019, 021, 028, 031, 043, 044 | 16 | Strongest; preferred — the hazard is designed out or the safe outcome is the default no fault/config/user can undo.
Protective measure (PROT, Level 2) | 002, 003, 005, 014, 015, 018, 020, 023, 024, 025, 026, 027, 029, 030, 032, 033, 034, 035, 036, 037, 038, 039, 042, 045, 047, 048, 049, 051 | 28 | Detect-and-act backstops, interlocks, limiters, safe-states, escalating alarms.
Information for safety (INFO, Level 3; all design-enforced except RCR-052) | 022, 040, 041, 046, 052 | 5 | The device forces the honest state; not a label/training the user must remember (except RCR-052 EOL).
Process / operational (PROC, Level 4) | 007, 016, 050, 053, 054, 055, 056, 057, 058 | 9 | Manufacturing, calibration, service, transport, disposal, governance controls.
Portfolio posture: inherent/protective-dominant — 44 of 58 controls at Level 1/2 (76%). This is the correct severity-first posture for a low-training, vulnerable last-mile user on an information device whose dominant harm is a wrong disposition decision. No S4/S5 safety risk is carried by information alone. The warnings-challenge (every use-related risk that would default to a label is instead carried by an inherent/protective control) is discharged in Control-Hierarchy-Analysis.md §3.
Release-critical set (highest-severity, single-fault / non-defeatable): RCR-001, 002, 019, 021, 022, 028, 029 — plus the two carried-Open D3 headline items RCR-001/RCR-005/RCR-006/RCR-017 (cold battery) and RCR-031 (RF-in-container), which have correct Level-1 placement but Evidence-Pending effectiveness (see §7).
---
5. Deduplication and reconciliation map (FMEA action ≡ HAL control)
The following FMEA recommended actions were found to be the same measure as a HAL control and are therefore carried as Source Failure Modes on the named RCR — not as separate requirements (this is the core dedup outcome):
RCR (single control) | FMEA recommended-action records folded in (= same measure) | HAL control it equals
RCR-001 | FA-HWFMEA-002/006, FA-SWFMEA-012 (charge-inhibit + FW invariant) | HW charge-inhibit <0 °C
RCR-002 | FA-HWFMEA-007/008, FA-HWFMEA-003, FA-DFMEA-009, FA-SYSFMEA-016 | BMS protection + secondary fuse
RCR-006 | FA-HWFMEA-001/012, FA-SWFMEA-021 (endurance + governor + gauge) | Energy governor + reserve
RCR-019 | FA-SWFMEA-001/002, FA-DFMEA-010, FA-SYSFMEA-001 | Two-sided detection
RCR-020 | FA-SWFMEA-007, FA-HWFMEA-017, FA-DFMEA-003/011 | Probe-fault/plausibility
RCR-021 | FA-SWFMEA-008/010/019, FA-SYSFMEA-005 | Verdict integrity (suppress false accept)
RCR-024 | FA-SWFMEA-018/019, FA-HWFMEA-033 | Power-fail-safe record
RCR-029 | FA-SWFMEA-023/024, FA-IFMEA-011, FA-SYSFMEA-002 | Alert-ahead + escalation
RCR-031 | FA-HWFMEA-029/030/032, FA-IFMEA-005 | RF-in-container / antenna
RCR-037 | FA-SWFMEA-026, FA-IFMEA-015 | Signed, rollback-safe OTA
RCR-045 | FA-SWFMEA-004/011, FA-USE-001 | Pre-arm readiness gate
RCR-053/054 | FA-HWFMEA-004, FA-SYSFMEA-009 | EOL calibration / correct-part
No FMEA safety-significant recommended action was left without an RCR, and no RCR was duplicated by a parallel FMEA requirement. Full per-record mapping is in each block's Source Failure Mode(s) and Dedup lines. PRD-SAFE reconciliation (each RCR → the PRD-SAFE it becomes/updates) is in PRD-SAFE-Reconciliation.md.
---
6. RCR register (RCR-001 … RCR-058), by subsystem
6.1 Battery / charge / thermal — physical-energy harm (severity-first, single-fault)
RCR-001 — Charge-inhibit below 0 °C by MCU-independent hardware interlock
Requirement (shall) | Charging shall be prevented below 0 °C (or the qualified cell datasheet lower charge limit, whichever is warmer) by an MCU-independent hardware interlock acting on cell-coupled temperature sensing; the inhibit shall persist in any firmware, configuration, or communications fault; firmware charge-inhibit shall be defense-in-depth only and shall not be able to defeat the hardware interlock.
Control Strategy | ISD (Level 1) primary; PROT + INFO(D-E) secondary. Non-defeatable, single-fault set.
Source Hazard(s) | HA-ENRG-001 (cold-charge plating → latent short), HA-ELEC-002, HA-USE-009; escalates to HA-TFC-002 (S5 aircraft-hold fire).
Source Failure Mode(s) | FA-HWFMEA-002 (cold-charge plating), FA-HWFMEA-006 (interlock fails), FA-SWFMEA-012 (FW charge-inhibit invariant violated) — = HAL control (same measure; folded, not duplicated).
Parent Requirement | PRD-SAFE-007 (no charge in unsafe/cold condition); SYS-091/094/095; PRD-ELEC-005/006.
Implementation (DO-) | DO-SAFE-001 → DO-EE-007 (HW interlock) + DO-FW-005 (FW defense-in-depth).
Verification (PV-) | PV-ENRG-001 (cold-charge fault-injection sweep, FI) + UN 38.3. Fault-injection mandatory (safety-critical): force FW fault/comms loss and confirm inhibit holds.
Validation (VAL-) | — (physical control; covered by verification + UN 38.3).
Independence / common-cause | HW interlock is independent of the MCU/FW; inhibit persists in any FW/comms fault and cannot be defeated by FW, config, or user (non-defeatable, HAL-2300 §8). Diversity of the NTC/interlock sensing path is a CCF item (FA-CCF-005).
New-risk (HAL-2700) | HA-NEW-001 — a device left cold cannot charge → under-charged/dead-on-arrival → loss-of-monitoring or discarded-as-dead. Balanced by RCR-045 (pre-arm SoC gate), RCR-006 (reserve), RCR-046 (charge-inhibited ≠ dead legibility), optional self-heat (RCR-005) / depot warm-before-charge.
Acceptance Criteria | Under a temperature sweep, charge current = 0 A at all cell temperatures below the qualified limit, including with FW halted/faulted and comms lost; inhibited state + cause reported on HMI/LED. The exact lower charge limit (0 °C vs the low-temp chemistry datasheet value) is Evidence-Pending on the final cell datasheet (closes with the battery selection + PV-ENRG-001).
Owner | Battery/Power (with Safety) · Status: Baselined (effectiveness on chemistry Open, O-2) · Notes: Primary D3 charge-safety control; one of the two D3 headline items validate-by-test.
RCR-002 — Autonomous cell protection + non-resettable secondary backstop
Requirement (shall) | The device shall provide autonomous over-voltage / under-voltage / over-current / over-temperature / short-circuit protection with a non-resettable secondary backstop (fuse and/or OVP) that is independent of the primary protection IC, such that the thermal hazard is terminated on a single fault (primary IC failure shall not reach the cell).
Control Strategy | PROT (Level 2) primary; ISD secondary. Single-fault set (S5).
Source Hazard(s) | HA-ENRG-002 (cell OV/OC/OT/short → runaway), HA-TFC-002 (S5), HA-TFC-003 (offgas), HA-ELEC-006.
Source Failure Mode(s) | FA-HWFMEA-007 (BMS/AFE fails to trip), FA-HWFMEA-008 (secondary fuse fails to clear), FA-HWFMEA-003 (cell internal short), FA-DFMEA-009 (no containment), FA-SYSFMEA-016 (system thermal event) — = HAL control.
Parent Requirement | PRD-SAFE-008 (no fire/thermal hazard, single-fault); SYS-096; PRD-ELEC-007/014.
Implementation (DO-) | DO-SAFE-002 → DO-EE-005 (BMS/AFE) + DO-EE-010 (secondary fuse/OVP).
Verification (PV-) | PV-ENRG-002 (protection fault-injection, DO-TST-009, FI) — force the primary IC failed and confirm the secondary backstop terminates the hazard; + UN 38.3.
Validation (VAL-) | —
Independence / common-cause | Secondary fuse/OVP is independent of the BMS IC (different device, different failure mechanism) — verified with the BMS forced-failed (HAL-2300 §8). Correct-rated/oriented parts enforced by RCR-054 (CCF).
New-risk (HAL-2700) | none material (the backstop is the balancing control); ties HA-NEW-010 (containment mass/material trade → RCR-008).
Acceptance Criteria | With the primary protection IC forced non-responsive, an induced OV/OC/OT/short is cleared by the secondary backstop before the cell reaches its thermal-runaway onset; no external flame propagation (with RCR-008). Runaway-onset temperature and clearing thresholds are per the qualified cell datasheet — Evidence-Pending until cell selection + PV-ENRG-002 / UN 38.3.
Owner | Battery/Power + Thermal/Battery-Safety · Status: Baselined · Notes: Release-critical; single-fault backstop is non-negotiable for the S5 path.
RCR-003 — Out-of-spec charge-source rejection / limiting
Requirement (shall) | The device shall reject or safely limit an out-of-spec charge source (reverse insertion, over-voltage, surge, short); the USB-PD requested PDO shall be capped to the qualified input envelope.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-ENRG-005 (out-of-spec source), HA-ELEC-002.
Source Failure Mode(s) | FA-HWFMEA-009 (out-of-spec charge source) — = HAL control.
Parent Requirement | PRD-SAFE-008; SYS-102; PRD-ELEC (USB-C sink).
Implementation (DO-) | DO-SAFE-003 → DO-EE-009 (USB-C sink) + DO-EE-010.
Verification (PV-) | PV-ELEC-002 (reverse / OVP / surge / short test).
Validation (VAL-) | —
Independence / common-cause | Input front-end and the BMS (RCR-002) are two independent layers — both must fail to reach the cell (HAL-2300 §8).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No cell/BMS damage or hazardous condition under reverse-insertion, OVP to the rated input ceiling, surge, and output short; PD PDO never exceeds the qualified envelope. PD cap value Evidence-Pending on the charge-path design (closes with DO-EE-009 + PV-ELEC-002).
Owner | Electrical/Power · Status: Baselined · Notes: —
RCR-004 — Cell retention (no crush/puncture through drop/vibration/transport)
Requirement (shall) | The cell shall be retained without crush or puncture through the specified drop/vibration/transport and reuse life, maintaining thermal and electrical clearance.
Control Strategy | ISD (Level 1) primary; PROT secondary. Single-fault set (S5 crush→short).
Source Hazard(s) | HA-ENRG-003 (crush on drop/transport), HA-MECH-009 (retention fails on shock), HA-TFC-002 (S5), HA-ENV-007 (altitude).
Source Failure Mode(s) | FA-DFMEA-008 (retention fails → crush), FA-SYSFMEA-016 — = HAL control.
Parent Requirement | PRD-SAFE-008/009; HWR-002; PRD-MECH.
Implementation (DO-) | DO-SAFE-010 → DO-MECH-012 (retention) + DO-EE-004 (pack).
Verification (PV-) | PV-MECH-004 (drop/vibe, FI-adjacent) + UN 38.3.
Validation (VAL-) | —
Independence / common-cause | Retention (mechanical) and containment (RCR-008 enclosure) are different physical mechanisms — no shared failure (HAL-2300 §8).
New-risk (HAL-2700) | ties HA-NEW-010 (enclosure mass/cold-brittleness trade) → controlled by material selection + RCR-017 range.
Acceptance Criteria | No cell deformation/puncture/short after the qualified drop/vibration profile and UN 38.3; clearance maintained. Drop height/vibration profile per IEC 60068 / UN 38.3 — profile to confirm in the test plan.
Owner | Mechanical + Battery/Power · Status: Baselined · Notes: —
RCR-005 — Self-heat (if adopted) bounded to enable-only, cannot bypass inhibit
Requirement (shall) | If a self-heat-before-charge feature is provided, it shall only enable the charge guard by warming the cell into the safe charge window; it shall not bypass the charge inhibit (RCR-001) and shall not exceed the safe upper thermal limit or a touchable-surface burn limit.
Control Strategy | PROT (Level 2), subordinate to RCR-001.
Source Hazard(s) | HA-ENRG-001, HA-TFC-001 (hot self-heat surface touched — burn).
Source Failure Mode(s) | covered under the charge-safety FTA (TOP-C) self-heat branch.
Parent Requirement | PRD-SAFE-007; D3; IEC 62368-1 touch-temperature.
Implementation (DO-) | DO-SAFE-001 (self-heat sub-function) → DO-EE-007.
Verification (PV-) | PV-TFC-001 (touch-temperature) + self-heat bound + inhibit-integrity test (FI).
Validation (VAL-) | VAL-TFC-001 (if self-heat adopted).
Independence / common-cause | Self-heat control cannot override the independent HW inhibit (RCR-001); bounded below the thermal/burn limit.
New-risk (HAL-2700) | HA-NEW-002 — self-heat is a new hot surface + energy source; balanced by the enable-only bound, touch-temp limit, and RCR-046 legibility. Conditional on self-heat adoption (CMP-013); N-A if not adopted.
Acceptance Criteria | Self-heat never enables charging below the safe window; element and any touchable surface stay below the IEC 62368-1 burn threshold; on any self-heat fault the charge inhibit still holds. Touch-temp limit value per IEC 62368-1; self-heat feasibility Evidence-Pending (CMP-013).
Owner | Thermal/Battery-Safety · Status: Baselined (conditional) · Notes: Archetype of one control (self-heat) needing another (touch-temp bound).
RCR-006 — Energy governor + reserve: no silent loss of monitoring
Requirement (shall) | The device shall implement an energy governor and reserve that prevent silent loss of monitoring: the reserve preserves record-commit and a low-battery alert; low-battery and dead-in-cold states are indicated honestly; the 90-day cold endurance shall be met.
Control Strategy | ISD (Level 1) primary; INFO(D-E) secondary.
Source Hazard(s) | HA-ENRG-006 (energy exhausted → silent gap reads uneventful), HA-SW-013.
Source Failure Mode(s) | FA-HWFMEA-001 (90-day cold endurance shortfall), FA-SWFMEA-021 (governor over-spend), FA-HWFMEA-012 (fuel-gauge over-reports SoC in cold) — = HAL control.
Parent Requirement | PRD-SAFE-012 (dead-in-cold as safety-relevant failure); PRD-REL-001; SYS-086/090/097.
Implementation (DO-) | DO-SAFE-012 → DO-FW-004 (governor) + DO-EE-008 (fuel gauge).
Verification (PV-) | PRD-VER-004 (90-day cold endurance) + fuel-gauge cold-accuracy test.
Validation (VAL-) | — (endurance is a performance verification; field endurance/gap monitoring per ISO 14971 §9).
Independence / common-cause | The governor + reserve is independent of the upload/comms path; the low-battery reserve is preserved before any discretionary spend. Shared cold-power is a CCF driver (FA-CCF-004).
New-risk (HAL-2700) | balances HA-NEW-001 (dead-on-arrival) with the reserve/honest indication; ties HA-NEW-011 (radio/time re-discipline energy is budgeted within the governor).
Acceptance Criteria | Over a 90-day −30 °C profile the device continues sampling+logging, commits the record, and raises a low-battery alert before exhaustion; the fuel gauge does not over-report SoC in cold beyond a bounded error. The 90-day margin, the cold SoC error bound, and the −30 °C duty profile are Evidence-Pending → PRD-VER-004 (O-2, the single largest program assumption).
Owner | Battery/Power + Firmware · Status: Open (D3 endurance, O-2) · Notes: D3 headline item; validate-by-test.
RCR-007 — UN 38.3 / IATA-DGR transport qualification
Requirement (shall) | The device and its battery shall be UN 38.3-tested and IATA-DGR-compliant, with transport documentation provided.
Control Strategy | PROC (Level 4).
Source Hazard(s) | HA-ENRG-003, HA-TFC-002 (S5), HA-MPE-005 (e-waste).
Source Failure Mode(s) | FA-SYSFMEA-016 (thermal event / transport).
Parent Requirement | PRD-SAFE-009 (safe/compliant Li transport); PRD-REG-006; PRD-ELEC-015.
Implementation (DO-) | DO-EE-015, DO-REG-006.
Verification (PV-) | UN 38.3 test suite (T.1–T.8) + IATA-DGR documentation review.
Validation (VAL-) | —
Independence / common-cause | Transport qualification is the process layer behind the physical controls (RCR-002/004/008).
New-risk (HAL-2700) | none.
Acceptance Criteria | Full UN 38.3 pass with report; IATA-DGR packing/marking/documentation in place. Test summary Evidence-Pending. No air-transport claim is made until this passes (release condition §7.3 of the Risk-Management Summary).
Owner | Battery/Power + Regulatory (EM-ENG-006) · Status: Baselined · Notes: —
RCR-008 — Fire-enclosure containment of a single-cell event
Requirement (shall) | The enclosure and internal materials shall contain a single-cell thermal event without external flame propagation and with low-toxicity offgassing, consistent with the IEC 62368-1 fire-enclosure model.
Control Strategy | ISD (Level 1). Single-fault backstop for the S5 crush/thermal path.
Source Hazard(s) | HA-TFC-002 (S5 aircraft-hold fire), HA-TFC-003 (toxic offgas in enclosed space).
Source Failure Mode(s) | FA-DFMEA-009 (no cell thermal containment), FA-SYSFMEA-016 — = HAL control.
Parent Requirement | PRD-SAFE-008; PRD-MECH; IEC 62368-1.
Implementation (DO-) | DO-MECH-005 (enclosure) + DO-PCB-010 (barrier).
Verification (PV-) | PV-TFC-002 (containment/fire-enclosure) + PV-TFC-003 (offgas characterisation) + UN 38.3.
Validation (VAL-) | —
Independence / common-cause | Containment is a different mechanism from retention (RCR-004) and protection (RCR-002) — the last-line backstop if both are defeated.
New-risk (HAL-2700) | HA-NEW-010 — containment adds mass/cost, constrains materials, and a vent path could conflict with the ingress seal (RCR-011); balanced by integrated vent-vs-seal design and cold-tough material choice.
Acceptance Criteria | An induced single-cell runaway is contained with no external flame; offgas toxicity/quantity characterised and consistent with enclosed-space (aircraft-hold) exposure limits. Offgas limits/quantities Evidence-Pending → PV-TFC-003. S5 residual stays A-High by rule with mandatory thermal-event monitoring.
Owner | Mechanical + Thermal/Battery-Safety · Status: Baselined (S5 residual A-High permanent) · Notes: —
6.2 Mechanical
RCR-009 — No accessible sharp/pinch/crush point
Requirement (shall) | The device shall present no accessible hazardous sharp edge, pinch, or crush point in normal handling or service.
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-MECH-001/002.
Source Failure Mode(s) | covered by DFMEA enclosure records (FAL-1000).
Parent Requirement | PRD-MECH; PRD-SAFE-005 (use-safety); IEC 62368-1.
Implementation (DO-) | DO-MECH-002/003.
Verification (PV-) | Edge/pinch inspection (IEC 62368-1 accessibility).
Validation (VAL-) | —
Independence / common-cause | Inherent geometry; no active control.
New-risk (HAL-2700) | ties HA-NEW-010 (handling ergonomics of the containment enclosure).
Acceptance Criteria | No accessible edge/pinch exceeds IEC 62368-1 accessibility limits in normal handling and serviceable-part access.
Owner | Mechanical · Status: Baselined · Notes: —
RCR-010 — Small-parts retention / choking prevention
Requirement (shall) | Small parts (probe tip, connector caps, any lanyard) shall be retained, or sized/warned, to prevent choking or strangulation of a child bystander.
Control Strategy | ISD (Level 1) primary; INFO secondary (warning only where retention infeasible).
Source Hazard(s) | HA-MECH-003 (detachable small part near a child).
Source Failure Mode(s) | DFMEA detachment records (FAL-1000).
Parent Requirement | PRD-MECH; PRD-SAFE-005; 16 CFR 1501.
Implementation (DO-) | DO-MECH-006 (retention) + DO-LBL (warning).
Verification (PV-) | PV-MECH-003 (small-parts / retention test).
Validation (VAL-) | —
Independence / common-cause | Retention preferred; warning is the residual only where retention is not feasible (warnings-challenge, HAL-2300 §7).
New-risk (HAL-2700) | none.
Acceptance Criteria | No detachable small part fits the small-parts cylinder without deliberate tooling; retention holds through drop/handling.
Owner | Mechanical + Human-Factors · Status: Baselined · Notes: —
RCR-011 — Enclosure/seal maintains ingress rating + cell containment through life
Requirement (shall) | The enclosure and seals shall maintain the ingress rating and cell containment through the specified drop/vibration/transport and reuse life.
Control Strategy | ISD (Level 1) primary; PROT secondary.
Source Hazard(s) | HA-MECH-004 (crack→ingress→data loss), HA-MECH-009, HA-ENV-003.
Source Failure Mode(s) | FA-DFMEA-001 (seal loses IP over cold cycling).
Parent Requirement | PRD-MECH; PRD-SAFE-010; HWR-038/040; IEC 60529.
Implementation (DO-) | DO-SAFE-010/011 → DO-MECH-005/012.
Verification (PV-) | PV-MECH (drop/vibe) + IP test (over reuse cycling).
Validation (VAL-) | —
Independence / common-cause | Ingress protection + containment are the mechanical barrier for record electronics (with RCR-015 coat).
New-risk (HAL-2700) | HA-NEW-009 — an external-antenna connector (if RCR-031 fallback adopted) is a new ingress path held to the same IP standard.
Acceptance Criteria | IP rating retained after the drop/vibration profile and N reuse thermal cycles; no cell containment loss. IP class and reuse-cycle count Evidence-Pending → PV-MECH / test plan.
Owner | Mechanical · Status: Baselined · Notes: —
RCR-012 — Fasteners/joints resist vibration/thermal loosening (no detachment)
Requirement (shall) | Fasteners and joints shall resist loosening under vibration and thermal cycling over the reuse life (no detachment).
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-MECH-006.
Source Failure Mode(s) | DFMEA fastener records (FAL-1000).
Parent Requirement | PRD-MECH; PRD-SAFE-010.
Implementation (DO-) | DO-MECH-004.
Verification (PV-) | Vibration + thermal-cycle inspection.
Validation (VAL-) | —
Independence / common-cause | Inherent joint robustness.
New-risk (HAL-2700) | none.
Acceptance Criteria | No fastener back-out or joint detachment after the qualified vibration + thermal-cycle profile over reuse life.
Owner | Mechanical · Status: Baselined · Notes: —
6.3 Electrical / EMC
RCR-013 — SELV only; no accessible hazardous voltage in single-fault
Requirement (shall) | All accessible circuits shall be SELV; no accessible hazardous voltage shall exist in normal or single-fault conditions.
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-ELEC-001.
Source Failure Mode(s) | HWFMEA electrical-safety records.
Parent Requirement | PRD-SAFE-008; PRD-ELEC; IEC 62368-1.
Implementation (DO-) | DO-EE (rails) + DO-PCB.
Verification (PV-) | IEC 62368-1 electrical-safety test.
Validation (VAL-) | —
Independence / common-cause | Inherent — shock hazard designed out.
New-risk (HAL-2700) | none.
Acceptance Criteria | No accessible point exceeds SELV limits in normal or single-fault per IEC 62368-1.
Owner | Electrical · Status: Baselined · Notes: —
RCR-014 — EMC immunity (no false reading/verdict/missed excursion) + emissions
Requirement (shall) | The device shall meet EMC immunity (ESD / RF / surge / EFT per IEC 61000) such that no interference produces a false reading, false verdict, or missed excursion, and shall meet emissions limits so it does not interfere with aircraft or medical equipment.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-ELEC-003/004 (EMC upset → false/missed excursion).
Source Failure Mode(s) | FA-HWFMEA-024 (EMC/ESD upset → false/missed) — = HAL control.
Parent Requirement | PRD-SAFE-001 (don't mislead) via data integrity; PRD-ELEC (EMC); IEC 61000.
Implementation (DO-) | DO-EE (EMC) + DO-PCB-010 + DO-EE-019.
Verification (PV-) | PV-ELEC-004 (EMC immunity + emissions, FI-adjacent) — immunity tested against the detection/verdict path integrity.
Validation (VAL-) | —
Independence / common-cause | Immunity designed so RF cannot corrupt the decision path; ties record-integrity chain (RCR-057).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No false/missed excursion or false verdict under the IEC 61000 immunity levels; emissions within FCC Part 15 / CE-RED and aircraft/medical limits. Immunity levels per IEC 61000 test plan.
Owner | Electrical/EMC · Status: Baselined · Notes: —
RCR-015 — Conformal coat + seal protect record electronics across −30↔+60 °C
Requirement (shall) | A conformal coat plus seal shall protect the record-integrity electronics from condensation, leakage, and corrosion across −30↔+60 °C cycling over the reuse life.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-ELEC-006/007, HA-ENV-003, HA-TFC-006 (condensation → record-electronics fault).
Source Failure Mode(s) | FA-DFMEA-001 (seal/coat over cold cycling).
Parent Requirement | PRD-SAFE-010; PRD-MECH/ELEC; HWR-038/040; IEC 60529.
Implementation (DO-) | DO-SAFE-011 → DO-MECH-005 (seal) + DO-PCB-010 (coat).
Verification (PV-) | PV-TFC-006 (condensation cycle + IP) + PV-ENV-003 + cleanliness.
Validation (VAL-) | —
Independence / common-cause | Barrier for the record-integrity electronics; complements RCR-011 enclosure.
New-risk (HAL-2700) | none.
Acceptance Criteria | No record-integrity fault after N cold↔warm condensation cycles; insulation resistance and cleanliness within limits. Cycle count / IR limits Evidence-Pending → PV-TFC-006.
Owner | Mechanical/PCB · Status: Baselined · Notes: —
RCR-016 — AOI + EOL screening of latent defects (no latent SPF on safety nets)
Requirement (shall) | Latent solder/PCB/component defects shall be screened by AOI + EOL test coverage such that no latent single-point failure escapes on safety-critical nets.
Control Strategy | PROC (Level 4).
Source Hazard(s) | HA-ELEC-008 (latent → intermittent → silent gap), HA-MFG-008.
Source Failure Mode(s) | FA-HWFMEA-023 (latent PCBA intermittent → gap) — = HAL control.
Parent Requirement | PRD-SAFE-010; PRD-REL; DfT plan.
Implementation (DO-) | DO-TST (AOI/ICT) + DO-PCB.
Verification (PV-) | PV-MFG-008 (AOI/ICT coverage review) + HASS.
Validation (VAL-) | —
Independence / common-cause | Screening of latent defects on safety-critical nets; ties record-integrity (RCR-024).
New-risk (HAL-2700) | none.
Acceptance Criteria | AOI/ICT coverage on safety-critical nets ≥ the defined coverage target; HASS applied; no latent SPF escape demonstrated by coverage analysis. Coverage target Evidence-Pending → DfT plan.
Owner | Manufacturing/Test · Status: Baselined · Notes: PFMEA (Stage 15) refines the process controls.
6.4 Environmental
RCR-017 — Operate + retain record across −30…+60 °C incl. 90-day cold endurance
Requirement (shall) | The device shall operate and retain the record across −30 °C…+60 °C, including the 90-day cold endurance and reserve.
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-ENV-001/002, HA-ENRG-006.
Source Failure Mode(s) | FA-HWFMEA-001 (endurance), FA-DFMEA-010 (internal-sensor thermal behaviour).
Parent Requirement | PRD-SAFE-012; PRD-REL-001; PRD-PERF; SYS environmental.
Implementation (DO-) | DO-SAFE-012 + DO-EE + DO-FW.
Verification (PV-) | IEC 60068 environmental + PRD-VER-004 (90-day cold endurance).
Validation (VAL-) | —
Independence / common-cause | Inherent environmental capability; the endurance sub-claim is shared with RCR-006.
New-risk (HAL-2700) | ties HA-NEW-010 (cold-tough material).
Acceptance Criteria | Full function and record retention across the range per IEC 60068; 90-day cold endurance met. Endurance margin Evidence-Pending → PRD-VER-004 (O-2).
Owner | Systems + Battery/Power · Status: Open (D3 endurance) · Notes: D3 headline; shares PRD-VER-004 with RCR-006.
RCR-018 — Tolerate cold↔warm condensation without record-integrity failure
Requirement (shall) | The device shall tolerate condensation from cold-to-warm cycling without record-integrity failure.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-ENV-003, HA-TFC-006.
Source Failure Mode(s) | FA-DFMEA-001.
Parent Requirement | PRD-SAFE-010; PRD-REL; IEC 60068.
Implementation (DO-) | DO-SAFE-011.
Verification (PV-) | PV-ENV-003 + PV-TFC-006 (condensation cycle).
Validation (VAL-) | —
Independence / common-cause | Protective barrier (with RCR-015).
New-risk (HAL-2700) | none.
Acceptance Criteria | No record-integrity loss after the condensation-cycle profile.
Owner | Mechanical/PCB · Status: Baselined · Notes: Verified jointly with RCR-015.
6.5 Software / detection / verdict / record / time — information-integrity harm (dominant class)
RCR-019 — Two-sided excursion detection, freeze reliability = heat, no undetected path
Requirement (shall) | The device shall perform two-sided (heat and freeze) product-stability excursion detection with freeze detection reliability equal to heat, with no undetected-excursion path.
Control Strategy | ISD (Level 1). Warnings-challenge carrier.
Source Hazard(s) | HA-SW-001 (missed excursion → compromised dose), HA-ENV-001 (missed freeze), HA-DATA-001.
Source Failure Mode(s) | FA-SWFMEA-001 (one-sided freeze miss), FA-SWFMEA-002 (cumulative criterion wrong), FA-DFMEA-010 (internal-sensor lag/bias), FA-SYSFMEA-001 (undetected excursion, S5 scale) — = HAL control.
Parent Requirement | PRD-SAFE-002 (reliable two-sided freeze detection) and PRD-SAFE-001 (don't mislead); SYS-149/150/010/014; PRD-FUN-006/007.
Implementation (DO-) | DO-SAFE-004 → DO-FW-006 (detection) + DO-EE-013 (probe front-end).
Verification (PV-) | PV-SW-001 (profiled excursions, FI) — freeze coverage measured equal to heat.
Validation (VAL-) | VAL-SW-001 (detection effectiveness against real product-stability profiles).
Independence / common-cause | Detection anchored to the product-stability profile (RCR-048); device+cloud recompute cross-check (RCR-021) provides a diverse second computation.
New-risk (HAL-2700) | HA-NEW-012 — over-sensitive detection → false-reject + alarm load; balanced by RCR-047 validated targets and profile-anchoring (RCR-048), never by weakening the safety bias.
Acceptance Criteria | Against a profiled excursion battery (heat and freeze, single and cumulative), missed-freeze rate ≤ missed-heat rate and both within the validated detection target; no input profile produces an undetected excursion. Detection target values Evidence-Pending → PV-SW-001 / VAL-SW-001; anchored to the WHO PQS accuracy target (O-7) once confirmed.
Owner | Embedded/Software Safety · Status: Baselined · Notes: Top HAL headline control; FTA TOP-A.
RCR-020 — Probe-fault / implausible-read detection; never a valid product temperature
Requirement (shall) | An absent, open, short, or wrong probe, and implausible sensor reads, shall be detected and flagged, and shall never be read as a valid product-zone temperature.
Control Strategy | PROT (Level 2). Warnings-challenge carrier.
Source Hazard(s) | HA-SW-008 (valid-looking temps not being measured), HA-USE-005 (probe in air/against wall), HA-DATA-001.
Source Failure Mode(s) | FA-SWFMEA-007 (plausibility fails to flag), FA-HWFMEA-017 (probe open/intermittent), FA-DFMEA-011 (placement unrepresentative), FA-DFMEA-003 (probe gland/strain-relief) — = HAL control.
Parent Requirement | PRD-SAFE-001; SYS-007/024/151; HWR-022/028.
Implementation (DO-) | DO-SAFE-005 → DO-FW-007 (drivers/plausibility) + DO-EE-013 (presence detect) + DO-IF-008.
Verification (PV-) | PV-SW-008 (probe-fault-injection, FI) — incl. internal-vs-external divergence.
Validation (VAL-) | VAL-USE-005 (placement-error remainder).
Independence / common-cause | Internal-vs-external divergence flags a probe fault rather than silently averaging (HA-NEW-005 safety-biased disagreement rule).
New-risk (HAL-2700) | HA-NEW-005 — disagreement logic; resolved safety-biased (any divergence → fault/qualify, never a clean accept).
Acceptance Criteria | Every injected probe fault (absent/open/short/wrong) and every implausible read is flagged and never scored as a valid product-zone temperature; divergence beyond the threshold raises a probe-fault flag. Plausibility/divergence thresholds Evidence-Pending → PV-SW-008.
Owner | Embedded/Software Safety + Sensing · Status: Baselined · Notes: FTA TOP-A.
RCR-021 — Verdict integrity: suppress false "accept"; device↔cloud consistent
Requirement (shall) | The verdict shall suppress a false "accept" on any data gap, sensor fault, or incomplete monitoring; the device and cloud verdicts shall be consistent, and the control shall not be bypassable by firmware fault, configuration, or user action.
Control Strategy | ISD (Level 1), non-defeatable. Warnings-challenge carrier; SF-equivalent.
Source Hazard(s) | HA-SW-002 (false accept on incomplete record), HA-SW-007, HA-DATA-004, HA-USE-007.
Source Failure Mode(s) | FA-SWFMEA-008 (false accept on incomplete record), FA-SWFMEA-010 (displayed verdict false/stale), FA-SWFMEA-019 (overwrite of unsent/excursion data), FA-SYSFMEA-005 (corrupted/false record) — = HAL control.
Parent Requirement | PRD-SAFE-001 (severity anchor) and PRD-SAFE-003 (no false reassurance); SYS-015/016/ 017/151; PRD-FUN-010/011.
Implementation (DO-) | DO-SAFE-006 → DO-FW-008 (verdict) + DO-CLD-006 (recompute) + DO-HF-002 (display).
Verification (PV-) | PV-SW-002 (gap-handling / verdict-consistency, FI) — attempt to force a clean accept over a gap and confirm it is impossible.
Validation (VAL-) | VAL-SW-002 (verdict integrity in workflow) + VAL-USE-001.
Independence / common-cause | Device and cloud verdicts cross-check (diverse second computation) so a single computation error cannot silently release a false accept (HAL-2300 §8, non-defeatability rule).
New-risk (HAL-2700) | HA-NEW-005 (disagreement logic → safety-biased) and HA-NEW-006 (conservative bias → wastage, the central benefit-risk trade), balanced by RCR-047/044/040, never by weakening the suppression.
Acceptance Criteria | No combination of gap/fault/config/user action produces an unqualified "accept" over an incomplete record; device and cloud verdicts agree or the disagreement resolves to qualify/reject (never a clean accept); the control cannot be disabled from FW/config/UI.
Owner | Embedded/Software Safety + Cloud · Status: Baselined · Notes: The single most important information-harm inherent control; release-critical; FTA TOP-A/D. Non-defeatable with RCR-022.
RCR-022 — Honest gap: monitoring gaps visible; incomplete record never a clean pass
Requirement (shall) | Any monitoring gap, outage, or lapse shall be made visible in the record and in on-device status; an incomplete record shall never show an unqualified clean pass.
Control Strategy | INFO (Level 3, design-enforced), non-defeatable. Warnings-challenge carrier.
Source Hazard(s) | HA-SW-013 (reset/brown-out gap reads uneventful), HA-DATA-004, HA-USE-004.
Source Failure Mode(s) | FA-SWFMEA-008 (false accept on incomplete), FA-SWFMEA-018 (torn/corrupt record) — = HAL control (design-enforced honesty half of RCR-021/024).
Parent Requirement | PRD-SAFE-003 (no false reassurance / honest gaps); SYS-160/168; PRD-FUN-017, PRD-REL-006/ 007, PRD-UX-008.
Implementation (DO-) | DO-SAFE-006/008.
Verification (PV-) | PV-SW-002 / PV-DATA-004 (gap-exposure test, FI) + PV-SW-013.
Validation (VAL-) | VAL-USE-004 (gap comprehension).
Independence / common-cause | Design-enforced — the device forces the honest state; not a label the user must read (HAL-2300 §7). Ties RCR-021 (verdict) + RCR-040 (surface qualifier).
New-risk (HAL-2700) | HA-NEW-006 (honest-gap → conservative reject → wastage), balanced ALARP both directions.
Acceptance Criteria | Every injected gap/outage/reset appears in the record and on-device status; no incomplete record renders as a clean pass on device or cloud.
Owner | Embedded/Software Safety · Status: Baselined · Notes: Non-defeatable with RCR-021; release-critical.
RCR-023 — Defined safe state on fault; watchdog-backed logged recovery
Requirement (shall) | On any unsafe or misleading fault the device shall enter a defined safe state (preserve the record, hold the cell safe, show honest status) with watchdog-backed, logged recovery.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-SW-009 (device believed monitoring but hung), HA-ELEC-008, HA-ENRG-006.
Source Failure Mode(s) | FA-HWFMEA-016 (watchdog fails to recover), FA-SWFMEA-028 (recovery assumes false continuity), FA-SWFMEA-029 (self-check fails) — = HAL control.
Parent Requirement | PRD-SAFE-010 (fail-safe/recover); SYS-047/099/167/169.
Implementation (DO-) | DO-SAFE-007 → DO-FW-003 (mode) + DO-FW-013 (fault handler).
Verification (PV-) | PV-SW-009 (fault-injection + watchdog, FI).
Validation (VAL-) | —
Independence / common-cause | Logging prevents a recovery masking a gap (HA-NEW-005); recovery must not assume false continuity (FA-SWFMEA-028).
New-risk (HAL-2700) | covered under HA-NEW-005 (recovery/continuity logic verified).
Acceptance Criteria | Every injected hang/fault transitions to the safe state without corrupting the committed log or armed state; watchdog recovers and the event is logged; no silent resumption assuming continuity.
Owner | Embedded/Software Safety · Status: Baselined · Notes: FTA TOP-A/D.
RCR-024 — Power-fail-safe storage; no torn/lost/fabricated record; tamper-evident
Requirement (shall) | Storage shall be power-fail-safe (atomic writes, brown-out barrier); there shall be no torn, lost, or fabricated record, no overwrite of unsent data, and the record shall be tamper-evident.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-SW-013 (brown-out gap), HA-ELEC-008/010 (cold-sag brown-out during write), HA-DATA-006.
Source Failure Mode(s) | FA-SWFMEA-018 (power-fail → torn/corrupt), FA-SWFMEA-019 (overwrite of unsent), FA-HWFMEA-033 (flash wear/retention loss) — = HAL control.
Parent Requirement | PRD-SAFE-010 and PRD-SAFE-003; SYS-018/021/028/099/160/168; ALCOA+.
Implementation (DO-) | DO-SAFE-008 → DO-FW-010 (storage) + DO-EE-002 (rail interlock).
Verification (PV-) | PV-SW-013 (power-fail-during-write, FI) + flash-endurance test.
Validation (VAL-) | —
Independence / common-cause | Brown-out barrier + atomic writes; independent rail interlock; ties record-integrity chain (RCR-057).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No torn/lost/fabricated record across repeated power loss during write (incl. cold sag); unsent/excursion data never overwritten; tamper-evidence intact; flash retention over reuse life met. Flash endurance / retention Evidence-Pending → endurance test.
Owner | Embedded/Software Safety · Status: Baselined · Notes: FTA TOP-D; core of the record-integrity spine.
RCR-025 — Trusted monotonic UTC time; drift bounded and disclosed when untrusted
Requirement (shall) | The device shall maintain trusted monotonic UTC time across sleep/brown-out/service; drift shall be bounded and disclosed when time is untrusted.
Control Strategy | PROT (Level 2) + design-enforced disclosure.
Source Hazard(s) | HA-SW-005 (disposition/audit on mis-timed record), HA-DATA-003, HA-COMM-005.
Source Failure Mode(s) | FA-HWFMEA-014 (RTC failure/drift), FA-SWFMEA-032 (auth/session failure) — = HAL control.
Parent Requirement | PRD-SAFE-003 (honest disclosure); SYS-006/118; PRD-FUN (time).
Implementation (DO-) | DO-EE-011 (RTC) + DO-FW.
Verification (PV-) | PV-SW-005 (clock-drift/discipline over cold).
Validation (VAL-) | —
Independence / common-cause | Time re-discipline opportunistic; untrusted-time state disclosed rather than silently trusted.
New-risk (HAL-2700) | HA-NEW-011 — re-discipline spends radio energy; budgeted within the governor (RCR-006).
Acceptance Criteria | Time stays monotonic and within the drift bound across sleep/brown-out/service over cold; on loss of trust the record and status disclose "untrusted time." Drift bound Evidence-Pending → PV-SW-005.
Owner | Firmware · Status: Baselined · Notes: FTA TOP-D.
RCR-026 — Calibration integrity; expired/invalid calibration blocks arming
Requirement (shall) | Calibration coefficients shall be integrity-protected; expired or invalid calibration shall block arming (pre-arm gate).
Control Strategy | PROT (Level 2) primary; PROC secondary.
Source Hazard(s) | HA-SW-014 (coefficient/units error), HA-SW-004 (biased conditioning), HA-MFG-001, HA-SVC-001, HA-DATA-001.
Source Failure Mode(s) | FA-SWFMEA-005 (coefficient/units error), FA-HWFMEA-018 (sensor drift over reuse), FA-DFMEA-010 (internal-sensor bias) — = HAL control.
Parent Requirement | PRD-SAFE-006 (prevent unmonitored/invalid trip) and PRD-SAFE-001; SYS-007/024; PRD-FUN (calibration).
Implementation (DO-) | DO-FW-007 (drivers/plausibility) + DO-SAFE-007 (gate) + DO-FW-009.
Verification (PV-) | PV-SW-004 (calibration integrity/expiry FI) + PV-DATA-001 (+ static analysis).
Validation (VAL-) | —
Independence / common-cause | Shared calibration/firmware is a fleet CCF driver (FA-CCF-001/002); enforced by the pre-arm gate (RCR-045) and EOL/depot coverage (RCR-053/049).
New-risk (HAL-2700) | HA-NEW-007 — a falsely-expired reading blocks a legitimate arming; balanced by justified interval margin + legible cause (RCR-046).
Acceptance Criteria | Corrupted/expired coefficients are detected and block arming; a units/coefficient error is caught by integrity check + static analysis; the fleet common-cause is bounded by coverage. Interval/coverage values Evidence-Pending → PV-SW-004 / PV-MFG-001.
Owner | Metrology + Firmware · Status: Baselined · Notes: FTA TOP-A (CCF); ties RCR-045/049/053.
RCR-027 — Availability self-check; flags "not monitoring" on outage
Requirement (shall) | An availability self-check shall verify the whole sensing→logging chain and flag the device not monitoring on an outage (silent-outage prevention).
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-SW-013 (silent gap), HA-SW-009, HA-ENRG-006.
Source Failure Mode(s) | FA-SWFMEA-029 (self-check fails), FA-HWFMEA-016 (watchdog) — = HAL control.
Parent Requirement | PRD-SAFE-003; SYS-169/170; PRD-REL-007.
Implementation (DO-) | DO-SAFE-006/007 → DO-FW-012.
Verification (PV-) | PV-SW-009 (self-check outage test, FI).
Validation (VAL-) | —
Independence / common-cause | Self-check independence from the monitored path is a CCF item (FA-CCF-013).
New-risk (HAL-2700) | covered under HA-NEW-005 (self-check must be independent, verified).
Acceptance Criteria | Every injected chain outage (sensor/log/timing) is detected and surfaced as "not monitoring"; no outage reads as healthy.
Owner | Embedded/Software Safety · Status: Baselined · Notes: FTA TOP-A.
6.6 Connectivity / cloud — alert delivery
RCR-028 — Every excursion committed locally, independent of connectivity
Requirement (shall) | Every excursion event shall be committed locally independent of connectivity, so no excursion depends on a live link to exist.
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-COMM-001 (held/administered before late alert), HA-COMM-002 (alert never communicated).
Source Failure Mode(s) | FA-SWFMEA-024 (S&F alert lost on reset), FA-IFMEA-011 (API accepts-but-never-delivers), FA-IFMEA-010 (ingest not idempotent), FA-SYSFMEA-002 (no/late alert) — = HAL control.
Parent Requirement | PRD-SAFE-003 (offline-first honesty); D2; SYS-111/113.
Implementation (DO-) | DO-SAFE-009 → DO-FW-011.
Verification (PV-) | PV-COMM-001/002 (offline-excursion / S&F test).
Validation (VAL-) | VAL-COMM-001.
Independence / common-cause | Offline-first invariant — device safety never depends on the cloud (HA-NEW-008); the cloud delays, never loses.
New-risk (HAL-2700) | none (this is the control that breaks the cloud dependence introduced by other controls).
Acceptance Criteria | An excursion during a total comms outage is committed locally and delivered alert-first on reconnection with no loss across resets.
Owner | Firmware + Connectivity · Status: Baselined · Notes: FTA TOP-B; the offline-first cornerstone.
RCR-029 — Alert-ahead-of-bulk, ack-required, escalate-if-unacknowledged
Requirement (shall) | Excursion alerts shall be sent ahead of bulk data, shall require acknowledgement, and shall escalate if unacknowledged; cloud-side consolidation shall avoid alarm fatigue.
Control Strategy | PROT (Level 2). Warnings-challenge carrier.
Source Hazard(s) | HA-COMM-001, HA-CLOUD-001 (responder unaware), HA-USE-002 (alarm fatigue).
Source Failure Mode(s) | FA-SWFMEA-023 (alert not generated/prioritised), FA-SWFMEA-024, FA-IFMEA-011, FA-SYSFMEA-002, FA-IFMEA-020 (real alert muted) — = HAL control.
Parent Requirement | PRD-SAFE-004 (alarm balance); SYS-011/013; PRD-CLOUD.
Implementation (DO-) | DO-SAFE-009 → DO-CLD-007 (rules) + DO-CLD-008 (notify).
Verification (PV-) | PV-SW-006 (alert-gen + escalation, FI) + PV-COMM-001.
Validation (VAL-) | VAL-SW-006 + VAL-USE-002 (alarm-management effectiveness).
Independence / common-cause | Escalation to a different responder rather than repeat-spamming the same one; cloud-side consolidation.
New-risk (HAL-2700) | HA-NEW-004 — escalation adds alarm load → fatigue; balanced by RCR-047 tuned targets + consolidation.
Acceptance Criteria | An excursion alert precedes bulk, requires ack, and escalates to an alternate responder within the target window if unacknowledged; consolidation holds alert volume within the validated fatigue-avoidance target. Latency/escalation/volume targets Evidence-Pending → PV-SW-006 / VAL-USE-002.
Owner | Firmware + Cloud + Human-Factors · Status: Baselined · Notes: Release-critical; FTA TOP-B.
RCR-030 — Degrade to store-and-forward on link loss; never record loss; alert-first on reconnect
Requirement (shall) | On connectivity loss the device shall degrade to store-and-forward buffering with never any record loss, and deliver alert-first on reconnection.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-COMM-001/003 (container blocks radio).
Source Failure Mode(s) | FA-SWFMEA-024, FA-HWFMEA-029, FA-IFMEA-005 (no in-container link) — = HAL control.
Parent Requirement | PRD-SAFE-003/010; D2; SYS-111.
Implementation (DO-) | DO-SAFE-009.
Verification (PV-) | PV-COMM-001 (store-and-forward test).
Validation (VAL-) | VAL-COMM-001.
Independence / common-cause | Buffering is local (independent of link); ties RCR-028.
New-risk (HAL-2700) | ties HA-NEW-011 (safety-plane backoff energy, budget-bounded).
Acceptance Criteria | Across arbitrary link outages the buffer never loses a record and delivers alert-first on reconnection; buffer depth covers the worst-case outage. Buffer depth / worst-case outage Evidence-Pending → PV-COMM-001 (bounded by PRD-VER-005 in-container result).
Owner | Firmware + Connectivity · Status: Baselined · Notes: The consequence-bounding control for O-1.
RCR-031 — RF closes the in-container budget, or IP-sealed external-antenna fallback
Requirement (shall) | The RF link (integral antenna) shall close the in-container link budget across the target markets, or an IP-sealed external-antenna fallback shall be provided; RF-in-container performance shall be qualified.
Control Strategy | ISD (Level 1).
Source Hazard(s) | HA-COMM-003 (excursion with no live notification), HA-COMM-006 (roaming gap).
Source Failure Mode(s) | FA-HWFMEA-029 (no in-container link), FA-HWFMEA-030 (cold RF margin loss), FA-HWFMEA-032 (SIM/roaming gap), FA-IFMEA-005 (no egress in container), FA-SYSFMEA-002 — = HAL control.
Parent Requirement | PRD-COMM; D2/D4; SYS-104/105/109.
Implementation (DO-) | DO-IF (antenna) + DO-EE-017/018.
Verification (PV-) | PRD-VER-005 (RF-in-container qualification) + PV-COMM-003.
Validation (VAL-) | VAL-COMM-003.
Independence / common-cause | Where the link cannot close, the consequence is bounded by store-and-forward (RCR-030) + on-receipt edge read; the real-time-alert claim is withheld on such lanes (release condition).
New-risk (HAL-2700) | HA-NEW-009 — external-antenna fallback adds a sealed-connector ingress path (RCR-011/015) and per-market config error path (RCR-056).
Acceptance Criteria | The integral antenna closes the in-container budget across the target markets/containers, or the external-antenna fallback does; where neither closes, store-and-forward + on-receipt verdict bound the consequence and the real-time claim is not made. In-container attenuation, link margin, and the integral-vs-external antenna decision are Evidence-Pending → PRD-VER-005 (O-1, largest connectivity unknown).
Owner | Connectivity + RF/Antenna + Regulatory · Status: Open (RF-in-container, O-1) · Notes: D3-class headline item; validate-by-test.
RCR-032 — Cloud never drops a safety event (dead-letter/retry, reconciliation)
Requirement (shall) | The cloud shall not drop a safety event: it shall dead-letter/retry on failure, never silently discard, and provide end-to-end delivery reconciliation.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CLOUD-001 (notify fail), HA-CLOUD-002, HA-CLOUD-004 (record lost), HA-COMM-002.
Source Failure Mode(s) | FA-IFMEA-011 (accepts-but-never-delivers), FA-IFMEA-010 (ingest not idempotent), FA-SWFMEA-032, FA-SYSFMEA-005 — = HAL control.
Parent Requirement | PRD-CLOUD; PRD-SAFE-003; SRS-CLOUD-023; SYS-176 (reconciliation).
Implementation (DO-) | DO-CLD-007 (SRS-CLOUD-023).
Verification (PV-) | PV-COMM-002 / PV-CLOUD-001 (ingest/notify fault-injection, FI) + PV-COMM-004.
Validation (VAL-) | VAL-COMM-002.
Independence / common-cause | Idempotent ingest + reconciliation; offline-first (RCR-028) means the device still holds the truth if the cloud fails.
New-risk (HAL-2700) | HA-NEW-008 — cloud is a fleet common-cause; balanced by offline-first + monitoring + DR (RCR-039).
Acceptance Criteria | Under injected ingest/notify failures no safety event is silently discarded; every event is dead-lettered/retried and reconciled end-to-end; duplicate ingest is idempotent.
Owner | Cloud · Status: Baselined · Notes: FTA TOP-B/D.
RCR-033 — Authoritative device↔trip↔owner↔responder binding
Requirement (shall) | The device↔trip↔owner↔responder binding shall be authoritative and correct; alerts shall route to the right responder; wrong-tenant binding shall be prevented.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CLOUD-003, HA-COMM-008 (real excursion invisible on its own record), HA-DATA-002 (another shipment's data).
Source Failure Mode(s) | FA-SYSFMEA-004 (record mis-attributed) — = HAL control.
Parent Requirement | PRD-CLOUD; PRD-SAFE-001; SYS-030/012/134.
Implementation (DO-) | DO-CLD (registry) + DO-SEC.
Verification (PV-) | PV-COMM-008 / PV-CLOUD-003 (association/routing test) + PV-MFG-005.
Validation (VAL-) | VAL-COMM-008 / VAL-USE-008.
Independence / common-cause | Unique device identity (RCR-055) anchors the binding.
New-risk (HAL-2700) | none material.
Acceptance Criteria | No cross-tenant mis-binding; every alert routes to the bound responder; mis-attribution attempts are rejected.
Owner | Cloud + Security · Status: Baselined · Notes: FTA TOP-D.
6.7 Cybersecurity-induced safety (independent adversarial verification)
RCR-034 — Capture-time authenticity/integrity; falsified data detectable/rejected
Requirement (shall) | Telemetry and the record shall carry capture-time authenticity/integrity (tamper-evidence); falsified or injected data shall be detectable and rejected.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CYBER-001 (false "clean" record from injected data), HA-DATA-001, HA-MPE-004.
Source Failure Mode(s) | FA-SWFMEA-031 (tamper-evidence not enforced), FA-SYSFMEA-005 — = HAL control.
Parent Requirement | PRD-SEC (safety-relevant) → PRD-SAFE-001; SYS-026/028.
Implementation (DO-) | DO-SEC (integrity) + DO-SAFE-008.
Verification (PV-) | PV-CYBER-001 (integrity / pen-test, independent adversarial — IV).
Validation (VAL-) | —
Independence / common-cause | Device-anchored authenticity; key lifecycle bounded by RCR-038/055/052.
New-risk (HAL-2700) | HA-NEW-003 — crypto adds latency/energy + a key-lifecycle single point; latency/energy is an accepted trade (budgeted vs PRD-VER-004/005); key lifecycle controlled.
Acceptance Criteria | Injected/falsified data is detected and rejected; capture-time authenticity holds under an independent pen-test. Pen-test scope Evidence-Pending → PV-CYBER-001 (IV).
Owner | Product-Security · Status: Baselined · Notes: FTA TOP-D; requires IV before residual acceptance.
RCR-035 — Alert path not silently suppressible; suppression raises an event
Requirement (shall) | The alarm/alert path shall not be silently suppressible; suppression or blocking shall itself raise a detectable event and escalate.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CYBER-002 (excursion hidden by muted/blocked alarm), HA-COMM-002.
Source Failure Mode(s) | FA-SWFMEA-033 (alarm suppression unauth) — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-004; DO-FW-016.
Implementation (DO-) | DO-SEC + DO-SAFE-009.
Verification (PV-) | PV-CYBER-002 (suppression pen-test, IV).
Validation (VAL-) | —
Independence / common-cause | Ties escalation (RCR-029) — a blocked path escalates rather than going silent.
New-risk (HAL-2700) | covered under HA-NEW-004 (alert load) — suppression detection does not add load, it raises a distinct event.
Acceptance Criteria | Any attempt to mute/block the alert path raises a distinct, escalated event; no silent suppression under pen-test.
Owner | Product-Security · Status: Baselined · Notes: IV required.
RCR-036 — Threshold/config changes authenticated, authorised, bounded, audited
Requirement (shall) | Excursion thresholds and configuration changes shall be authenticated, authorised, bounded, and audited; unauthorised change shall be prevented.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CYBER-003 (detection defeated by altered thresholds), HA-USE-003, HA-CLOUD-002.
Source Failure Mode(s) | FA-IFMEA-016 (BLE unauth pairing → rogue config) — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-001; SYS-107/110.
Implementation (DO-) | DO-SEC + DO-CLD (audit).
Verification (PV-) | PV-CYBER-003 (config-tamper pen-test, IV).
Validation (VAL-) | —
Independence / common-cause | Bounds prevent an out-of-range threshold even by an authorised actor; ties profile validation (RCR-048).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No unauthorised or out-of-bounds threshold/config change succeeds; every change is authenticated, bounded, and audited under pen-test.
Owner | Product-Security + Cloud · Status: Baselined · Notes: IV required; FTA TOP-A/D.
RCR-037 — Signed, rollback-protected, fail-safe A/B OTA; never bricks; windows disclosed
Requirement (shall) | OTA shall be signed, integrity- and rollback-protected, fail-safe A/B with automatic rollback; it shall never brick a device; unmonitored update windows shall be disclosed.
Control Strategy | PROT (Level 2) (S5 via HA-CYBER-005).
Source Hazard(s) | HA-SW-010 (bad update → wrong logic/brick), HA-CYBER-005 (malicious firmware → fleet unsafe), HA-CLOUD-007.
Source Failure Mode(s) | FA-SWFMEA-026 (bad OTA wrong-logic/brick), FA-IFMEA-015 (unsigned image applied) — = HAL control.
Parent Requirement | PRD-SAFE-010 (fail-safe incl. OTA) + PRD-SEC-008; SYS-032/114/139.
Implementation (DO-) | DO-SAFE-014 → DO-FW-014/015 + DO-CLD-010.
Verification (PV-) | PV-SW-010 (rollback + version-skew) + PV-CYBER-005 (OTA pen-test, IV).
Validation (VAL-) | —
Independence / common-cause | HA-NEW-008 — OTA is a fleet common-cause; controlled by staged/canary rollout + automatic rollback + halt-on-regression (CCF FA-CCF-007).
New-risk (HAL-2700) | HA-NEW-003 (crypto), HA-NEW-008 (fleet common-cause) — both controlled by rollout discipline + offline-first.
Acceptance Criteria | An unsigned/corrupt image is rejected; a bad image auto-rolls-back with trip data preserved; no brick under version-skew; unmonitored window disclosed; staged rollout halts on regression. Pen-test confirms signing/rollback cannot be defeated. Pen-test scope Evidence-Pending → PV-CYBER-005 (IV).
Owner | Firmware + Cloud + Product-Security · Status: Baselined · Notes: Secure signed OTA + rollback is a dominant SCM control; IV required.
RCR-038 — Secure boot to HW root of trust; debug ports locked; tamper-evident
Requirement (shall) | The device shall secure-boot against a hardware root of trust; production debug ports shall be locked/disabled; the enclosure shall be tamper-evident.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-CYBER-006 (forged records / cloned device), HA-CYBER-008, HA-SVC-006.
Source Failure Mode(s) | FA-IFMEA-015, FA-IFMEA-016, FA-SWFMEA-034 (debug mode enabled in field) — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-010; SYS-032/114.
Implementation (DO-) | DO-SEC-006 + DO-FW.
Verification (PV-) | PV-CYBER-006 (secure-boot / debug-lock pen-test, IV).
Validation (VAL-) | —
Independence / common-cause | Root of trust anchors OTA (RCR-037) and identity (RCR-055).
New-risk (HAL-2700) | HA-NEW-003 — the key/secure-element is a single point; key lifecycle bounded (RCR-055 identity, RCR-052 erasure).
Acceptance Criteria | Only signed firmware boots; debug ports are non-functional in production; tamper is evident; no bypass under pen-test.
Owner | Product-Security · Status: Baselined · Notes: IV required.
RCR-039 — Cloud resists compromise (RBAC/SSO, monitoring, tested DR)
Requirement (shall) | The cloud shall resist compromise (RBAC/SSO, monitoring of auth/integrity/backlog events, tested backup/restore) so a platform incident cannot blind the fleet or corrupt records silently.
Control Strategy | PROT (Level 2) primary; PROC secondary (S8 fleet-blinding).
Source Hazard(s) | HA-CYBER-009 (ransomware → fleet blind + records altered), HA-CLOUD-004/006/008.
Source Failure Mode(s) | FA-SWFMEA-032, FA-SYSFMEA-005 (cloud side) — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-011; PRD-CLOUD.
Implementation (DO-) | DO-CLD + DO-SEC.
Verification (PV-) | PV-CYBER-009 (cloud security review + DR test, IV).
Validation (VAL-) | —
Independence / common-cause | HA-NEW-008 — cloud is the fleet common-cause; offline-first (RCR-028) is the independence guarantee; DR-tested restore bounds a platform incident.
New-risk (HAL-2700) | HA-NEW-008 (fleet common-cause + connectivity dependence) — controlled by offline-first + rollout/monitoring discipline.
Acceptance Criteria | RBAC/SSO enforced; auth/integrity/backlog events monitored; backup/restore demonstrated within the DR objective; a simulated compromise cannot silently blind the fleet or corrupt records. DR RTO/RPO Evidence-Pending → PV-CYBER-009 (IV).
Owner | Cloud + Product-Security · Status: Baselined · Notes: IV + DR test required.
6.8 Data & decision
RCR-040 — Disposition surface shows incomplete-monitoring qualifier; never a clean pass on a gap
Requirement (shall) | The disposition surface shall show the incomplete-monitoring qualifier and shall never present a gap or fault as a clean pass (ties RCR-021/022).
Control Strategy | INFO (Level 3, design-enforced). Warnings-challenge carrier.
Source Hazard(s) | HA-DATA-004 (incomplete record as clean pass), HA-DATA-005, HA-USE-007.
Source Failure Mode(s) | FA-SWFMEA-008, FA-SYSFMEA-008 (reject overridden) — decision-support half of RCR-021/022.
Parent Requirement | PRD-SAFE-003 and PRD-SAFE-005; PRD-UX; SYS (disposition surface).
Implementation (DO-) | DO-SAFE-006 + DO-HF-002.
Verification (PV-) | PV-DATA-004 (gap-exposure) + PV-DATA-005 (visualisation review).
Validation (VAL-) | VAL-USE-004 (decision-support usability) + VAL-USE-005.
Independence / common-cause | Design-enforced surface — cannot render a clean pass over a gap (ties RCR-021/022).
New-risk (HAL-2700) | HA-NEW-006 — the honest qualifier drives conservative rejects; the qualifier tells the human why so an override (RCR-044) can be proportionate.
Acceptance Criteria | No incomplete/faulted trip renders as a clean pass on the disposition surface; the qualifier is present and comprehended in summative usability. Comprehension target Evidence-Pending → VAL-USE-004.
Owner | Human-Factors + Cloud · Status: Baselined · Notes: Usability-gated.
RCR-041 — Estimated/interpolated/last-known data flagged, never presented as measured
Requirement (shall) | Estimated, interpolated, or last-known data (e.g., GNSS, gap-fill) shall be flagged as such and shall never be presented as measured or authoritative.
Control Strategy | INFO (Level 3, design-enforced). Warnings-challenge carrier.
Source Hazard(s) | HA-DATA-007 (estimate believed a measurement), HA-COMM-003.
Source Failure Mode(s) | provenance/visualisation records (FAL-1400/1500).
Parent Requirement | PRD-SAFE-003; PRD-UX; ALCOA+.
Implementation (DO-) | DO-CLD + DO-HF.
Verification (PV-) | PV-DATA-007 (provenance/visualisation review).
Validation (VAL-) | VAL-USE-005 (estimate-vs-measured comprehension).
Independence / common-cause | Provenance carried through export (RCR-042).
New-risk (HAL-2700) | none material.
Acceptance Criteria | Every estimated/interpolated/last-known value is visibly flagged; users distinguish estimate from measurement in summative usability.
Owner | Human-Factors + Cloud · Status: Baselined · Notes: —
RCR-042 — Export/API preserves attribution, time, gaps, tamper-evidence; no silent loss
Requirement (shall) | Export and API shall preserve attribution, time, gaps, and tamper-evidence in a documented format, with no silent data loss.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-DATA-006 (record missing evidence), HA-MPE-004 (data-integrity at scale).
Source Failure Mode(s) | FA-SYSFMEA-005 (corrupted/false record) — = HAL control.
Parent Requirement | PRD-CLOUD; PRD-SAFE-003; SYS (API/audit); ALCOA+ / Part-11-style.
Implementation (DO-) | DO-CLD (API/audit).
Verification (PV-) | PV-DATA-006 (power-fail + export) + traceability test.
Validation (VAL-) | —
Independence / common-cause | Ties the record-integrity spine (RCR-057); export cannot strip attribution/gaps.
New-risk (HAL-2700) | none material.
Acceptance Criteria | Exports preserve attribution, time, gaps, and tamper-evidence in the documented format; no silent loss on export; format audit passes.
Owner | Cloud + Quality/Regulatory · Status: Baselined · Notes: Part-11-style audit trail / data authenticity.
6.9 Use-related / human factors — warnings-challenge decisive
RCR-043 — Legible, language-independent verdict (gloved, low-light, −30 °C)
Requirement (shall) | The accept/reject/quarantine verdict shall be legible and language-independent (icon+colour+word), readable gloved, in low light, and at −30 °C within the responsiveness budget.
Control Strategy | ISD (Level 1). Warnings-challenge carrier.
Source Hazard(s) | HA-USE-001 (verdict misread), HA-USE-008, HA-ENV-002.
Source Failure Mode(s) | FA-SWFMEA-010 (displayed verdict false/stale), FA-IFMEA-019 (verdict misread) — = HAL control.
Parent Requirement | PRD-SAFE-005 (use-error protection) and PRD-SAFE-001; PRD-UX-001/007; SYS-015/041.
Implementation (DO-) | DO-SAFE-013 + DO-HF-002/005.
Verification (PV-) | PV-USE-001 (cold-display) + display-integrity (FI, with RCR-021).
Validation (VAL-) | VAL-USE-001 (summative usability).
Independence / common-cause | The verdict is designed readable by the vulnerable user — not "trained to interpret" (HAL-2300 §7).
New-risk (HAL-2700) | none material.
Acceptance Criteria | In summative usability with gloved users in low light at cold, the verdict is read correctly within the responsiveness budget at the target success rate. Success-rate/responsiveness targets Evidence-Pending → VAL-USE-001.
Owner | Human-Factors · Status: Baselined · Notes: Usability-gated; release-relevant use control.
RCR-044 — Verdict-override requires deliberate, confirmed, recorded action
Requirement (shall) | A verdict override shall require a deliberate, confirmed, recorded action; accidental acceptance of a compromised dose shall be prevented.
Control Strategy | ISD (Level 1). Warnings-challenge carrier.
Source Hazard(s) | HA-USE-007 (reject overridden under pressure), HA-IUE-007, HA-MPE-001.
Source Failure Mode(s) | FA-SYSFMEA-008 (reject overridden), FA-USE-006 — = HAL control.
Parent Requirement | PRD-SAFE-005 and PRD-SAFE-001; PRD-UX; SYS (override); DO-CLD audit.
Implementation (DO-) | DO-HF + DO-CLD (audit).
Verification (PV-) | PV-DATA-008 (recommendation/override workflow).
Validation (VAL-) | VAL-USE-007 (override-workflow usability).
Independence / common-cause | Override is hard and audited (not discouraged by a note); the honest qualifier (RCR-040) informs the decision.
New-risk (HAL-2700) | balances HA-NEW-006 — an audited override exists so a genuinely-good marginal lot is not blindly destroyed (wastage control), without weakening the safety bias.
Acceptance Criteria | Accidental acceptance is not possible; a deliberate override is captured with actor, reason, and time; override rate + reasons are monitorable. Workflow effectiveness Evidence-Pending → VAL-USE-007.
Owner | Human-Factors + Cloud · Status: Baselined · Notes: Usability-gated; ties the wastage trade (HA-MPE-001/003).
RCR-045 — Pre-arm readiness gate (SoC / calibration / firmware / profile / probe)
Requirement (shall) | A pre-arm readiness gate shall block arming unless SoC-for-trip, valid/unexpired calibration, valid firmware/profile, and probe-present all pass; an un-armed/void state shall be unmistakable.
Control Strategy | PROT (Level 2), non-defeatable. Warnings-challenge carrier.
Source Hazard(s) | HA-USE-003/004 (forgot-to-arm / void trip; wrong profile), HA-SW-003, HA-SVC-005.
Source Failure Mode(s) | FA-SWFMEA-004 (wrong/default profile), FA-SWFMEA-011 (runs unmonitored), FA-HWFMEA-017 (probe open), FA-USE-001 (forgot-to-arm), FA-DFMEA-011/003 (placement/gland) — = HAL control.
Parent Requirement | PRD-SAFE-006 (prevent unmonitored trip start); SYS-023/024; DO-FW-032.
Implementation (DO-) | DO-SAFE-007 + DO-FW-032.
Verification (PV-) | PV-SW-003 (pre-arm gate + arming/void, FI).
Validation (VAL-) | VAL-USE-004 (pre-arm/void comprehension).
Independence / common-cause | The "is it ready?" judgement is taken off the low-training user and enforced by the device; non-defeatable (HAL-2300 §7).
New-risk (HAL-2700) | HA-NEW-007 — an over-strict gate blocks a legitimate arming → availability/stock-out; balanced by justified margin + legible cause (RCR-046) + spare logistics.
Acceptance Criteria | No trip arms with a failing readiness item; the un-armed/void state is unmistakable; a blocked arming shows its cause. Gate thresholds set with justified margin (not hair-trigger).
Owner | Firmware + Human-Factors · Status: Baselined · Notes: Non-defeatable; carries multiple warnings-challenge use hazards.
RCR-046 — Charge-inhibited/fault/unready states legible with cause (cold ≠ dead)
Requirement (shall) | Charge-inhibited, fault, and unready states shall be made legible with their cause (LED/OLED); a cold, charge-inhibited device shall not be mistaken for a dead/discardable device.
Control Strategy | INFO (Level 3, design-enforced). Warnings-challenge carrier.
Source Hazard(s) | HA-USE-009 (cold charge-inhibited device discarded as dead), HA-ENRG-001.
Source Failure Mode(s) | state-legibility records (FAL-1600).
Parent Requirement | PRD-SAFE-005/007; PRD-UX; SYS-024/094/100.
Implementation (DO-) | DO-SAFE-013 → DO-HF-005 + DO-EE-008.
Verification (PV-) | PV-USE (state-legibility test).
Validation (VAL-) | VAL-USE (discard-as-dead comprehension).
Independence / common-cause | Design-enforced legibility; the charge safety itself is inherent HW (RCR-001).
New-risk (HAL-2700) | balances HA-NEW-001 (dead-on-arrival) and HA-NEW-007 (blocked arming shows its cause so the depot can remedy).
Acceptance Criteria | Users correctly distinguish charge-inhibited/unready from dead in summative usability; cause is shown; discard-as-dead rate is monitorable.
Owner | Human-Factors · Status: Baselined · Notes: —
RCR-047 — Alerting tuned to validated false-alarm/missed-excursion targets (no fatigue)
Requirement (shall) | Alerting shall be tuned to validated false-alarm and missed-excursion targets that avoid alarm fatigue muting a real excursion; consolidation shall be applied.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-USE-002 (real excursion muted), HA-SW-006, HA-MPE-003 (false-reject stock-out).
Source Failure Mode(s) | FA-SYSFMEA-013 (false-reject stock-out), FA-IFMEA-020 (real alert muted) — = HAL control.
Parent Requirement | PRD-SAFE-004 (alarm balance); PRD-CLOUD-003; PRD-FUN-007.
Implementation (DO-) | DO-SAFE-004 + DO-CLD-003.
Verification (PV-) | PV-USE-002 (alarm-management).
Validation (VAL-) | VAL-USE-002 (false-alarm/false-reject targets — ALARP both directions).
Independence / common-cause | The single control that keeps the safety-vs-availability trade balanced.
New-risk (HAL-2700) | balances HA-NEW-004 (alarm fatigue), HA-NEW-006 (wastage), HA-NEW-012 (over-sensitive detection) — the central benefit-risk tuning control.
Acceptance Criteria | False-alarm and missed-excursion rates meet the validated targets in both directions; consolidation holds volume within the fatigue-avoidance target. Targets Evidence-Pending → VAL-USE-002; anchored to the product-stability profile and WHO PQS target (O-7).
Owner | Human-Factors + Cloud + System-Safety · Status: Baselined · Notes: Validate-by-test; carries the accepted public-health trade (HA-MPE-003).
RCR-048 — Provisioning validates trip profile against the product-stability library
Requirement (shall) | Provisioning shall validate the selected trip profile against the product-stability library and flag an implausible or mismatched profile.
Control Strategy | PROT (Level 2).
Source Hazard(s) | HA-USE-003 (wrong thresholds), HA-SW-003.
Source Failure Mode(s) | FA-SWFMEA-002 (cumulative criterion), FA-SWFMEA-004 (wrong/default profile) — = HAL control.
Parent Requirement | PRD-SAFE-006 and PRD-SAFE-002; PRD-APP/CLOUD; SYS-154.
Implementation (DO-) | DO-APP + DO-CLD.
Verification (PV-) | PV-SW-003 (profile-validation test).
Validation (VAL-) | VAL-USE-003.
Independence / common-cause | Profile-anchoring makes detection thresholds the vaccine's real stability limits (ties RCR-019 / HA-NEW-012).
New-risk (HAL-2700) | ties HA-NEW-007 (a valid profile falsely flagged mismatched → blocked arming).
Acceptance Criteria | An implausible/mismatched profile is flagged and blocks arming (via RCR-045); a correct profile validates against the stability library.
Owner | Firmware + Cloud + Metrology · Status: Baselined · Notes: —
6.10 Service / maintenance
RCR-049 — Recalibration interval enforced by the pre-arm expiry gate
Requirement (shall) | The recalibration interval shall be enforced by the pre-arm expiry gate (RCR-026); an out-of-calibration device shall not be armed.
Control Strategy | PROT (Level 2) primary; PROC secondary.
Source Hazard(s) | HA-SVC-001 (depot recal escape → systematic bias), HA-MFG-001.
Source Failure Mode(s) | FA-SYSFMEA-009 (depot recal escape, fleet CCF) — = HAL control.
Parent Requirement | PRD-SAFE-006; PRD-SVC; SYS-024.
Implementation (DO-) | DO-FW-009 + DO-SVC.
Verification (PV-) | PV-SVC-001 (interval enforcement) + PV-SW-004.
Validation (VAL-) | —
Independence / common-cause | Depot recal is a fleet CCF driver (FA-CCF-001/002); enforced device-side by the gate.
New-risk (HAL-2700) | ties HA-NEW-007 (a marginally-valid calibration read as expired blocks arming).
Acceptance Criteria | No device with expired calibration can be armed; depot recal-escape is caught by the gate; the fleet escape rate is monitorable.
Owner | Metrology + Service · Status: Baselined · Notes: FTA TOP-A (CCF).
RCR-050 — Service restores seal/config/firmware; post-service BIST + pre-arm gate
Requirement (shall) | Service shall restore the seal, configuration, and firmware; a post-service BIST + pre-arm gate shall verify integrity before redeploy.
Control Strategy | PROC (Level 4) primary; PROT secondary.
Source Hazard(s) | HA-SVC-003/005 (service leaves device unready), HA-SVC-002 (wrong firmware).
Source Failure Mode(s) | service-restore records (FAL-1400); FA-SWFMEA-026 (wrong firmware) via RCR-037.
Parent Requirement | PRD-SAFE-010; PRD-SVC.
Implementation (DO-) | DO-SVC + DO-TST.
Verification (PV-) | PV-SVC (post-service BIST test).
Validation (VAL-) | —
Independence / common-cause | Post-service BIST + pre-arm gate (RCR-045) is the protective verification behind the service process.
New-risk (HAL-2700) | none material.
Acceptance Criteria | After service, seal/config/firmware are restored and verified by BIST + pre-arm gate before redeploy; a mis-serviced device cannot arm.
Owner | Service + Quality · Status: Baselined · Notes: —
RCR-051 — Debug/dev modes locked before deployment; unexited mode detected
Requirement (shall) | Diagnostic/developer/debug modes shall be locked/disabled before field deployment; an unexited service mode shall be detected.
Control Strategy | PROT (Level 2) primary; PROC secondary.
Source Hazard(s) | HA-SVC-006 (debug mode left enabled), HA-CYBER-006.
Source Failure Mode(s) | FA-SWFMEA-034 (debug/dev mode enabled in field) — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-010; PRD-SVC; DO-EE-012/DO-FW-017.
Implementation (DO-) | DO-SEC + DO-SVC.
Verification (PV-) | PV-CYBER-006 (mode-lock verification, IV).
Validation (VAL-) | —
Independence / common-cause | Ties secure boot / debug-lock (RCR-038).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No field device boots with debug/dev mode active; an unexited mode is detected and flagged; confirmed under pen-test.
Owner | Product-Security + Service · Status: Baselined · Notes: IV required.
RCR-052 — Safe end-of-life instructions (battery/data/key erasure) + design support
Requirement (shall) | Safe end-of-life instructions (battery removal/recycling, data/key erasure) shall be provided and supported by design.
Control Strategy | INFO (Level 3) + PROC — the one accepted non-design-enforced information control (no higher control removes the disposal hazard).
Source Hazard(s) | HA-SVC-007 (improper Li disposal), HA-CYBER-008 (key/data remanence), HA-MPE-005 (e-waste).
Source Failure Mode(s) | disposal records (FAL-1400).
Parent Requirement | PRD-SVC/REG; PRD-SEC (key erasure).
Implementation (DO-) | DO-SVC + DO-LBL + DO-SEC.
Verification (PV-) | Disposal-procedure + key-erasure review.
Validation (VAL-) | —
Independence / common-cause | Design support (battery removal, key/data erasure) so the instruction is followable.
New-risk (HAL-2700) | bounds the key-lifecycle single point (HA-NEW-003) at decommission.
Acceptance Criteria | Disposal instructions exist and are supported by design; key/data erasure is demonstrable at decommission. Justified as information-only because no inherent/protective measure removes the disposal hazard (HAL-2300 §7).
Owner | Service + Product-Security + Regulatory · Status: Baselined · Notes: The single justified information-only control (see Control-Hierarchy-Analysis §4).
6.11 Manufacturing / supply-chain
RCR-053 — EOL calibration coverage/accuracy + birth record; miscalibrated unit fails EOL
Requirement (shall) | EOL calibration shall meet the defined coverage and accuracy and record a birth record; a miscalibrated unit shall not pass EOL.
Control Strategy | PROC (Level 4).
Source Hazard(s) | HA-MFG-001 (miscalibrated unit passes EOL), HA-SVC-001.
Source Failure Mode(s) | FA-SYSFMEA-009 (calibration escape, fleet CCF) — = HAL control.
Parent Requirement | PRD-SAFE-001 (accuracy basis); PRD-MFG; SYS-019/020.
Implementation (DO-) | DO-TST + DO-MFG.
Verification (PV-) | PV-MFG-001 (EOL calibration coverage audit).
Validation (VAL-) | —
Independence / common-cause | EOL coverage bounds the fleet miscalibration CCF; birth record anchors traceability.
New-risk (HAL-2700) | none material.
Acceptance Criteria | EOL calibration coverage/accuracy meets the target; every unit has a birth record; a miscalibrated unit fails EOL. Coverage/accuracy target Evidence-Pending → PV-MFG-001; anchored to the WHO PQS accuracy target (O-7).
Owner | Metrology + Manufacturing · Status: Baselined · Notes: FTA TOP-A (CCF).
RCR-054 — Correct cell chemistry + correct-rated/oriented protection parts (poka-yoke + AOI)
Requirement (shall) | Correct cell chemistry and correct-rated/oriented protection components shall be verified (poka-yoke + AOI); counterfeit or wrong parts shall be prevented.
Control Strategy | PROC (Level 4) primary; PROT secondary (poka-yoke = error-proofing).
Source Hazard(s) | HA-MFG-002 (wrong/counterfeit chemistry), HA-MFG-003 (reversed/wrong-rated protection part), HA-ENRG-002.
Source Failure Mode(s) | FA-HWFMEA-004 (wrong chemistry), FA-HWFMEA-008 (secondary fuse), FA-HWFMEA-003 — = HAL control.
Parent Requirement | PRD-SAFE-008; PRD-MFG; SYS (incoming/AOI); DO-CFG-004.
Implementation (DO-) | DO-MFG + DO-TST.
Verification (PV-) | PV-MFG-002 (incoming + AOI) + PV-MFG-003 (poka-yoke + AOI).
Validation (VAL-) | —
Independence / common-cause | Correct chemistry/part underpins RCR-001/002 (charge safety) — a wrong part defeats them (CCF).
New-risk (HAL-2700) | none material.
Acceptance Criteria | No wrong-chemistry cell or wrong-rated/reversed protection part passes incoming + AOI; poka-yoke prevents mis-orientation.
Owner | Manufacturing + Supply-Chain · Status: Baselined · Notes: FTA TOP-C; underpins battery safety.
RCR-055 — Unique cryptographic identity + serial; duplicates prevented
Requirement (shall) | Each device shall carry a unique cryptographic identity and serial; duplicate identity/serial shall be prevented.
Control Strategy | PROC (Level 4).
Source Hazard(s) | HA-MFG-005 (wrong/duplicate identity), HA-CYBER-006.
Source Failure Mode(s) | FA-SYSFMEA-004 (mis-attributed), FA-IFMEA-016 — = HAL control.
Parent Requirement | PRD-SEC → PRD-SAFE-001 (attribution); PRD-MFG; SYS-016.
Implementation (DO-) | DO-MFG (provisioning) + DO-SEC.
Verification (PV-) | PV-MFG-005 (identity-uniqueness check).
Validation (VAL-) | —
Independence / common-cause | Anchors the device↔trip binding (RCR-033) and secure boot/OTA (RCR-037/038).
New-risk (HAL-2700) | bounds the key-lifecycle (HA-NEW-003).
Acceptance Criteria | No duplicate identity/serial is provisioned; each device is cryptographically unique; duplicates are detected.
Owner | Manufacturing + Product-Security · Status: Baselined · Notes: —
RCR-056 — Correct regional/firmware configuration provisioned + verified at EOL
Requirement (shall) | Correct regional and firmware configuration (band plan, market profile) shall be provisioned and verified at EOL.
Control Strategy | PROC (Level 4).
Source Hazard(s) | HA-MFG-006 (wrong region/firmware config), HA-COMM-006 (roaming/coverage).
Source Failure Mode(s) | FA-HWFMEA-032 (SIM/roaming gap) — = HAL control.
Parent Requirement | PRD-MFG/COMM; PRD-SAFE-010 (config integrity); SYS-027/035/115/116.
Implementation (DO-) | DO-MFG + DO-FW.
Verification (PV-) | PV-MFG (config-verification) + PV-COMM-006 (roaming + config).
Validation (VAL-) | VAL-COMM-006.
Independence / common-cause | Config poka-yoke bounds the per-market variant error (HA-NEW-009).
New-risk (HAL-2700) | HA-NEW-009 — per-market antenna/band variant config-error path; controlled by EOL verification + poka-yoke.
Acceptance Criteria | Every unit ships with the correct regional/firmware config, verified at EOL; a wrong-region variant is caught.
Owner | Manufacturing + Connectivity · Status: Baselined · Notes: —
6.12 Overarching governance
RCR-057 — Record-integrity chain preserves mission evidence (aggregate)
Requirement (shall) | The record-integrity chain (detection→verdict→record→export) shall preserve mission evidence such that a wrong disposition or lot-loss cannot arise from data loss or corruption (aggregate of RCR-019…024 and RCR-042).
Control Strategy | PROC (Level 4) governance aggregate; carries Level-1/2 inherent/protective members.
Source Hazard(s) | HA-MPE-001 (valid product destroyed), HA-MPE-002 (mass administration, S5), HA-MPE-004 (data-integrity at scale).
Source Failure Mode(s) | FA-SYSFMEA-012 (mass administration, S5), FA-SYSFMEA-005 — aggregate spine.
Parent Requirement | PRD-SAFE-001 and PRD-SAFE-003; SYS spine.
Implementation (DO-) | DO-SAFE-004…008 (the record-integrity spine).
Verification (PV-) | Traceability + integration test (aggregate of the spine PVs: PV-SW-001/002/008/013, PV-DATA-004/006).
Validation (VAL-) | aggregate of VAL-SW-001/002, VAL-USE-001/004.
Independence / common-cause | Aggregates the inherent (RCR-019/021), protective (RCR-020/023/024), and export (RCR-042) record controls; the two independent choke points (honest verdict + supported decision).
New-risk (HAL-2700) | HA-NEW-006 (the aggregate safety-vs-availability trade) — balanced ALARP both ways.
Acceptance Criteria | No single data-loss/corruption path produces a wrong disposition across the integrated detection→verdict→record→export chain; the S5 mass-administration top event is bounded to residual A-High with mandatory monitoring.
Owner | Chief Systems (EM-ENG-001) + System-Safety · Status: Baselined (S5 residual A-High) · Notes: The mission-integrity spine; release-critical aggregate.
RCR-058 — Overall residual + benefit-risk evaluated before release
Requirement (shall) | Overall residual risk and benefit-risk shall be evaluated per the ISO 14971 risk-management plan before release (HAL-2900).
Control Strategy | PROC (Level 4) governance gate.
Source Hazard(s) | HA-ORES-001 (overall residual), all top hazards in aggregate.
Source Failure Mode(s) | — (governance; not a failure-mode control).
Parent Requirement | PRD-SAFE-011 (residual-risk acceptability); ISO 14971.
Implementation (DO-) | Risk-Management File (HAL-2900 + Risk-Management-Summary).
Verification (PV-) | Risk-management review (design review / release board).
Validation (VAL-) | — (benefit-risk analysis).
Independence / common-cause | The overall gate that no individual control replaces.
New-risk (HAL-2700) | aggregates all accepted trades (HA-NEW-004/006/007/011/012 + latency/energy of HA-NEW-003) into the benefit-risk statement.
Acceptance Criteria | Overall residual risk is evaluated and the benefit-risk conclusion signed off per ISO 14971; the eight release conditions (Risk-Management-Summary §7) are met before release. S4/S5 region-A residuals require EM-ENG-006 + EM-ENG-001 + benefit-risk sign-off; the S5 aircraft-fire residual requires documented benefit-risk + post-market monitoring.
Owner | Quality/Regulatory (EM-ENG-006) + Chief Systems (EM-ENG-001) · Status: Baselined (conditional per Risk-Management-Summary §8) · Notes: Release gate.
---
7. Carried-Open items (validate-by-test, release-gating)
Two program-top items have correct Level-1 control placement but Evidence-Pending effectiveness and gate release (Risk-Management-Summary §3, O-1/O-2). They are carried explicitly here with test-based acceptance:
Open | RCR(s) | Why Open | Closes on | Acceptance (validate-by-test)
D3 low-temp battery (−30 °C / 90-day endurance + cold-charge safety + fade) | RCR-001, RCR-005, RCR-006, RCR-017 | Endurance margin, cold-charge boundary, and reuse-fade Evidence-Pending on the final cell datasheet + test | PRD-VER-004 + PV-ENRG-001 + UN 38.3 | 90-day cold endurance met with reserve; charge current = 0 A below the qualified limit under FW/comms fault; fade within reuse-life budget — all demonstrated by test, not asserted
RF-in-container (in-container link margin + antenna decision) | RCR-031 (+ RCR-030 consequence-bound) | In-container attenuation, cold RF margin, and integral-vs-external antenna undecided | PRD-VER-005 + PV-COMM-003 | Link closes in-container across target markets/containers, or the external-antenna fallback does; where it cannot, store-and-forward + on-receipt verdict bound the consequence and the real-time claim is withheld
No RCR asserts an unevidence-backed quantity. Every temperature limit, endurance margin, RF margin, detection target, alarm target, calibration coverage, drift bound, DR objective, and pen-test result is marked Evidence-Pending with its named closing test above and in the blocks.
8. Coverage and closure
• 58 RCRs, no duplicates. Every HAL control (HAL-2400 RCR-001…058) is carried; every safety-significant FMEA recommended action (FMEA-to-Requirements-Hazard-Matrix) is folded into an RCR as a Source Failure Mode (§5); every preliminary PRD-SAFE-001…012 intent is reconciled to ≥1 RCR (PRD-SAFE-Reconciliation.md).
• Every High/unacceptable HAL hazard has ≥1 RCR (see Risk-Control-Traceability.md); every RCR has an implementation reference (DO-) and a verification target (PV-/VAL-) — no control is on paper only (HA-CIR / HAL-2500 posture preserved).
• Fault-injection is marked for every safety-critical control (PV-…, FI) and independent adversarial verification for every security control (PV-CYBER-…, IV), per HAL-2600.
• New risks (HAL-2700) are noted on every RCR that introduces or balances one; the dominant safety-vs-availability trade (HA-NEW-006) and the fleet common-cause/connectivity trade (HA-NEW-008) are carried ALARP both directions (RCR-047/044/040; RCR-028 offline-first + RCR-037/039 rollout discipline).
Feeds | Stage 13 (Verification, PV-) — the PV- column is the verification demand set with FI/IV flags; Stage 14 (Validation, VAL-) — the VAL- column and the usability-gated use controls; Design/Product Requirements — the re-baselined PRD-SAFE (PRD-SAFE-Reconciliation.md); and ORCH-018 / EM-ENG-012 — this register is the Stage-12 segment of the end-to-end traceability spine (Risk-Control-Traceability.md).
9. Review history
• 2026-07-19 v1.0 — Initial consolidated RCR register (Stage 12, ORCH-018 / EM-ENG-012). Reconciled HAL-2400 + FMEA-to-Requirements-Hazard-Matrix + PRD-SAFE into 58 deduplicated RCRs. All PV-/VAL- Evidence-Pending (Stages 13–14); D3 battery and RF-in-container carried Open.
Risk-Control Traceability — Smart Cold Chain Monitor (SCM)
Stage 12 (ORCH-018 / EM-ENG-012) · Owner: EM-ENG-012 (Traceability) with Chief Systems (EM-ENG-001), System-Safety, and Quality/Regulatory (EM-ENG-006). The single-row-per-RCR chain table for the SCM: each Risk-Control Requirement traced through Source Hazard(s) → Source Failure Mode(s) → Design Output → Verification → Validation → Residual Risk. This is the Stage-12 segment of the end-to-end traceability spine consumed by ORCH-018 / EM-ENG-012, and the row-per-RCR complement to the Hazard-Traceability-Matrix (row-per-hazard) and the FMEA-to-Requirements-Hazard-Matrix (row-per-failure). All PV-/VAL- and residuals are Evidence-Pending (Stages 13–14); severity-first per CLAUDE.md rule 3. Regions per HAL-2200: R unacceptable · A ALARP (A-High for S5) · B broadly acceptable. FI = fault-injection (safety); IV = independent adversarial (security).
---
1. How to read
• RCR — the consolidated risk-control requirement (RCR-1000-register.md); Cls = control strategy (1 ISD / 2 PROT / 3 INFO-D-E / 4 PROC).
• Hazard (HA-) — the HAL hazard(s) it controls; ★ = S4/S5 spine hazard.
• Failure (FA-) — the ERFKL failure-mode record(s) whose recommended action this RCR is (FMEA-to-Requirements-Hazard-Matrix).
• Design Output (DO-) — the controlled implementation (DOL-1200 + disciplines).
• Verify (PV-) / Validate (VAL-) — Stage 13 / Stage 14 targets; FI/IV flagged.
• Residual — projected region/AP after control (HAL-2800), Evidence-Pending. Open where a load-bearing value is unproven (D3, RF-in-container).
2. RCR → Hazard → Failure → Design Output → Verification → Validation → Residual
2.1 Battery / charge / thermal
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
001 | 1 | ENRG-001★, ELEC-002, USE-009, →TFC-002★★ | HWFMEA-002/006, SWFMEA-012 | SAFE-001→EE-007+FW-005 | ENRG-001 (FI)+UN 38.3 | — | A Med (Open D3)
002 | 2 | ENRG-002, TFC-002★★, TFC-003 | HWFMEA-007/008/003, DFMEA-009, SYSFMEA-016 | SAFE-002→EE-005+EE-010 | ENRG-002 (FI, DO-TST-009) | — | A Med; S5→A-High
003 | 2 | ENRG-005, ELEC-002 | HWFMEA-009 | SAFE-003→EE-009+EE-010 | ELEC-002 | — | B/A
004 | 1 | ENRG-003, MECH-009, TFC-002★★, ENV-007 | DFMEA-008, SYSFMEA-016 | SAFE-010→MECH-012+EE-004 | MECH-004+UN 38.3 | — | A Med
005 | 2 | ENRG-001★, TFC-001 | (charge-safety FTA TOP-C) | SAFE-001(self-heat)→EE-007 | TFC-001+inhibit-integrity (FI) | TFC-001 | B (conditional)
006 | 1 | ENRG-006★, SW-013 | HWFMEA-001/012, SWFMEA-021 | SAFE-012→FW-004+EE-008 | PRD-VER-004+gauge | — | A Med (Open D3)
007 | 4 | ENRG-003, TFC-002★★, MPE-005 | SYSFMEA-016 | EE-015, REG-006 | UN 38.3 (T.1–T.8) | — | A Med; S5→A-High
008 | 1 | TFC-002★★, TFC-003 | DFMEA-009, SYSFMEA-016 | MECH-005, PCB-010 | TFC-002/003+UN 38.3 | — | A-High (S5)
2.2 Mechanical / Electrical / Environmental
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
009 | 1 | MECH-001/002 | DFMEA (enclosure) | MECH-002/003 | edge/pinch (IEC 62368-1) | — | B
010 | 1 | MECH-003 | DFMEA (detachment) | MECH-006+LBL | MECH-003 | — | B
011 | 1 | MECH-004/009, ENV-003 | DFMEA-001 | SAFE-010/011→MECH-005/012 | MECH(drop/vibe)+IP | — | A Med
012 | 1 | MECH-006 | DFMEA (fastener) | MECH-004 | vibe+thermal-cycle | — | B
013 | 1 | ELEC-001 | HWFMEA (elec-safety) | EE(rails), PCB | IEC 62368-1 elec | — | B
014 | 2 | ELEC-003/004 | HWFMEA-024 | EE(EMC)+PCB-010+EE-019 | ELEC-004 (immunity/emissions) | — | A Med
015 | 2 | ELEC-006/007, ENV-003, TFC-006 | DFMEA-001 | SAFE-011→MECH-005+PCB-010 | TFC-006+ENV-003 | — | A Med
016 | 4 | ELEC-008, MFG-008 | HWFMEA-023 | TST(AOI/ICT), PCB | MFG-008+HASS | — | A Med
017 | 1 | ENV-001/002, ENRG-006★ | HWFMEA-001, DFMEA-010 | SAFE-012, EE, FW | IEC 60068+PRD-VER-004 | — | A Med (Open D3)
018 | 2 | ENV-003, TFC-006 | DFMEA-001 | SAFE-011 | ENV-003+TFC-006 | — | A Med
2.3 Software / detection / verdict / record / time
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
019 | 1 | SW-001★, ENV-001, DATA-001★ | SWFMEA-001/002, DFMEA-010, SYSFMEA-001 | SAFE-004→FW-006+EE-013 | SW-001 (FI) | VAL-SW-001 | A Med
020 | 2 | SW-008★, USE-005, DATA-001★ | SWFMEA-007, HWFMEA-017, DFMEA-011/003 | SAFE-005→FW-007+EE-013 | SW-008 (FI) | VAL-USE-005 | A Med
021 | 1 | SW-002★, SW-007, DATA-004★, USE-007★ | SWFMEA-008/010/019, SYSFMEA-005 | SAFE-006→FW-008+CLD-006+HF-002 | SW-002 (FI) | VAL-SW-002, VAL-USE-001 | A Med
022 | 3 | SW-013★, DATA-004★, USE-004★ | SWFMEA-008/018 | SAFE-006/008 | SW-002/DATA-004 (FI)+SW-013 | VAL-USE-004 | A Med
023 | 2 | SW-009 | HWFMEA-016, SWFMEA-028/029 | SAFE-007→FW-003/013 | SW-009 (FI) | — | A Med
024 | 2 | SW-013★, ELEC-008/010, DATA-006 | SWFMEA-018/019, HWFMEA-033 | SAFE-008→FW-010+EE-002 | SW-013 (FI)+endurance | — | A Med
025 | 2 | SW-005, DATA-003, COMM-005 | HWFMEA-014, SWFMEA-032 | EE-011(RTC)+FW | SW-005 (drift/cold) | — | A Med
026 | 2 | SW-014, SW-004, MFG-001, SVC-001, DATA-001★ | SWFMEA-005, HWFMEA-018, DFMEA-010 | FW-007+SAFE-007+FW-009 | SW-004 (FI)+DATA-001 | — | A Med (CCF)
027 | 2 | SW-013★, SW-009, ENRG-006★ | SWFMEA-029, HWFMEA-016 | SAFE-006/007→FW-012 | SW-009 (FI) | — | A Med
2.4 Connectivity / cloud
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
028 | 1 | COMM-001★, COMM-002★ | SWFMEA-024, IFMEA-011/010, SYSFMEA-002 | SAFE-009→FW-011 | COMM-001/002 | VAL-COMM-001 | A Med
029 | 2 | COMM-001★, CLOUD-001, USE-002★ | SWFMEA-023/024, IFMEA-011/020, SYSFMEA-002 | SAFE-009→CLD-007/008 | SW-006 (FI)+COMM-001 | VAL-SW-006, VAL-USE-002 | A Med
030 | 2 | COMM-001★, COMM-003★ | SWFMEA-024, HWFMEA-029, IFMEA-005 | SAFE-009 | COMM-001 | VAL-COMM-001 | A Med
031 | 1 | COMM-003★, COMM-006 | HWFMEA-029/030/032, IFMEA-005, SYSFMEA-002 | IF(antenna)+EE-017/018 | PRD-VER-005+COMM-003 | VAL-COMM-003 | A Med (Open RF)
032 | 2 | CLOUD-001/002/004, COMM-002★ | IFMEA-011/010, SWFMEA-032, SYSFMEA-005 | CLD-007 (SRS-CLOUD-023) | COMM-002/CLOUD-001 (FI) | VAL-COMM-002 | A Med
033 | 2 | CLOUD-003, COMM-008, DATA-002 | SYSFMEA-004 | CLD(registry)+SEC | COMM-008/CLOUD-003+MFG-005 | VAL-COMM-008/USE-008 | A Med
2.5 Cybersecurity-induced safety (IV)
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
034 | 2 | CYBER-001★, DATA-001★, MPE-004 | SWFMEA-031, SYSFMEA-005 | SEC(integrity)+SAFE-008 | CYBER-001 (IV) | — | A Med
035 | 2 | CYBER-002★, COMM-002★ | SWFMEA-033 | SEC+SAFE-009 | CYBER-002 (IV) | — | A Med
036 | 2 | CYBER-003, USE-003, CLOUD-002 | IFMEA-016 | SEC+CLD(audit) | CYBER-003 (IV) | — | A Med
037 | 2 | SW-010★, CYBER-005★, CLOUD-007 | SWFMEA-026, IFMEA-015 | SAFE-014→FW-014/015+CLD-010 | SW-010+CYBER-005 (IV) | — | A Med/High
038 | 2 | CYBER-006, CYBER-008, SVC-006 | IFMEA-015/016, SWFMEA-034 | SEC-006+FW | CYBER-006 (IV) | — | A Med
039 | 2 | CYBER-009★, CLOUD-004/006/008 | SWFMEA-032, SYSFMEA-005 | CLD+SEC | CYBER-009 (IV)+DR | — | A Med
2.6 Data & decision / human factors
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
040 | 3 | DATA-004★, DATA-005, USE-007★ | SWFMEA-008, SYSFMEA-008 | SAFE-006+HF-002 | DATA-004/005 | VAL-USE-004, VAL-USE-005 | A Med
041 | 3 | DATA-007, COMM-003★ | (provenance recs) | CLD+HF | DATA-007 | VAL-USE-005 | A Med
042 | 2 | DATA-006, MPE-004 | SYSFMEA-005 | CLD(API/audit) | DATA-006+traceability | — | A Med
043 | 1 | USE-001★, USE-008, ENV-002 | SWFMEA-010, IFMEA-019 | SAFE-013+HF-002/005 | USE-001+display (FI) | VAL-USE-001 | A Med
044 | 1 | USE-007★, IUE-007, MPE-001 | SYSFMEA-008, USE-006 | HF+CLD(audit) | DATA-008 | VAL-USE-007 | A Med
045 | 2 | USE-003/004★, SW-003, SVC-005 | SWFMEA-004/011, HWFMEA-017, USE-001, DFMEA-011/003 | SAFE-007+FW-032 | SW-003 (FI) | VAL-USE-004 | A Med
046 | 3 | USE-009, ENRG-001★ | (state-legibility recs) | SAFE-013→HF-005+EE-008 | USE(state-legibility) | VAL-USE | A Med
047 | 2 | USE-002★, SW-006, MPE-003★ | SYSFMEA-013, IFMEA-020 | SAFE-004+CLD-003 | USE-002 | VAL-USE-002 | A Med (accepted trade)
048 | 2 | USE-003, SW-003 | SWFMEA-002/004 | APP+CLD | SW-003 | VAL-USE-003 | A Med
2.7 Service / manufacturing / governance
RCR | Cls | Hazard (HA-) | Failure (FA-) | Design Output (DO-) | Verify (PV-) | Validate | Residual
049 | 2 | SVC-001★, MFG-001 | SYSFMEA-009 | FW-009+SVC | SVC-001+SW-004 | — | A Med (CCF)
050 | 4 | SVC-003/005, SVC-002 | (service-restore recs), SWFMEA-026 | SVC+TST | SVC (post-service BIST) | — | A Med
051 | 2 | SVC-006, CYBER-006 | SWFMEA-034 | SEC+SVC | CYBER-006 (IV) | — | A Med
052 | 3 | SVC-007, CYBER-008, MPE-005 | (disposal recs) | SVC+LBL+SEC | disposal+key-erasure review | — | A Med
053 | 4 | MFG-001★, SVC-001 | SYSFMEA-009 | TST+MFG | MFG-001 (cal coverage) | — | A Med (CCF)
054 | 4 | MFG-002/003, ENRG-002 | HWFMEA-004/008/003 | MFG+TST+CFG-004 | MFG-002/003 (AOI/poka-yoke) | — | A Med
055 | 4 | MFG-005, CYBER-006 | SYSFMEA-004, IFMEA-016 | MFG(provisioning)+SEC | MFG-005 (identity) | — | A Med
056 | 4 | MFG-006, COMM-006 | HWFMEA-032 | MFG+FW | MFG(config)+COMM-006 | VAL-COMM-006 | A Med
057 | 4 | MPE-001, MPE-002★★(S5), MPE-004 | SYSFMEA-012, SYSFMEA-005 | SAFE-004…008 (spine) | aggregate spine PVs+integration | aggregate VAL-SW/USE | A-High (S5)
058 | 4 | ORES-001, all top | — (governance) | Risk-Management File (HAL-2900) | risk-management review | benefit-risk | conditional
3. Residual-risk roll-up (from HAL-2800/2900, projected)
Class | RCRs | Projected residual | Note
Physical energy/battery/thermal | 001–008, 054 | A Med; S5 (HA-TFC-002 aircraft-hold fire) → A-High permanent | Single-fault-safe, non-defeatable HW; UN 38.3; mandatory thermal-event monitoring
Information pathway (dominant) | 019–027, 040, 043–048, 057 | A Med (≈20 S4 residuals) | Two independent choke points (honest verdict + supported decision); not yet demonstrated — closes on PV-SW-002/PV-DATA-004 + VAL-USE-001/004
Connectivity/cloud | 028–033 | A Med; RF-in-container Open | Local-commit + S&F + ack/escalation; O-1 driver = PRD-VER-005
Security→safety | 034–039, 051, 055 | A Med (A-Med/High for OTA S5) | Device-anchored authenticity, non-suppressible alarm, secure OTA/boot; pending IV
Use/human-factors | 043–048, 040 | A Med, usability-gated | Closable only by summative usability (VAL-USE-001/002/004/007)
Accepted public-health trade | 044, 047, 057 | A Med, balanced not eliminated | False-reject/wastage/stock-out (HA-MPE-003) monitored as a first-class metric (HA-NEW-006)
Physical lower-severity (S1–S2) | 009–013 | B (broadly acceptable) | Inherent/protective
D3 battery Open | 001, 005, 006, 017 | A Med Open | Endurance/charge-safety/fade → PRD-VER-004 + PV-ENRG-001
4. Coverage assertions (release-gate contribution)
• Every High/unacceptable HAL hazard maps to ≥1 RCR. The ~60 top hazard rows of the Hazard-Traceability-Matrix each appear in §2 under their controlling RCR(s); no R-region hazard is left without a control.
• Every RCR has a Design Output and a Verification target (columns above); no control is on paper only.
• Fault-injection is planned for every safety-critical RCR (FI marked) and independent adversarial verification for every security RCR (IV marked) — the HAL-2600 discipline carried into Stage 12.
• Two items carried Open (mirrors HAL/ERFKL): RF-in-container (RCR-031 → PRD-VER-005) and D3 cold endurance (RCR-006/017 → PRD-VER-004); the S5 rows (RCR-008/057) stay A-High by rule with mandatory monitoring.
• All residuals are projected / Evidence-Pending until Stages 13–14 return evidence.
5. Handoff
Stage 13 (Verification, PV-) | the Verify column (with FI/IV flags) is the verification demand set.
Stage 14 (Validation, VAL-) | the Validate column and the usability-gated use controls.
Product/Design Requirements | the re-baselined PRD-SAFE (PRD-SAFE-Reconciliation.md).
ORCH-018 / EM-ENG-012 | this table is the Stage-12 segment of the end-to-end spine (Requirement → Function → Failure Mode → Cause → Risk → Mitigation/RCR → Design Output → Verification → Validation → Residual → Field), joining the Hazard-Traceability-Matrix (Stage 10) and the FMEA-to-Requirements-Hazard-Matrix (Stage 11).
Review history | 2026-07-19 v1.0 initial (Stage 12).

## 13 Verification
Verification
The plan to prove, by test and analysis, that the product meets every promise.
↩ Contents
In this section:
• Product Verification Library (PVL) — Smart Cold Chain Monitor (SCM)
• PVL-400 (PV-ANAL) — Analysis Verification — Smart Cold Chain Monitor (SCM)
• PVL-1500 (PV-COMP) — Compliance Verification — Smart Cold Chain Monitor (SCM)
• PVL-1700 (PV-DEF) — Defect Management — Smart Cold Chain Monitor (SCM)
• PVL-800 (PV-EMC) — EMC & Radio Verification — Smart Cold Chain Monitor (SCM)
• PVL-700 (PV-ENV) — Environmental Verification — Smart Cold Chain Monitor (SCM)
• PVL-1800 (PV-EVID) — Verification Evidence — Smart Cold Chain Monitor (SCM)
• PVL-500 (PV-FUNC) — Functional Verification — Smart Cold Chain Monitor (SCM)
• PVL-000 (PV-GOV) — Verification Governance — Smart Cold Chain Monitor (SCM)
• PVL-1200 (PV-HW) — Hardware Verification — Smart Cold Chain Monitor (SCM)
• PVL-300 (PV-INSP) — Inspection Verification — Smart Cold Chain Monitor (SCM)
• PVL-1300 (PV-MFG) — Manufacturing Verification — Smart Cold Chain Monitor (SCM)
• PVL-600 (PV-PERF) — Performance Verification — Smart Cold Chain Monitor (SCM)
• PVL-100 (PV-PLAN) — Verification Planning — Smart Cold Chain Monitor (SCM)
• PVL-1400 (PV-REL) — Reliability Verification — Smart Cold Chain Monitor (SCM)
• PVL-200 (PV-REQ) — Requirements Verification — Smart Cold Chain Monitor (SCM)
• PVL-900 (PV-SAFE) — Safety Verification — Smart Cold Chain Monitor (SCM)
• PVL-1000 (PV-SEC) — Cybersecurity Verification — Smart Cold Chain Monitor (SCM)
• PVL-1900 (PV-SUM) — Verification Summary — Smart Cold Chain Monitor (SCM)
• PVL-1100 (PV-SW) — Software Verification — Smart Cold Chain Monitor (SCM)
• PVL-1600 (PV-TRACE) — Traceability & Coverage — Smart Cold Chain Monitor (SCM)
• Verification Plan — Smart Cold Chain Monitor (SCM)
• Verification Traceability & Coverage Matrix — Smart Cold Chain Monitor (SCM)
Product Verification Library (PVL) — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 13 — Verification (ORCH-014) · Owner/Lead: Principal Verification/Reliability/Test/Quality Engineer (EM-ENG-006) · Integrator: Chief Systems (EM-ENG-001) · Traceability: ORCH-018 / EM-ENG-012 · Status: Authored playbook — all activities Planned / Evidence-Pending (pre-execution; no results fabricated).
This library is the SCM's verification playbook: it defines everything required to prove the design outputs satisfy the design inputs — method (Test/Analysis/Inspection/Demonstration), objective acceptance criteria, traceability, independence, and evidence required — but it does not store executed test reports (those are verification evidence, referenced from PVL-1800). It follows 00-governance/product-verification-standard.md and the 20-module map, and is based on 00-concept/concept-and-decisions.md (D1–D4), the Stage-05/06/07 requirement baselines, the Stage-10 Hazard Analysis, the Stage-11 ERFKL, and the Stage-12 Risk-Control Requirements (RCR-001…058).
---
Product in one line
A reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shippers to log temperature/humidity/shock, detect two-sided excursions, alert in time, give an offline accept/reject verdict at the door, and provide an auditable, tamper-evident cold-chain record. Operating range −30…+60 °C; endurance target 90 days. The two top program risks — the D3 −30 °C/90-day battery tension and the RF-in-container link — are carried as explicit, release-gating validate-by-test items.
---
Module table (20 modules · 195 verification activities)
Module | Code | File | Activities | Focus
PVL-000 | PV-GOV | [PV-GOV-000-governance.md](PV-GOV-000-governance.md) | 6 | Policy, independence, config control, regression, deviation, gates
PVL-100 | PV-PLAN | [PV-PLAN-100-planning.md](PV-PLAN-100-planning.md) | 6 | Strategy, coverage matrix, phasing, resources, dependencies
PVL-200 | PV-REQ | [PV-REQ-200-requirements.md](PV-REQ-200-requirements.md) | 6 | Requirement testability, completeness, allocation, RCR coverage
PVL-300 | PV-INSP | [PV-INSP-300-inspection.md](PV-INSP-300-inspection.md) | 10 | Drawing/PCB/code/label/packaging/supplier inspection
PVL-400 | PV-ANAL | [PV-ANAL-400-analysis.md](PV-ANAL-400-analysis.md) | 14 | Tolerance/thermal/power/battery/RF/worst-case/reliability/MSA/threat
PVL-500 | PV-FUNC | [PV-FUNC-500-functional.md](PV-FUNC-500-functional.md) | 30 | Detection/verdict, alarm+delivery, logging, OTA, comms, S&F, data, use, service (hosts PV-COMM/CLOUD/DATA/USE/SVC)
PVL-600 | PV-PERF | [PV-PERF-600-performance.md](PV-PERF-600-performance.md) | 12 | Accuracy, latency, responsiveness, boot, storage, throughput, charge time
PVL-700 | PV-ENV | [PV-ENV-700-environmental.md](PV-ENV-700-environmental.md) | 13 | −30…+60, thermal/condensation cycling, IP67, drop/shock/vibe, altitude, UV (hosts PV-MECH)
PVL-800 | PV-EMC | [PV-EMC-800-emc.md](PV-EMC-800-emc.md) | 6 | Emissions, immunity (decision-path integrity), coexistence, radio FCC/CE (hosts PV-ELEC-004)
PVL-900 | PV-SAFE | [PV-SAFE-900-safety.md](PV-SAFE-900-safety.md) | 11 | HAL risk-control verification, battery/thermal/charge FI, single-fault, record-integrity spine (hosts PV-ENRG/PV-TFC)
PVL-1000 | PV-SEC | [PV-SEC-1000-cybersecurity.md](PV-SEC-1000-cybersecurity.md) | 11 | Independent pen-tests (auth, OTA, secure boot, cloud/DR, integrity), TLS, secure storage, SBOM, keys (hosts PV-CYBER)
PVL-1100 | PV-SW | [PV-SW-1100-software.md](PV-SW-1100-software.md) | 15 | Safety software fault-injection (detection/verdict/gap/safe-state/power-fail/cal/time/alert/pre-arm/OTA), static/coverage/timing/soak
PVL-1200 | PV-HW | [PV-HW-1200-hardware.md](PV-HW-1200-hardware.md) | 9 | Rails/charge/protection/sensors/connectors/SI-PI/observability (hosts PV-ELEC-002)
PVL-1300 | PV-MFG | [PV-MFG-1300-manufacturing.md](PV-MFG-1300-manufacturing.md) | 8 | EOL calibration, incoming/AOI, poka-yoke, identity, config, latent-defect coverage, FAI/Cpk
PVL-1400 | PV-REL | [PV-REL-1400-reliability.md](PV-REL-1400-reliability.md) | 10 | D3 90-day cold endurance (PRD-VER-004) + RF-in-container (PRD-VER-005), HALT/HASS, battery aging, flash endurance, MTBF
PVL-1500 | PV-COMP | [PV-COMP-1500-compliance.md](PV-COMP-1500-compliance.md) | 12 | WHO PQS, EU GDP/Part 11, US FDA, UN 38.3, IATA, radio, IEC 62368-1/62304, EMC, IP, RoHS/REACH, ISO 14971 file
PVL-1600 | PV-TRACE | [PV-TRACE-1600-traceability.md](PV-TRACE-1600-traceability.md) | 4 | Coverage chains + gap analysis + metrics
PVL-1700 | PV-DEF | [PV-DEF-1700-defect-management.md](PV-DEF-1700-defect-management.md) | 4 | Defect capture, RCA/CAPA, regression, waiver/closure
PVL-1800 | PV-EVID | [PV-EVID-1800-evidence.md](PV-EVID-1800-evidence.md) | 4 | Protocols, raw data, calibration certs, review/approval/retention
PVL-1900 | PV-SUM | [PV-SUM-1900-summary.md](PV-SUM-1900-summary.md) | 4 | Coverage/pass-rate, open issues, residual risk, release recommendation
Total |  | 20 modules | 195
Companion documents: [Verification-Plan.md](Verification-Plan.md) (strategy, DVT/PVT phasing, sequencing, resources, gates) · [Verification-Traceability-Coverage-Matrix.md](Verification-Traceability-Coverage-Matrix.md) (Requirement/RCR → PV → method → acceptance → status; gap analysis).
---
ID scheme
Verification activities use PV-<CODE>-NNN. To honour the Stage-12 RCR verification references, the thematic IDs those RCRs already cite are preserved as the actual activity IDs and hosted in the module whose scope matches: PV-ENRG-/PV-TFC- (battery/thermal safety) in PV-SAFE; PV-SW- in PV-SW; PV-CYBER- in PV-SEC; PV-COMM-/PV-CLOUD-/PV-DATA-/PV-USE-/PV-SVC- in PV-FUNC; PV-MECH- in PV-ENV; PV-ELEC-002 in PV-HW / PV-ELEC-004 in PV-EMC; PV-MFG- in PV-MFG; the D3/RF headline tests (PRD-VER-004/005) as PV-REL-001/PV-REL-005. Each host module documents the IDs it carries.
---
Applicability Register
All 20 modules are Applicable — the SCM is a globally distributed, regulated, reusable, connected, safety- significant IoT device with device + cloud + mobile tiers. No module is excluded.
Module | Status | Note
All 20 (PV-GOV … PV-SUM) | Applicable | Every module carries real, product-specific verification activities.
PV-SAFE (900) | Applicable — reconcile w/ HAL | Each activity is a HAL hazard-record Verification Reference (PRD-VER-007); gates residual-risk closure.
PV-COMP (1500) | Applicable — Evidence-Pending items | WHO PQS category, US classification, target-market set, test houses Evidence-Pending (concept §7 / D4).
PV-REL (1400) | Applicable — 2 Open items | D3 90-day endurance (O-2) and RF-in-container (O-1) carried Open; release-gating.
Borderline within modules (confirm, not excluded): self-heat FI (PV-TFC-001) conditional on CMP-013 adoption; RF-exposure likely MPE not SAR (PV-EMC-005); several numeric limits Evidence-Pending with named closing tests.
---
Coverage note
• 58/58 RCRs (Stage 12) have ≥1 matching PV activity — no control on paper only (Verification-Traceability-Coverage-Matrix §2).
• Fault injection is provided for every safety-critical control (PV-ENRG/PV-TFC/PV-SW FI set, PV-SAFE-003/004); independent adversarial verification for every security control (PV-CYBER-).
• PRD-VER-004 (D3 endurance/charge) → PV-REL-001 + PV-ENRG-001; PRD-VER-005 (RF-in-container) → PV-REL-005 — both carried Open and release-gating.
• Every applicable PRD/SYS/SRS/HWR/MER/FWR/IFR requirement maps to a PV module (matrix §3). No unverified applicable requirement; Evidence-Pending numeric targets carry their named closing test and are never invented.
Evidence & honesty
Per evidence-standard.md / hallucination-policy.md: standards are cited by designation, revision to confirm (no clause asserted as fact); acceptance criteria are objective/measurable; unresolved quantities are Evidence-Pending with the named closing test; and no results are recorded — this is a pre-execution playbook.
Chain position & feeds
Flow: PRD (05) → SYS (06) → Subsystem/SRS (07) → Design Outputs (09) → Hazard Analysis (10) / ERFKL (11) → Risk-Control Requirements (12) → Verification (13, this library) → Validation (14) → … → Traceability (18). Feeds Validation (Stage 14, PVaL / ORCH-015) (VAL- targets, usability-gated use controls), Release (Stage 16) (coverage + open-item evidence), and ORCH-018 / EM-ENG-012 (Stage 18) (the verification segment of the end-to-end traceability spine). Safety verification reconciles with HAL (Stage 10).
Convening: ORCH-014 (Verification). Lead: EM-ENG-006. Integrator: EM-ENG-001. Discipline owners per module.
PVL-400 (PV-ANAL) — Analysis Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with the discipline principals (Power, Thermal, RF, PCB, Reliability, Firmware). Verifies by calculation, modelling, worst-case, and simulation those characteristics that are proven by analysis (often as a predicate to, or in place of where justified, a physical test). Analysis supporting a safety-critical claim is always confirmed by a downstream Test (severity-first, PRD-VER-003). All entries Planned/Evidence-Pending.
1. Purpose
Verify by analysis: tolerance stack, thermal, power/energy budget, worst-case circuit, battery (endurance, cold-capacity, fade, charge-safety envelope), RF/network link budget, storage-capacity sizing, reliability prediction, EMC prediction, and measurement-system capability (MSA).
2. Scope
System- and subsystem-level analyses across S1–S9, especially the D3 energy/battery tension and the RF-in-container link budget (both closed by test — PRD-VER-004/005). Excludes the physical tests those analyses predict (PVL-600/700/800/1400).
3. Applicable Standards & References
IEC 62368-1 (thermal/touch-temp limits); IEC 60068 (environmental analysis basis); IEC 61000 (EMC prediction); UN 38.3 (battery-safety basis); ISO/IEC 17025 & AIAG/VDA MSA (measurement capability); reliability-prediction methods (e.g., Telcordia SR-332 / MIL-HDBK-217-class, method to confirm); friis/link- budget methods for RF. Designations; revisions to confirm.
4. Governance & Method
Method is Analysis (calculation/model/worst-case/Monte-Carlo/simulation). Each analysis states its assumptions, inputs (datasheet / released requirement), method, and result with margin. Analyses feeding a safety claim name the confirming Test.
5. Verification Categories
Tolerance · thermal · power/energy budget · worst-case circuit · battery/endurance · RF/link budget · storage/capacity · reliability prediction (MTBF/FIT) · EMC prediction · MSA/measurement capability.
6. Verification Register (this module)
PV-ANAL-001 — Mechanical tolerance-stack analysis (seal / probe / fit)
Objective | Confirm by tolerance-stack analysis that seal groove, probe gland, HMI window, and enclosure interfaces maintain fit, sealing, and clearance across the tolerance range and −30…+60 °C.
Method | Analysis
I/O Trace | MER-mechanical; DO-MECH-018 (stack); DO-VER-002
Risk-Control Trace | RCR-011 (IP/containment), RCR-015 (seal)
Acceptance Criteria | Worst-case stack maintains seal compression and clearance with positive margin across tolerance and thermal expansion; result feeds the IP test (PV-ENV-001).
Level | Subsystem
Sample/Config | GD&T model + material CTE data
Facility/Equipment/Tools | Tolerance-analysis tool; CAD
Independence | Mechanical (EM-ENG-002); Quality reviews
Standard/Compliance Reference | ASME Y14.5; IEC 60529 (target)
Evidence Required | Tolerance-stack report
Owner | Mechanical (EM-ENG-002)
Status | Planned
Notes | Confirmed by PV-ENV-001 (IP) and PV-ENV drop/vibe.
PV-ANAL-002 — System thermal analysis (−30…+60 °C, self-heating, touch temperature)
Objective | Confirm by thermal model that all components stay within rated limits and touchable surfaces (incl. any self-heat element) stay below the IEC 62368-1 burn threshold across the range.
Method | Analysis
I/O Trace | SYS-069/071; DO-MECH-017; DO-SAFE-001 (self-heat)
Risk-Control Trace | RCR-005 (self-heat touch-temp bound), RCR-002 (thermal margins)
Acceptance Criteria | All parts within datasheet temperature limits at +60 °C worst case; touch temperatures below IEC 62368-1 limits; self-heat element bounded; feeds PV-TFC-001 (touch-temp test).
Level | System
Sample/Config | Thermal model + power map
Facility/Equipment/Tools | Thermal-simulation tool
Independence | Thermal/Mechanical
Standard/Compliance Reference | IEC 62368-1 touch-temperature
Evidence Required | Thermal-analysis report
Owner | Thermal/Battery-Safety
Status | Planned
Notes | Confirmed by PV-TFC-001.
PV-ANAL-003 — Power / energy budget analysis for 90-day cold endurance (D3)
Objective | Confirm by energy-budget model that the duty-cycled design meets ≥90 days armed operation from a full charge under the worst-case −30 °C cold-duty profile, with documented margin and reserve.
Method | Analysis
I/O Trace | SYS-045/046/086/090; PRD-PERF-005; DO-SYS-007 (energy); DO-FW-004 (governor); D3
Risk-Control Trace | RCR-006 (energy governor + reserve), RCR-017 (endurance)
Acceptance Criteria | Modelled endurance ≥ 90 days with reserve preserved for record-commit + low-battery alert under the worst-case cold-duty and connectivity profile; margin % stated. The result is a prediction — the load-bearing claim is Evidence-Pending and closed by test PV-REL-001 / PRD-VER-004.
Level | System
Sample/Config | Energy model (cell datasheet + duty profile)
Facility/Equipment/Tools | Energy-budget spreadsheet/model
Independence | Power (with Firmware); Reliability reviews
Standard/Compliance Reference | cell datasheet; D3
Evidence Required | Power-budget analysis report
Owner | Battery/Power (with Firmware EM-ENG-007)
Status | Evidence-Pending
Notes | D3 headline; analysis predicts, PV-REL-001 (90-day) confirms. Final cell datasheet pending.
PV-ANAL-004 — Low-temperature battery capacity & cycle-fade analysis
Objective | Confirm by analysis of the selected chemistry the usable capacity at −30 °C and capacity fade over the reuse-cycle life, feeding the endurance budget and reserve.
Method | Analysis
I/O Trace | SYS-046; PRD-REL-001/002/003; DO-EE-003 (cell); D3
Risk-Control Trace | RCR-006/017 (cold endurance), RCR-001 (charge boundary)
Acceptance Criteria | Predicted −30 °C usable capacity and end-of-life fade keep the endurance budget positive; the qualified lower charge limit is derived. Evidence-Pending on final cell datasheet; confirmed by PV-REL-004 (battery aging) and PV-REL-001.
Level | Subsystem
Sample/Config | Cell datasheet + cycle model
Facility/Equipment/Tools | Battery model
Independence | Battery/Power; Reliability reviews
Standard/Compliance Reference | cell datasheet; D3
Evidence Required | Battery capacity/fade analysis
Owner | Battery/Power
Status | Evidence-Pending
Notes | Underpins RCR-001 charge-limit value.
PV-ANAL-005 — Worst-case circuit analysis (charge path, protection, rails)
Objective | Confirm by worst-case (tolerance/temperature/aging) circuit analysis that the charge path, protection thresholds, SELV rails, and brown-out barrier meet limits under all conditions.
Method | Analysis
I/O Trace | HWR (charge/protection/rails); DO-EE-005/007/009/010
Risk-Control Trace | RCR-002/003 (protection/charge-source), RCR-013 (SELV), RCR-024 (brown-out barrier)
Acceptance Criteria | Protection trip thresholds, charge inhibit boundary, SELV limits, and brown-out barrier hold across worst-case tolerance/temperature/aging with margin; feeds PV-ENRG-001/002, PV-ELEC-002.
Level | Subsystem
Sample/Config | Circuit models + component tolerances
Facility/Equipment/Tools | SPICE / worst-case tool
Independence | Electrical (EM-ENG-003)
Standard/Compliance Reference | IEC 62368-1
Evidence Required | Worst-case analysis report
Owner | Electrical/Power (EM-ENG-003)
Status | Planned
Notes | Predicate for the battery-safety fault-injection tests.
PV-ANAL-006 — RF / network in-container link-budget analysis
Objective | Confirm by link-budget analysis the achievable LTE-M/Wi-Fi/BLE/GNSS margins inside a representative metallized/insulated shipper, and whether an external antenna is required.
Method | Analysis
I/O Trace | SYS-080/104/109/117; PRD-CONN-007; DO-EE-017/018 (antenna); AD-1
Risk-Control Trace | RCR-031 (RF closes budget or external-antenna fallback)
Acceptance Criteria | Predicted in-container link margin per band across target containers; integral-vs- external antenna recommendation derived. Evidence-Pending — the principal open architectural risk; closed by test PV-REL-005 / PRD-VER-005.
Level | System
Sample/Config | RF model + container attenuation estimates
Facility/Equipment/Tools | Link-budget model; antenna sim
Independence | RF/Antenna + Connectivity
Standard/Compliance Reference | Friis/link-budget; FCC/CE band plans
Evidence Required | Link-budget analysis
Owner | RF/Antenna + Connectivity
Status | Evidence-Pending
Notes | RF-in-container Open item (O-1); analysis predicts, PRD-VER-005 confirms.
PV-ANAL-007 — Log-storage capacity sizing analysis (offline 90-day trip)
Objective | Confirm by sizing analysis that non-volatile storage holds a full 90-day multi-channel trip at the minimum interval with ≥50 % spare, no overwrite of unsent data.
Method | Analysis
I/O Trace | SYS-043/127; PRD-PERF-006; DO-FW-010; SRS-DATA (900)
Risk-Control Trace | RCR-024 (no overwrite of unsent), RCR-030 (buffer depth)
Acceptance Criteria | Computed worst-case record size ≤ 50 % of NVM capacity; buffer depth covers the worst-case offline duration. Storage sizing Evidence-Pending (SYS-127); confirmed by PV-PERF-007.
Level | Subsystem
Sample/Config | Record-format spec + channel count
Facility/Equipment/Tools | Sizing model
Independence | Firmware; Systems reviews
Standard/Compliance Reference | —
Evidence Required | Storage-sizing analysis
Owner | Firmware (EM-ENG-007)
Status | Evidence-Pending
Notes | —
PV-ANAL-008 — Reliability prediction (MTBF / FIT) and allocation
Objective | Predict system MTBF/FIT and allocate reliability budgets to subsystems as the basis for the reliability-demonstration plan.
Method | Analysis
I/O Trace | PRD-REL; SYS-163…176; FAL-1900 (reliability analysis)
Risk-Control Trace | RCR-016 (latent-defect basis), RCR-023 (fault handling)
Acceptance Criteria | Predicted MTBF meets the reliability target with allocation to subsystems; dominant failure contributors identified for HALT/HASS focus. Target Evidence-Pending; confirmed by PV-REL-006.
Level | System
Sample/Config | BOM + stress data
Facility/Equipment/Tools | Reliability-prediction tool
Independence | Reliability; Quality reviews
Standard/Compliance Reference | reliability-prediction method (to confirm)
Evidence Required | Reliability-prediction report
Owner | Reliability
Status | Evidence-Pending
Notes | —
PV-ANAL-009 — EMC pre-compliance / emissions prediction analysis
Objective | Predict conducted/radiated emissions and identify immunity risk areas (detection/verdict path) before formal EMC test.
Method | Analysis
I/O Trace | HWR (EMC); DO-EE (EMC), DO-PCB-010
Risk-Control Trace | RCR-014 (EMC immunity — no false/missed excursion; emissions limits)
Acceptance Criteria | Predicted emissions within FCC Part 15 / CE-RED margins; immunity-sensitive nets (probe front-end, detection path) identified and mitigated; feeds PV-EMC / PV-ELEC-004.
Level | Subsystem
Sample/Config | PCB layout + models
Facility/Equipment/Tools | EMC-simulation / pre-scan
Independence | Electrical/EMC
Standard/Compliance Reference | CISPR/IEC 61000; FCC Part 15; CE RED
Evidence Required | EMC prediction / pre-scan report
Owner | Electrical/EMC
Status | Planned
Notes | —
PV-ANAL-010 — Measurement-system analysis (MSA) for temperature/RH verification
Objective | Confirm the verification measurement systems (temperature/RH references, chambers) are capable of resolving the ±0.5 °C / ±5 %RH acceptance limits.
Method | Analysis
I/O Trace | DO-VER-003/008; SYS-033/036; PRD-PERF-001/008
Risk-Control Trace | RCR-053 (calibration accuracy basis)
Acceptance Criteria | Measurement-system uncertainty is small relative to the ±0.5 °C / ±5 %RH tolerance (gauge R&R / uncertainty budget within the accepted fraction); references ISO/IEC 17025-traceable.
Level | System
Sample/Config | Reference standards + chamber
Facility/Equipment/Tools | Calibrated references; MSA method
Independence | Metrology/Quality
Standard/Compliance Reference | AIAG/VDA MSA; ISO/IEC 17025
Evidence Required | MSA / uncertainty-budget report
Owner | Quality/Metrology
Status | Planned
Notes | Underwrites every accuracy test's credibility.
PV-ANAL-011 — Single-fault / FMEA-consistency safety analysis
Objective | Confirm by analysis that no single fault defeats a safety function (charge-inhibit, protection, verdict integrity, alert path) and that the design matches the FMEA/FTA assumptions.
Method | Analysis
I/O Trace | SYS-safety; FAL-1800 (FTA), FAL-1300/1400 (FMEA)
Risk-Control Trace | RCR-001/002/021/022 (single-fault sets), RCR-057 (record-integrity spine)
Acceptance Criteria | Every S4/S5 safety function is shown single-fault-tolerant or has an independent backstop; analysis is consistent with the FTA top events; findings feed the fault-injection tests.
Level | System
Sample/Config | FTA/FMEA + design
Facility/Equipment/Tools | FTA tool
Independence | System-Safety; independent of design authors
Standard/Compliance Reference | ISO 14971; IEC 62368-1 single-fault
Evidence Required | Single-fault analysis report
Owner | System-Safety (EM-ENG-006)
Status | Planned
Notes | Defines the fault set that PV-ENRG/PV-SW fault-injection tests must force.
PV-ANAL-012 — Cybersecurity threat / attack-surface analysis
Objective | Confirm by threat modelling that the attack surface (BLE, cellular, cloud, OTA, debug) is enumerated and each safety-relevant threat has a control, feeding the pen-test scope.
Method | Analysis
I/O Trace | SRS-SEC (600); PRD-SEC; DO-SEC-
Risk-Control Trace | RCR-034…039, RCR-051 (all security controls)
Acceptance Criteria | Threat model covers all interfaces; every safety-relevant threat maps to a control and a PV-CYBER pen-test; residual threats documented. Feeds independent pen-test scoping (PV-CYBER).
Level | System
Sample/Config | Architecture + interfaces
Facility/Equipment/Tools | Threat-modelling method (e.g., STRIDE)
Independence | Product-Security (EM-ENG-008)
Standard/Compliance Reference | IEC 62443 concepts; OWASP
Evidence Required | Threat model / attack-surface report
Owner | Product-Security (EM-ENG-008)
Status | Planned
Notes | Scopes the IV pen-tests (PV-CYBER-001…009).
PV-ANAL-013 — Detection-algorithm worst-case / boundary analysis
Objective | Confirm by analysis that the two-sided excursion-detection and cumulative-exposure logic has no undetected-excursion boundary case and that freeze reliability equals heat.
Method | Analysis
I/O Trace | SYS-149/150; PRD-FUN-006/007; DO-FW-006; DO-VER-007 (reference profiles)
Risk-Control Trace | RCR-019 (two-sided detection), RCR-048 (profile anchoring)
Acceptance Criteria | No input profile in the analysed space yields an undetected excursion; cumulative criterion correct at boundaries; feeds the profiled-excursion test PV-SW-001.
Level | Subsystem
Sample/Config | Algorithm spec + stability profiles
Facility/Equipment/Tools | Model/simulation
Independence | Embedded SW + Metrology
Standard/Compliance Reference | WHO PQS accuracy basis (to confirm)
Evidence Required | Algorithm boundary-analysis report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | Analysis predicate to PV-SW-001 (FI).
PV-ANAL-014 — Timing / real-time schedulability & clock-drift analysis
Objective | Confirm by analysis that the acquire-settle-stamp cycle fits within the logging interval and that RTC drift stays within the disclosed bound over cold/sleep.
Method | Analysis
I/O Trace | SYS-037/038/118; PRD-FUN (time); DO-EE-011 (RTC), DO-FW
Risk-Control Trace | RCR-025 (trusted time / drift bound)
Acceptance Criteria | Worst-case task timing fits the interval with margin; predicted RTC drift within the bound across −30 °C/sleep; feeds PV-SW-005 (clock-drift test).
Level | Subsystem
Sample/Config | Task model + RTC spec
Facility/Equipment/Tools | Schedulability analysis
Independence | Firmware
Standard/Compliance Reference | —
Evidence Required | Timing/drift analysis report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | Confirmed by PV-SW-005 and PV-PERF timing tests.
7. Traceability Notes
Analyses either stand as the verification (where a test is not feasible/necessary) or predict the outcome a downstream Test confirms (safety-critical claims always confirmed by test). Each safety analysis names its confirming test in Notes; the D3 (PV-ANAL-003/004) and RF (PV-ANAL-006) analyses are predicates to the program-top tests PRD-VER-004/005.
8. Coverage/Completion Notes
Complete when all analyses are documented with margins and the safety-critical ones are confirmed by their named tests. D3 and RF analyses are Evidence-Pending (closed by test). All Planned/Evidence-Pending.
PVL-1500 (PV-COMP) — Compliance Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Quality/Regulatory (EM-ENG-006) with the discipline principals. Verifies the SCM meets the compliance obligations of its multi-market regime (D4): WHO PQS performance, EU GDP data-integrity, US FDA distribution-monitoring expectations, air/ground transport (UN 38.3 / IATA DGR), radio (FCC/CE-RED), safety (IEC 62368-1), EMC (CISPR/IEC 61000), environmental/ingress (IEC 60068 / 60529), software lifecycle (IEC 62304), and materials (RoHS/REACH). Research-driven: exact WHO PQS category and US classification are Evidence-Pending (concept §7). Standards cited by designation — no clause asserted as fact. All entries Planned/Evidence-Pending.
1. Purpose
Verify each market-access and regulatory obligation is met with the evidence a certification body/auditor requires, and map each obligation to the technical verification that demonstrates it.
2. Scope
Compliance verification across all markets in D4. This module largely aggregates and audits the evidence produced by the technical modules (safety, EMC, environmental, reliability, security, software) against each standard; it does not duplicate the tests.
3. Applicable Standards & References
WHO PQS E006 (temperature-monitoring device performance, category to confirm); EU GDP (data-integrity for medicinal-product distribution); 21 CFR Part 11 / ALCOA+ (electronic records); GAMP 5 (CSV); IEC 62304 (software lifecycle); ISO 14971 (risk management); UN 38.3 & IATA DGR (Li transport); FCC Part 15 / CE-RED (radio); IEC 62368-1 (safety); CISPR/IEC 61000 (EMC); IEC 60068/60529 (environmental/IP); RoHS/REACH; ISO 13485 (QMS, if applicable to classification). All designations; revisions & applicability to confirm.
4. Governance & Method
Method is Inspection (dossier/evidence audit against each standard) and Test (where a compliance test is the direct evidence, e.g., UN 38.3). Each obligation names the technical PV activity that produces its evidence.
5. Verification Categories
WHO PQS performance · EU GDP / data-integrity (Part 11 / ALCOA+) · US FDA distribution-monitoring · UN 38.3 / IATA DGR transport · radio type-approval · product safety · EMC · environmental/IP · IEC 62304 SW lifecycle · RoHS/REACH · ISO 14971 risk-management file.
6. Verification Register (this module)
PV-COMP-001 — WHO PQS performance-specification conformance
Objective | Confirm the SCM meets the applicable WHO PQS temperature-monitoring device performance specification (accuracy class, alarm behaviour, logging).
Method | Inspection, Test
I/O Trace | PRD-REG; SYS-177…190; concept D4; DO-VER-003
Risk-Control Trace | RCR-019/053 (accuracy/detection anchor to PQS)
Acceptance Criteria | Performance evidence (PV-PERF-001/002, PV-SW-001) meets the confirmed WHO PQS category requirements; the exact category and spec revision are confirmed. PQS category Evidence-Pending (concept §7).
Level | System
Sample/Config | DVT evidence set
Facility/Equipment/Tools | PQS test protocol; accredited lab
Independence | Regulatory (EM-ENG-006)
Standard/Compliance Reference | WHO PQS E006 (category/revision to confirm)
Evidence Required | PQS conformance dossier
Owner | Regulatory (EM-ENG-006)
Status | Evidence-Pending
Notes | Anchors detection/accuracy targets (O-7).
PV-COMP-002 — EU GDP / data-integrity (Part 11 / ALCOA+) conformance
Objective | Confirm the record and audit trail meet EU GDP data-integrity and ALCOA+/Part-11-style expectations for medicinal-product distribution.
Method | Inspection
I/O Trace | PRD-REG; PRD-CLOUD (audit); DO-CLD (audit); ALCOA+
Risk-Control Trace | RCR-042 (export integrity), RCR-034 (authenticity), RCR-057
Acceptance Criteria | Records are Attributable, Legible, Contemporaneous, Original, Accurate (+ Complete, Consistent, Enduring, Available); audit trail evidence (PV-DATA-006, PV-SEC-003) satisfies GDP expectations.
Level | System
Sample/Config | Record/audit evidence
Facility/Equipment/Tools | Data-integrity audit checklist
Independence | Quality/Regulatory
Standard/Compliance Reference | EU GDP; 21 CFR Part 11; ALCOA+; GAMP 5
Evidence Required | Data-integrity compliance record
Owner | Regulatory + Cloud
Status | Planned
Notes | CSV per GAMP 5 for the cloud record system.
PV-COMP-003 — US FDA distribution-monitoring / classification confirmation
Objective | Confirm the SCM's US regulatory position (accessory/information tool vs device classification) and that distribution-monitoring/data-integrity expectations are met.
Method | Inspection
I/O Trace | PRD-REG; concept §7/D4
Risk-Control Trace | RCR-058 (regulatory readiness)
Acceptance Criteria | US classification/pathway confirmed; applicable data-integrity expectations met. Classification Evidence-Pending (concept §7).
Level | System
Sample/Config | Regulatory dossier
Facility/Equipment/Tools | Regulatory assessment
Independence | Regulatory (EM-ENG-006)
Standard/Compliance Reference | FDA distribution/data-integrity (to confirm)
Evidence Required | US regulatory position memo
Owner | Regulatory (EM-ENG-006)
Status | Evidence-Pending
Notes | ISO 13485 applicability depends on classification.
PV-COMP-004 — UN 38.3 lithium-battery transport test conformance
Objective | Confirm the battery/device passes the UN 38.3 test suite (T.1–T.8) with report.
Method | Test
I/O Trace | PRD-SAFE-009; PRD-REG-006; DO-EE-015; PRD-ELEC-015
Risk-Control Trace | RCR-007; HA-TFC-002, HA-ENRG-003
Acceptance Criteria | Full UN 38.3 T.1–T.8 pass with report; no air-transport claim until this passes.
Level | Subsystem/System
Sample/Config | Production-representative cells/packs/devices
Facility/Equipment/Tools | Accredited UN 38.3 lab
Independence | Accredited test lab
Standard/Compliance Reference | UN 38.3 (Manual of Tests and Criteria)
Evidence Required | UN 38.3 test summary
Owner | Battery/Power + Regulatory
Status | Evidence-Pending
Notes | RCR-007 verification reference; release condition.
PV-COMP-005 — IATA DGR packing/marking/documentation conformance
Objective | Confirm Li-battery air-transport packing, marking, and documentation conform to IATA DGR.
Method | Inspection
I/O Trace | PRD-REG-006; DO-EE-015, DO-REG-006
Risk-Control Trace | RCR-007 (transport)
Acceptance Criteria | Packing, marking, SoC, and documentation conform to IATA DGR; conformance recorded.
Level | Production
Sample/Config | Production packaging
Facility/Equipment/Tools | DGR checklist
Independence | Regulatory/Quality
Standard/Compliance Reference | IATA DGR
Evidence Required | DGR conformance record
Owner | Regulatory (EM-ENG-006)
Status | Planned
Notes | Complements PV-INSP-007.
PV-COMP-006 — Radio type-approval (FCC / CE-RED) certification
Objective | Confirm each radio is certified for the target markets.
Method | Test, Inspection
I/O Trace | PRD-REG (radio); PV-EMC-004
Risk-Control Trace | RCR-056 (regional config)
Acceptance Criteria | FCC/CE-RED certifications obtained for the target markets; grant/DoC on file. Market set Evidence-Pending (D4).
Level | System
Sample/Config | Regional variants
Facility/Equipment/Tools | Accredited radio lab
Independence | Accredited lab
Standard/Compliance Reference | FCC Part 15; CE RED
Evidence Required | Radio certifications
Owner | Regulatory + RF
Status | Evidence-Pending
Notes | Uses PV-EMC-004 evidence.
PV-COMP-007 — Product-safety (IEC 62368-1) conformance
Objective | Confirm the device meets IEC 62368-1 (energy sources, single-fault, touch-temperature, fire enclosure, electrical safety).
Method | Test, Inspection
I/O Trace | PRD-SAFE-008; PV-SAFE-001/002, PV-TFC-001/002, PV-ENRG-002
Risk-Control Trace | RCR-002/008/013/005
Acceptance Criteria | IEC 62368-1 requirements met per the safety-module evidence; certification obtainable.
Level | System
Sample/Config | DVT evidence set
Facility/Equipment/Tools | Accredited safety lab
Independence | Accredited lab
Standard/Compliance Reference | IEC 62368-1
Evidence Required | Safety compliance dossier
Owner | Safety + Regulatory
Status | Planned
Notes | Aggregates PV-SAFE/PV-TFC/PV-ENRG evidence.
PV-COMP-008 — EMC directive conformance (CISPR / IEC 61000)
Objective | Confirm EMC conformance for the target markets.
Method | Test, Inspection
I/O Trace | PV-EMC-001/PV-ELEC-004/PV-EMC-003
Risk-Control Trace | RCR-014
Acceptance Criteria | Emissions and immunity meet the market EMC requirements; DoC obtainable.
Level | System
Sample/Config | DVT evidence set
Facility/Equipment/Tools | Accredited EMC lab
Independence | Accredited lab
Standard/Compliance Reference | CISPR/IEC 61000; CE EMC
Evidence Required | EMC compliance dossier
Owner | EMC + Regulatory
Status | Planned
Notes | Uses PVL-800 evidence.
PV-COMP-009 — Environmental / ingress (IEC 60068 / 60529) conformance
Objective | Confirm the environmental and IP claims are substantiated for the markets.
Method | Inspection
I/O Trace | PVL-700 evidence
Risk-Control Trace | RCR-011/017/018
Acceptance Criteria | IEC 60068 exposures and IEC 60529 IP class substantiated at the confirmed severities.
Level | System
Sample/Config | PVL-700 evidence
Facility/Equipment/Tools | Dossier audit
Independence | Quality/Regulatory
Standard/Compliance Reference | IEC 60068; IEC 60529
Evidence Required | Environmental compliance record
Owner | Regulatory + Reliability
Status | Planned
Notes | —
PV-COMP-010 — IEC 62304 software-lifecycle compliance
Objective | Confirm the software lifecycle (planning, requirements, architecture, verification, config/ problem management) meets IEC 62304 at the assigned safety class.
Method | Inspection
I/O Trace | SRS library; PVL-1100 evidence; IEC 62304
Risk-Control Trace | RCR-019/021/037 (software safety controls)
Acceptance Criteria | IEC 62304 lifecycle artefacts complete for the assigned class; verification/coverage evidence (PV-SW-014/015) present; software safety classification documented.
Level | System
Sample/Config | Software lifecycle file
Facility/Equipment/Tools | Compliance checklist
Independence | Quality + Embedded SW
Standard/Compliance Reference | IEC 62304
Evidence Required | IEC 62304 compliance record
Owner | Quality + Embedded SW (EM-ENG-007)
Status | Planned
Notes | Software safety class to confirm.
PV-COMP-011 — RoHS / REACH materials conformance
Objective | Confirm materials/BOM conform to RoHS/REACH.
Method | Inspection, Analysis
I/O Trace | PRD-REG; PV-INSP-008; DO-CFG
Risk-Control Trace | RCR-052 (EOL/materials)
Acceptance Criteria | RoHS/REACH declarations complete for the BOM; no restricted substance above limits.
Level | Production
Sample/Config | BOM + declarations
Facility/Equipment/Tools | Materials-compliance database
Independence | Quality/Supply-Chain
Standard/Compliance Reference | RoHS; REACH
Evidence Required | Materials-compliance record
Owner | Quality + Supply-Chain
Status | Planned
Notes | —
PV-COMP-012 — ISO 14971 risk-management-file completeness
Objective | Confirm the risk-management file is complete and the risk controls are verified/effective per ISO 14971 before release.
Method | Inspection
I/O Trace | HAL (Stage 10); RCR register (Stage 12); PV-SAFE-005
Risk-Control Trace | RCR-058 (overall residual/benefit-risk gate)
Acceptance Criteria | Risk-management file complete; every control verified (PV-SAFE-005); overall residual and benefit-risk signed off; the eight release conditions met.
Level | System
Sample/Config | Risk-management file
Facility/Equipment/Tools | ISO 14971 review
Independence | Quality/Regulatory + System-Safety
Standard/Compliance Reference | ISO 14971
Evidence Required | Risk-management review record
Owner | Regulatory (EM-ENG-006) + System-Safety
Status | Planned
Notes | RCR-058 verification reference; the release gate; feeds PVL-1900.
7. Traceability Notes
Compliance activities aggregate and audit the technical evidence (safety/EMC/env/reliability/security/software) against each standard/regime; UN 38.3 (PV-COMP-004) is the RCR-007 reference and ISO 14971 file (PV-COMP-012) is the RCR-058 reference. Feeds Release (Stage 16) and certification.
8. Coverage/Completion Notes
Complete when every applicable regime is substantiated. WHO PQS category, US classification, target-market set, and test houses are Evidence-Pending (concept §7 / D4). All Planned/Evidence-Pending.
PVL-1700 (PV-DEF) — Defect Management — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Quality (EM-ENG-006) with the discipline principals. Governs how verification defects/nonconformances are captured, triaged by severity, root-caused, corrected, retested, and closed — so no verification failure is lost and no safety failure closes without disposition. All entries Planned.
1. Purpose
Define and verify the defect-management loop for verification failures: capture, severity/priority, root cause, corrective action, regression/retest, deviation/waiver, and closure.
2. Scope
All verification nonconformances across the 20 modules. Interfaces with PV-GOV-005 (deviation/waiver) and PVL- 1900 (open-issue summary). Field CAPA (Stage 17) is downstream.
3. Applicable Standards & References
ISO 13485 §8.3 (nonconforming product) / §8.5 (CAPA); ISO 14971 (risk re-assessment on safety defects); IEC 62304 (software problem resolution). Designations; revisions to confirm.
4. Governance & Method
Method is Inspection of the defect process/records. Severity is severity-first (CLAUDE.md rule 3): safety/security defects dominate and block release until closed.
5. Verification Categories
Defect capture · severity/priority triage · root-cause analysis · corrective action · regression/retest · deviation/waiver · closure/verification.
6. Verification Register (this module)
PV-DEF-001 — Defect capture & severity/priority triage process
Objective | Confirm every verification failure is captured with a severity/priority and routed, with safety/security defects escalated.
Method | Inspection
I/O Trace | PV-GOV-005; PRD-VER-001
Risk-Control Trace | RCR-058 (no unresolved safety defect at release)
Acceptance Criteria | 100 % of verification failures captured with severity/priority; safety/security defects escalated to System-Safety/Security; no failure silently closed.
Level | System
Sample/Config | Defect log
Facility/Equipment/Tools | QMS defect/CAPA module
Independence | Quality
Standard/Compliance Reference | ISO 13485 §8.3
Evidence Required | Defect-log + triage records
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | —
PV-DEF-002 — Root-cause & corrective-action verification
Objective | Confirm each significant defect has a root-cause analysis and a corrective action verified effective.
Method | Inspection
I/O Trace | ISO 13485 §8.5; FAL-2000 (RCA/CAPA)
Risk-Control Trace | RCR-058; affected RCRs
Acceptance Criteria | Every significant/safety defect has RCA and a verified-effective corrective action; no recurrence in retest.
Level | System
Sample/Config | CAPA records
Facility/Equipment/Tools | QMS CAPA
Independence | Quality + discipline owner
Standard/Compliance Reference | ISO 13485 §8.5; ISO 14971
Evidence Required | RCA/CAPA records
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | —
PV-DEF-003 — Regression / retest after correction
Objective | Confirm corrected defects are retested and the change-impacted PV activities re-run (per PV-GOV-004).
Method | Inspection
I/O Trace | PV-GOV-004 (regression); PV-FUNC-015
Risk-Control Trace | RCR-019/021/037 (safety/verdict/OTA regression)
Acceptance Criteria | Every correction is retested; impacted PV activities re-run; safety changes trigger the full safety regression; results recorded.
Level | System
Sample/Config | Retest records
Facility/Equipment/Tools | CI regression harness
Independence | Quality + Embedded SW
Standard/Compliance Reference | IEC 62304
Evidence Required | Retest/regression records
Owner | Quality + Embedded SW (EM-ENG-007)
Status | Planned
Notes | —
PV-DEF-004 — Deviation/waiver & closure verification
Objective | Confirm deviations/waivers are dispositioned and every defect is formally closed, with safety/ security waivers requiring risk re-assessment and sign-off.
Method | Inspection
I/O Trace | PV-GOV-005; RCR-058
Risk-Control Trace | RCR-058 (release gate)
Acceptance Criteria | No defect open at release without a signed disposition; safety/security waivers carry a risk re-assessment and EM-ENG-006 + EM-ENG-001 sign-off; closure verified.
Level | System
Sample/Config | Deviation/waiver + closure log
Facility/Equipment/Tools | QMS
Independence | Quality + System-Safety
Standard/Compliance Reference | ISO 13485; ISO 14971
Evidence Required | Closure/waiver records
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | Feeds PVL-1900 open-issue list.
7. Traceability Notes
Defect management links every verification failure to its disposition and to the affected requirement/RCR; open safety/security defects gate release (RCR-058). Feeds PVL-1900 and Stage-17 field CAPA.
8. Coverage/Completion Notes
Complete when the defect loop is in force and, at release, no safety/security defect is open without a signed disposition. All Planned.
PVL-800 (PV-EMC) — EMC & Radio Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-003 (PCB/Electronics/EMC) with Connectivity/RF and Regulatory. Qualifies emissions, immunity (with the safety-relevant requirement that no interference produces a false reading/verdict or missed excursion), wireless coexistence, and radio type-approval (FCC/CE/RED). Hosts the EMC-immunity test PV-ELEC-004 referenced by Stage-12. CISPR/IEC 61000 and FCC Part 15 / CE RED govern; levels set in the Verification Plan. All entries Planned/Evidence-Pending.
1. Purpose
Qualify the SCM's electromagnetic compatibility and radio compliance so it neither interferes with aircraft/ medical equipment nor is upset (false/missed excursion, false verdict) by its EM environment, and so its radios are type-approved for its markets.
2. Scope
Radiated/conducted emissions, ESD, radiated/conducted RF immunity, EFT/burst, surge, magnetic field, power- frequency, wireless coexistence, RF exposure, and radio type-approval across S2–S7. Excludes RF-in-container link performance (PVL-1400 PV-REL-005 / PVL-500 PV-COMM-003) and functional comms (PVL-500).
3. Applicable Standards & References
CISPR 11 / CISPR 32 (emissions); IEC 61000-4-2 (ESD) / -4-3 (radiated immunity) / -4-4 (EFT) / -4-5 (surge) / -4-6 (conducted immunity) / -4-8 (magnetic field); FCC Part 15 (subparts B/C); CE RED (EN 300/301 series); RF-exposure (SAR/MPE where applicable). Designations; revisions/levels to confirm.
4. Governance & Method
Method is Test at an accredited EMC/radio facility. Immunity is tested against the integrity of the detection/verdict path (not just "no reset"): the device is exercised through detection/verdict during immunity exposure. This is FI-adjacent — interference is the injected disturbance.
5. Verification Categories
Radiated/conducted emissions · ESD · radiated/conducted RF immunity · EFT/burst · surge · magnetic-field · wireless coexistence · RF exposure · radio type-approval.
6. Verification Register (this module)
PV-EMC-001 — Radiated & conducted emissions
Objective | Confirm radiated and conducted emissions are within limits so the device does not interfere with aircraft or medical equipment.
Method | Test
I/O Trace | SYS (EMC); PRD-ELEC (EMC); DO-EE (EMC)
Risk-Control Trace | RCR-014 (emissions limits)
Acceptance Criteria | Emissions within CISPR/FCC Part 15 / CE-RED limits across the operating modes, including radio transmit.
Level | System
Sample/Config | DVT units, all radios active
Facility/Equipment/Tools | Accredited EMC chamber
Independence | Accredited test house (independent)
Standard/Compliance Reference | CISPR 11/32; FCC Part 15B; CE RED
Evidence Required | Emissions test report
Owner | EMC (EM-ENG-003) + Regulatory
Status | Planned
Notes | —
PV-ELEC-004 — EMC immunity vs detection/verdict integrity (ESD/RF/EFT/surge)
Objective | Confirm no immunity disturbance (ESD, radiated/conducted RF, EFT, surge, magnetic field) produces a false reading, false verdict, or missed excursion, and that the device recovers to a defined state.
Method | Test (FI-adjacent — interference is the injected fault)
I/O Trace | SYS (EMC); PRD-SAFE-001 (don't mislead) via data integrity; DO-EE (EMC), DO-PCB-010, DO-EE-019
Risk-Control Trace | RCR-014 (immunity — no false/missed excursion)
Acceptance Criteria | Under the IEC 61000 immunity levels, while exercising detection/verdict: 0 false readings, 0 false verdicts, 0 missed excursions; any upset self-recovers to the defined state with the event logged. Immunity levels per IEC 61000 test plan.
Level | System
Sample/Config | DVT units exercised through detection/verdict
Facility/Equipment/Tools | Accredited EMC chamber + stimulus injector
Independence | Accredited test house; System-Safety reviews criteria
Standard/Compliance Reference | IEC 61000-4-2/-4-3/-4-4/-4-5/-4-6/-4-8
Evidence Required | Immunity test report with detection/verdict integrity results
Owner | EMC (EM-ENG-003) + System-Safety
Status | Planned
Notes | RCR-014 verification reference; the safety-critical EMC test — pass criterion is data-decision integrity, not merely "no reset".
PV-EMC-003 — Wireless coexistence (LTE-M / Wi-Fi / BLE / GNSS)
Objective | Confirm the co-located radios coexist without self-desensitisation that degrades alert delivery or GNSS below spec.
Method | Test
I/O Trace | SYS-104…118; PRD-CONN; DO-EE-017/018
Risk-Control Trace | RCR-031 (link availability), RCR-030
Acceptance Criteria | With all radios active, each meets its required sensitivity/throughput; no coexistence-induced loss of alert delivery.
Level | System
Sample/Config | DVT units, all radios
Facility/Equipment/Tools | RF conducted/radiated test setup
Independence | RF/Connectivity + test house
Standard/Compliance Reference | relevant coexistence practice
Evidence Required | Coexistence test report
Owner | RF/Connectivity (EM-ENG-003)
Status | Planned
Notes | —
PV-EMC-004 — Radio type-approval (FCC / CE-RED) transmitter compliance
Objective | Confirm each intentional radiator meets transmitter type-approval requirements for the target markets.
Method | Test
I/O Trace | PRD-REG (radio); SYS-177…190; PRD-CONN
Risk-Control Trace | RCR-056 (correct regional config), RCR-014
Acceptance Criteria | Each radio passes FCC Part 15C / CE-RED transmitter measurements for its band plan; certification obtainable. Target-market set per regulatory stage.
Level | System
Sample/Config | DVT units (per regional variant)
Facility/Equipment/Tools | Accredited radio test house
Independence | Accredited test house
Standard/Compliance Reference | FCC Part 15C; CE RED (EN 300/301 series)
Evidence Required | Radio type-approval reports
Owner | Regulatory (EM-ENG-006) + RF
Status | Evidence-Pending
Notes | Target markets Evidence-Pending (D4).
PV-EMC-005 — RF exposure (SAR / MPE) evaluation
Objective | Confirm RF-exposure compliance for the device's transmit power and use position.
Method | Test, Analysis
I/O Trace | PRD-REG (RF exposure); PRD-CONN
Risk-Control Trace | RCR-014 (emissions/exposure)
Acceptance Criteria | RF exposure within the applicable SAR/MPE limits for the intended use position.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | SAR/MPE test setup
Independence | Accredited test house
Standard/Compliance Reference | FCC/CE RF-exposure limits
Evidence Required | RF-exposure evaluation
Owner | Regulatory + RF
Status | Planned
Notes | Likely MPE (low power, not body-worn) — confirm.
PV-EMC-006 — Electrical fast-transient / surge on the USB-C charge port
Objective | Confirm the wired charge/data port withstands EFT/burst and surge without damage or hazardous condition.
Method | Test
I/O Trace | HWR (USB-C); PRD-ELEC; DO-EE-009/010
Risk-Control Trace | RCR-003 (charge-source robustness), RCR-014
Acceptance Criteria | No damage or hazardous condition after EFT/burst and surge on the charge port; port remains functional or fails safe.
Level | Subsystem
Sample/Config | DVT units on charger
Facility/Equipment/Tools | EFT/surge generators
Independence | EMC test house
Standard/Compliance Reference | IEC 61000-4-4 / -4-5
Evidence Required | Port EFT/surge report
Owner | Electrical/EMC (EM-ENG-003)
Status | Planned
Notes | Complements PV-ELEC-002 (charge-source).
7. Traceability Notes
EMC activities verify RCR-014 (immunity/emissions) and support RCR-031/056 (radio availability/config). The safety-critical entry is PV-ELEC-004 — pass/fail is decision-path integrity. Radio type-approval feeds the compliance module (PVL-1500).
8. Coverage/Completion Notes
Complete when emissions, immunity (with decision-path integrity), coexistence, radio type-approval, and RF exposure pass at accredited facilities for the target markets. Levels/markets Evidence-Pending. All Planned.
PVL-700 (PV-ENV) — Environmental Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with Mechanical (EM-ENG-002), Materials, Reliability. Qualifies the SCM against its environmental and mechanical-environment envelope: −30…+60 °C operating (−40…+70 °C storage), thermal cycling, humidity/condensation, ingress (IP67), drop/shock/vibration, altitude, transport, UV, and cleaning. Hosts the mechanical-environment tests (PV-MECH-###) referenced by Stage-12. IEC 60068 series and IEC 60529 govern; exact severities are set in the Verification Plan. All entries Planned/Evidence-Pending.
1. Purpose
Qualify survival, function, accuracy retention, and record integrity across the full environmental envelope of the cold-chain journey and multi-trip reuse.
2. Scope
Environmental exposures and mechanical-environment tests at Subsystem/System level (S1 dominant, all subsystems exposed). Excludes performance measurement (PVL-600), long-life reliability (PVL-1400), and EMC (PVL-800).
3. Applicable Standards & References
IEC 60068-2-1 (cold) / -2-2 (dry heat) / -2-14 (change of temperature) / -2-30 & -2-38 (damp-heat/temp-humidity cycling) / -2-78 (damp heat steady) / -2-64 (random vibration) / -2-6 (sinusoidal) / -2-27 (shock) / -2-31 (free-fall); IEC 60529 (IP); UN 38.3 T.1 (altitude); IEC 60068-2-5 / ISO 4892 (solar/UV). Designations; revisions & severities to confirm in the Verification Plan.
4. Governance & Method
Method is Test (with Inspection for IP/dimensional confirmation). Sequencing: non-destructive before destructive; accuracy/functional checks before and after each exposure. Reuse-life exposures (thermal/ condensation cycling, cleaning) run to the confirmed cycle count.
5. Verification Categories
Cold/heat operating · storage/transport temperature · thermal cycling · humidity · condensation cycling · cold display legibility · vibration · shock · drop · IP/ingress · cleaning/chemical · altitude/pressure · UV/ weathering.
6. Verification Register (this module)
PV-ENV-001 — Operating temperature range −30…+60 °C (function + accuracy)
Objective | Confirm full function (sensing, logging, verdict, alert) and retained accuracy across continuous operation at −30 °C and +60 °C.
Method | Test
I/O Trace | SYS-069/071; PRD-ENV-001; DO-VER-009
Risk-Control Trace | RCR-017 (operate + retain record across range)
Acceptance Criteria | Full function and record retention at both extremes; accuracy within ±0.5 °C (PV-PERF-001). (90-day cold endurance = PV-REL-001.)
Level | System
Sample/Config | DVT units in chamber
Facility/Equipment/Tools | Climatic chamber; temperature reference
Independence | Reliability/Quality
Standard/Compliance Reference | IEC 60068-2-1 / -2-2
Evidence Required | Cold/heat operating test report
Owner | Reliability (EM-ENG-006)
Status | Planned
Notes | RCR-017 verification reference (function); endurance = PV-REL-001.
PV-ENV-002 — Storage / transport temperature survival (−40…+70 °C)
Objective | Confirm survival without permanent degradation of function/accuracy/seal/data at the non-operating extremes.
Method | Test
I/O Trace | SYS-070; PRD-ENV-002
Risk-Control Trace | RCR-011 (seal), RCR-017
Acceptance Criteria | After exposure at −40 °C and +70 °C (confirm vs cell datasheet), function/accuracy/ seal/data intact; no permanent degradation.
Level | System
Sample/Config | DVT units (powered-off)
Facility/Equipment/Tools | Climatic chamber
Independence | Reliability
Standard/Compliance Reference | IEC 60068-2-1 / -2-2
Evidence Required | Storage-temperature survival report
Owner | Reliability
Status | Evidence-Pending
Notes | Upper/lower bounds Evidence-Pending vs cell datasheet (SYS-070).
PV-ENV-003 — Condensation & cold↔warm thermal cycling (reuse life)
Objective | Confirm tolerance of repeated condensation and cold-to-warm cycling over reuse life without record-integrity failure, corrosion, seal, or accuracy loss.
Method | Test
I/O Trace | SYS-073; PRD-ENV-004; PRD-MECH-005/009; DO-SAFE-011
Risk-Control Trace | RCR-018 (condensation tolerance), RCR-015 (coat/seal)
Acceptance Criteria | After N condensation/thermal cycles (count Evidence-Pending) no record-integrity fault, corrosion, seal degradation, or accuracy shift; insulation resistance within limits.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Temperature-humidity cycling chamber
Independence | Reliability/Materials
Standard/Compliance Reference | IEC 60068-2-30 / -2-38
Evidence Required | Condensation-cycle report
Owner | Reliability/Materials
Status | Evidence-Pending
Notes | RCR-018 verification reference; joint with PV-TFC-006; cycle count Evidence-Pending.
PV-ENV-004 — Humidity operating range (10–95 %RH non-condensing)
Objective | Confirm function and accuracy across the humidity operating range.
Method | Test
I/O Trace | SYS-072; PRD-ENV-004
Risk-Control Trace | RCR-015 (record electronics protection)
Acceptance Criteria | Function and accuracy retained across 10–95 %RH non-condensing.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Humidity chamber
Independence | Reliability
Standard/Compliance Reference | IEC 60068-2-78
Evidence Required | Damp-heat steady report
Owner | Reliability
Status | Planned
Notes | —
PV-MECH-001 — Transport vibration survival (random + sinusoidal)
Objective | Confirm survival of combined air/road/rail vibration without loss of function, data, joint fatigue failure, or probe/device displacement.
Method | Test
I/O Trace | SYS-076; PRD-ENV-005; PRD-MECH-003/007
Risk-Control Trace | RCR-012 (fasteners resist loosening), RCR-004 (cell retention)
Acceptance Criteria | After the qualified random+sinusoidal profile: full function, 0 data loss, no connector/solder fatigue failure, no probe/device position shift. Profile Evidence-Pending (SYS-076).
Level | System
Sample/Config | DVT units on shaker
Facility/Equipment/Tools | Vibration table; accelerometers
Independence | Mechanical/Reliability
Standard/Compliance Reference | IEC 60068-2-64 / -2-6
Evidence Required | Vibration test report
Owner | Mechanical (EM-ENG-002)
Status | Evidence-Pending
Notes | Profile severity set in Verification Plan.
PV-MECH-002 — Mechanical shock survival
Objective | Confirm survival of handling shock without loss of function/data, seal breach, or accuracy shift.
Method | Test
I/O Trace | SYS-077; PRD-ENV-006; PRD-MECH-007
Risk-Control Trace | RCR-011 (seal), RCR-004 (cell retention)
Acceptance Criteria | After the qualified shock pulse: full function, 0 data loss, seal intact, accuracy unchanged. Shock pulse Evidence-Pending (SYS-077).
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Shock machine
Independence | Mechanical/Reliability
Standard/Compliance Reference | IEC 60068-2-27
Evidence Required | Shock test report
Owner | Mechanical (EM-ENG-002)
Status | Evidence-Pending
Notes | —
PV-MECH-003 — Free-fall drop & small-parts / retention test
Objective | Confirm free-fall drop survival (function/data/seal/accuracy) and that no detachable small part is released (choking prevention).
Method | Test
I/O Trace | SYS-077; PRD-ENV-006; PRD-MECH; 16 CFR 1501 (small parts)
Risk-Control Trace | RCR-010 (small-parts retention), RCR-011 (seal)
Acceptance Criteria | After the qualified drop height: full function, 0 data loss, seal intact, accuracy unchanged; no detachable small part fits the small-parts cylinder without tooling. Drop height Evidence-Pending (SYS-077).
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Drop tester; small-parts cylinder
Independence | Mechanical + Human-Factors
Standard/Compliance Reference | IEC 60068-2-31; 16 CFR 1501
Evidence Required | Drop / small-parts report
Owner | Mechanical (EM-ENG-002)
Status | Evidence-Pending
Notes | RCR-010 verification reference.
PV-MECH-004 — Cell retention through drop/vibration/transport (no crush/puncture)
Objective | Confirm the cell is retained without crush/puncture through the drop/vibration/transport profile and reuse life, maintaining thermal/electrical clearance.
Method | Test
I/O Trace | SYS-076/077; PRD-SAFE-008/009; HWR-002; DO-MECH-012, DO-EE-004
Risk-Control Trace | RCR-004 (cell retention) — FI-adjacent (mechanical single-fault)
Acceptance Criteria | No cell deformation/puncture/short after the qualified drop/vibration profile and UN 38.3; clearance maintained. Profile per IEC 60068 / UN 38.3.
Level | Subsystem
Sample/Config | DVT packs/units
Facility/Equipment/Tools | Drop/vibration; post-test CT/teardown
Independence | Mechanical + Battery/Power
Standard/Compliance Reference | IEC 60068-2-27/-2-31/-2-64; UN 38.3
Evidence Required | Cell-retention test report
Owner | Mechanical + Battery/Power
Status | Planned
Notes | RCR-004 verification reference; safety single-fault path (with PV-TFC-002).
PV-ENV-005 — Ingress protection IP67 (over reuse cycling)
Objective | Confirm the enclosure and closed charge/service port meet IP67 (dust-tight; temporary immersion) and retain it after reuse cycling.
Method | Test, Inspection
I/O Trace | SYS-078; PRD-ENV-007; PRD-MECH-005/006; DO-MECH-005
Risk-Control Trace | RCR-011 (ingress + containment through life)
Acceptance Criteria | IP6X dust + IPX7 immersion pass per IEC 60529 initially and after N reuse/thermal cycles; no ingress; IP class & cycle count Evidence-Pending.
Level | System
Sample/Config | DVT units (pre/post cycling)
Facility/Equipment/Tools | Dust chamber; immersion tank
Independence | Mechanical/Quality
Standard/Compliance Reference | IEC 60529
Evidence Required | IP test report
Owner | Mechanical (EM-ENG-002)
Status | Evidence-Pending
Notes | RCR-011 verification reference; confirm IP67 vs cleaning method (SYS-078).
PV-ENV-006 — Cleaning & decontamination chemical compatibility (reuse life)
Objective | Confirm external surfaces, seals, window, controls, and labels withstand the specified cleaning agents over reuse life without crazing, seal/IP loss, or legibility loss.
Method | Test
I/O Trace | SYS-079; PRD-ENV-008; PRD-MECH-009/010; PRD-SVC-002
Risk-Control Trace | RCR-011 (seal/IP maintained), RCR-052 (surface durability)
Acceptance Criteria | After the agent list and cycle count (Evidence-Pending): no material crazing, seal degradation, IP loss, or label/display illegibility.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Chemical exposure rig
Independence | Materials/Reliability
Standard/Compliance Reference | RoHS/REACH (materials); manufacturer agent list
Evidence Required | Cleaning-compatibility report
Owner | Materials/Reliability
Status | Evidence-Pending
Notes | Agent list/cycle count Evidence-Pending (SYS-079).
PV-ENV-007 — Altitude / cargo-hold pressure operation
Objective | Confirm operation and safety through reduced/cycling cargo-hold pressure (≥4,570 m cabin altitude) without loss of function, seal failure, or battery hazard.
Method | Test
I/O Trace | SYS-082; PRD-ENV-012; PRD-REG-006; D4
Risk-Control Trace | RCR-011 (seal under Δp), RCR-004 (battery)
Acceptance Criteria | Function and safety maintained through the altitude profile and pressure-change rate; seal integrity retained; consistent with UN 38.3 T.1 altitude simulation.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | Altitude/pressure chamber
Independence | Reliability/Mechanical
Standard/Compliance Reference | UN 38.3 T.1; IATA DGR / DO-160 low-pressure concept
Evidence Required | Altitude test report
Owner | Reliability/Mechanical
Status | Planned
Notes | Battery altitude also covered under UN 38.3 (PV-COMP/PV-SAFE).
PV-ENV-008 — UV / solar / weathering & marking durability
Objective | Confirm enclosure materials and required markings withstand UV/sunlight/outdoor staging over reuse life without embrittlement, discolouration impairing the display, or loss of marking legibility.
Method | Test
I/O Trace | SYS-085; PRD-MECH-010; PRD-ENV-008
Risk-Control Trace | RCR-052 (marking/durability), RCR-055 (identity legibility)
Acceptance Criteria | After the UV/weathering exposure (profile Evidence-Pending): no embrittlement, display-impairing discolouration, or loss of identification/mandatory-marking legibility.
Level | Subsystem
Sample/Config | DVT enclosures/labels
Facility/Equipment/Tools | UV/weathering chamber
Independence | Materials
Standard/Compliance Reference | IEC 60068-2-5; ISO 4892
Evidence Required | UV/weathering report
Owner | Materials
Status | Evidence-Pending
Notes | Exposure profile Evidence-Pending (SYS-085).
PV-ENV-009 — Transport (ISTA-class) packaged-product test
Objective | Confirm the packaged product survives a representative distribution/transport simulation (drop, vibration, compression) without damage or data loss.
Method | Test
I/O Trace | SYS-076/077; PRD-ENV-005/006; PRD-REG-006
Risk-Control Trace | RCR-004/011 (retention/seal), RCR-007 (transport)
Acceptance Criteria | After the transport-simulation sequence: packaging protects the product; no damage, data loss, or seal breach.
Level | System
Sample/Config | Packaged DVT units
Facility/Equipment/Tools | Transport-simulation lab
Independence | Packaging/Quality
Standard/Compliance Reference | ISTA-series (to confirm); IEC 60068
Evidence Required | Transport-simulation report
Owner | Packaging/Quality
Status | Planned
Notes | Complements UN 38.3 battery transport (PV-SAFE/PV-COMP).
7. Traceability Notes
Environmental activities qualify the SYS-069…085 environmental requirements and provide the environmental half of RCR-004/010/011/012/015/017/018. The −30…+60 °C chamber and representative shipper come from DO-VER-009. The battery-specific transport/altitude safety is shared with PV-SAFE/PV-COMP (UN 38.3).
8. Coverage/Completion Notes
Complete when the full IEC 60068 / 60529 suite passes at the confirmed severities on production-representative units, including reuse-life cycling. Severities and cycle counts Evidence-Pending (Verification Plan). All Planned/Evidence-Pending.
PVL-1800 (PV-EVID) — Verification Evidence — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Quality (EM-ENG-006). Governs the verification evidence/records — protocols, raw data, reports, logs, photos/videos, calibration certificates, review records, approvals, signatures — that each PV activity's "Evidence Required" field references. The PVL defines what to verify and how; this module governs where the executed evidence lives and how it stays attributable and controlled. No evidence exists yet (pre- execution). All entries Planned.
1. Purpose
Define the evidence package structure, control, and integrity for the SCM verification program so every result is attributable, reproducible, and auditable (ALCOA+), and referenced from each PV activity.
2. Scope
Evidence lifecycle for all 20 modules. Excludes authoring the results (produced at execution) and the playbook definitions (the other modules).
3. Applicable Standards & References
ISO 13485 §4.2.5 (records); ISO/IEC 17025 (calibration certificates); ALCOA+ / 21 CFR Part 11 (record integrity); GAMP 5 (CSV for the evidence system). Designations; revisions to confirm.
4. Governance & Method
Method is Inspection of the evidence structure/control. Evidence uses the DO-VER-005 verification logging schema (stimulus, measured value, limit, verdict, unit serial, timestamp, config/version).
5. Verification Categories
Protocol control · raw-data capture · report structure · calibration certificates · review/approval records · photo/video/log capture · evidence integrity (ALCOA+) · retention.
6. Verification Register (this module)
PV-EVID-001 — Protocol control & approval
Objective | Confirm every executed PV activity runs against an approved, version-controlled protocol.
Method | Inspection
I/O Trace | DO-VER-010 (pass/fail mapping); PV-GOV-006
Risk-Control Trace | RCR-058 (evidence quality)
Acceptance Criteria | Every PV activity has an approved protocol before execution; protocol version recorded in the result; no ad-hoc execution.
Level | System
Sample/Config | Protocol library
Facility/Equipment/Tools | PLM/QMS
Independence | Quality
Standard/Compliance Reference | ISO 13485; GAMP 5
Evidence Required | Approved protocol set
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | —
PV-EVID-002 — Raw-data & report capture per the logging schema
Objective | Confirm raw data and reports capture stimulus, measured value, limit, verdict, unit serial, timestamp, and config/version per DO-VER-005 (ALCOA+).
Method | Inspection
I/O Trace | DO-VER-005 (schema); PV-GOV-003 (config)
Risk-Control Trace | RCR-057 (integrity discipline mirrored in evidence)
Acceptance Criteria | Each result is attributable and reconstructable per the schema; raw data retained alongside the report.
Level | System
Sample/Config | Evidence records
Facility/Equipment/Tools | Verification data system
Independence | Quality
Standard/Compliance Reference | ALCOA+; 21 CFR Part 11
Evidence Required | Report + raw-data package
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | —
PV-EVID-003 — Calibration certificates & equipment traceability
Objective | Confirm every measurement uses equipment with a current ISO/IEC 17025 calibration certificate on file.
Method | Inspection
I/O Trace | DO-VER-008; PV-PLAN-004; PV-ANAL-010 (MSA)
Risk-Control Trace | RCR-053 (accuracy truth)
Acceptance Criteria | All verification equipment calibrated and traceable; certificates referenced from the affected results; no result on out-of-cal equipment.
Level | System
Sample/Config | Calibration register
Facility/Equipment/Tools | Metrology system
Independence | Metrology/Quality
Standard/Compliance Reference | ISO/IEC 17025
Evidence Required | Calibration certificates
Owner | Quality/Metrology
Status | Planned
Notes | —
PV-EVID-004 — Review, approval & signature records + retention
Objective | Confirm each verification result is reviewed, approved, and signed (independent where required) and retained per the records policy.
Method | Inspection
I/O Trace | PV-GOV-002 (independence); ISO 13485 §4.2.5
Risk-Control Trace | RCR-058 (release evidence)
Acceptance Criteria | Every result carries review/approval/signature (independent for safety/security) and is retained per policy; the evidence index is complete for release.
Level | System
Sample/Config | Evidence index
Facility/Equipment/Tools | QMS/eQMS
Independence | Quality
Standard/Compliance Reference | ISO 13485 §4.2.5; 21 CFR Part 11 (signatures)
Evidence Required | Review/approval/signature records
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | Feeds PVL-1900 and Release (Stage 16).
7. Traceability Notes
Every PV activity's "Evidence Required" field points here; this module keeps the playbook and executed evidence distinct (per the verification standard) and feeds the release evidence package (Stage 16).
8. Coverage/Completion Notes
Complete when protocols, raw data, calibration certificates, and review/approval records exist for every executed PV activity and the evidence index is release-ready. No evidence exists yet (pre-execution). All Planned.
PVL-500 (PV-FUNC) — Functional Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-007 (Embedded SW) + EM-ENG-009 (Cloud/Mobile) with System-Safety and Human-Factors. Verifies the SCM's functional behaviour end-to-end: capture → detect → warn → decide → prove. This module hosts the functional device tests (PV-FUNC-###) and, to honour the Stage-12 verification references, the connectivity (PV-COMM-###), cloud (PV-CLOUD-###), data/decision (PV-DATA-###), use/HMI (PV-USE-###), and service-mode (PV-SVC-###) functional activities. Safety-critical software fault-injection lives in PVL-1100 (PV-SW); this module cross-references it. All entries Planned/Evidence-Pending.
1. Purpose
Verify each functional capability performs correctly and its interfaces behave as specified: excursion detection & verdict, alarm generation + delivery, gap-free logging, OTA, communications, offline store-and-forward, cloud ingest/notify/reconciliation, mobile provisioning, and factory/service/diagnostic modes.
2. Scope
Functional and interface behaviour at Unit → Integration → System levels across S3–S9. Excludes performance timing/margins (PVL-600), fault-injection safety proofs (PVL-1100 PV-SW), and usability effectiveness (Validation, Stage 14).
3. Applicable Standards & References
IEC 62304 (software functional verification); ISO 13485 §7.3.6; the device↔cloud ICD (IFR); ALCOA+ (record integrity); WHO PQS E006 (detection/accuracy basis, to confirm). Designations; revisions to confirm.
4. Governance & Method
Method is Test (exercise + observe) and Demonstration (observe intended operation). Reference stimuli (excursion/shock/energy profiles, radio-loopback) come from DO-VER-007; observability from DO-VER-004/006. Functional tests establish correctness; the FI safety proofs in PV-SW establish robustness under fault.
5. Verification Categories
Detection & verdict · alarm gen + delivery · logging & records · OTA · communications & store-and-forward · cloud ingest/notify/reconciliation · mobile provisioning · data provenance & export · HMI/use · factory/ service/diagnostic modes · error handling & recovery · regression.
6. Verification Register (this module)
6.1 Detection & verdict (functional correctness)
PV-FUNC-001 — Two-sided excursion detection functional correctness
Objective | Confirm the device detects heat and freeze single-point and cumulative excursions against the active profile and registers them at the sampling cadence, independent of connectivity.
Method | Test
I/O Trace | SYS-010/014/038/149/150; PRD-FUN-006/007; DO-FW-006
Risk-Control Trace | RCR-019 (two-sided detection), HA-SW-001
Acceptance Criteria | Against the DO-VER-007 profiled-excursion battery, every seeded heat/freeze/ cumulative excursion is detected within one logging interval; 0 missed excursions in the functional set. (Freeze-equal-heat coverage and undetected-path proof are the FI test PV-SW-001.)
Level | Subsystem
Sample/Config | DVT units + reference temperature profiles
Facility/Equipment/Tools | Climatic chamber; temperature reference; profile injector
Independence | Embedded SW; System-Safety reviews
Standard/Compliance Reference | WHO PQS E006 (to confirm)
Evidence Required | Detection functional-test report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | Functional correctness; robustness = PV-SW-001 (FI).
PV-FUNC-002 — Offline accept/reject/quarantine verdict computation & display
Objective | Confirm the device computes and displays the accept/reject/quarantine verdict from the stored trip record against the active profile with no network/phone/login.
Method | Test, Demonstration
I/O Trace | SYS-015/042/151; PRD-FUN-010; DO-FW-008, DO-HF-002
Risk-Control Trace | RCR-021 (verdict integrity), RCR-043 (legible verdict)
Acceptance Criteria | For known trip records the on-device verdict matches the expected verdict and renders without any network/phone/login; full-record verdict computes within the responsiveness budget (measured in PV-PERF-004). Gap-suppression robustness is PV-SW-002 (FI).
Level | Subsystem
Sample/Config | DVT units + known trip datasets
Facility/Equipment/Tools | Reference record injector
Independence | Embedded SW; Human-Factors reviews
Standard/Compliance Reference | —
Evidence Required | Verdict functional-test report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | —
PV-FUNC-003 — Probe-fault / implausible-read functional flagging
Objective | Confirm absent/open/short/wrong probe and implausible reads are functionally detected and flagged, and internal-vs-external divergence raises a probe-fault flag.
Method | Test
I/O Trace | SYS-007/024/151; HWR-022/028; DO-FW-007, DO-EE-013
Risk-Control Trace | RCR-020 (probe-fault detection), HA-SW-008
Acceptance Criteria | Each induced probe condition is flagged and never scored as a valid product-zone temperature; divergence beyond threshold flags a fault. (Full fault-injection sweep is PV-SW-008.)
Level | Subsystem
Sample/Config | DVT units + probe fault harness
Facility/Equipment/Tools | Probe-emulation fixture
Independence | Embedded SW + Sensing
Standard/Compliance Reference | —
Evidence Required | Probe-fault functional report
Owner | Embedded SW + Sensing
Status | Planned
Notes | Robustness = PV-SW-008 (FI).
PV-FUNC-004 — Pre-arm readiness gate & arming/void behaviour
Objective | Confirm arming is blocked unless SoC-for-trip, valid calibration, valid firmware/profile, and probe-present all pass, and the un-armed/void state is unmistakable.
Method | Test, Demonstration
I/O Trace | SYS-023/024; PRD-SAFE-006; DO-FW-032, DO-SAFE-007
Risk-Control Trace | RCR-045 (pre-arm gate), RCR-048 (profile validation)
Acceptance Criteria | No trip arms with any failing readiness item; blocked arming shows its cause; the void state is unmistakable on HMI. (Non-defeatability under fault is PV-SW-003.)
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Test-mode injector
Independence | Firmware + Human-Factors
Standard/Compliance Reference | —
Evidence Required | Pre-arm gate functional report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | —
6.2 Logging & records
PV-FUNC-005 — Gap-free interval logging & event capture
Objective | Confirm the device logs every scheduled sample at the configured interval and captures shock/ event records without gaps or drops over a full trip.
Method | Test
I/O Trace | SYS-001/002/044; PRD-FUN-001/004; DO-FW-010
Risk-Control Trace | RCR-024 (no lost record), RCR-022 (honest gap)
Acceptance Criteria | Over an extended run the sample count equals expected with 0 gaps; concurrent shock events captured without dropping base samples. (Power-fail robustness is PV-SW-013.)
Level | Subsystem
Sample/Config | DVT units + long-run harness
Facility/Equipment/Tools | Automated logging harness; shock stimulator
Independence | Firmware
Standard/Compliance Reference | ALCOA+
Evidence Required | Logging-integrity report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | —
PV-DATA-001 — Calibration coefficient application & data-integrity functional check
Objective | Confirm calibration coefficients are correctly applied to raw readings and that a units/ coefficient error is caught by the integrity check (functional path).
Method | Test
I/O Trace | SYS-007/024; PRD-FUN (calibration); DO-FW-007/009
Risk-Control Trace | RCR-026 (calibration integrity), HA-SW-014
Acceptance Criteria | Applied readings match the reference within the accuracy budget for known coefficients; a corrupted/invalid coefficient set is detected. (Expiry/blocks-arming FI is PV-SW-004.)
Level | Subsystem
Sample/Config | DVT units + reference bath
Facility/Equipment/Tools | Temperature reference; coefficient injector
Independence | Metrology + Firmware
Standard/Compliance Reference | ISO/IEC 17025 (traceability)
Evidence Required | Calibration-application report
Owner | Metrology + Firmware
Status | Planned
Notes | RCR-026 verification reference (functional half; FI half = PV-SW-004).
PV-DATA-004 — Monitoring-gap exposure in record & status (functional)
Objective | Confirm any monitoring gap/outage/reset is made visible in the record and on-device status and never rendered as a clean pass.
Method | Test
I/O Trace | SYS-160/168; PRD-SAFE-003; DO-SAFE-006/008
Risk-Control Trace | RCR-022 (honest gap), RCR-040 (disposition qualifier)
Acceptance Criteria | Each induced gap appears in the record and status; no incomplete record renders as a clean pass on device or cloud. (Fault-forced version is PV-SW-002/013.)
Level | Integration
Sample/Config | DVT device + cloud
Facility/Equipment/Tools | Gap-injection harness
Independence | Firmware + Cloud
Standard/Compliance Reference | ALCOA+
Evidence Required | Gap-exposure functional report
Owner | Firmware + Cloud
Status | Planned
Notes | RCR-022/040 verification reference.
PV-DATA-005 — Disposition-surface visualisation review (no gap shown as pass)
Objective | Confirm the disposition surface shows the incomplete-monitoring qualifier and never presents a gap/fault as a clean pass.
Method | Inspection, Demonstration
I/O Trace | PRD-UX; PRD-SAFE-003/005; DO-HF-002, DO-CLD
Risk-Control Trace | RCR-040 (disposition qualifier)
Acceptance Criteria | Review of every disposition-surface state confirms a gap/fault always carries the qualifier and cannot render as a clean pass. (Comprehension is validated in VAL-USE-004.)
Level | System
Sample/Config | Cloud/mobile UI build
Facility/Equipment/Tools | UI review harness
Independence | Human-Factors + Cloud
Standard/Compliance Reference | IEC 62366-1 (design); ALCOA+
Evidence Required | Visualisation review record
Owner | Human-Factors + Cloud
Status | Planned
Notes | RCR-040 verification reference.
PV-DATA-006 — Export/API preserves attribution, time, gaps, tamper-evidence
Objective | Confirm export and API preserve attribution, time, gaps, and tamper-evidence with no silent loss, including after a power-fail during commit.
Method | Test
I/O Trace | SYS (API/audit); PRD-CLOUD; PRD-SAFE-003; DO-CLD (API/audit)
Risk-Control Trace | RCR-042 (export integrity), RCR-057 (record spine)
Acceptance Criteria | Exported/queried records preserve attribution, trusted time, gap markers, and tamper-evidence in the documented format; byte-level comparison shows no silent loss; format audit passes.
Level | Integration
Sample/Config | Device + cloud + export
Facility/Equipment/Tools | Export/API test harness
Independence | Cloud + Quality/Regulatory
Standard/Compliance Reference | ALCOA+ / 21 CFR Part 11-style
Evidence Required | Export-integrity report
Owner | Cloud + Quality (EM-ENG-006)
Status | Planned
Notes | RCR-042 verification reference.
PV-DATA-007 — Estimated/interpolated/last-known data flagged as non-measured
Objective | Confirm estimated/interpolated/last-known values (GNSS, gap-fill) are flagged and never presented as measured/authoritative.
Method | Inspection, Test
I/O Trace | PRD-SAFE-003; PRD-UX; DO-CLD, DO-HF
Risk-Control Trace | RCR-041 (provenance)
Acceptance Criteria | Every estimated/interpolated/last-known value is visibly flagged in UI and export; provenance carried through export.
Level | System
Sample/Config | Cloud/mobile + export
Facility/Equipment/Tools | Provenance test harness
Independence | Human-Factors + Cloud
Standard/Compliance Reference | ALCOA+
Evidence Required | Provenance-flag report
Owner | Human-Factors + Cloud
Status | Planned
Notes | RCR-041 verification reference.
PV-DATA-008 — Verdict-override workflow (deliberate, confirmed, recorded)
Objective | Confirm a verdict override requires a deliberate, confirmed, recorded action and that accidental acceptance of a compromised dose is not possible.
Method | Test, Demonstration
I/O Trace | PRD-UX; PRD-SAFE-005; DO-HF, DO-CLD (audit)
Risk-Control Trace | RCR-044 (override control)
Acceptance Criteria | No accidental acceptance path exists; a deliberate override captures actor, reason, and time in the audit trail; override is monitorable. (Effectiveness validated in VAL-USE-007.)
Level | System
Sample/Config | Cloud/mobile build
Facility/Equipment/Tools | Workflow test harness
Independence | Human-Factors + Cloud
Standard/Compliance Reference | IEC 62366-1 (design); ALCOA+
Evidence Required | Override-workflow test report
Owner | Human-Factors + Cloud
Status | Planned
Notes | RCR-044 verification reference.
6.3 Alarm generation & delivery / communications / store-and-forward
PV-COMM-001 — Offline-excursion capture & store-and-forward, alert-first on reconnect
Objective | Confirm an excursion during a total comms outage is committed locally and delivered alert-first on reconnection with no record loss.
Method | Test
I/O Trace | SYS-013/040/111; PRD-FUN-008; D2; DO-FW-011
Risk-Control Trace | RCR-028 (offline-first commit), RCR-030 (store-and-forward), RCR-029 (alert-ahead)
Acceptance Criteria | Across induced link outages the buffer loses 0 records and delivers the alert ahead of bulk on reconnection; buffer depth covers the worst-case outage.
Level | Integration
Sample/Config | Device + emulated network + cloud
Facility/Equipment/Tools | Network emulator; cloud test tenant
Independence | Firmware + Connectivity + Cloud
Standard/Compliance Reference | D2 (store-and-forward)
Evidence Required | Store-and-forward test report
Owner | Firmware + Connectivity
Status | Planned
Notes | RCR-028/030 verification reference; buffer depth bounded by PRD-VER-005 result.
PV-COMM-002 — Cloud never drops a safety event (dead-letter/retry/reconciliation)
Objective | Confirm no safety event is silently discarded under ingest/notify failure — every event is dead-lettered/retried and reconciled end-to-end, with idempotent duplicate ingest.
Method | Test
I/O Trace | SYS-176; PRD-CLOUD; PRD-SAFE-003; DO-CLD-007 (SRS-CLOUD-023)
Risk-Control Trace | RCR-032 (no dropped event), HA-CLOUD-001
Acceptance Criteria | Under injected ingest/notify failures, 0 safety events lost; all reconciled; duplicate ingest is idempotent. (Fault-injection depth shared with PV-CLOUD-001.)
Level | Integration
Sample/Config | Cloud test tenant + fault injector
Facility/Equipment/Tools | Chaos/fault-injection harness
Independence | Cloud; Quality reviews
Standard/Compliance Reference | —
Evidence Required | Delivery-reconciliation report
Owner | Cloud (EM-ENG-009)
Status | Planned
Notes | RCR-032 verification reference (FI at cloud tier).
PV-COMM-003 — RF-in-container data-egress / alert-delivery functional test
Objective | Confirm data egress and alert delivery from inside a representative metallized/insulated shipper across the target radios/bands (functional companion to the reliability qualification PV-REL-005).
Method | Test
I/O Trace | SYS-080/104/109/117; PRD-CONN-007; PRD-VER-005; DO-VER-009
Risk-Control Trace | RCR-031 (RF closes budget or external-antenna fallback)
Acceptance Criteria | With the device inside the representative shipper, buffered data and queued alerts are delivered whenever a usable link is available; the integral or external-antenna configuration is identified. In-container margin/antenna decision Evidence-Pending → PRD-VER-005 / PV-REL-005 (O-1).
Level | System
Sample/Config | DVT device inside representative shipper(s)
Facility/Equipment/Tools | Representative shippers; RF test rig; live/field networks
Independence | Connectivity + RF/Antenna
Standard/Compliance Reference | FCC/CE band plans; PRD-VER-005
Evidence Required | RF-in-container delivery report
Owner | Connectivity + RF/Antenna
Status | Evidence-Pending
Notes | RCR-031 verification reference; Open item O-1.
PV-COMM-004 — Multi-bearer failover & opportunistic offload (LTE-M/Wi-Fi/BLE)
Objective | Confirm the device selects LTE-M primary and opportunistically offloads via Wi-Fi/BLE, and fails over without record loss.
Method | Test
I/O Trace | SYS-111/113; PRD-CONN; DO-FW-011, DO-IF (RF)
Risk-Control Trace | RCR-030 (degrade without loss), RCR-032
Acceptance Criteria | Bearer selection follows the specified priority; failover between bearers loses 0 records; offload occurs when a usable bearer appears.
Level | Integration
Sample/Config | Device + multi-bearer emulator
Facility/Equipment/Tools | Cellular/Wi-Fi/BLE emulation
Independence | Connectivity + Firmware
Standard/Compliance Reference | —
Evidence Required | Failover test report
Owner | Connectivity + Firmware
Status | Planned
Notes | —
PV-COMM-006 — Regional/band configuration & roaming behaviour
Objective | Confirm the device operates on the correct regional band plan and roams/attaches per market configuration, with wrong-region config caught.
Method | Test
I/O Trace | SYS-027/035/115/116; PRD-COMM; DO-FW, DO-MFG (config)
Risk-Control Trace | RCR-056 (correct regional/firmware config), HA-COMM-006
Acceptance Criteria | Correctly-configured units attach/roam on the target bands; a wrong-region variant is detected/blocked (config verified at EOL by PV-MFG-006).
Level | System
Sample/Config | Regional variant units
Facility/Equipment/Tools | Multi-band network emulation; live roaming
Independence | Connectivity + Manufacturing
Standard/Compliance Reference | regional band plans
Evidence Required | Roaming/config test report
Owner | Connectivity + Manufacturing
Status | Planned
Notes | RCR-056 verification reference.
PV-COMM-008 — Device↔trip↔owner↔responder association & alert routing
Objective | Confirm alerts route to the correct bound responder and no cross-tenant mis-binding occurs.
Method | Test
I/O Trace | SYS-012/030/134; PRD-CLOUD; PRD-SAFE-001; DO-CLD (registry), DO-SEC
Risk-Control Trace | RCR-033 (authoritative binding), RCR-055 (unique identity)
Acceptance Criteria | Every alert routes to the bound responder; cross-tenant mis-binding attempts are rejected; unique identity anchors the binding.
Level | Integration
Sample/Config | Multi-tenant cloud test
Facility/Equipment/Tools | Cloud test tenants
Independence | Cloud + Security
Standard/Compliance Reference | —
Evidence Required | Association/routing test report
Owner | Cloud + Security
Status | Planned
Notes | RCR-033 verification reference.
PV-CLOUD-001 — Cloud ingest/notify pipeline functional & fault-injection
Objective | Confirm the cloud ingest→process→notify pipeline functions correctly and, under injected component failures, does not drop or corrupt safety events.
Method | Test
I/O Trace | SRS-CLOUD (400); PRD-CLOUD; DO-CLD-007/008
Risk-Control Trace | RCR-032 (no dropped event), HA-CLOUD-001/002
Acceptance Criteria | Pipeline processes events correctly; under injected failures (queue, worker, notifier) 0 safety events lost; alerts still delivered/escalated.
Level | Integration
Sample/Config | Cloud test tenant
Facility/Equipment/Tools | Chaos/fault-injection harness
Independence | Cloud; Quality reviews
Standard/Compliance Reference | —
Evidence Required | Cloud pipeline test report
Owner | Cloud (EM-ENG-009)
Status | Planned
Notes | RCR-032 verification reference (cloud FI).
PV-CLOUD-003 — Cloud verdict recompute & device↔cloud consistency
Objective | Confirm the cloud independently recomputes the verdict from the record and agrees with the device, or the disagreement resolves to qualify/reject (never a clean accept).
Method | Test
I/O Trace | SYS-015/016/017; PRD-CLOUD; DO-CLD-006
Risk-Control Trace | RCR-021 (verdict integrity — diverse second computation)
Acceptance Criteria | For a corpus of trip records the cloud recompute matches the device verdict, or any disagreement resolves conservatively; no gap yields a clean accept on either side.
Level | Integration
Sample/Config | Device + cloud + record corpus
Facility/Equipment/Tools | Record corpus; cloud tenant
Independence | Cloud + Embedded SW
Standard/Compliance Reference | —
Evidence Required | Recompute-consistency report
Owner | Cloud + Embedded SW
Status | Planned
Notes | Diverse-computation half of RCR-021 (device FI half is PV-SW-002).
6.4 OTA
PV-FUNC-010 — OTA update functional flow (download, apply, A/B, resume)
Objective | Confirm the OTA flow downloads, verifies, applies to the inactive A/B slot, switches, and resumes normal operation without losing trip data.
Method | Test
I/O Trace | SYS-032/114/139; SRS-OTA (700); DO-FW-014/015, DO-CLD-010
Risk-Control Trace | RCR-037 (signed, fail-safe A/B OTA)
Acceptance Criteria | A valid signed image is downloaded, applied, and activated; trip data preserved; unmonitored update window disclosed. (Rollback/version-skew robustness = PV-SW-010; signing pen-test = PV-CYBER-005.)
Level | Integration
Sample/Config | Device + cloud OTA service
Facility/Equipment/Tools | OTA test harness
Independence | Firmware + Cloud
Standard/Compliance Reference | IEC 62304 (change)
Evidence Required | OTA functional report
Owner | Firmware + Cloud
Status | Planned
Notes | Robustness/security split to PV-SW-010 / PV-CYBER-005.
6.5 Mobile provisioning
PV-FUNC-011 — BLE provisioning & trip profile selection (mobile)
Objective | Confirm the mobile app provisions the device over BLE, selects a trip profile validated against the stability library, and arms via the readiness gate.
Method | Test, Demonstration
I/O Trace | SRS-MOB (500); PRD-MOB; DO-APP, DO-IF (BLE-GATT)
Risk-Control Trace | RCR-048 (profile validation), RCR-036 (authenticated config)
Acceptance Criteria | Provisioning completes over BLE; an implausible/mismatched profile is flagged and blocks arming; only authenticated changes succeed. (Auth robustness = PV-CYBER-003.)
Level | Integration
Sample/Config | Device + mobile app
Facility/Equipment/Tools | Mobile test devices
Independence | Mobile + Firmware
Standard/Compliance Reference | —
Evidence Required | Provisioning test report
Owner | Mobile (EM-ENG-009) + Firmware
Status | Planned
Notes | —
6.6 HMI / use (functional legibility)
PV-USE-001 — On-device verdict/status cold-display legibility (functional)
Objective | Confirm the verdict and status are legibly rendered (icon+colour+word) and update within the responsiveness budget down to the −30 °C operating extreme.
Method | Test
I/O Trace | SYS-041/074; PRD-UX-001/007; PRD-ENV-003; DO-HF-002/005, DO-SAFE-013
Risk-Control Trace | RCR-043 (legible verdict)
Acceptance Criteria | Display renders the verdict legibly with adequate contrast and update latency within the responsiveness budget at −30 °C; measured contrast/latency meet the DO-VER-003 thresholds. (Gloved low-light summative success rate is validated in VAL-USE-001.)
Level | Subsystem
Sample/Config | DVT units in cold chamber
Facility/Equipment/Tools | Climatic chamber; photometer; timer
Independence | Human-Factors + Display/HMI
Standard/Compliance Reference | IEC 62366-1 (design)
Evidence Required | Cold-legibility test report
Owner | Human-Factors + HMI
Status | Evidence-Pending
Notes | RCR-043 verification reference; cold-legibility threshold Evidence-Pending (SYS-074).
PV-USE-002 — Alarm-management / consolidation functional behaviour
Objective | Confirm alert consolidation and escalation behave per the alarm-management design so a real excursion is not muted by volume.
Method | Test
I/O Trace | SYS-011/013; PRD-CLOUD-003; PRD-SAFE-004; DO-CLD-003, DO-SAFE-004
Risk-Control Trace | RCR-047 (alarm tuning), RCR-029 (escalation)
Acceptance Criteria | Consolidation holds alert volume within the configured limit while a real excursion always produces a distinct, escalating alert; no real alert is suppressed by consolidation. (False-alarm/ false-reject targets are validated in VAL-USE-002.)
Level | Integration
Sample/Config | Cloud tenant + alert scenarios
Facility/Equipment/Tools | Alert-scenario harness
Independence | Human-Factors + Cloud
Standard/Compliance Reference | —
Evidence Required | Alarm-management functional report
Owner | Human-Factors + Cloud
Status | Planned
Notes | RCR-047 verification reference (functional half; targets validated Stage 14).
PV-USE-003 — State-legibility: charge-inhibited/fault/unready (cold ≠ dead)
Objective | Confirm charge-inhibited, fault, and unready states are shown with their cause on LED/OLED so a cold, charge-inhibited device is not mistaken for dead.
Method | Test, Demonstration
I/O Trace | SYS-024/094/100; PRD-SAFE-005/007; DO-HF-005, DO-EE-008
Risk-Control Trace | RCR-046 (state legibility)
Acceptance Criteria | Each state renders a distinct, cause-carrying indication; the charge-inhibited state is visibly distinct from a dead device. (Comprehension validated in VAL-USE.)
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | State injector
Independence | Human-Factors
Standard/Compliance Reference | —
Evidence Required | State-legibility report
Owner | Human-Factors
Status | Planned
Notes | RCR-046 verification reference.
6.7 Factory / service / diagnostic modes
PV-SVC-001 — Recalibration-interval enforcement (pre-arm expiry gate)
Objective | Confirm an out-of-calibration (expired) device cannot be armed — the interval is enforced by the pre-arm expiry gate.
Method | Test
I/O Trace | SYS-024; PRD-SVC; PRD-SAFE-006; DO-FW-009, DO-SVC
Risk-Control Trace | RCR-049 (interval enforcement), RCR-026
Acceptance Criteria | A device past its calibration expiry is blocked from arming with the cause shown; the depot recal-escape is caught by the gate. (Integrity/expiry FI = PV-SW-004.)
Level | Subsystem
Sample/Config | DVT units with set calibration dates
Facility/Equipment/Tools | Time/calibration injector
Independence | Metrology + Service
Standard/Compliance Reference | —
Evidence Required | Interval-enforcement report
Owner | Metrology + Service
Status | Planned
Notes | RCR-049 verification reference.
PV-SVC-002 — Post-service BIST & restore-to-ready verification
Objective | Confirm service restores seal/config/firmware and a post-service BIST + pre-arm gate verifies integrity before redeploy; a mis-serviced device cannot arm.
Method | Test, Demonstration
I/O Trace | PRD-SVC; PRD-SAFE-010; DO-SVC, DO-TST
Risk-Control Trace | RCR-050 (service restore + BIST)
Acceptance Criteria | After the service procedure, seal/config/firmware are verified restored by BIST + pre-arm gate; a deliberately mis-serviced unit is blocked from arming.
Level | Subsystem
Sample/Config | Serviced DVT units
Facility/Equipment/Tools | Service fixture; BIST
Independence | Service + Quality
Standard/Compliance Reference | —
Evidence Required | Post-service BIST report
Owner | Service + Quality
Status | Planned
Notes | RCR-050 verification reference.
PV-FUNC-012 — Factory-mode & diagnostic-observability functional check
Objective | Confirm factory/verification test modes (forced excursion, forced cold-cell, sensor injection, radio-loopback) expose acceptance-critical behaviour and are governed.
Method | Test
I/O Trace | DO-VER-006 (test modes), DO-FW-012 (BIST); SRS-DIAG (1400)
Risk-Control Trace | RCR-051 (mode lock), DO-VER-006 (no fielded bypass)
Acceptance Criteria | Each test mode exposes the intended behaviour for verification; modes are authenticated and cannot silently defeat a safety control in the field. (Field-lock pen-test = PV-CYBER-006.)
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Test-mode interface
Independence | Firmware + Quality
Standard/Compliance Reference | —
Evidence Required | Test-mode functional report
Owner | Firmware + Quality
Status | Planned
Notes | Enables many other PV activities; security-locked in field (PV-CYBER-006).
PV-FUNC-013 — Factory-reset & re-provisioning functional check
Objective | Confirm a controlled factory reset clears trip/config state and re-provisioning restores a known-good baseline (reusable-device lifecycle, D1).
Method | Test
I/O Trace | PRD-SVC; SRS-DEV/FLEET; DO-FW, DO-CLD
Risk-Control Trace | RCR-050 (restore), RCR-052 (data/key erasure at EOL)
Acceptance Criteria | Factory reset clears prior trip/config; re-provisioning yields the known-good baseline; no residual prior-trip data leaks into a new trip.
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Provisioning tools
Independence | Firmware + Cloud
Standard/Compliance Reference | —
Evidence Required | Reset/re-provision report
Owner | Firmware + Cloud
Status | Planned
Notes | —
6.8 Error handling & regression
PV-FUNC-014 — Functional error-handling & graceful-degradation check
Objective | Confirm defined error paths (sensor error, storage-near-full, low battery, link loss) degrade gracefully and surface honest status rather than failing silently.
Method | Test
I/O Trace | SYS (error handling); SRS-DEV/REL; DO-FW-013
Risk-Control Trace | RCR-023 (safe state), RCR-027 (availability self-check)
Acceptance Criteria | Each induced error condition produces the specified graceful behaviour and honest status; no silent failure. (Fault-forced safe-state proof = PV-SW-009.)
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Fault harness
Independence | Firmware
Standard/Compliance Reference | IEC 62304
Evidence Required | Error-handling report
Owner | Firmware
Status | Planned
Notes | —
PV-FUNC-015 — Functional regression suite (device + cloud + mobile)
Objective | Confirm a repeatable regression suite re-verifies core functional behaviour after any change.
Method | Test
I/O Trace | PV-GOV-004 (regression policy); all functional requirements
Risk-Control Trace | RCR-037 (change discipline), RCR-019/021 (detection/verdict across change)
Acceptance Criteria | The regression suite covers detection, verdict, logging, alarm delivery, OTA, and comms; it runs in CI and gates release; safety-relevant changes trigger the full safety regression.
Level | System
Sample/Config | Automated regression harness
Facility/Equipment/Tools | CI + device farm
Independence | Quality + Embedded SW
Standard/Compliance Reference | IEC 62304
Evidence Required | Regression-suite definition + run records
Owner | Quality + Embedded SW
Status | Planned
Notes | Ties PV-GOV-004.
7. Traceability Notes
Functional activities verify SYS/SRS/PRD functional requirements and provide the functional half of many RCR controls whose robustness is proven by the fault-injection tests in PVL-1100 (PV-SW). The RCR verification references PV-COMM/PV-CLOUD/PV-DATA/PV-USE/PV-SVC are hosted here. Feeds Validation (Stage 14) which proves these functions meet user needs in representative use.
8. Coverage/Completion Notes
Complete when every functional requirement and the hosted RCR references pass at the appropriate level, with the RF-in-container functional test (PV-COMM-003) closed by PRD-VER-005. All Planned/Evidence-Pending.
PVL-000 (PV-GOV) — Verification Governance — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 13 — Verification (ORCH-014) · Owner: Principal Verification/Reliability/Test/Quality Engineer (EM-ENG-006) with Chief Systems (EM-ENG-001) and the discipline principals. Governance module of the Product Verification Library (PVL). Based on 00-governance/product-verification-standard.md, the concept D1–D4, and the Stage-05/06/07 requirement baselines. Every activity is a playbook entry (planned method + objective acceptance + evidence required); no results are recorded — all evidence is Planned / Evidence-Pending until executed.
1. Purpose
Define how the SCM's verification is governed: the philosophy ("did we build the product correctly?" — each design output satisfies its design input), the method taxonomy, independence rules, acceptance policy, configuration control of the item-under-test, regression policy, deviation handling, and readiness/completion gates. This module is the constitution the other 19 PVL modules obey.
2. Scope
All SCM verification across device (S1–S7), cloud (S8), and mobile (S9); Unit → Subsystem → Integration → System → Production levels; Test / Analysis / Inspection / Demonstration methods. Excludes validation (intended-use effectiveness — Stage 14, PVaL/ORCH-015) and the executed evidence itself (protocols, raw data, reports — referenced from PVL-1800).
3. Applicable Standards & References
ISO 13485:2016 §7.3.6 (design verification); FDA design controls 21 CFR 820.30; ISO/IEC/IEEE 29148 (requirements verification); IEC 62304 (software lifecycle & verification); ISO 14971 (risk-control verification); ISO/IEC 17025 (measurement traceability); GAMP 5 (CSV); AIAG/VDA MSA (measurement-system capability). All cited by designation; revisions to confirm — no clause asserted as fact.
4. Governance & Method
• Method taxonomy. Every activity names one primary method: Test (exercise + measure), Analysis (calculation/model/worst-case/simulation), Inspection (examine drawing/code/hardware/document), Demonstration (observe intended operation). Method allocation is risk-first (PRD-VER-003): safety- and function-critical requirements get Test wherever feasible; Analysis/Inspection in their place is justified in the entry Notes.
• Fault injection (FI). Every safety-critical risk control (RCR marked FI in Stage 12) is verified by fault injection — the fault is forced and the control is proven to hold. FI is mandatory, not optional (HAL-2600).
• Independent adversarial verification (IV). Every security control (RCR-034…039, RCR-051 marked IV) is verified by an independent party (adversarial pen-test / third-party review), not by the implementing engineer.
• Objective acceptance. No subjective safety verdict. Each activity carries a measurable pass/fail rule (DO-VER-010). Where a target value is not yet evidence-backed it is Evidence-Pending with the named closing test — never invented.
5. Verification Categories (governance activities)
Philosophy & policy · independence · acceptance · configuration management of the UUT · regression · deviation/waiver · readiness · completion & release recommendation.
6. Verification Register (this module)
PV-GOV-001 — Verification policy and method-allocation rule established
Objective | Confirm the SCM verification program applies the method taxonomy and severity-first method allocation to every requirement and risk control.
Method | Inspection
I/O Trace | PRD-VER-001, PRD-VER-003; SRS-VER (1500); DO-VER-010
Risk-Control Trace | governs all RCR-### verification; ISO 14971 method rigor
Acceptance Criteria | 100 % of requirements and 58/58 RCRs carry a named method; every FI-flagged RCR has a Test with fault injection and every IV-flagged RCR has an independent activity; no safety-critical requirement is verified by Inspection/Analysis without a documented justification.
Level | System
Sample/Config | Verification plan + coverage matrix (this library)
Facility/Equipment/Tools | PLM/QMS; requirements/trace tool
Independence | Quality (EM-ENG-006) reviews; independent of design authors
Standard/Compliance Reference | ISO 13485 §7.3.6; ISO/IEC/IEEE 29148; ISO 14971
Evidence Required | Approved verification plan; method-allocation audit record
Owner | Quality/Verification (EM-ENG-006)
Status | Planned
Notes | The governing gate for PV-REQ and PV-TRACE.
PV-GOV-002 — Independence of verification demonstrated (author ≠ verifier)
Objective | Confirm each verification is performed/reviewed by a party independent of the design author, with security controls verified by an external adversarial party.
Method | Inspection
I/O Trace | PRD-VER-001; DO-VER-006 (test modes not a fielded bypass)
Risk-Control Trace | RCR-034…039, RCR-051 (IV); all safety RCRs (independent review)
Acceptance Criteria | Every safety-critical and security activity records an independent reviewer/executor; security pen-tests (PV-CYBER-###) are executed by a party independent of firmware/cloud development; independence exceptions are documented and approved.
Level | System
Sample/Config | Verification records + independence register
Facility/Equipment/Tools | QMS; third-party test-house/pen-test engagement
Independence | By definition — this activity audits independence
Standard/Compliance Reference | ISO 13485; IEC 62443 concepts (independent assessment)
Evidence Required | Independence register; third-party engagement records
Owner | Quality/Verification (EM-ENG-006)
Status | Planned
Notes | Test-house and pen-test vendor selection is Evidence-Pending (concept §7).
PV-GOV-003 — Configuration control of the unit-under-test (UUT) baseline
Objective | Ensure every verification is executed against a controlled, identified hardware/firmware/ cloud/mobile configuration baseline so results are attributable and reproducible.
Method | Inspection
I/O Trace | DOL-100 (config baseline); DO-VER-005 (logging schema); SRS-VER
Risk-Control Trace | RCR-055 (unique identity), RCR-056 (correct config)
Acceptance Criteria | Each verification record captures unit serial, HW rev, FW version, cloud/mobile build, calibration state, and profile config (ALCOA+); no verification is accepted against an unidentified/uncontrolled build.
Level | System
Sample/Config | DVT/PVT build records; software build manifests
Facility/Equipment/Tools | PLM/QMS; CI build system; SBOM (PV-SEC)
Independence | Quality/Config management
Standard/Compliance Reference | ISO 13485 §7.5; GAMP 5; ALCOA+
Evidence Required | UUT configuration record per test; build manifest
Owner | Quality/Config Management
Status | Planned
Notes | Ties DO-VER-005 verification logging schema.
PV-GOV-004 — Regression and change-impact verification policy
Objective | Confirm a defined regression policy re-verifies impacted requirements/controls after any design, firmware, or cloud change (especially OTA and detection/verdict logic).
Method | Inspection
I/O Trace | PRD-VER-002 (traceability enables impact analysis); SRS-OTA (700); DO-FW-014/015
Risk-Control Trace | RCR-037 (OTA), RCR-019/021 (detection/verdict integrity across change)
Acceptance Criteria | A documented change-impact procedure maps each change to the requirements/RCRs it touches and mandates re-run of the affected PV activities; safety/verdict/detection changes trigger the full safety regression suite; staged OTA halts on regression (RCR-037).
Level | System
Sample/Config | Change records + regression suite definition
Facility/Equipment/Tools | Trace tool; CI regression harness
Independence | Quality
Standard/Compliance Reference | IEC 62304 (change control); ISO 13485
Evidence Required | Regression policy; change-impact records
Owner | Quality + Embedded SW (EM-ENG-007)
Status | Planned
Notes | Regression scope for OTA is release-critical.
PV-GOV-005 — Deviation, nonconformance, and waiver handling
Objective | Confirm every deviation from a protocol or acceptance criterion is recorded, dispositioned, and (for safety/security) escalated before any release decision.
Method | Inspection
I/O Trace | PRD-VER-001; DO-VER-010; SRS-VER
Risk-Control Trace | RCR-058 (overall residual/benefit-risk gate)
Acceptance Criteria | No verification failure is closed without a recorded root cause and disposition (PVL-1700); no safety/security acceptance criterion is waived without EM-ENG-006 + EM-ENG-001 sign-off and a risk re-assessment (ISO 14971).
Level | System
Sample/Config | Deviation/waiver log
Facility/Equipment/Tools | QMS CAPA/deviation module
Independence | Quality; safety waivers require System-Safety
Standard/Compliance Reference | ISO 13485 §8.3/8.5; ISO 14971
Evidence Required | Deviation/waiver records with dispositions
Owner | Quality/Verification (EM-ENG-006)
Status | Planned
Notes | Feeds PVL-1700 defect management and PVL-1900 summary.
PV-GOV-006 — Verification readiness and completion gates defined
Objective | Establish the entry (readiness) and exit (completion) gates for each verification phase (DVT, PVT) so verification neither starts prematurely nor closes with open safety items.
Method | Inspection
I/O Trace | PRD-VER-001; Verification-Plan.md; DO-VER-009 (chamber/RF config)
Risk-Control Trace | RCR-058; the two Open items (D3, RF-in-container)
Acceptance Criteria | Readiness gate requires released requirements, controlled UUT, calibrated equipment, and approved protocols; completion gate requires all applicable PV activities executed or formally deferred, zero open safety/security nonconformances, and a PVL-1900 summary; the D3 and RF-in-container Open items are explicitly gated.
Level | System
Sample/Config | Phase gate checklists
Facility/Equipment/Tools | QMS; design-review/release board
Independence | Quality + Chief Systems
Standard/Compliance Reference | ISO 13485 design-review; ISO 14971
Evidence Required | Gate review minutes; phase readiness/completion checklists
Owner | Quality (EM-ENG-006) + Chief Systems (EM-ENG-001)
Status | Planned
Notes | Sequencing detailed in Verification-Plan.md.
7. Traceability Notes
Governance activities have no single parent requirement; they enforce the whole-program rules that make every other module's PV-<CODE>-NNN traceable, independent, controlled, and objectively judged. Feeds every module and the PVL-1600 traceability + PVL-1900 summary.
8. Coverage/Completion Notes
Governance is complete when policy, independence, configuration control, regression, deviation handling, and readiness/completion gates are approved and in force before DVT begins. All entries Planned (pre-execution playbook).
PVL-1200 (PV-HW) — Hardware Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: PCB/Electronics (EM-ENG-003) with Power, Sensing, RF. Verifies the electrical/ electronic hardware: rails and power integrity, charge path, protection, sensors and signal chain, connectors, signal integrity, and mechanical-electrical interfaces. Hosts the charge-source robustness test PV-ELEC-002 referenced by Stage-12. Battery-safety fault-injection is in PVL-900; EMC in PVL-800. All entries Planned.
1. Purpose
Verify hardware meets its electrical requirements: power/rail integrity, charge-path behaviour and robustness, sensor/signal-chain accuracy support, connector durability, signal/power integrity, and HW test-point observability.
2. Scope
S2–S7 hardware at Subsystem level. Excludes battery-safety FI (PVL-900), EMC (PVL-800), and long-life reliability/aging (PVL-1400).
3. Applicable Standards & References
IEC 62368-1 (electrical safety limits); USB-C / USB-PD specification (charge-source); IPC (assembly); component datasheets; ISO/IEC 17025 (references). Designations; revisions to confirm.
4. Governance & Method
Method is Test (bench measurement) and Inspection (test-point/observability). Worst-case predictions come from PV-ANAL-005; safety-critical charge/protection behaviour is confirmed by FI in PV-ENRG-001/002.
5. Verification Categories
Rails/power integrity · charge path · protection thresholds · sensor/signal chain · connectors · signal integrity · thermal (bench) · test-point observability.
6. Verification Register (this module)
PV-HW-001 — Power-rail integrity & brown-out barrier
Objective | Confirm all rails meet voltage/ripple/sequencing limits across load and temperature, and the brown-out barrier asserts at the specified threshold.
Method | Test
I/O Trace | HWR (rails); DO-EE (rails), DO-EE-002 (rail interlock)
Risk-Control Trace | RCR-024 (brown-out barrier), RCR-013 (SELV)
Acceptance Criteria | Rails within tolerance/ripple/sequence across load and −30…+60 °C; brown-out barrier asserts at the designed threshold, protecting the write path.
Level | Subsystem
Sample/Config | DVT PCBAs
Facility/Equipment/Tools | Bench PSU, scope, electronic load, chamber
Independence | Electrical (EM-ENG-003)
Standard/Compliance Reference | IEC 62368-1 (SELV)
Evidence Required | Rail-integrity report
Owner | Electrical (EM-ENG-003)
Status | Planned
Notes | Confirms PV-ANAL-005; ties PV-SW-013.
PV-ELEC-002 — Charge-source robustness (reverse / OVP / surge / short) & PD cap
Objective | Confirm the device rejects/safely limits an out-of-spec charge source (reverse, over-voltage, surge, short) and the USB-PD request is capped to the qualified input envelope.
Method | Test
I/O Trace | SYS-102; PRD-ELEC (USB-C sink); DO-EE-009/010
Risk-Control Trace | RCR-003; HA-ENRG-005, HA-ELEC-002
Acceptance Criteria | No cell/BMS damage or hazardous condition under reverse-insertion, OVP to the rated ceiling, surge, and output short; PD PDO never exceeds the qualified envelope. PD cap value Evidence-Pending → DO-EE-009.
Level | Subsystem
Sample/Config | DVT units + programmable source
Facility/Equipment/Tools | Programmable USB-C/PD source; surge/short rig
Independence | Electrical/Power
Standard/Compliance Reference | USB-PD spec; IEC 62368-1
Evidence Required | Charge-source robustness report
Owner | Electrical/Power (EM-ENG-003)
Status | Evidence-Pending
Notes | RCR-003 verification reference; complements PV-EMC-006.
PV-HW-003 — Charge-path functional & thermal behaviour (permitted window)
Objective | Confirm the charge path charges correctly within the permitted temperature window and manages thermal behaviour.
Method | Test
I/O Trace | HWR (charge); PRD-ELEC-004; DO-EE-007
Risk-Control Trace | RCR-001 (inhibit boundary), RCR-002
Acceptance Criteria | Correct charge profile within the permitted window; charge inhibited below the qualified limit (FI proof = PV-ENRG-001); charge-path temperatures within limits.
Level | Subsystem
Sample/Config | DVT units in chamber
Facility/Equipment/Tools | Charger; chamber; thermocouples
Independence | Power
Standard/Compliance Reference | cell datasheet; IEC 62368-1
Evidence Required | Charge-path report
Owner | Battery/Power
Status | Planned
Notes | Safety boundary FI = PV-ENRG-001.
PV-HW-004 — Protection-threshold bench verification (OV/UV/OC/OT/short)
Objective | Confirm the primary protection IC trips at the designed OV/UV/OC/OT/short thresholds across temperature.
Method | Test
I/O Trace | HWR (protection); PRD-ELEC-007/014; DO-EE-005
Risk-Control Trace | RCR-002 (autonomous protection)
Acceptance Criteria | Each protection function trips within tolerance of the designed threshold across temperature; recovery per spec. (Single-fault backstop FI = PV-ENRG-002.)
Level | Subsystem
Sample/Config | DVT packs
Facility/Equipment/Tools | Source-measure; chamber
Independence | Battery/Power
Standard/Compliance Reference | IEC 62368-1; cell datasheet
Evidence Required | Protection-threshold report
Owner | Battery/Power
Status | Planned
Notes | —
PV-HW-005 — Sensor & signal-chain bench accuracy support
Objective | Confirm the sensor front-ends and signal chain support the ±0.5 °C / ±5 %RH accuracy at the hardware level (noise, offset, gain, self-heating).
Method | Test
I/O Trace | HWR-022/028 (sensors); SYS-033/036; DO-EE-013
Risk-Control Trace | RCR-019/020 (trustworthy sensing), RCR-053
Acceptance Criteria | Signal-chain noise/offset/gain and self-heating within the error budget allocated to hardware; probe front-end supports presence/fault detection.
Level | Subsystem
Sample/Config | DVT PCBAs + reference sources
Facility/Equipment/Tools | Precision references; chamber
Independence | Sensing/Electrical
Standard/Compliance Reference | ISO/IEC 17025
Evidence Required | Signal-chain report
Owner | Sensing/Electrical
Status | Planned
Notes | Feeds PV-PERF-001 accuracy.
PV-HW-006 — Connector & USB-C port electrical / durability verification
Objective | Confirm the USB-C and probe/service connectors meet electrical spec and mating-cycle durability.
Method | Test
I/O Trace | HWR (connectors); PRD-ELEC; DO-EE-009
Risk-Control Trace | RCR-011 (port ingress), RCR-003
Acceptance Criteria | Connectors meet electrical spec and survive the rated mating cycles without contact degradation; port retains IP when closed (IP test = PV-ENV-005).
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Mating-cycle rig; contact-resistance meter
Independence | Electrical/Mechanical
Standard/Compliance Reference | USB-C connector spec
Evidence Required | Connector durability report
Owner | Electrical/Mechanical (EM-ENG-003)
Status | Planned
Notes | Durability life also in PV-REL.
PV-HW-007 — Fuel-gauge / SoC accuracy bench (cold behaviour)
Objective | Confirm the fuel gauge reports SoC within a bounded error and does not over-report SoC in cold.
Method | Test
I/O Trace | SYS-097; PRD-ELEC; DO-EE-008
Risk-Control Trace | RCR-006 (honest energy state)
Acceptance Criteria | SoC error within the bound across temperature; no over-report in cold beyond the bound. Cold SoC error bound Evidence-Pending → PV-REL-001.
Level | Subsystem
Sample/Config | DVT units in chamber
Facility/Equipment/Tools | Chamber; coulomb-counter reference
Independence | Power
Standard/Compliance Reference | cell/gauge datasheet
Evidence Required | Fuel-gauge accuracy report
Owner | Battery/Power
Status | Evidence-Pending
Notes | Supports RCR-006 honest indication.
PV-HW-008 — Signal & power integrity (high-speed, radio, sensor nets)
Objective | Confirm signal/power integrity on high-speed, radio, and sensor nets meets design margins.
Method | Test, Analysis
I/O Trace | HWR; DO-PCB-010; DO-EE-019
Risk-Control Trace | RCR-014 (clean decision path)
Acceptance Criteria | Eye/margin, PDN impedance, and coupling within limits; sensor nets clean enough to support the accuracy budget.
Level | Subsystem
Sample/Config | DVT PCBAs
Facility/Equipment/Tools | Scope/VNA/TDR
Independence | PCB/Electronics (EM-ENG-003)
Standard/Compliance Reference | —
Evidence Required | SI/PI report
Owner | PCB/Electronics (EM-ENG-003)
Status | Planned
Notes | —
PV-HW-009 — Test-point & diagnostic-observability verification
Objective | Confirm physical test points and firmware-exposed diagnostics make every acceptance-critical net/state observable for verification and production test.
Method | Inspection, Test
I/O Trace | DO-VER-004; DO-PCB-011; DO-FW-012
Risk-Control Trace | enables verification of RCR-001/002/020/024 (safety nets/states)
Acceptance Criteria | All acceptance-critical nets/states (charge-inhibit, protection, probe-presence, safe-state) are observable via test point or diagnostic; coverage reviewed.
Level | Subsystem
Sample/Config | DVT PCBAs
Facility/Equipment/Tools | Bed-of-nails / diagnostic interface
Independence | Quality/Electrical
Standard/Compliance Reference | —
Evidence Required | Observability coverage record
Owner | Electrical + Quality
Status | Planned
Notes | Serves production test (PVL-1300) too.
7. Traceability Notes
Hardware activities verify the HWR/electrical requirements and provide the bench basis for battery-safety FI (PV-ENRG), EMC (PV-EMC), and accuracy (PV-PERF). PV-ELEC-002 is the RCR-003 verification reference. Feeds manufacturing test (PVL-1300) and reliability (PVL-1400).
8. Coverage/Completion Notes
Complete when all rails/charge/protection/sensor/connector/SI-PI tests pass on production-representative hardware. PD cap and cold SoC bound Evidence-Pending. All Planned/Evidence-Pending.
PVL-300 (PV-INSP) — Inspection Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with Mechanical (EM-ENG-002), PCB/Electronics (EM-ENG-003), Embedded SW (EM-ENG-007). Verifies conformance by examining drawings, hardware, PCB, code, documents, labels, packaging, and supplier documentation — where the truth is established by examination rather than by exercise. All entries Planned.
1. Purpose
Verify by inspection those characteristics best proven by examination: mechanical dimensions/GD&T, PCB fabrication and assembly conformance, code review against coding standards, documentation/label/packaging correctness, and supplier documentation completeness.
2. Scope
Enclosure/probe/HMI drawings (S1), PCBA (S2–S7), firmware/software source, labels & transport markings, packaging, and incoming supplier documentation. Excludes functional exercise (PVL-500), performance measurement (PVL-600), and production-line inspection setup (PVL-1300).
3. Applicable Standards & References
IPC-A-610 / IPC-A-600 (PCB/PCBA acceptance); ASME Y14.5 (GD&T); IEC 62368-1 (markings/accessibility); UN 38.3 / IATA DGR (transport marking); IEC 62304 (code review); RoHS/REACH (materials declarations). Designations; revisions to confirm.
4. Governance & Method
Method is Inspection throughout; sampling per an AQL/coverage plan for production-representative units. Safety-relevant inspections (label warnings, transport marking, containment features) are 100 %.
5. Verification Categories
Drawing/GD&T · mechanical · electrical/PCB · assembly · code review · documentation · label & packaging · supplier documentation.
6. Verification Register (this module)
PV-INSP-001 — Enclosure & probe critical-dimension / GD&T conformance
Objective | Confirm as-built enclosure, seal groove/gland, probe tip, HMI window, and connector apertures conform to the released critical-dimension & tolerance schedule.
Method | Inspection
I/O Trace | MER-mechanical (seal/gland); DO-MECH-002/011/018; DO-VER-002
Risk-Control Trace | RCR-011 (IP/containment dims), RCR-015 (seal)
Acceptance Criteria | All critical dimensions within tolerance per DO-VER-002; seal-groove and probe-tip dimensions within limits; CMM/gauge measurement traceable.
Level | Subsystem
Sample/Config | DVT enclosures (n per plan)
Facility/Equipment/Tools | CMM, calibrated gauges, optical comparator
Independence | Quality/Metrology (independent of design)
Standard/Compliance Reference | ASME Y14.5; ISO/IEC 17025
Evidence Required | Dimensional report vs schedule
Owner | Quality/Mechanical (EM-ENG-002)
Status | Planned
Notes | Feeds IP test (PV-ENV) and production dimensional inspection.
PV-INSP-002 — PCB fabrication conformance (bare board)
Objective | Confirm bare-board fabrication (stackup, impedance, plating, clearances, IPC class) conforms.
Method | Inspection
I/O Trace | HWR (PCB); DO-PCB-001…011
Risk-Control Trace | RCR-013 (SELV clearances), RCR-016 (latent-defect basis)
Acceptance Criteria | Board meets the specified IPC-A-600 class; controlled-impedance coupons within tolerance; safety-critical clearances/creepage meet IEC 62368-1.
Level | Subsystem
Sample/Config | Fab coupons + sample boards
Facility/Equipment/Tools | Cross-section microscopy, TDR, microscope
Independence | Quality/PCB (EM-ENG-003)
Standard/Compliance Reference | IPC-A-600; IEC 62368-1 spacings
Evidence Required | Fab conformance + coupon report
Owner | PCB/Electronics (EM-ENG-003)
Status | Planned
Notes | —
PV-INSP-003 — PCBA assembly workmanship & safety-net conformance
Objective | Confirm assembled PCBA workmanship (solder, orientation, correct parts on safety nets) conforms to IPC-A-610 target class.
Method | Inspection
I/O Trace | HWR; DO-PCB; DO-EE-005/010 (protection parts)
Risk-Control Trace | RCR-054 (correct-rated/oriented protection parts), RCR-002
Acceptance Criteria | IPC-A-610 class met; protection/fuse/BMS parts correct value and orientation; 0 defects on safety-critical nets.
Level | Subsystem/Production
Sample/Config | DVT/PVT PCBAs
Facility/Equipment/Tools | AOI, X-ray (BGA), microscope
Independence | Quality/Manufacturing
Standard/Compliance Reference | IPC-A-610
Evidence Required | AOI/X-ray + workmanship report
Owner | Manufacturing + PCB (EM-ENG-003)
Status | Planned
Notes | AOI coverage on safety nets ties PV-MFG-008 / RCR-016.
PV-INSP-004 — Firmware / device-software code review vs coding standard
Objective | Confirm firmware and device software conform to the adopted coding standard, with safety- and security-critical modules given focused review.
Method | Inspection
I/O Trace | FWR; SRS-DEV (200); DO-FW-; IEC 62304
Risk-Control Trace | RCR-019/021/023/024 (detection/verdict/safe-state/storage code)
Acceptance Criteria | Coding-standard compliance recorded; all safety/security modules peer-reviewed; review findings dispositioned; 0 open critical review findings on safety code.
Level | Unit
Sample/Config | Frozen source baseline
Facility/Equipment/Tools | Code-review tooling; static analyzer (PV-SW)
Independence | Reviewer ≠ author
Standard/Compliance Reference | IEC 62304; MISRA-style guideline (to confirm)
Evidence Required | Code-review records
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | Complements static analysis PV-SW-014.
PV-INSP-005 — Cloud / mobile code and configuration review
Objective | Confirm cloud and mobile source and infrastructure config are reviewed against secure-coding and platform standards.
Method | Inspection
I/O Trace | SRS-CLOUD (400), SRS-MOB (500); DO-CLD-, DO-APP-
Risk-Control Trace | RCR-032/039 (cloud integrity/DR), RCR-042 (export)
Acceptance Criteria | Secure-coding review complete; IaC/config reviewed; safety-relevant cloud services (ingest, notify, reconciliation) peer-reviewed; findings dispositioned.
Level | Unit
Sample/Config | Cloud/mobile build baseline
Facility/Equipment/Tools | Code-review + SAST tooling
Independence | Reviewer ≠ author; security consulted
Standard/Compliance Reference | OWASP ASVS concepts; IEC 62443 concepts
Evidence Required | Cloud/mobile review records
Owner | Cloud/Mobile (EM-ENG-009) + Security (EM-ENG-008)
Status | Planned
Notes | —
PV-INSP-006 — Label, marking & warning inspection (device + battery)
Objective | Confirm device labels, safety/warning markings, regulatory marks, and battery/transport markings are present, correct, legible, and durable.
Method | Inspection
I/O Trace | DO-LBL; PRD-REG; PRD-MECH-010; SYS-085 (marking durability)
Risk-Control Trace | RCR-010 (small-parts warning), RCR-052 (EOL/disposal), RCR-007 (UN 38.3 marking)
Acceptance Criteria | All required marks present and correct; content matches the regulatory dossier; markings survive the cleaning/UV durability checks (with PV-ENV); 100 % inspection of safety/transport marks.
Level | Production
Sample/Config | Production-representative labels
Facility/Equipment/Tools | Reference artwork; abrasion/solvent rub test
Independence | Quality/Regulatory (EM-ENG-006)
Standard/Compliance Reference | IEC 62368-1; UN 38.3 / IATA DGR; regional marks
Evidence Required | Label conformance report
Owner | Regulatory/Quality (EM-ENG-006)
Status | Evidence-Pending
Notes | Exact regional mark set follows the regulatory stage.
PV-INSP-007 — Packaging and transport-configuration inspection
Objective | Confirm shipping packaging and Li-battery transport configuration (state-of-charge, packing, documentation) conform to IATA DGR / UN 38.3 packing requirements.
Method | Inspection
I/O Trace | PRD-REG-006; DO-EE-015; PRD-ELEC-015
Risk-Control Trace | RCR-007 (transport qualification)
Acceptance Criteria | Packaging, marking, and DGR documentation conform; ship SoC within limit; no nonconformance on transport-safety items.
Level | Production
Sample/Config | Production packaging
Facility/Equipment/Tools | DGR checklist; SoC meter
Independence | Quality/Regulatory
Standard/Compliance Reference | IATA DGR; UN 38.3
Evidence Required | Packaging/DGR conformance record
Owner | Regulatory/Quality (EM-ENG-006)
Status | Planned
Notes | Complements the UN 38.3 test (PV-COMP / PV-SAFE).
PV-INSP-008 — Supplier documentation & materials-declaration review
Objective | Confirm incoming supplier documentation (cell datasheet, protection-IC, RoHS/REACH declarations, CoC) is complete and matches the approved BOM.
Method | Inspection
I/O Trace | HWR (BOM); PRD-REG (RoHS/REACH); DO-CFG-004
Risk-Control Trace | RCR-054 (correct chemistry/parts), RCR-055 (identity)
Acceptance Criteria | Every safety-critical component has a current datasheet/declaration on file; cell chemistry & protection ratings match the approved design; RoHS/REACH declarations complete.
Level | Production
Sample/Config | Incoming lots
Facility/Equipment/Tools | QMS supplier records
Independence | Quality/Supply-Chain
Standard/Compliance Reference | RoHS/REACH; ISO 13485 §7.4
Evidence Required | Supplier documentation review record
Owner | Quality + Supply-Chain
Status | Evidence-Pending
Notes | Final cell datasheet Evidence-Pending (concept §7) — gates RCR-001/054 acceptance values.
PV-INSP-009 — Design-output completeness & change-class inspection
Objective | Confirm the design-output package (DOL) is complete, released, and change-controlled before verification closes.
Method | Inspection
I/O Trace | Design-Output-Register; DOL-100 config baseline
Risk-Control Trace | RCR-058 (release readiness)
Acceptance Criteria | All Critical design outputs Released (or deviation-tracked); revision control intact; the verification-enabling outputs (DO-VER-001/003/006/010) are Released.
Level | System
Sample/Config | DOL register
Facility/Equipment/Tools | PLM
Independence | Quality
Standard/Compliance Reference | ISO 13485 §7.3.6/§7.3.7
Evidence Required | Design-output completeness review
Owner | Quality + Chief Systems (EM-ENG-001)
Status | Planned
Notes | —
PV-INSP-010 — Tamper-evidence & serviceable-access feature inspection
Objective | Confirm tamper-evident features and locked debug/service access are physically present and correct on production-representative units.
Method | Inspection
I/O Trace | DO-SEC-006; DO-EE-012 (debug port); DO-MECH (tamper)
Risk-Control Trace | RCR-038 (tamper-evident/debug lock), RCR-051 (mode lock)
Acceptance Criteria | Tamper evidence present; production debug/service ports physically locked or disabled; no accessible unlocked programming interface. (Functional lock verified by PV-CYBER-006.)
Level | Production
Sample/Config | Production-representative units
Facility/Equipment/Tools | Visual + probe inspection
Independence | Quality/Security
Standard/Compliance Reference | IEC 62443 concepts
Evidence Required | Tamper/access inspection record
Owner | Product-Security (EM-ENG-008)
Status | Planned
Notes | Physical half of RCR-038; functional half is PV-CYBER-006 (IV).
7. Traceability Notes
Inspection activities verify conformance of design outputs to drawings/standards/documents; they feed production inspection (PVL-1300) and the safety/compliance modules. Label/transport inspections tie the compliance module.
8. Coverage/Completion Notes
Complete when all drawing/PCB/code/label/packaging/supplier inspections pass on production-representative configuration. All Planned/Evidence-Pending.
PVL-1300 (PV-MFG) — Manufacturing Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Manufacturing/Test with Quality (EM-ENG-006), Metrology, Supply-Chain, Product- Security. Verifies the production processes and end-of-line (EOL) test coverage that keep the fielded fleet safe and accurate: EOL calibration + birth record, incoming/AOI screening, poka-yoke, unique identity, regional/firmware config, and latent-defect screening. Hosts the PV-MFG-### activities referenced by Stage-12 (RCR-016/053/054/055/056). PFMEA (Stage 15) refines the process controls. All entries Planned/Evidence-Pending.
1. Purpose
Verify manufacturing/EOL processes prevent the fleet-scale common-cause failures (miscalibration, wrong chemistry/part, duplicate identity, wrong regional config, latent defect) that individual controls depend on.
2. Scope
Assembly, EOL test/calibration, incoming inspection, AOI/ICT, serialization/identity, config provisioning, and production software at Production level. Excludes design verification of the parts (PVL-1200) and the transfer/ process-validation activities of Stage 15.
3. Applicable Standards & References
ISO 13485 §7.5 (production/process control); IPC-A-610 (assembly); ISO/IEC 17025 (calibration traceability); AIAG/VDA PFMEA & MSA; poka-yoke/error-proofing practice. Designations; revisions to confirm.
4. Governance & Method
Method is Test (EOL functional/calibration), Inspection (AOI/incoming), and Analysis (coverage). Safety-critical EOL screens are 100 %. Fleet common-cause controls (calibration, chemistry, identity, config) are verified for coverage and escape rate.
5. Verification Categories
EOL calibration & birth record · incoming inspection & AOI · poka-yoke error-proofing · unique identity/serial · regional/firmware config · AOI/ICT latent-defect coverage · production-software/programming.
6. Verification Register (this module)
PV-MFG-001 — EOL calibration coverage/accuracy + birth record
Objective | Confirm EOL calibration meets the defined coverage and accuracy, records a birth record, and that a miscalibrated unit cannot pass EOL.
Method | Test
I/O Trace | SYS-019/020; PRD-SAFE-001; PRD-MFG; DO-TST, DO-MFG
Risk-Control Trace | RCR-053, RCR-026, RCR-049; HA-MFG-001, FTA TOP-A (CCF)
Acceptance Criteria | EOL calibration coverage/accuracy meets the target; every unit has a birth record; a seeded miscalibrated unit fails EOL. Coverage/accuracy target Evidence-Pending → anchored to WHO PQS accuracy target (O-7).
Level | Production
Sample/Config | EOL line + reference standards
Facility/Equipment/Tools | Calibration bath/reference (ISO/IEC 17025); EOL tester
Independence | Metrology + Quality
Standard/Compliance Reference | ISO/IEC 17025; WHO PQS E006 (to confirm)
Evidence Required | EOL calibration coverage audit + birth-record sample
Owner | Metrology + Manufacturing
Status | Evidence-Pending
Notes | RCR-053 verification reference; bounds the fleet miscalibration CCF.
PV-MFG-002 — Incoming inspection: correct cell chemistry & protection parts
Objective | Confirm incoming inspection prevents wrong-chemistry cells and wrong-rated protection parts.
Method | Inspection, Test
I/O Trace | PRD-SAFE-008; PRD-MFG; DO-CFG-004
Risk-Control Trace | RCR-054; HA-MFG-002/003, FTA TOP-C
Acceptance Criteria | No wrong-chemistry cell or wrong-rated protection part passes incoming; supplier CoC/datasheet verified (PV-INSP-008); counterfeit screening applied.
Level | Production
Sample/Config | Incoming lots
Facility/Equipment/Tools | Incoming test/inspection; supplier records
Independence | Quality + Supply-Chain
Standard/Compliance Reference | ISO 13485 §7.4
Evidence Required | Incoming-inspection records
Owner | Quality + Supply-Chain
Status | Planned
Notes | RCR-054 verification reference; underpins battery safety.
PV-MFG-003 — Poka-yoke + AOI for orientation/value of safety parts
Objective | Confirm poka-yoke error-proofing and AOI prevent mis-orientation or wrong value of protection/ safety components.
Method | Inspection
I/O Trace | PRD-MFG; DO-MFG, DO-TST
Risk-Control Trace | RCR-054, RCR-010 (retention assembly)
Acceptance Criteria | Poka-yoke prevents mis-orientation; AOI catches wrong/reversed protection parts; challenge units are caught 100 %.
Level | Production
Sample/Config | Line + challenge samples
Facility/Equipment/Tools | AOI; poka-yoke fixtures
Independence | Manufacturing + Quality
Standard/Compliance Reference | IPC-A-610; error-proofing practice
Evidence Required | Poka-yoke/AOI verification
Owner | Manufacturing (with Supply-Chain)
Status | Planned
Notes | RCR-054 verification reference.
PV-MFG-005 — Unique cryptographic identity & serial provisioning
Objective | Confirm each device is provisioned with a unique cryptographic identity and serial, and duplicates are prevented/detected.
Method | Test
I/O Trace | SYS-016; PRD-MFG; PRD-SEC → PRD-SAFE-001; DO-MFG (provisioning), DO-SEC
Risk-Control Trace | RCR-055, RCR-033; HA-MFG-005
Acceptance Criteria | No duplicate identity/serial provisioned; each device cryptographically unique; duplicate attempts detected/rejected.
Level | Production
Sample/Config | Provisioning line
Facility/Equipment/Tools | Provisioning station; key-management
Independence | Manufacturing + Product-Security
Standard/Compliance Reference | IEC 62443 concepts
Evidence Required | Identity-uniqueness verification
Owner | Manufacturing + Product-Security
Status | Planned
Notes | RCR-055 verification reference; anchors device↔trip binding.
PV-MFG-006 — Regional / firmware configuration provisioned & verified at EOL
Objective | Confirm each unit ships with the correct regional/firmware config (band plan, market profile) and a wrong-region variant is caught at EOL.
Method | Test
I/O Trace | SYS-027/035/115/116; PRD-MFG/COMM; PRD-SAFE-010; DO-MFG, DO-FW
Risk-Control Trace | RCR-056; HA-MFG-006, HA-COMM-006, HA-NEW-009
Acceptance Criteria | Every unit ships with the correct regional/firmware config, verified at EOL; a wrong-region variant is caught. (Functional roaming = PV-COMM-006.)
Level | Production
Sample/Config | EOL line (regional variants)
Facility/Equipment/Tools | EOL config verifier
Independence | Manufacturing + Connectivity
Standard/Compliance Reference | regional band plans
Evidence Required | Config-verification record
Owner | Manufacturing + Connectivity
Status | Planned
Notes | RCR-056 verification reference.
PV-MFG-007 — Production programming & software-load verification
Objective | Confirm production programming loads the correct signed firmware/config and verifies the load.
Method | Test
I/O Trace | PRD-MFG; SRS-OTA/DEV; DO-MFG, DO-FW
Risk-Control Trace | RCR-037 (signed image), RCR-050 (restore)
Acceptance Criteria | Only the correct signed firmware/config is loaded; load verified (checksum/ signature) before the unit passes; wrong/unsigned load rejected.
Level | Production
Sample/Config | Programming station
Facility/Equipment/Tools | Programmer; signature verifier
Independence | Manufacturing + Security
Standard/Compliance Reference | IEC 62304 (config)
Evidence Required | Programming verification record
Owner | Manufacturing + Firmware
Status | Planned
Notes | —
PV-MFG-008 — AOI / ICT coverage & latent-defect screening (safety nets)
Objective | Confirm AOI + ICT/EOL coverage on safety-critical nets is sufficient that no latent single- point failure escapes, and HASS is applied.
Method | Analysis, Inspection
I/O Trace | PRD-SAFE-010; PRD-REL; DfT plan; DO-TST (AOI/ICT), DO-PCB
Risk-Control Trace | RCR-016; HA-ELEC-008, HA-MFG-008
Acceptance Criteria | AOI/ICT coverage on safety-critical nets ≥ the defined coverage target; HASS applied; coverage analysis shows no latent SPF escape. Coverage target Evidence-Pending → DfT plan.
Level | Production
Sample/Config | Line coverage data
Facility/Equipment/Tools | AOI/ICT; HASS chamber
Independence | Manufacturing/Test + Quality
Standard/Compliance Reference | IPC; DfT practice
Evidence Required | Coverage analysis + HASS record
Owner | Manufacturing/Test
Status | Evidence-Pending
Notes | RCR-016 verification reference; PFMEA (Stage 15) refines.
PV-MFG-009 — First-article & process-capability (Cpk) on critical characteristics
Objective | Confirm first-article inspection and process capability on critical-to-quality/function characteristics before production ramp.
Method | Inspection, Analysis
I/O Trace | DO-VER-001/002 (CTQ/CTF, dimensions); PRD-QUAL
Risk-Control Trace | RCR-011/015 (seal/coat quality), RCR-053
Acceptance Criteria | First article conforms to all CTQ/CTF; process capability meets the Cpk target on critical characteristics.
Level | Production
Sample/Config | First-article + ramp samples
Facility/Equipment/Tools | CMM/gauges; SPC
Independence | Quality
Standard/Compliance Reference | AIAG/VDA; ISO 13485
Evidence Required | FAI + capability study
Owner | Quality (EM-ENG-006)
Status | Planned
Notes | Feeds Stage-15 process validation.
7. Traceability Notes
Manufacturing activities verify the process/EOL controls that bound the fleet common-cause failures (RCR-016/053/054/055/056) — the PV-MFG-### verification references. They feed Stage-15 (transfer/PFMEA) and the supply-chain inspections (PVL-300).
8. Coverage/Completion Notes
Complete when EOL calibration/coverage, incoming/AOI, poka-yoke, identity, config, and latent-defect screening are verified on the production line. Coverage/accuracy targets Evidence-Pending. All Planned/Evidence-Pending.
PVL-600 (PV-PERF) — Performance Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with Sensing/Metrology, Firmware, RF, Power. Verifies the measurable performance targets that make the SCM's verdict trustworthy and its record complete: measurement accuracy, detection/alert latency, responsiveness, boot, storage capacity, throughput, and charge time. The 90-day cold endurance (the D3 headline) is a reliability test (PV-REL-001 / PRD-VER-004); this module carries the performance-timing and capacity set. All entries Planned/Evidence-Pending.
1. Purpose
Verify accuracy (±0.5 °C / 0.1 °C res / ±5 %RH), detection latency, connected/offline alert latency, HMI/verdict responsiveness (3 s incl. cold/low-battery), boot/wake time, storage capacity (90-day + 50 % spare), write throughput, data-delivery success, and charge time.
2. Scope
Performance/timing at Subsystem/System level across S3–S8. Excludes long-duration endurance (PVL-1400) and qualification exposures (PVL-700).
3. Applicable Standards & References
WHO PQS E006 (accuracy class, to confirm); ISO/IEC 17025 (reference traceability); IEC 60068 (test exposure during performance measurement); AIAG/VDA MSA. Designations; revisions to confirm.
4. Governance & Method
Method is Test (measure against a traceable reference). Accuracy tests run across the full −30…+60 °C range, not extrapolated. Evidence-Pending numeric targets (carried Judgment from PRD) are confirmed here.
5. Verification Categories
Measurement accuracy/resolution · detection latency · alert latency (connected/offline) · responsiveness · boot/wake · storage capacity · write throughput · delivery success · charge time.
6. Verification Register (this module)
PV-PERF-001 — Temperature-measurement accuracy across −30…+60 °C
Objective | Confirm each temperature channel measures within ±0.5 °C (or the confirmed WHO PQS class) across the full operating range against a traceable reference.
Method | Test
I/O Trace | SYS-033/071; PRD-PERF-001; PRD-ENV-001; DO-VER-003
Risk-Control Trace | RCR-053 (accuracy basis), RCR-019 (trustworthy detection)
Acceptance Criteria | Measured error ≤ ±0.5 °C (target; confirm vs WHO PQS E006) at multiple setpoints across −30…+60 °C for all channels; MSA (PV-ANAL-010) supports the measurement.
Level | Subsystem
Sample/Config | DVT units (n per plan) in chamber
Facility/Equipment/Tools | Climatic chamber; ISO/IEC 17025 temperature reference
Independence | Metrology/Quality
Standard/Compliance Reference | WHO PQS E006 (to confirm); ISO/IEC 17025
Evidence Required | Accuracy test report across range
Owner | Sensing/Metrology
Status | Evidence-Pending
Notes | Accuracy class Evidence-Pending (SYS-033); anchors detection targets (O-7).
PV-PERF-002 — Temperature resolution & humidity accuracy
Objective | Confirm 0.1 °C stored resolution and ±5 %RH humidity accuracy across 10–95 %RH non-condensing.
Method | Test
I/O Trace | SYS-035/036/072; PRD-PERF-003/008; PRD-ENV-004
Risk-Control Trace | RCR-019 (detection resolution basis)
Acceptance Criteria | Stored resolution ≤ 0.1 °C; RH error ≤ ±5 %RH across 10–95 %RH non-condensing at reference conditions.
Level | Subsystem
Sample/Config | DVT units in humidity chamber
Facility/Equipment/Tools | Humidity chamber; RH reference
Independence | Metrology/Quality
Standard/Compliance Reference | ISO/IEC 17025
Evidence Required | Resolution/RH accuracy report
Owner | Sensing/Metrology
Status | Planned
Notes | —
PV-PERF-003 — On-device excursion-detection latency
Objective | Confirm the device registers a detected excursion within one logging interval of the qualifying sample, independent of connectivity.
Method | Test
I/O Trace | SYS-038; PRD-PERF-004; DO-FW-006
Risk-Control Trace | RCR-019 (prompt detection)
Acceptance Criteria | Detection registered within one logging interval for all seeded excursions; measured at −30 °C and low battery worst case.
Level | Subsystem
Sample/Config | DVT units + profile injector
Facility/Equipment/Tools | Chamber; timing harness
Independence | Embedded SW; Quality
Standard/Compliance Reference | —
Evidence Required | Detection-latency report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | —
PV-PERF-004 — Connected & offline alert latency; verdict/HMI responsiveness
Objective | Confirm connected alert delivery ≤ 5 min of detection, offline alert-first delivery within the bounded time of reconnection, and verdict/HMI response within 3 s incl. −30 °C and low battery.
Method | Test
I/O Trace | SYS-039/040/041/042; PRD-PERF-004/007; DO-FW-008, DO-HF-002
Risk-Control Trace | RCR-029 (alert-ahead), RCR-043 (responsive verdict)
Acceptance Criteria | Connected detection→cloud alert ≤ 5 min (target; confirm); offline alert delivered ahead of bulk within the bound on reconnection; verdict/status renders ≤ 3 s at −30 °C/low battery.
Level | Integration/System
Sample/Config | Device + network emulator + cloud; chamber
Facility/Equipment/Tools | Network emulator; chamber; timing harness
Independence | Connectivity + Firmware; Quality
Standard/Compliance Reference | D2
Evidence Required | Latency/responsiveness report
Owner | Connectivity + Firmware
Status | Evidence-Pending
Notes | 5-min target Evidence-Pending (SYS-039); in-container latency bounded by PRD-VER-005.
PV-PERF-005 — Boot / wake / cold-start time
Objective | Confirm boot and low-power wake times meet the timing budget so scheduled samples are not delayed, including at −30 °C.
Method | Test
I/O Trace | SYS-037; PRD-PERF; DO-FW-003 (power-state governor)
Risk-Control Trace | RCR-006 (energy governor timing)
Acceptance Criteria | Wake-to-sample and cold-start times within the budget with margin at −30 °C; no sample delayed into the next interval.
Level | Subsystem
Sample/Config | DVT units in chamber
Facility/Equipment/Tools | Chamber; timing harness
Independence | Firmware
Standard/Compliance Reference | —
Evidence Required | Boot/wake timing report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | —
PV-PERF-006 — Full-record verdict computation time (worst case)
Objective | Confirm verdict computation over a full 90-day record completes within the responsiveness budget on the embedded MCU at cold/low-battery worst case.
Method | Test
I/O Trace | SYS-042; PRD-PERF-007; DO-FW-008
Risk-Control Trace | RCR-021 (verdict availability), RCR-043
Acceptance Criteria | Verdict computes and renders within the responsiveness budget for a full-length record at −30 °C/low battery.
Level | Subsystem
Sample/Config | DVT units + full-length record
Facility/Equipment/Tools | Chamber; timing harness
Independence | Embedded SW
Standard/Compliance Reference | —
Evidence Required | Verdict-timing report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | —
PV-PERF-007 — Log-storage capacity for a full offline 90-day trip (+50 % spare)
Objective | Confirm NVM holds a full 90-day multi-channel trip at the minimum interval with ≥50 % spare, without overwriting unsent data.
Method | Test
I/O Trace | SYS-043/127; PRD-PERF-006; DO-FW-010
Risk-Control Trace | RCR-024 (no overwrite), RCR-030 (buffer depth)
Acceptance Criteria | A full-length trip is stored with ≥50 % spare capacity remaining; no unsent data overwritten; predicted by PV-ANAL-007.
Level | Subsystem
Sample/Config | DVT units + long-run/simulated fill
Facility/Equipment/Tools | Long-run harness
Independence | Firmware
Standard/Compliance Reference | —
Evidence Required | Storage-capacity report
Owner | Firmware (EM-ENG-007)
Status | Evidence-Pending
Notes | Storage sizing Evidence-Pending (SYS-127).
PV-PERF-008 — Sustained logging write throughput (base + event burst)
Objective | Confirm the write path commits base logging plus concurrent shock-event capture at peak rate without dropping/corrupting records.
Method | Test
I/O Trace | SYS-044; PRD-PERF-006; DO-FW-010
Risk-Control Trace | RCR-024 (no lost record)
Acceptance Criteria | At the peak combined rate over sustained operation, 0 dropped/overwritten/corrupted records.
Level | Subsystem
Sample/Config | DVT units + burst injector
Facility/Equipment/Tools | Shock-event stimulator; throughput harness
Independence | Firmware
Standard/Compliance Reference | —
Evidence Required | Throughput report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | —
PV-PERF-009 — Data-delivery success rate over a representative trip
Objective | Confirm the device delivers the target fraction of buffered data over a representative trip whenever connectivity occurs, in time order without loss.
Method | Test
I/O Trace | SYS-049; PRD-PERF-011; PRD-CONN-007
Risk-Control Trace | RCR-030 (degrade without loss), RCR-032
Acceptance Criteria | Delivered fraction meets the target over a representative connectivity profile with intact time order and 0 data loss. Fraction Evidence-Pending; bounded by RF-in-container PRD-VER-005.
Level | System
Sample/Config | Device + representative network profile
Facility/Equipment/Tools | Network emulator; field lanes
Independence | Connectivity + Cloud
Standard/Compliance Reference | —
Evidence Required | Delivery-success report
Owner | Connectivity + Cloud
Status | Evidence-Pending
Notes | Success fraction Evidence-Pending (SYS-049).
PV-PERF-010 — GNSS location-fix cadence within energy budget
Objective | Confirm location fixes at the configured cadence within the location energy allocation without breaching the 90-day endurance.
Method | Test
I/O Trace | SYS-048; PRD-PERF-010; PRD-CONN-005
Risk-Control Trace | RCR-006 (energy budget)
Acceptance Criteria | Achieved fix cadence meets the target within the location energy allocation; energy draw within budget. GNSS-in-container cadence Evidence-Pending (SYS-048).
Level | Subsystem
Sample/Config | DVT units (in/out of shipper)
Facility/Equipment/Tools | GNSS test; power analyzer
Independence | Connectivity + Power
Standard/Compliance Reference | —
Evidence Required | GNSS cadence/energy report
Owner | Connectivity + Power
Status | Evidence-Pending
Notes | —
PV-PERF-011 — Charge time to trip-readiness (permitted charge temperature)
Objective | Confirm recharge from empty to trip-readiness meets the charge-time target at a permitted charge temperature (charging inhibited below 0 °C).
Method | Test
I/O Trace | SYS-050; PRD-PERF-012; PRD-ELEC-004; DO-EE-007
Risk-Control Trace | RCR-001 (charge inhibit boundary)
Acceptance Criteria | Charge to trip-readiness within the target time at the permitted temperature; charge correctly inhibited below 0 °C. Charge-time target Evidence-Pending (SYS-050).
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Charger; chamber; power analyzer
Independence | Power
Standard/Compliance Reference | —
Evidence Required | Charge-time report
Owner | Battery/Power
Status | Evidence-Pending
Notes | Charge-safety boundary verified in PV-ENRG-001.
PV-PERF-012 — Cloud/API throughput & scalability (fleet)
Objective | Confirm the cloud sustains fleet-scale ingest/query/notify throughput within response-time targets.
Method | Test
I/O Trace | SRS-CLOUD/PERF (400/1000); PRD-CLOUD
Risk-Control Trace | RCR-032 (no dropped event under load)
Acceptance Criteria | At the projected fleet load, ingest/query/notify meet response-time SLOs with 0 dropped safety events.
Level | System
Sample/Config | Cloud load-test tenant
Facility/Equipment/Tools | Load-test harness
Independence | Cloud; Quality
Standard/Compliance Reference | —
Evidence Required | Load/scalability report
Owner | Cloud (EM-ENG-009)
Status | Planned
Notes | —
7. Traceability Notes
Performance activities close the Evidence-Pending numeric targets carried Judgment from the PRD/SYS (±0.5 °C, 5-min, 3-s, 50 % spare, delivery fraction). The 90-day endurance is verified in PV-REL-001. Feeds Validation (trust in the verdict) and the release summary.
8. Coverage/Completion Notes
Complete when all performance targets are measured within limits (or the Evidence-Pending targets are set and met by test). All Planned/Evidence-Pending.
PVL-100 (PV-PLAN) — Verification Planning — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with EM-ENG-001. Strategy, scope, matrix, schedule, resources, facilities, dependencies, milestones, and deliverables for SCM verification. The full narrative strategy and phase sequencing live in the companion Verification-Plan.md; this module registers the plan as verifiable governance items. All entries Planned/Evidence-Pending.
1. Purpose
Register the planning artefacts that scope and sequence SCM verification: the strategy, the requirement→PV coverage matrix, the DVT/PVT phase plan, the resource/facility/equipment plan, and the dependency/risk plan.
2. Scope
Program-level planning for all subsystems and tiers; the two program-top qualification tests (D3 cold endurance PRD-VER-004; RF-in-container PRD-VER-005); certification-house dependencies (Evidence-Pending).
3. Applicable Standards & References
ISO 13485 §7.3.6; ISO/IEC/IEEE 29148; IEC 60068 / 60529 / 62368-1 / 61000 / UN 38.3 (qualification test scheduling); ISO/IEC 17025 (facility/reference traceability); GAMP 5. Designations; revisions to confirm.
4. Governance & Method
Planning is verified by Inspection of the plan artefacts against the requirement/RCR baseline. The plan is risk-first: the D3 battery and RF-in-container Open items (Stage 12 §7) are scheduled earliest with the longest lead, because they gate architecture and release.
5. Verification Categories
Strategy · scope/objectives · coverage matrix · schedule/phasing · resources/personnel · facilities/ equipment/tools · dependencies/risks · milestones/deliverables.
6. Verification Register (this module)
PV-PLAN-001 — Verification strategy approved
Objective | Confirm an approved verification strategy exists defining levels, methods, phases, and the severity-first ordering of the SCM verification program.
Method | Inspection
I/O Trace | PRD-VER-001..008; SYS/PRD/SRS baselines; DO-VER-001/003/010
Risk-Control Trace | RCR-058; all RCR PV- targets
Acceptance Criteria | Strategy covers Unit→Production levels, all four methods, DVT/PVT phases, and names the D3/RF-in-container Open items as gating; approved by EM-ENG-006 + EM-ENG-001.
Level | System
Sample/Config | Verification-Plan.md
Facility/Equipment/Tools | PLM/QMS
Independence | Quality authors; Chief Systems approves
Standard/Compliance Reference | ISO 13485 §7.3.6
Evidence Required | Signed verification plan
Owner | EM-ENG-006
Status | Planned
Notes | See Verification-Plan.md.
PV-PLAN-002 — Requirement-to-PV coverage matrix complete
Objective | Confirm every requirement (PRD/SYS/SRS/HWR/MER/FWR/IFR) and every RCR maps to ≥1 PV activity.
Method | Inspection
I/O Trace | whole requirement baseline; RCR-1000 register
Risk-Control Trace | all 58 RCRs
Acceptance Criteria | 0 unverified applicable requirements; 58/58 RCRs have a matching PV activity; every FI/IV flag preserved; documented in Verification-Traceability-Coverage-Matrix.md.
Level | System
Sample/Config | Coverage matrix
Facility/Equipment/Tools | Trace tool
Independence | Quality; Chief Systems reconciles
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Coverage matrix with gap analysis
Owner | EM-ENG-006 + EM-ENG-001
Status | Planned
Notes | The matrix is the demand set for scheduling.
PV-PLAN-003 — DVT / PVT phase plan and sequencing
Objective | Confirm a phased test plan (Design Verification Test → Process/Production Verification Test) with sequencing that respects sample maturity and destructive-test ordering.
Method | Inspection
I/O Trace | PRD-VER-006; DO-VER-009; IEC 60068 / UN 38.3
Risk-Control Trace | RCR-007 (UN 38.3), RCR-001/002/008 (battery safety), RCR-017 (env)
Acceptance Criteria | Phase plan sequences non-destructive before destructive tests, schedules long-lead qualifications (90-day endurance, UN 38.3, EMC/radio) with margin, and defines sample allocation per phase.
Level | System
Sample/Config | DVT/PVT schedule
Facility/Equipment/Tools | Program schedule; test-house calendar
Independence | Quality + Program
Standard/Compliance Reference | IEC 60068; UN 38.3; CISPR/IEC 61000
Evidence Required | Approved phase schedule
Owner | EM-ENG-006
Status | Planned
Notes | 90-day endurance is the schedule critical path.
PV-PLAN-004 — Facilities, equipment, and calibrated-reference plan
Objective | Confirm the facilities and traceable references needed (climatic chambers, IP/leak rigs, vibration/shock/drop, EMC chamber, RF test rig, temperature/RH references, load banks) are identified and calibration-traceable.
Method | Inspection
I/O Trace | DO-VER-008/009; PRD-VER-006; ISO/IEC 17025
Risk-Control Trace | enables RCR-017/031/001 verification
Acceptance Criteria | Each planned test names its facility/equipment with required accuracy and ISO/IEC 17025 calibration traceability; the −30…+60 °C chamber and representative in-container RF rig are specified (DO-VER-009).
Level | System
Sample/Config | Facility/equipment register
Facility/Equipment/Tools | Internal lab + accredited test houses
Independence | Metrology/Quality
Standard/Compliance Reference | ISO/IEC 17025; IEC 60068/60529
Evidence Required | Facility/equipment plan; calibration certificates (at execution)
Owner | Quality/Metrology
Status | Evidence-Pending
Notes | Test houses & representative shipper Evidence-Pending (concept §7, DO-VER-009).
PV-PLAN-005 — Resources, personnel, and responsibilities allocated
Objective | Confirm each PV activity has a named owner discipline and the personnel/skills (incl. independent pen-test) are resourced.
Method | Inspection
I/O Trace | PRD-VER-001; PV-GOV-002 (independence)
Risk-Control Trace | RCR-034…039 (IV resourcing)
Acceptance Criteria | Every PV activity has an owner; independent security testers and accredited test houses are contracted or planned; no activity is unassigned.
Level | System
Sample/Config | RACI / resource plan
Facility/Equipment/Tools | Program resource plan
Independence | Quality + Program
Standard/Compliance Reference | ISO 13485
Evidence Required | Resource/RACI plan
Owner | EM-ENG-006
Status | Planned
Notes | —
PV-PLAN-006 — Dependencies, verification risks, and milestones tracked
Objective | Confirm cross-stage dependencies (final cell datasheet, antenna decision, WHO PQS category, test-house lead) and verification risks are tracked with milestones and deliverables.
Method | Inspection
I/O Trace | concept §7; PRD-VER-004/005; SYS-046/080
Risk-Control Trace | Open items D3 (RCR-001/005/006/017) and RF-in-container (RCR-031)
Acceptance Criteria | Each Evidence-Pending input has a named closing test and a milestone; verification risks (long-lead qual, sample availability, cold-chamber time) have mitigations; deliverables (protocols, reports, summary) scheduled.
Level | System
Sample/Config | Dependency/risk/milestone register
Facility/Equipment/Tools | Program schedule
Independence | Quality + Program
Standard/Compliance Reference | ISO 13485
Evidence Required | Dependency & milestone register
Owner | EM-ENG-006 + EM-ENG-001
Status | Planned
Notes | D3 and RF-in-container are the top verification-program risks.
7. Traceability Notes
Planning items scope and schedule the demand set produced by PV-REQ/PV-TRACE and consumed by every technical module. The matrix (PV-PLAN-002) is authoritative for coverage.
8. Coverage/Completion Notes
Planning is complete when strategy, coverage matrix, phase plan, facility/resource plans, and dependency register are approved before DVT. All Planned/Evidence-Pending.
PVL-1400 (PV-REL) — Reliability Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Reliability with Battery/Power, RF, Firmware, Quality. Verifies the SCM's reliability and life: HALT/HASS, accelerated life, battery aging/cycle-life, flash/storage endurance, connector durability, and the MTBF demonstration. This module owns the two program-top validate-by-test items: the D3 −30 °C / 90-day cold-endurance test (PV-REL-001 = PRD-VER-004) and the RF-in-container qualification test (PV-REL-005 = PRD-VER-005). Both are carried Open (Stage-12 §7) and gate release. All entries Planned/ Evidence-Pending.
1. Purpose
Verify the SCM meets its reliability and life targets over multi-trip reuse, and close the D3 battery and RF-in-container Open items by test — not by analysis alone.
2. Scope
Reliability/life/aging/endurance at Subsystem/System level over the reuse life. Excludes single-exposure environmental qualification (PVL-700) and functional comms (PVL-500), which it draws on.
3. Applicable Standards & References
IEC 60068 (accelerated/aging exposures); UN 38.3 (battery); HALT/HASS practice; flash-endurance datasheet methods; reliability-demonstration (e.g., Telcordia/MIL-HDBK-217-class prediction as basis, method to confirm). Designations; revisions to confirm.
4. Governance & Method
Method is Test (long-duration/accelerated). The D3 endurance and RF-in-container tests use the DO-VER-009 cold-chamber and representative-shipper configuration. Reliability confirms the analyses (PV-ANAL-003/004/006/
008) by physical test — the safety-critical claims are not accepted on analysis alone.
5. Verification Categories
90-day cold endurance (D3) · RF-in-container qualification (O-1) · HALT · HASS · accelerated life/aging · battery cycle-life & fade · flash/storage endurance · connector/switch durability · MTBF demonstration.
6. Verification Register (this module)
PV-REL-001 — 90-day −30 °C cold-endurance test (D3 headline; = PRD-VER-004)
Objective | Prove by test ≥90 days of continuous armed operation from a full charge under the worst-case −30 °C cold-duty profile, retaining sensing/logging/verdict/alert with reserve preserved and honest low- battery indication.
Method | Test
I/O Trace | SYS-045/046/086/090; PRD-PERF-005; PRD-REL-001; PRD-VER-004; DO-SAFE-012, DO-FW-004, DO-VER-009; D3
Risk-Control Trace | RCR-006, RCR-017; HA-ENRG-006; O-2 (largest program assumption)
Acceptance Criteria | Over a 90-day −30 °C profile the device continues sampling+logging, commits the record, and raises a low-battery alert before exhaustion; the fuel gauge does not over-report SoC in cold beyond the bound; documented energy margin > 0. 90-day margin, cold SoC error bound, and −30 °C duty profile Evidence-Pending → this test (O-2).
Level | System
Sample/Config | DVT units, final chemistry, in cold chamber (n per plan)
Facility/Equipment/Tools | −30 °C climatic chamber (DO-VER-009); power logging; long-run harness
Independence | Reliability; Battery/Power + System-Safety review
Standard/Compliance Reference | D3; IEC 60068-2-1; cell datasheet
Evidence Required | 90-day cold-endurance test report
Owner | Reliability + Battery/Power
Status | Open / Evidence-Pending (D3, O-2)
Notes | PRD-VER-004; the schedule critical path and top program risk; confirms PV-ANAL-003/004.
PV-REL-002 — Cold-capacity & reserve verification at −30 °C
Objective | Prove usable capacity and reserve at −30 °C are sufficient to preserve record-commit + low- battery alert, confirming the reserve design.
Method | Test
I/O Trace | SYS-046; PRD-REL-001/002; DO-EE-003, DO-FW-004
Risk-Control Trace | RCR-006 (reserve)
Acceptance Criteria | Measured −30 °C usable capacity supports the endurance budget; reserve preserves record-commit + alert before exhaustion.
Level | Subsystem
Sample/Config | DVT packs in chamber
Facility/Equipment/Tools | Chamber; battery cycler
Independence | Battery/Power + Reliability
Standard/Compliance Reference | cell datasheet; D3
Evidence Required | Cold-capacity/reserve report
Owner | Battery/Power
Status | Evidence-Pending
Notes | Confirms PV-ANAL-004; supports PV-REL-001.
PV-REL-003 — Battery cycle-life & capacity-fade over reuse
Objective | Prove capacity fade over the reuse-cycle life keeps the endurance budget positive to end-of-life.
Method | Test
I/O Trace | PRD-REL-003; SYS-165-line; DO-EE-003; D3
Risk-Control Trace | RCR-006/017 (endurance to EOL)
Acceptance Criteria | After the qualified number of charge/discharge cycles, retained capacity keeps the 90-day cold-endurance budget positive; fade within the reuse-life budget. Cycle count Evidence-Pending.
Level | Subsystem
Sample/Config | DVT packs on cycler
Facility/Equipment/Tools | Battery cycler; chamber
Independence | Battery/Power + Reliability
Standard/Compliance Reference | cell datasheet
Evidence Required | Cycle-life/fade report
Owner | Battery/Power
Status | Evidence-Pending
Notes | Confirms PV-ANAL-004; part of the D3 close-out set.
PV-REL-004 — Charge-safety over aging / degraded-cell behaviour
Objective | Prove the charge-inhibit and protection controls remain effective as the cell ages/degrades over reuse.
Method | Test (with fault injection on aged cells)
I/O Trace | PRD-REL-003; PRD-SAFE-007/008; DO-EE-007, DO-SAFE-001/002
Risk-Control Trace | RCR-001/002 (charge/protection effectiveness over life)
Acceptance Criteria | Charge inhibit and protection hold on aged/degraded cells across the cycle life; no charge below limit and no runaway path emerges with aging.
Level | Subsystem
Sample/Config | Aged DVT packs
Facility/Equipment/Tools | Cycler; abuse rig; chamber
Independence | Battery-Safety + Reliability
Standard/Compliance Reference | IEC 62368-1; UN 38.3
Evidence Required | Aged-cell charge-safety report
Owner | Battery/Power + System-Safety
Status | Evidence-Pending
Notes | Life-companion to PV-ENRG-001/002.
PV-REL-005 — RF-in-container qualification test (O-1; = PRD-VER-005)
Objective | Prove by test the in-container link margin (LTE-M/Wi-Fi/BLE/GNSS) inside representative metallized/insulated shippers across target markets, and settle the integral-vs-external antenna decision.
Method | Test
I/O Trace | SYS-080/104/109/117; PRD-CONN-007; PRD-ENV-009; PRD-VER-005; DO-EE-017/018, DO-VER-009
Risk-Control Trace | RCR-031 (+ RCR-030 consequence-bound); HA-COMM-003; O-1
Acceptance Criteria | The integral antenna closes the in-container budget across the target markets/ containers, or the external-antenna fallback does; where neither closes, store-and-forward + on-receipt verdict bound the consequence and the real-time claim is withheld on those lanes. In-container attenuation, cold RF margin, and antenna decision Evidence-Pending → this test (O-1).
Level | System
Sample/Config | DVT devices inside representative shippers (multiple SKUs); cold + ambient
Facility/Equipment/Tools | Representative shippers; RF test rig; live/field networks; chamber
Independence | RF/Antenna + Connectivity + Regulatory
Standard/Compliance Reference | PRD-VER-005; FCC/CE band plans
Evidence Required | RF-in-container qualification report
Owner | RF/Antenna + Connectivity
Status | Open / Evidence-Pending (O-1)
Notes | PRD-VER-005; principal open architectural risk; confirms PV-ANAL-006; functional companion PV-COMM-003.
PV-REL-006 — HALT (highly accelerated life test) — margin discovery
Objective | Discover design margins and precipitate latent failure modes by stepping thermal/vibration stress beyond spec.
Method | Test
I/O Trace | PRD-REL; FAL-1900; PV-ANAL-008
Risk-Control Trace | RCR-016 (latent-defect discovery), RCR-023
Acceptance Criteria | Operating and destruct margins characterised; failure modes root-caused and dispositioned; margin adequate above the operating envelope.
Level | System
Sample/Config | DVT units
Facility/Equipment/Tools | HALT chamber (thermal + vibration)
Independence | Reliability
Standard/Compliance Reference | HALT practice
Evidence Required | HALT report
Owner | Reliability
Status | Planned
Notes | Sets HASS screen limits (PV-REL-007).
PV-REL-007 — HASS + flash/storage-endurance & retention
Objective | Establish the HASS production screen and prove flash/storage endurance and retention over the reuse life at temperature.
Method | Test
I/O Trace | PRD-REL-005; SYS-127; DO-FW-010; PV-REL-006
Risk-Control Trace | RCR-016 (screen), RCR-024 (record retention)
Acceptance Criteria | HASS screen precipitates infant-mortality without over-stress; flash endurance and retention meet the reuse-life record-retention requirement at temperature. Endurance/retention Evidence-Pending.
Level | Subsystem/Production
Sample/Config | DVT/PVT units
Facility/Equipment/Tools | HASS chamber; flash-endurance harness; chamber
Independence | Reliability + Manufacturing
Standard/Compliance Reference | HASS practice; flash datasheet
Evidence Required | HASS + flash-endurance report
Owner | Reliability + Firmware
Status | Evidence-Pending
Notes | Supports RCR-024 (PV-SW-013 tests the write-atomicity; this tests wear/retention).
PV-REL-008 — Connector / port / switch durability life
Objective | Prove USB-C port, probe connector, and any button/switch survive the rated actuation/mating cycles over reuse life.
Method | Test
I/O Trace | HWR (connectors/switch); PRD-REL; PRD-SVC
Risk-Control Trace | RCR-011 (port ingress retained), RCR-050
Acceptance Criteria | Connectors/switches meet electrical/mechanical spec after the rated cycles; port retains IP after cycling (with PV-ENV-005).
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Mating/actuation cycle rigs
Independence | Mechanical/Reliability
Standard/Compliance Reference | connector/switch spec
Evidence Required | Durability-life report
Owner | Mechanical + Reliability
Status | Planned
Notes | Reuse-life companion to PV-HW-006.
PV-REL-009 — Reliability / MTBF demonstration
Objective | Demonstrate the field reliability (MTBF/availability) target by an accelerated demonstration test, confirming the PV-ANAL-008 prediction.
Method | Test, Analysis
I/O Trace | PRD-REL; SYS-163…176; PV-ANAL-008
Risk-Control Trace | RCR-023/027 (availability)
Acceptance Criteria | Demonstrated MTBF/availability meets the target at the specified confidence; failure modes root-caused. Target/confidence Evidence-Pending.
Level | System
Sample/Config | DVT/PVT population
Facility/Equipment/Tools | Life-test racks; chambers
Independence | Reliability + Quality
Standard/Compliance Reference | reliability-demonstration method (to confirm)
Evidence Required | MTBF demonstration report
Owner | Reliability
Status | Evidence-Pending
Notes | Long-lead; feeds the release summary.
PV-REL-010 — Multi-trip reuse-life / accelerated-aging system test
Objective | Prove the whole device survives the multi-trip reuse life (thermal cycling + cleaning + handling + charge cycles) without accuracy, seal, or record-integrity degradation.
Method | Test
I/O Trace | PRD-REL; SYS-073/079; D1 (reusable)
Risk-Control Trace | RCR-011/015/018 (seal/coat/condensation over life), RCR-053 (accuracy)
Acceptance Criteria | After the accelerated reuse-life sequence, accuracy within ±0.5 °C, seal/IP retained, and record integrity intact.
Level | System
Sample/Config | DVT units through combined-stress sequence
Facility/Equipment/Tools | Chambers; cleaning/handling rigs; cycler
Independence | Reliability
Standard/Compliance Reference | IEC 60068; D1
Evidence Required | Reuse-life system report
Owner | Reliability
Status | Evidence-Pending
Notes | Integrates the reuse-life claims across modules.
7. Traceability Notes
Reliability owns the two carried-Open validate-by-test items (PV-REL-001 = PRD-VER-004; PV-REL-005 = PRD-VER-005) and confirms the D3/RF analyses (PV-ANAL-003/004/006). Flash endurance supports RCR-024; aging-charge-safety supports RCR-001/002 over life. Feeds the release summary (both Open items gate release).
8. Coverage/Completion Notes
Complete when HALT/HASS, life/aging, flash endurance, durability, and MTBF pass, and the two Open items (90-day cold endurance; RF-in-container) close by test. Both are release-gating. All Planned/Evidence-Pending.
PVL-200 (PV-REQ) — Requirements Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-001 (integrator) with EM-ENG-006. Confirms the SCM requirement set is itself verifiable — each requirement is atomic, testable, allocated, has an acceptance criterion, and is covered by ≥1 PV activity — before and during execution. This is the "verify the requirements are verifiable" module. All entries Planned.
1. Purpose
Verify requirement quality and coverage: every PRD/SYS/SRS/HWR/MER/FWR/IFR requirement is unambiguous, testable, allocated to a verification method and level, and traced to ≥1 PV activity; every RCR has a PV target; no orphan or conflict escapes.
2. Scope
The full Stage-05 (220 PRD), Stage-06 (190 SYS), Stage-07 (180 subsystem + 308 SRS) requirement baseline and the 58 RCRs (Stage 12). Excludes authoring the requirements (Stages 05–07) and the trace-matrix mechanics (PVL-1600).
3. Applicable Standards & References
ISO/IEC/IEEE 29148 (requirements engineering & verification); ISO 13485 §7.3; ISO 14971 (risk-control traceability). Designations; revisions to confirm.
4. Governance & Method
Method is Inspection (requirement review) and Analysis (coverage/gap computation). A requirement is "verifiable" only if it has a single testable "shall", an objective acceptance criterion, an allocated method, and a level.
5. Verification Categories
Ambiguity/testability review · completeness · conflict detection · allocation & method check · derived- requirement coverage · acceptance-criteria objectivity · RCR coverage · closure.
6. Verification Register (this module)
PV-REQ-001 — Requirement testability and ambiguity review
Objective | Confirm each requirement is a single, atomic, measurable "shall" with no ambiguous or unverifiable language.
Method | Inspection
I/O Trace | all PRD/SYS/SRS/HWR/MER/FWR/IFR; PRD-VER-001
Risk-Control Trace | RCR-058 (defensibility of the requirement set)
Acceptance Criteria | 100 % of requirements reviewed; every non-testable/ambiguous requirement is flagged and corrected or given an objective acceptance criterion; 0 requirements without a verification method.
Level | System
Sample/Config | Requirement baseline (frozen rev)
Facility/Equipment/Tools | Requirements/trace tool
Independence | Quality reviews; Chief Systems adjudicates
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Requirement-quality review record
Owner | EM-ENG-001 + EM-ENG-006
Status | Planned
Notes | Evidence-Pending numeric targets (±0.5 °C, 5-min, 90-day, IP67) are acceptable if the closing test is named.
PV-REQ-002 — Requirement completeness and conflict review
Objective | Confirm the requirement set is complete for the intended use and internally non-conflicting (especially the D3 endurance vs charge-safety and safety-vs-availability tensions).
Method | Inspection
I/O Trace | PRD library index; SYS decomposition-allocation-matrix; STN-conflicts-and-tradeoffs
Risk-Control Trace | RCR-006/017 (endurance), RCR-047 (alarm balance), HA-NEW-006 trade
Acceptance Criteria | No intended-use function lacks a requirement; identified conflicts (D3, alarm balance) are resolved or carried as explicit balanced requirements; 0 unresolved conflicts at baseline.
Level | System
Sample/Config | Requirement baseline
Facility/Equipment/Tools | Trace tool
Independence | Chief Systems + Quality
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Completeness/conflict review record
Owner | EM-ENG-001
Status | Planned
Notes | —
PV-REQ-003 — Allocation and verification-level assignment check
Objective | Confirm each requirement is allocated to a subsystem (S1–S9/SYS) and a verification level (Unit/Subsystem/Integration/System/Production).
Method | Inspection
I/O Trace | SYS decomposition-allocation-matrix; DO-VER-010
Risk-Control Trace | RCR allocation (Stage 12 §6 by subsystem)
Acceptance Criteria | Every requirement has an allocation target and a level; no requirement is verified only at a level too high to isolate its failure (e.g., safe-state verified at Unit + System).
Level | System
Sample/Config | Allocation matrix
Facility/Equipment/Tools | Trace tool
Independence | Quality
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Allocation/level review record
Owner | EM-ENG-001
Status | Planned
Notes | —
PV-REQ-004 — Acceptance-criteria objectivity review
Objective | Confirm each requirement's PV activity carries an objective, measurable pass/fail rule with a limit and tolerance (no subjective safety verdict).
Method | Inspection
I/O Trace | DO-VER-003 (limits), DO-VER-010 (pass/fail); PRD-VER-001
Risk-Control Trace | all safety RCRs (zero-ambiguity acceptance)
Acceptance Criteria | 100 % of safety/function-critical PV activities have a numeric or binary pass/fail; Evidence-Pending limits are flagged with the closing test; 0 subjective safety criteria.
Level | System
Sample/Config | PV register (this library)
Facility/Equipment/Tools | QMS
Independence | Quality
Standard/Compliance Reference | ISO 13485; AIAG/VDA MSA (measurement behind each limit)
Evidence Required | Acceptance-criteria audit
Owner | EM-ENG-006
Status | Planned
Notes | —
PV-REQ-005 — Derived-requirement coverage review
Objective | Confirm engineering-derived requirements (not explicit in a parent) are still traced and verified.
Method | Inspection
I/O Trace | SYS "Type: Derived" set; SRS derived items
Risk-Control Trace | RCR-023/027 (derived safe-state/self-check)
Acceptance Criteria | Every Derived requirement traces upward to a need/decision and forward to a PV activity; 0 derived requirements without verification.
Level | System
Sample/Config | Requirement baseline
Facility/Equipment/Tools | Trace tool
Independence | Chief Systems
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Derived-coverage review record
Owner | EM-ENG-001
Status | Planned
Notes | —
PV-REQ-006 — Risk-control (RCR) coverage and FI/IV-flag review
Objective | Confirm every RCR (Stage 12) has ≥1 matching PV activity here, with fault-injection preserved for safety-critical and independent adversarial verification for security controls.
Method | Inspection
I/O Trace | RCR-1000 register (all 58); HAL HA-### verification references
Risk-Control Trace | RCR-001…058
Acceptance Criteria | 58/58 RCRs mapped to a PV activity; every FI-flagged RCR (RCR-001/002/019/020/021/ 022/023/024/026/029/037/045 …) maps to a Test with fault injection; every IV-flagged RCR (RCR-034/035/036/ 037/038/039/051) maps to an independent pen-test/review; 0 controls "on paper only".
Level | System
Sample/Config | RCR register + PV register
Facility/Equipment/Tools | Trace tool
Independence | Quality + System-Safety
Standard/Compliance Reference | ISO 14971; IEC 62443 concepts
Evidence Required | RCR→PV coverage table (Verification-Traceability-Coverage-Matrix.md)
Owner | EM-ENG-006 + System-Safety
Status | Planned
Notes | The core Stage-12→Stage-13 handshake; see the coverage matrix.
7. Traceability Notes
This module proves the requirement set is verifiable and fully covered; it feeds PV-PLAN scheduling and PVL-1600 traceability. Every downstream module inherits the acceptance-criteria objectivity rule (PV-REQ-004).
8. Coverage/Completion Notes
Complete when the requirement baseline is frozen, quality-reviewed, and 100 % covered (incl. 58/58 RCRs) with no orphan requirement. All Planned.
PVL-900 (PV-SAFE) — Safety Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: System-Safety with EM-ENG-006 and Battery/Thermal, Electrical, Mechanical. This module verifies each HAL risk control was implemented correctly and is effective — every activity is the "Verification Reference" in the HAL hazard record (HA-<CODE>-NNN) and the RCR register (Stage 12). Every S4/S5 / high-severity safety control is verified by fault injection (FI): the fault is forced and the control is proven to hold. No residual risk (HAL-2800/2900) can be closed until its PV-SAFE activity passes. This module hosts the battery/charge/thermal fault-injection (PV-ENRG-###, PV-TFC-###) and the system-safety single-fault/safe-state proofs (PV-SAFE-###). Software-control FI lives in PVL-1100 (PV-SW); security-control IV in PVL-1000 (PV-SEC) — both cross-referenced in §6.3. All entries Planned/Evidence-Pending.
1. Purpose
Verify that every accepted risk control (RCR-001…058) is implemented and effective, with fault injection for every safety-critical control and single-fault-tolerance demonstrated for the S5 physical-energy and information-harm paths.
2. Scope
Battery/charge/thermal safety, mechanical/electrical safety, safe-state/fault handling, and the system-level single-fault and record-integrity-spine proofs. Cross-references the software (PV-SW) and security (PV-CYBER) control verifications. Excludes the environmental exposures those controls survive (PVL-700) and validation of use-effectiveness (Stage 14).
3. Applicable Standards & References
ISO 14971 (risk-control verification & effectiveness); IEC 62368-1 (energy sources, single-fault, touch- temperature, fire enclosure); UN 38.3 (battery safety); IEC 60068 (exposure). Designations; revisions to confirm.
4. Governance & Method
Method is Test with fault injection (FI) for all FI-flagged RCRs; Inspection/Analysis only where a physical test cannot isolate the control (justified in Notes). Each entry states the forced fault and the "control holds" pass criterion. S5 residuals (e.g., aircraft-hold fire) remain A-High by rule with mandatory monitoring even after the control passes (Risk-Management-Summary).
5. Verification Categories
Charge-safety FI · cell-protection FI · thermal/fire containment · touch-temperature · condensation/record electronics · safe-state/fault handling (→PV-SW) · single-fault tolerance · record-integrity spine · HAL risk-control verification reference index.
6. Verification Register (this module)
6.1 Battery / charge / thermal — physical-energy harm (FI mandatory)
PV-ENRG-001 — Cold-charge inhibit fault-injection sweep (charge = 0 A below limit)
Objective | Prove charging is prevented below the qualified cell lower charge limit by the MCU-independent hardware interlock, and that the inhibit persists under forced firmware/comms fault.
Method | Test — fault injection (FI)
I/O Trace | SYS-091/094/095; PRD-SAFE-007; PRD-ELEC-005/006; DO-SAFE-001, DO-EE-007, DO-FW-005
Risk-Control Trace | RCR-001; HA-ENRG-001, HA-ELEC-002, HA-USE-009 → HA-TFC-002 (S5)
Acceptance Criteria | Over a temperature sweep, charge current = 0 A at all cell temperatures below the qualified limit — including with firmware halted/faulted and comms lost; inhibited state + cause reported on HMI/LED. Exact lower charge limit Evidence-Pending on final cell datasheet.
Level | Subsystem
Sample/Config | DVT packs/units, HW interlock populated
Facility/Equipment/Tools | Climatic chamber; source-measure; fault-injection harness (halt FW / cut comms)
Independence | System-Safety executes; independent of firmware author
Standard/Compliance Reference | IEC 62368-1; UN 38.3; cell datasheet
Evidence Required | Cold-charge FI sweep report
Owner | Battery/Power + System-Safety
Status | Evidence-Pending (D3 charge boundary)
Notes | Primary D3 charge-safety control; one of the two D3 headline validate-by-test items.
PV-ENRG-002 — Cell-protection fault-injection (primary IC failed → secondary backstop)
Objective | Prove that with the primary protection IC forced non-responsive, an induced OV/OC/OT/short is terminated by the independent non-resettable secondary backstop before thermal-runaway onset.
Method | Test — fault injection (FI)
I/O Trace | SYS-096; PRD-SAFE-008; PRD-ELEC-007/014; DO-SAFE-002, DO-EE-005/010
Risk-Control Trace | RCR-002; HA-ENRG-002, HA-TFC-002 (S5), HA-TFC-003, HA-ELEC-006
Acceptance Criteria | With the primary IC forced failed, induced OV/OC/OT/short is cleared by the secondary backstop before the cell reaches runaway onset; no external flame propagation (with PV-TFC-002). Onset/clearing thresholds per qualified cell datasheet — Evidence-Pending.
Level | Subsystem
Sample/Config | DVT packs, secondary backstop populated
Facility/Equipment/Tools | Abuse-test rig; source-measure; thermal instrumentation
Independence | System-Safety + Battery-Safety; independent execution
Standard/Compliance Reference | IEC 62368-1; UN 38.3
Evidence Required | Protection FI report (DO-TST-009)
Owner | Battery/Power + Thermal/Battery-Safety
Status | Evidence-Pending
Notes | Release-critical; single-fault backstop non-negotiable for the S5 path.
PV-TFC-001 — Self-heat / touch-temperature bound & inhibit-integrity (if adopted)
Objective | Prove any self-heat feature only enables charging within the safe window, never bypasses the charge inhibit, and keeps element/touchable surfaces below the IEC 62368-1 burn threshold.
Method | Test — fault injection (FI on self-heat control)
I/O Trace | PRD-SAFE-007; DO-SAFE-001 (self-heat), DO-EE-007
Risk-Control Trace | RCR-005; HA-ENRG-001, HA-TFC-001 (burn); HA-NEW-002
Acceptance Criteria | Self-heat never enables charge below the safe window; element/touchable surface stay below the IEC 62368-1 burn limit; on any self-heat fault the charge inhibit still holds. Conditional on self-heat adoption (CMP-013); N/A if not adopted.
Level | Subsystem
Sample/Config | DVT units with self-heat (if adopted)
Facility/Equipment/Tools | Chamber; thermocouples; fault harness
Independence | Thermal/Battery-Safety
Standard/Compliance Reference | IEC 62368-1 touch-temperature
Evidence Required | Self-heat / touch-temp FI report
Owner | Thermal/Battery-Safety
Status | Evidence-Pending
Notes | RCR-005 verification reference; feasibility Evidence-Pending (CMP-013).
PV-TFC-002 — Fire-enclosure containment of a single-cell thermal event
Objective | Prove the enclosure/internal materials contain an induced single-cell runaway without external flame propagation.
Method | Test
I/O Trace | PRD-SAFE-008; PRD-MECH; DO-MECH-005, DO-PCB-010
Risk-Control Trace | RCR-008; HA-TFC-002 (S5), HA-TFC-003
Acceptance Criteria | An induced single-cell runaway is contained with no external flame; consistent with the IEC 62368-1 fire-enclosure model. (S5 residual remains A-High with mandatory thermal-event monitoring.)
Level | Subsystem
Sample/Config | DVT enclosures + induced-runaway cell
Facility/Equipment/Tools | Abuse/fire-safety test cell
Independence | Thermal/Battery-Safety + Mechanical; independent execution
Standard/Compliance Reference | IEC 62368-1 fire enclosure; UN 38.3
Evidence Required | Containment test report
Owner | Mechanical + Thermal/Battery-Safety
Status | Planned
Notes | RCR-008 verification reference; last-line backstop for the S5 crush/thermal path.
PV-TFC-003 — Offgas characterisation (enclosed-space / aircraft-hold exposure)
Objective | Characterise the toxicity/quantity of offgas from an induced single-cell event against enclosed-space exposure limits.
Method | Test
I/O Trace | PRD-SAFE-008; DO-MECH-005
Risk-Control Trace | RCR-008; HA-TFC-003 (toxic offgas)
Acceptance Criteria | Offgas toxicity/quantity characterised and consistent with aircraft-hold enclosed- space exposure limits. Offgas limits/quantities Evidence-Pending.
Level | Subsystem
Sample/Config | Induced-runaway units in gas-capture rig
Facility/Equipment/Tools | Gas-analysis / offgas-capture setup
Independence | Battery-Safety + third-party lab
Standard/Compliance Reference | enclosed-space exposure limits (to confirm)
Evidence Required | Offgas characterisation report
Owner | Thermal/Battery-Safety
Status | Evidence-Pending
Notes | RCR-008 verification reference; supports the S5 aircraft-hold benefit-risk.
PV-TFC-006 — Condensation-cycle record-electronics integrity (safety-relevant)
Objective | Prove the conformal-coat + seal protect record-integrity electronics through cold↔warm condensation cycling over reuse life.
Method | Test
I/O Trace | PRD-SAFE-010; HWR-038/040; DO-SAFE-011, DO-MECH-005, DO-PCB-010
Risk-Control Trace | RCR-015/018; HA-TFC-006, HA-ELEC-006/007, HA-ENV-003
Acceptance Criteria | After N cold↔warm condensation cycles: no record-integrity fault; insulation resistance and cleanliness within limits. Cycle count/IR limits Evidence-Pending.
Level | Subsystem
Sample/Config | DVT PCBAs/units
Facility/Equipment/Tools | Temp-humidity cycling chamber; IR meter
Independence | Mechanical/PCB
Standard/Compliance Reference | IEC 60068-2-30/-2-38; IEC 60529
Evidence Required | Condensation-cycle IR report
Owner | Mechanical/PCB (EM-ENG-002/003)
Status | Evidence-Pending
Notes | RCR-015/018 verification reference; joint with PV-ENV-003.
6.2 Mechanical / electrical safety
PV-SAFE-001 — SELV / no accessible hazardous voltage (single-fault)
Objective | Prove no accessible point exceeds SELV limits in normal or single-fault conditions.
Method | Test
I/O Trace | PRD-SAFE-008; PRD-ELEC; DO-EE (rails), DO-PCB
Risk-Control Trace | RCR-013; HA-ELEC-001
Acceptance Criteria | No accessible point exceeds SELV limits in normal or single-fault per IEC 62368-1.
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Electrical-safety test set
Independence | Electrical + Safety
Standard/Compliance Reference | IEC 62368-1
Evidence Required | Electrical-safety test report
Owner | Electrical (EM-ENG-003)
Status | Planned
Notes | RCR-013 verification reference.
PV-SAFE-002 — No accessible sharp / pinch / crush point
Objective | Prove no accessible hazardous edge/pinch/crush point in normal handling or service.
Method | Inspection
I/O Trace | PRD-MECH; PRD-SAFE-005; DO-MECH-002/003
Risk-Control Trace | RCR-009; HA-MECH-001/002
Acceptance Criteria | No accessible edge/pinch exceeds IEC 62368-1 accessibility limits in handling and serviceable-part access.
Level | Subsystem
Sample/Config | DVT units
Facility/Equipment/Tools | Accessibility probes
Independence | Mechanical + Safety
Standard/Compliance Reference | IEC 62368-1 accessibility
Evidence Required | Edge/pinch inspection record
Owner | Mechanical (EM-ENG-002)
Status | Planned
Notes | RCR-009 verification reference; small-parts = PV-MECH-003.
6.3 Safe-state / fault-handling & information-harm controls (cross-reference index)
The following safety-critical controls are verified by fault injection in PVL-1100 (PV-SW) or independent adversarial verification in PVL-1000 (PV-SEC); they are indexed here because each is a HAL risk-control Verification Reference and must pass before its residual risk closes:
RCR | Control | Verified by | FI/IV
RCR-019 | Two-sided detection, no undetected path | PV-SW-001 | FI
RCR-020 | Probe-fault / implausible-read | PV-SW-008 | FI
RCR-021 | Verdict integrity (suppress false accept) | PV-SW-002 (+PV-CLOUD-003) | FI
RCR-022 | Honest gap / no clean pass on incomplete | PV-SW-002 / PV-DATA-004 | FI
RCR-023 | Safe state on fault + watchdog recovery | PV-SW-009 | FI
RCR-024 | Power-fail-safe storage (no torn/lost record) | PV-SW-013 | FI
RCR-025 | Trusted monotonic time / drift bound | PV-SW-005 | Test
RCR-026 | Calibration integrity / expiry blocks arming | PV-SW-004 | FI
RCR-027 | Availability self-check (flags "not monitoring") | PV-SW-009 | FI
RCR-029 | Alert-ahead + escalation | PV-SW-006 | FI
RCR-037 | Signed, rollback-safe A/B OTA | PV-SW-010 + PV-CYBER-005 | FI+IV
RCR-045 | Pre-arm readiness gate | PV-SW-003 | FI
RCR-034…039, 051 | Security-induced-safety controls | PV-CYBER-001…009 | IV
PV-SAFE-003 — System single-fault-tolerance integration proof
Objective | Prove at system integration that no single fault across the physical-energy and information- harm paths defeats a safety function, confirming the PV-ANAL-011 analysis by test at the integrated level.
Method | Test — fault injection (FI)
I/O Trace | SYS-safety; FAL-1800 (FTA top events); DO-SAFE-001…014
Risk-Control Trace | RCR-001/002/021/022/057 (single-fault sets); FTA TOP-A/B/C/D
Acceptance Criteria | For each FTA top event, forcing any single credible fault at the integrated system does not produce the top event (charge below limit, cell runaway, false accept, undelivered alert, lost record); backstops engage as designed.
Level | System
Sample/Config | Integrated DVT system (device+cloud)
Facility/Equipment/Tools | System fault-injection harness
Independence | System-Safety; independent of design authors
Standard/Compliance Reference | ISO 14971; IEC 62368-1 single-fault
Evidence Required | System single-fault FI report
Owner | System-Safety (EM-ENG-006)
Status | Planned
Notes | The integrated confirmation that the unit/subsystem FI proofs hold together.
PV-SAFE-004 — Record-integrity spine integration proof (no wrong disposition from data loss)
Objective | Prove across the integrated detection→verdict→record→export chain that no single data-loss/ corruption path produces a wrong disposition.
Method | Test — fault injection (FI)
I/O Trace | SYS spine; PRD-SAFE-001/003; DO-SAFE-004…008
Risk-Control Trace | RCR-057 (aggregate); RCR-019/021/022/024/042; HA-MPE-001/002/004
Acceptance Criteria | No single data-loss/corruption path across the integrated chain yields a wrong disposition; the S5 mass-administration top event is bounded to residual A-High with mandatory monitoring.
Level | System
Sample/Config | Integrated DVT system
Facility/Equipment/Tools | Fault-injection + record-corpus harness
Independence | Chief Systems (EM-ENG-001) + System-Safety
Standard/Compliance Reference | ISO 14971; ALCOA+
Evidence Required | Record-integrity spine FI report
Owner | Chief Systems + System-Safety
Status | Planned
Notes | RCR-057 verification reference; release-critical aggregate.
PV-SAFE-005 — Residual-risk & risk-control effectiveness review (per control)
Objective | Confirm every HAL risk control has a passed verification reference and its residual risk is acceptable, feeding HAL-2800/2900.
Method | Inspection, Analysis
I/O Trace | HAL-2600 (risk-control verification), HAL-2800/2900 (residual)
Risk-Control Trace | all RCR-001…058
Acceptance Criteria | 58/58 RCRs have a passed (or explicitly deferred) verification reference; every S4/S5 residual is reviewed; no residual closes without its PV evidence; open items (D3, RF) are gated.
Level | System
Sample/Config | RCR register + PV evidence index
Facility/Equipment/Tools | QMS; risk-management file
Independence | System-Safety + Quality/Regulatory
Standard/Compliance Reference | ISO 14971 §7/§8
Evidence Required | Risk-control effectiveness / residual review
Owner | System-Safety + EM-ENG-006
Status | Planned
Notes | Feeds RCR-058 release gate and PVL-1900 summary.
7. Traceability Notes
Every PV-SAFE/PV-ENRG/PV-TFC activity is a HAL hazard-record "Verification Reference" and an RCR verification target; the §6.3 index maps the software/security-verified controls. This module reconciles with HAL Stage 10 (HAL-2600) and gates residual-risk closure (HAL-2800/2900) and the RCR-058 release gate.
8. Coverage/Completion Notes
Complete when every safety control's verification reference passes (FI for safety-critical, IV for security) and the residual-risk review closes — except the two carried-Open items (D3 battery charge/endurance; RF-in-container), which gate release. All Planned/Evidence-Pending.
PVL-1000 (PV-SEC) — Cybersecurity Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Product-Security (EM-ENG-008) with independent testers. Verifies the SCM's security controls — authentication, authorisation, encryption, secure boot, OTA security, tamper/integrity, secure storage, audit logging, SBOM/dependency, key lifecycle — with every security control verified by independent adversarial verification (IV / penetration testing), not by the implementer. Hosts the PV-CYBER-### pen-tests referenced by Stage-12 (RCR-034…039, RCR-051). Security controls here are also safety controls (cybersecurity-induced safety). All entries Planned/Evidence-Pending.
1. Purpose
Verify that the SCM resists compromise that could falsify records, suppress alerts, defeat detection, brick or subvert firmware, clone identity, or blind the fleet — each proven by an independent adversarial test.
2. Scope
Device (S3), connectivity (S5), cloud (S8), mobile (S9), and OTA (S3/S8) security. Includes SBOM/dependency verification and key lifecycle. Excludes secure-coding review (PVL-300) and functional auth flows (PVL-500).
3. Applicable Standards & References
IEC 62443 concepts (industrial/IoT security assurance); OWASP ASVS / MASVS (application/mobile); NIST SP 800- series concepts; CycloneDX/SPDX (SBOM); TLS best practice. Designations; revisions to confirm.
4. Governance & Method
Method is Test via independent adversarial verification (IV) for the safety-relevant controls, plus Analysis/Inspection for SBOM/dependency. IV is executed by a party independent of firmware/cloud/mobile development (PV-GOV-002). Pen-test scope derives from the threat model (PV-ANAL-012).
5. Verification Categories
Capture-time integrity/authenticity · alert-path non-suppression · config/threshold authorisation · secure OTA · secure boot / debug lock · cloud resilience & DR · TLS/transport · secure storage · audit logging · SBOM/ dependency · key rotation/lifecycle.
6. Verification Register (this module)
PV-CYBER-001 — Capture-time authenticity/integrity pen-test (falsified data rejected)
Objective | Prove telemetry/records carry capture-time authenticity and that falsified/injected data is detected and rejected under an independent pen-test.
Method | Test — independent adversarial (IV)
I/O Trace | SYS-026/028; PRD-SEC → PRD-SAFE-001; DO-SEC (integrity), DO-SAFE-008
Risk-Control Trace | RCR-034; HA-CYBER-001, HA-DATA-001
Acceptance Criteria | Injected/falsified data is detected and rejected; capture-time authenticity holds under the independent pen-test; no undetected forgery path. Pen-test scope Evidence-Pending.
Level | System
Sample/Config | DVT device + cloud
Facility/Equipment/Tools | Independent pen-test lab
Independence | External pen-tester (independent of dev)
Standard/Compliance Reference | IEC 62443; OWASP
Evidence Required | Pen-test report (integrity)
Owner | Product-Security (EM-ENG-008)
Status | Evidence-Pending
Notes | RCR-034 verification reference; required before residual acceptance.
PV-CYBER-002 — Alert-path non-suppression pen-test
Objective | Prove the alert/alarm path cannot be silently suppressed — any suppression/blocking raises a detectable, escalated event.
Method | Test — IV
I/O Trace | PRD-SEC → PRD-SAFE-004; DO-SEC, DO-SAFE-009, DO-FW-016
Risk-Control Trace | RCR-035; HA-CYBER-002, HA-COMM-002
Acceptance Criteria | Any attempt to mute/block the alert path raises a distinct escalated event; no silent suppression under pen-test.
Level | System
Sample/Config | DVT device + cloud
Facility/Equipment/Tools | Independent pen-test lab
Independence | External pen-tester
Standard/Compliance Reference | IEC 62443
Evidence Required | Suppression pen-test report
Owner | Product-Security (EM-ENG-008)
Status | Evidence-Pending
Notes | RCR-035 verification reference.
PV-CYBER-003 — Threshold/config-tamper authorisation pen-test
Objective | Prove excursion-threshold and config changes are authenticated, authorised, bounded, and audited; no unauthorised or out-of-bounds change succeeds.
Method | Test — IV
I/O Trace | SYS-107/110; PRD-SEC → PRD-SAFE-001; DO-SEC, DO-CLD (audit)
Risk-Control Trace | RCR-036; HA-CYBER-003, HA-USE-003
Acceptance Criteria | No unauthorised/out-of-bounds threshold/config change succeeds; every change is authenticated, bounded, and audited under pen-test (incl. BLE unauthorised pairing).
Level | System
Sample/Config | DVT device + mobile + cloud
Facility/Equipment/Tools | Independent pen-test lab
Independence | External pen-tester
Standard/Compliance Reference | IEC 62443; OWASP MASVS
Evidence Required | Config-tamper pen-test report
Owner | Product-Security + Cloud
Status | Evidence-Pending
Notes | RCR-036 verification reference.
PV-CYBER-005 — Secure-OTA pen-test (signing/rollback cannot be defeated)
Objective | Prove OTA signing, integrity, and rollback protection cannot be defeated — unsigned/corrupt images are rejected, a bad image auto-rolls-back, and no device bricks.
Method | Test — IV
I/O Trace | SYS-032/114/139; PRD-SAFE-010 + PRD-SEC-008; DO-SAFE-014, DO-FW-014/015, DO-CLD-010
Risk-Control Trace | RCR-037; HA-SW-010, HA-CYBER-005 (S5)
Acceptance Criteria | Unsigned/corrupt image rejected; bad image auto-rolls-back with trip data preserved; no brick under version-skew; staged rollout halts on regression; signing/rollback not defeatable under pen-test. (Functional rollback = PV-SW-010.) Pen-test scope Evidence-Pending.
Level | System
Sample/Config | DVT device + OTA service
Facility/Equipment/Tools | Independent pen-test lab
Independence | External pen-tester
Standard/Compliance Reference | IEC 62443; secure-boot/OTA practice
Evidence Required | OTA pen-test report
Owner | Product-Security + Firmware + Cloud
Status | Evidence-Pending
Notes | RCR-037 verification reference; dominant SCM control.
PV-CYBER-006 — Secure-boot / debug-lock / field-mode-lock pen-test
Objective | Prove only signed firmware boots, production debug/dev/service ports are non-functional in the field, and no unexited service mode can be exploited.
Method | Test — IV
I/O Trace | SYS-032/114; PRD-SEC → PRD-SAFE-010; DO-SEC-006, DO-EE-012, DO-FW-017
Risk-Control Trace | RCR-038, RCR-051; HA-CYBER-006/008, HA-SVC-006
Acceptance Criteria | Only signed firmware boots; debug/dev/service ports non-functional in production; tamper evident; no field device boots with debug/dev mode active; no bypass under pen-test.
Level | System
Sample/Config | Production-representative units
Facility/Equipment/Tools | Independent pen-test lab; JTAG/SWD probes
Independence | External pen-tester
Standard/Compliance Reference | IEC 62443
Evidence Required | Secure-boot/mode-lock pen-test report
Owner | Product-Security (EM-ENG-008)
Status | Evidence-Pending
Notes | RCR-038/051 verification reference; physical lock inspected in PV-INSP-010.
PV-CYBER-009 — Cloud security & disaster-recovery test (fleet cannot be silently blinded)
Objective | Prove the cloud enforces RBAC/SSO, monitors auth/integrity/backlog events, and has a tested backup/restore so a platform incident cannot silently blind the fleet or corrupt records.
Method | Test — IV (+ DR drill)
I/O Trace | PRD-SEC → PRD-SAFE-011; PRD-CLOUD; DO-CLD, DO-SEC
Risk-Control Trace | RCR-039; HA-CYBER-009, HA-CLOUD-004/006/008
Acceptance Criteria | RBAC/SSO enforced; auth/integrity/backlog monitoring in place; backup/restore demonstrated within the DR objective; a simulated compromise cannot silently blind the fleet or corrupt records. DR RTO/RPO Evidence-Pending.
Level | System
Sample/Config | Cloud production-representative environment
Facility/Equipment/Tools | Independent security assessment; DR drill
Independence | External assessor
Standard/Compliance Reference | IEC 62443; SOC-2/ISO 27001 concepts
Evidence Required | Cloud security assessment + DR-test report
Owner | Cloud + Product-Security
Status | Evidence-Pending
Notes | RCR-039 verification reference; offline-first (RCR-028) is the independence guarantee.
PV-SEC-001 — TLS / transport-security verification (device↔cloud, mobile↔cloud)
Objective | Confirm all links use current TLS with correct certificate validation and no downgrade.
Method | Test
I/O Trace | SRS-SEC (600); PRD-SEC; DO-SEC, DO-IF (protocol)
Risk-Control Trace | RCR-034 (integrity in transit), RCR-036
Acceptance Criteria | TLS configuration meets the adopted baseline; certificate validation enforced; no downgrade/weak-cipher path; scanned clean.
Level | Integration
Sample/Config | Device/mobile ↔ cloud
Facility/Equipment/Tools | TLS scanner; MITM proxy
Independence | Security
Standard/Compliance Reference | TLS best practice; OWASP
Evidence Required | TLS/transport test report
Owner | Product-Security (EM-ENG-008)
Status | Planned
Notes | —
PV-SEC-002 — Secure storage (keys / credentials / record at rest)
Objective | Confirm keys, credentials, and records at rest are protected on device and in cloud.
Method | Test, Inspection
I/O Trace | SRS-SEC; PRD-SEC; DO-SEC, DO-FW-010
Risk-Control Trace | RCR-034 (integrity), RCR-038 (root of trust)
Acceptance Criteria | Keys/credentials stored in protected storage (secure element/keystore); records at rest protected; no plaintext secret recoverable under inspection/extraction.
Level | Subsystem
Sample/Config | DVT device + cloud
Facility/Equipment/Tools | Extraction tools; storage inspection
Independence | Security
Standard/Compliance Reference | IEC 62443; OWASP MASVS
Evidence Required | Secure-storage test report
Owner | Product-Security
Status | Planned
Notes | —
PV-SEC-003 — Audit-log integrity & completeness verification
Objective | Confirm security/safety-relevant events (config change, override, suppression, auth failure) are logged completely and tamper-evidently.
Method | Test
I/O Trace | SRS-SEC; PRD-SEC; DO-CLD (audit)
Risk-Control Trace | RCR-036 (audited config), RCR-044 (override audit)
Acceptance Criteria | All defined security/safety events appear in the audit log with actor/time; log is tamper-evident; no gap under test.
Level | Integration
Sample/Config | Cloud + device
Facility/Equipment/Tools | Audit-log test harness
Independence | Security + Quality
Standard/Compliance Reference | ALCOA+; 21 CFR Part 11-style
Evidence Required | Audit-log verification report
Owner | Product-Security + Cloud
Status | Planned
Notes | —
PV-SEC-004 — SBOM generation & dependency-vulnerability verification
Objective | Confirm a complete SBOM is produced for device/cloud/mobile and dependencies are scanned for known vulnerabilities with a disposition process.
Method | Analysis, Inspection
I/O Trace | SRS-SEC; PRD-SEC; DO-CFG (build manifests)
Risk-Control Trace | RCR-037 (supply-chain integrity), RCR-039
Acceptance Criteria | SBOM (CycloneDX/SPDX) generated for each build; dependency scan run; no unmitigated critical vulnerability at release; disposition recorded.
Level | System
Sample/Config | Release builds
Facility/Equipment/Tools | SBOM/SCA tooling
Independence | Security
Standard/Compliance Reference | CycloneDX/SPDX; NIST concepts
Evidence Required | SBOM + vulnerability report
Owner | Product-Security (EM-ENG-008)
Status | Planned
Notes | Ties PV-GOV-003 configuration control.
PV-SEC-005 — Key-lifecycle / rotation & erasure verification
Objective | Confirm device keys are provisioned uniquely, rotatable, and erasable at decommission (bounds the key-lifecycle single point).
Method | Test
I/O Trace | PRD-SEC; DO-SEC, DO-MFG (provisioning)
Risk-Control Trace | RCR-055 (unique identity), RCR-052 (key erasure at EOL), RCR-038
Acceptance Criteria | Keys unique per device; rotation works without service loss; key/data erasure demonstrable at decommission; no key reuse across devices.
Level | System
Sample/Config | DVT devices
Facility/Equipment/Tools | Provisioning + key-management tools
Independence | Security + Manufacturing
Standard/Compliance Reference | IEC 62443; NIST key-management concepts
Evidence Required | Key-lifecycle test report
Owner | Product-Security + Manufacturing
Status | Planned
Notes | Erasure is the RCR-052 verification (with disposal review).
7. Traceability Notes
Every PV-CYBER activity is the verification reference for a security-induced-safety RCR (RCR-034…039, 051) and is executed IV per HAL-2600. Security controls also appear in the PV-SAFE §6.3 index (they gate residual-risk closure). SBOM/key-lifecycle support supply-chain and EOL controls. Feeds the compliance and summary modules.
8. Coverage/Completion Notes
Complete when all IV pen-tests pass with findings dispositioned, SBOM is clean at release, and key lifecycle is demonstrated. Pen-test scope/DR objectives Evidence-Pending. All Planned/Evidence-Pending.
PVL-1900 (PV-SUM) — Verification Summary — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Quality/Verification (EM-ENG-006) + Chief Systems (EM-ENG-001). The verification conclusion: coverage, pass rate, open issues, residual risks, and the release recommendation — assembled after execution. This is a playbook: no results exist yet; the summary is a planned template with the acceptance gates it will apply. All entries Planned; the two carried-Open items (D3, RF-in-container) are explicitly release-gating.
1. Purpose
Define the verification summary and release-recommendation the program will produce when the PVL is executed: coverage, pass rate, open issues, residual-risk status, and the go/no-go recommendation into Validation (Stage 14) and Release (Stage 16).
2. Scope
Program-level roll-up of all 20 modules. Consumes PVL-1600 coverage, PVL-1700 open issues, PVL-900/1000 safety/ security closure, and PVL-1400 D3/RF results.
3. Applicable Standards & References
ISO 13485 §7.3.6 (verification conclusion); ISO 14971 (residual-risk/benefit-risk); the Risk-Management-Summary release conditions. Designations; revisions to confirm.
4. Governance & Method
Method is Inspection/Analysis (roll-up + gate application). The summary applies severity-first gates: no safety/security nonconformance open; the two Open items closed by test; residual risk accepted per ISO 14971.
5. Verification Categories
Coverage summary · pass-rate · open-issue register · residual-risk status · conclusions · release recommendation.
6. Verification Register (this module)
PV-SUM-001 — Verification coverage & pass-rate summary
Objective | Summarise coverage (requirement/RCR/FI/IV %) and pass rate across all modules.
Method | Inspection, Analysis
I/O Trace | PVL-1600 coverage; all module registers
Risk-Control Trace | RCR-058
Acceptance Criteria | Coverage = 100 % of applicable requirements and 58/58 RCRs; pass rate reported; every FI/IV activity accounted for; exceptions listed.
Level | System
Sample/Config | All PV results (at execution)
Facility/Equipment/Tools | Trace tool; QMS
Independence | Quality + Chief Systems
Standard/Compliance Reference | ISO 13485 §7.3.6
Evidence Required | Coverage/pass-rate summary
Owner | EM-ENG-006 + EM-ENG-001
Status | Planned
Notes | —
PV-SUM-002 — Open-issue & residual-risk status
Objective | Roll up open verification issues and the residual-risk status, flagging any safety/security item and the two Open program items.
Method | Inspection, Analysis
I/O Trace | PVL-1700 defects; PV-SAFE-005 residual; PVL-1400 D3/RF
Risk-Control Trace | RCR-058; O-1 (RF), O-2 (D3)
Acceptance Criteria | No safety/security issue open without disposition; residual-risk status per control; D3 (PV-REL-001) and RF-in-container (PV-REL-005) status explicit.
Level | System
Sample/Config | Defect + residual registers
Facility/Equipment/Tools | QMS
Independence | Quality + System-Safety
Standard/Compliance Reference | ISO 14971
Evidence Required | Open-issue/residual summary
Owner | EM-ENG-006 + System-Safety
Status | Planned
Notes | —
PV-SUM-003 — Verification conclusion & release recommendation
Objective | State the verification conclusion (design outputs satisfy design inputs) and a go/no-go recommendation into Validation and Release.
Method | Inspection, Analysis
I/O Trace | all modules; PRD-VER-001…008; RCR-058
Risk-Control Trace | RCR-058 (release gate); the eight release conditions
Acceptance Criteria | Verification conclusion supported by evidence; the eight release conditions (Risk-Management-Summary §7) evaluated; the two Open items closed or the release is conditional/blocked; recommendation signed by EM-ENG-006 + EM-ENG-001 (+ benefit-risk for S5 residual).
Level | System
Sample/Config | Full verification package
Facility/Equipment/Tools | Release/design-review board
Independence | Quality + Chief Systems + System-Safety
Standard/Compliance Reference | ISO 13485; ISO 14971
Evidence Required | Verification summary report + release recommendation
Owner | EM-ENG-006 + EM-ENG-001
Status | Planned
Notes | Feeds Validation (Stage 14) and Release (Stage 16).
PV-SUM-004 — Lessons-learned & feed to ORCH-018 / Stage 14/16
Objective | Capture verification lessons-learned and hand the verified baseline + open items to Validation (Stage 14), Release (Stage 16), and ORCH-018 / EM-ENG-012.
Method | Inspection
I/O Trace | PVL-1600; PVaL (Stage 14); Release (Stage 16); ORCH-018
Risk-Control Trace | RCR-058
Acceptance Criteria | Lessons-learned recorded; the verified baseline and residual open items handed off; the Stage-13 segment integrated into the lifecycle traceability spine.
Level | System
Sample/Config | Summary package
Facility/Equipment/Tools | PLM/QMS
Independence | Quality + EM-ENG-012
Standard/Compliance Reference | ISO 13485
Evidence Required | Lessons-learned + handoff record
Owner | EM-ENG-006 + EM-ENG-012
Status | Planned
Notes | —
7. Traceability Notes
The summary rolls up all 20 modules and applies the RCR-058 release gate; it feeds Validation (Stage 14), Release (Stage 16), and ORCH-018 / EM-ENG-012 (Stage 18).
8. Coverage/Completion Notes
Complete when the verification is executed, coverage/pass-rate/open-issues/residual-risk are summarised, and a signed release recommendation exists. The D3 and RF-in-container Open items gate a clean recommendation. All Planned (pre-execution).
PVL-1100 (PV-SW) — Software Verification — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Embedded SW (EM-ENG-007) with Cloud/Mobile and System-Safety. Verifies the SCM software (device firmware, cloud, mobile) by unit/integration test, static & dynamic analysis, coverage, timing, and recovery — and hosts the safety-critical software fault-injection (FI) activities that are the verification references for the information-integrity RCRs (RCR-019…027, RCR-029, RCR-037, RCR-045). Under IEC 62304 the software safety classification drives rigor. All entries Planned/Evidence-Pending.
1. Purpose
Verify software correctness and robustness: the detection/verdict/record/time/alert logic under fault injection, plus unit/integration/static-analysis/coverage/timing/recovery across firmware, cloud, and mobile.
2. Scope
Device firmware (S3, SRS-DEV 200), cloud (S8, SRS-CLOUD 400), mobile (S9, SRS-MOB 500), OTA (S3/S8, SRS-OTA 700), plus data/reliability/safety software modules. Excludes functional black-box flows (PVL-500, cross- referenced) and security pen-tests (PVL-1000).
3. Applicable Standards & References
IEC 62304 (software lifecycle, unit/integration verification, class-driven rigor); the adopted coding standard (MISRA-style, to confirm); coverage criteria (statement/branch/MC-DC as classified); ISO 14971 (software risk controls). Designations; revisions to confirm.
4. Governance & Method
Method is Test (unit/integration/fault-injection) and Analysis (static/coverage/timing). Every FI-flagged software RCR is verified by forcing the fault and proving the control holds. Coverage targets are set by the IEC 62304 software safety class of each module. Fault-injection tests here are the robustness proofs behind the functional-correctness tests in PVL-500.
5. Verification Categories
Detection FI · probe-fault FI · verdict/gap FI · safe-state/watchdog FI · power-fail-storage FI · calibration- integrity FI · clock/time · alert-gen/escalation FI · pre-arm-gate FI · OTA rollback/version-skew · unit/ integration · static analysis · coverage · timing/schedulability · recovery/logging.
6. Verification Register (this module)
6.1 Safety-critical software fault-injection (RCR verification references)
PV-SW-001 — Profiled-excursion detection FI (freeze reliability = heat; no undetected path)
Objective | Prove two-sided detection with freeze reliability equal to heat and no undetected-excursion path, by injecting a full profiled-excursion battery.
Method | Test — fault injection (FI)
I/O Trace | SYS-010/014/149/150; PRD-SAFE-002/001; PRD-FUN-006/007; DO-FW-006, DO-EE-013
Risk-Control Trace | RCR-019; HA-SW-001, HA-ENV-001, FTA TOP-A
Acceptance Criteria | Missed-freeze rate ≤ missed-heat rate and both within the validated detection target; no injected input profile produces an undetected excursion. Target values Evidence-Pending → anchored to WHO PQS accuracy target (O-7).
Level | Subsystem
Sample/Config | DVT firmware + profile injector (DO-VER-007)
Facility/Equipment/Tools | Detection test harness; reference profiles
Independence | Embedded SW; System-Safety reviews
Standard/Compliance Reference | IEC 62304; WHO PQS E006 (to confirm)
Evidence Required | Detection FI report
Owner | Embedded SW (EM-ENG-007)
Status | Evidence-Pending
Notes | RCR-019 verification reference; top HAL headline; validated in VAL-SW-001.
PV-SW-002 — Gap-handling / verdict-consistency FI (no false clean accept)
Objective | Prove no combination of data gap, sensor fault, config, or user action produces an unqualified "accept" over an incomplete record, and device↔cloud verdicts stay consistent.
Method | Test — FI
I/O Trace | SYS-015/016/017/151; PRD-SAFE-001/003; DO-FW-008, DO-CLD-006, DO-HF-002
Risk-Control Trace | RCR-021, RCR-022; HA-SW-002/007, HA-DATA-004, FTA TOP-A/D
Acceptance Criteria | No gap/fault/config/user action yields an unqualified accept; device and cloud agree or resolve to qualify/reject; the control cannot be disabled from FW/config/UI.
Level | Integration
Sample/Config | DVT device + cloud + fault injector
Facility/Equipment/Tools | Gap/fault-injection harness
Independence | Embedded SW + Cloud; System-Safety reviews
Standard/Compliance Reference | IEC 62304; ISO 14971
Evidence Required | Verdict-integrity FI report
Owner | Embedded SW + Cloud
Status | Planned
Notes | RCR-021/022 verification reference; most important information-harm control.
PV-SW-003 — Pre-arm readiness-gate FI (non-defeatable)
Objective | Prove arming is blocked unless SoC/calibration/firmware/profile/probe all pass, and the gate is non-defeatable under fault, and validate profile-mismatch flagging.
Method | Test — FI
I/O Trace | SYS-023/024/154; PRD-SAFE-006; DO-SAFE-007, DO-FW-032
Risk-Control Trace | RCR-045, RCR-048; HA-USE-003/004, HA-SW-003, FTA
Acceptance Criteria | No trip arms with any failing readiness item even under injected fault; blocked arming shows its cause; an implausible/mismatched profile is flagged and blocks arming.
Level | Subsystem
Sample/Config | DVT firmware + injector
Facility/Equipment/Tools | Fault-injection harness
Independence | Firmware; System-Safety reviews
Standard/Compliance Reference | IEC 62304
Evidence Required | Pre-arm gate FI report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | RCR-045/048 verification reference; non-defeatable.
PV-SW-004 — Calibration-integrity / expiry FI (blocks arming)
Objective | Prove corrupted/expired calibration is detected and blocks arming, and a units/coefficient error is caught by integrity check + static analysis.
Method | Test — FI (+ static analysis)
I/O Trace | SYS-007/024; PRD-SAFE-006/001; DO-FW-007/009, DO-SAFE-007
Risk-Control Trace | RCR-026, RCR-049; HA-SW-014/004, FTA TOP-A (CCF)
Acceptance Criteria | Corrupted/expired coefficients are detected and block arming; a units/coefficient error is caught. Interval/coverage values Evidence-Pending → PV-MFG-001.
Level | Subsystem
Sample/Config | DVT firmware + coefficient injector
Facility/Equipment/Tools | Fault harness; static analyzer
Independence | Metrology + Firmware
Standard/Compliance Reference | IEC 62304; ISO/IEC 17025
Evidence Required | Calibration-integrity FI report
Owner | Metrology + Firmware
Status | Planned
Notes | RCR-026/049 verification reference.
PV-SW-005 — Trusted-time / clock-drift verification over cold/sleep
Objective | Prove time stays monotonic and within the drift bound across sleep/brown-out/service over cold, and that loss of trust is disclosed.
Method | Test
I/O Trace | SYS-006/118; PRD-SAFE-003; DO-EE-011 (RTC), DO-FW
Risk-Control Trace | RCR-025; HA-SW-005, HA-DATA-003, FTA TOP-D
Acceptance Criteria | Time monotonic and within the drift bound across sleep/brown-out/service over cold; on loss of trust the record/status disclose "untrusted time." Drift bound Evidence-Pending.
Level | Subsystem
Sample/Config | DVT units in chamber
Facility/Equipment/Tools | Chamber; time reference
Independence | Firmware
Standard/Compliance Reference | IEC 62304
Evidence Required | Clock-drift report
Owner | Firmware (EM-ENG-007)
Status | Evidence-Pending
Notes | RCR-025 verification reference; predicted by PV-ANAL-014.
PV-SW-006 — Alert-generation & escalation FI (alert-ahead, ack, escalate)
Objective | Prove an excursion alert is generated and prioritised ahead of bulk, requires ack, and escalates to an alternate responder if unacknowledged, under injected loss/fault.
Method | Test — FI
I/O Trace | SYS-011/013; PRD-SAFE-004; DO-SAFE-009, DO-CLD-007/008
Risk-Control Trace | RCR-029; HA-COMM-001, HA-CLOUD-001, FTA TOP-B
Acceptance Criteria | Alert precedes bulk, requires ack, escalates to an alternate responder within the target window if unacknowledged, under injected drops; consolidation holds volume within target. Latency/ escalation targets Evidence-Pending → VAL-USE-002.
Level | Integration
Sample/Config | Device + cloud + fault injector
Facility/Equipment/Tools | Alert/fault harness
Independence | Firmware + Cloud; Human-Factors reviews
Standard/Compliance Reference | IEC 62304
Evidence Required | Alert-escalation FI report
Owner | Firmware + Cloud
Status | Planned
Notes | RCR-029 verification reference.
PV-SW-008 — Probe-fault-injection sweep (never a valid product temperature)
Objective | Prove every injected probe fault (absent/open/short/wrong) and every implausible read is flagged and never scored as a valid product-zone temperature, incl. internal-vs-external divergence.
Method | Test — FI
I/O Trace | SYS-007/024/151; PRD-SAFE-001; HWR-022/028; DO-FW-007, DO-EE-013, DO-IF-008
Risk-Control Trace | RCR-020; HA-SW-008, HA-USE-005, FTA TOP-A
Acceptance Criteria | Every injected probe fault/implausible read is flagged and never scored valid; divergence beyond threshold raises a probe-fault flag. Thresholds Evidence-Pending.
Level | Subsystem
Sample/Config | DVT + probe-emulation fixture
Facility/Equipment/Tools | Probe-fault injector
Independence | Embedded SW + Sensing
Standard/Compliance Reference | IEC 62304
Evidence Required | Probe-fault FI report
Owner | Embedded SW + Sensing
Status | Planned
Notes | RCR-020 verification reference.
PV-SW-009 — Safe-state / watchdog & availability self-check FI
Objective | Prove any injected hang/fault transitions to the safe state without corrupting the committed log or armed state, the watchdog recovers with the event logged, and every chain outage is surfaced as "not monitoring."
Method | Test — FI
I/O Trace | SYS-047/099/167/169/170; PRD-SAFE-010; DO-SAFE-006/007, DO-FW-003/012/013
Risk-Control Trace | RCR-023, RCR-027; HA-SW-009/013, FTA TOP-A
Acceptance Criteria | Every injected hang/fault → safe state without corrupting the log/armed state; watchdog recovers with the event logged; no silent resumption assuming continuity; every injected chain outage surfaced as "not monitoring."
Level | Subsystem
Sample/Config | DVT firmware + injector
Facility/Equipment/Tools | Fault-injection + watchdog harness
Independence | Embedded SW; System-Safety reviews
Standard/Compliance Reference | IEC 62304; ISO 14971
Evidence Required | Safe-state/self-check FI report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | RCR-023/027 verification reference.
PV-SW-010 — OTA rollback / version-skew FI (never bricks)
Objective | Prove a bad image auto-rolls-back with trip data preserved and no brick under version-skew; unmonitored update window disclosed.
Method | Test — FI
I/O Trace | SYS-032/114/139; PRD-SAFE-010; SRS-OTA (700); DO-FW-014/015, DO-CLD-010
Risk-Control Trace | RCR-037; HA-SW-010, FTA
Acceptance Criteria | Corrupt/bad image auto-rolls-back with trip data preserved; no brick under version-skew; unmonitored window disclosed. (Signing/rollback-defeat = PV-CYBER-005 IV.)
Level | Integration
Sample/Config | Device + OTA service
Facility/Equipment/Tools | OTA fault harness
Independence | Firmware + Cloud
Standard/Compliance Reference | IEC 62304
Evidence Required | OTA rollback FI report
Owner | Firmware + Cloud
Status | Planned
Notes | RCR-037 verification reference (functional/robustness half).
PV-SW-013 — Power-fail-during-write FI (no torn/lost/fabricated record)
Objective | Prove no torn/lost/fabricated record across repeated power loss during write (incl. cold sag) and that unsent/excursion data is never overwritten and tamper-evidence stays intact.
Method | Test — FI
I/O Trace | SYS-018/021/028/099/160/168; PRD-SAFE-010/003; DO-SAFE-008, DO-FW-010, DO-EE-002
Risk-Control Trace | RCR-024, RCR-022; HA-SW-013, HA-ELEC-008/010, FTA TOP-D
Acceptance Criteria | No torn/lost/fabricated record across repeated power loss during write (incl. cold sag); unsent/excursion data never overwritten; tamper-evidence intact. (Flash retention = PV-REL-007.)
Level | Subsystem
Sample/Config | DVT units + power-cycling rig (cold)
Facility/Equipment/Tools | Programmable power cut rig; chamber
Independence | Embedded SW; Quality reviews
Standard/Compliance Reference | IEC 62304; ALCOA+
Evidence Required | Power-fail-write FI report
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | RCR-024/022 verification reference; core of the record-integrity spine.
6.2 General software verification
PV-SW-014 — Static analysis & compiler-warning verification
Objective | Confirm static analysis and zero-warning builds across firmware/cloud/mobile, with safety classes held to the stricter ruleset.
Method | Analysis
I/O Trace | all SRS software modules; IEC 62304
Risk-Control Trace | RCR-019/021/024 (defect-free safety code)
Acceptance Criteria | Static analysis run on every build; no unresolved high-severity finding on safety code; builds warning-clean or warnings dispositioned.
Level | Unit
Sample/Config | Source baselines
Facility/Equipment/Tools | SAST / static analyzer; CI
Independence | Embedded SW + Quality
Standard/Compliance Reference | IEC 62304; coding standard
Evidence Required | Static-analysis reports
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | Complements code review PV-INSP-004/005.
PV-SW-015 — Unit-test coverage verification (class-driven)
Objective | Confirm unit tests meet the coverage target set by each module's IEC 62304 software safety class.
Method | Test, Analysis
I/O Trace | SRS software modules; IEC 62304
Risk-Control Trace | RCR-019/021/023 (safety code coverage)
Acceptance Criteria | Coverage meets the per-class target (e.g., statement/branch, MC-DC for highest class); safety modules at the highest target; gaps justified.
Level | Unit
Sample/Config | Source + test suites
Facility/Equipment/Tools | Coverage tool; CI
Independence | Embedded SW + Quality
Standard/Compliance Reference | IEC 62304
Evidence Required | Coverage reports
Owner | Embedded SW (EM-ENG-007)
Status | Planned
Notes | Coverage target per software safety classification (to confirm).
PV-SW-016 — Software integration verification (device / cloud / mobile interfaces)
Objective | Confirm integrated software components interact correctly across the device↔cloud↔mobile interfaces per the ICD.
Method | Test
I/O Trace | IFR (interfaces); SRS-CTX/CONN; DO-IF-
Risk-Control Trace | RCR-032 (cloud), RCR-033 (binding)
Acceptance Criteria | All interface contracts pass integration test; error/edge cases handled; ICD conformance demonstrated.
Level | Integration
Sample/Config | Integrated device+cloud+mobile
Facility/Equipment/Tools | Integration test harness
Independence | Embedded SW + Cloud + Mobile
Standard/Compliance Reference | IEC 62304
Evidence Required | Integration test report
Owner | Embedded SW + Cloud
Status | Planned
Notes | —
PV-SW-017 — Real-time timing / schedulability verification
Objective | Confirm by test that the acquire-settle-stamp cycle and safety tasks meet their deadlines, confirming PV-ANAL-014.
Method | Test
I/O Trace | SYS-037/038; DO-FW-003; SRS-PERF (1000)
Risk-Control Trace | RCR-006 (governor timing), RCR-025 (time)
Acceptance Criteria | Measured worst-case task timing meets deadlines with margin at −30 °C; no missed sample.
Level | Subsystem
Sample/Config | DVT firmware in chamber
Facility/Equipment/Tools | Timing instrumentation
Independence | Firmware
Standard/Compliance Reference | IEC 62304
Evidence Required | Timing verification report
Owner | Firmware (EM-ENG-007)
Status | Planned
Notes | Confirms PV-ANAL-014.
PV-SW-018 — Memory-leak / resource-exhaustion / long-run stability
Objective | Confirm no memory leak, resource exhaustion, or degradation over a long-run soak (device and cloud).
Method | Test
I/O Trace | SRS-REL (1100); PRD-REL
Risk-Control Trace | RCR-023 (stable operation), RCR-027
Acceptance Criteria | No memory growth/leak or resource exhaustion over the soak duration; heap/handles stable; no crash.
Level | Subsystem/System
Sample/Config | DVT device + cloud soak
Facility/Equipment/Tools | Soak harness; memory profiler
Independence | Embedded SW + Cloud
Standard/Compliance Reference | IEC 62304
Evidence Required | Soak/leak report
Owner | Embedded SW + Cloud
Status | Planned
Notes | Complements reliability soak (PV-REL).
7. Traceability Notes
The §6.1 fault-injection tests are the verification references for the information-integrity RCRs and are indexed in PV-SAFE §6.3; §6.2 provides the IEC 62304 software rigor. Functional-correctness counterparts are in PVL-500. Feeds Validation (detection/verdict effectiveness) and the summary.
8. Coverage/Completion Notes
Complete when all FI safety proofs pass, static analysis/coverage meet class targets, and integration/timing/ soak pass. Detection targets and thresholds Evidence-Pending (closed by test/validation). All Planned.
PVL-1600 (PV-TRACE) — Traceability & Coverage — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 with Chief Systems (EM-ENG-001) and EM-ENG-012 (lifecycle traceability). Verifies the coverage chains — requirement→test, risk→test, hazard→test, design-output→test, subsystem→test — and computes the gap analysis. The detailed matrix is the companion Verification-Traceability-Coverage-Matrix.md. All entries Planned.
1. Purpose
Prove every applicable requirement, risk control, hazard, and critical design output maps to ≥1 verification activity, and surface any gap (unverified requirement, control without a PV, safety control without fault injection, security control without IV).
2. Scope
Coverage/gap analysis across PRD/SYS/SRS/HWR/MER/FWR/IFR requirements, RCR-001…058, HAL HA-###, and DO- design outputs. Feeds ORCH-018 / EM-ENG-012 (Stage 18 end-to-end matrix).
3. Applicable Standards & References
ISO/IEC/IEEE 29148; ISO 13485 §7.3; ISO 14971 (risk-control coverage). Designations; revisions to confirm.
4. Governance & Method
Method is Inspection/Analysis (coverage computation). Coverage is bidirectional; every gap is a nonconformance dispositioned via PVL-1700.
5. Verification Categories
Requirement→test · risk-control→test · hazard→test · design-output→test · subsystem→test · FI/IV-flag coverage · gap analysis · coverage metrics.
6. Verification Register (this module)
PV-TRACE-001 — Requirement→test coverage & gap analysis
Objective | Confirm every applicable requirement (PRD/SYS/SRS/HWR/MER/FWR/IFR) maps to ≥1 PV activity and list any gap.
Method | Inspection, Analysis
I/O Trace | whole requirement baseline; PRD-VER-001/002
Risk-Control Trace | RCR-058 (defensible coverage)
Acceptance Criteria | 0 applicable requirements without a PV activity; any exception recorded with rationale; coverage % reported.
Level | System
Sample/Config | Coverage matrix
Facility/Equipment/Tools | Trace tool
Independence | Quality + Chief Systems
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Requirement-coverage report
Owner | EM-ENG-006 + EM-ENG-001
Status | Planned
Notes | See Verification-Traceability-Coverage-Matrix.md.
PV-TRACE-002 — Risk-control (RCR) & hazard→test coverage with FI/IV flags
Objective | Confirm every RCR (58) and every High/unacceptable HAL hazard maps to a PV activity, with fault injection preserved for safety-critical and IV for security controls.
Method | Inspection, Analysis
I/O Trace | RCR-1000 register; HAL Hazard-Traceability-Matrix
Risk-Control Trace | RCR-001…058; all High/unacceptable HA-###
Acceptance Criteria | 58/58 RCRs mapped; every FI-flagged RCR → a Test with fault injection; every IV-flagged RCR → an independent activity; 0 controls "on paper only"; the two Open items flagged.
Level | System
Sample/Config | RCR/HAL + PV register
Facility/Equipment/Tools | Trace tool
Independence | Quality + System-Safety
Standard/Compliance Reference | ISO 14971; IEC 62443 concepts
Evidence Required | RCR/hazard-coverage report
Owner | EM-ENG-006 + System-Safety
Status | Planned
Notes | Core Stage-12→13 handshake; matrix §2.
PV-TRACE-003 — Design-output→test & subsystem→test coverage
Objective | Confirm every critical design output (DO-) and every subsystem (S1–S9) is covered by verification.
Method | Inspection, Analysis
I/O Trace | Design-Output-Register; SYS allocation matrix
Risk-Control Trace | RCR implementation DO- references
Acceptance Criteria | Every Critical DO- and every subsystem maps to ≥1 PV activity; gaps recorded.
Level | System
Sample/Config | DOL register + PV register
Facility/Equipment/Tools | Trace tool
Independence | Quality + Chief Systems
Standard/Compliance Reference | ISO 13485 §7.3.6
Evidence Required | DO-/subsystem coverage report
Owner | EM-ENG-001
Status | Planned
Notes | —
PV-TRACE-004 — Coverage metrics & Stage-18 handoff
Objective | Compute coverage metrics and hand the Stage-13 segment to the ORCH-018 / EM-ENG-012 end-to-end traceability spine.
Method | Analysis
I/O Trace | PRD-VER-002; ORCH-018 traceability
Risk-Control Trace | RCR-058
Acceptance Criteria | Coverage metrics (requirement %, RCR %, FI/IV %) computed and reported; the Stage-13 verification segment integrated into the lifecycle matrix.
Level | System
Sample/Config | Coverage matrix
Facility/Equipment/Tools | Trace tool
Independence | EM-ENG-012
Standard/Compliance Reference | ISO/IEC/IEEE 29148
Evidence Required | Coverage-metrics report
Owner | EM-ENG-012 + EM-ENG-006
Status | Planned
Notes | Feeds Stage 18.
7. Traceability Notes
This module is the coverage authority; the detailed row-by-row matrix is in Verification-Traceability-Coverage-Matrix.md. Feeds ORCH-018 / EM-ENG-012 (Stage 18) and the release summary.
8. Coverage/Completion Notes
Complete when all coverage chains show 0 gaps (or documented exceptions) and the metrics are reported. All Planned.
Verification Plan — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: Principal Verification/Reliability/Test/Quality Engineer (EM-ENG-006) with Chief Systems (EM-ENG-001) and the discipline principals. The strategy, phasing, sequencing, resources, and gates for executing the Product Verification Library (PVL, 20 modules, 195 activities). Companion to the PVL modules, the Verification-Traceability-Coverage-Matrix, and index.md. Based on the verification standard, D1–D4, and the Stage-05/06/07 baselines. Pre-execution: all activities Planned/Evidence-Pending.
---
1. Purpose & scope
Verification answers "did we build the SCM correctly?" — that each design output satisfies its design input(s). This plan governs how the 195 verification activities are executed across Unit → Subsystem → Integration → System → Production levels, using Test / Analysis / Inspection / Demonstration, with objective acceptance criteria and controlled evidence (ISO 13485 §7.3.6). Validation ("did we build the right SCM?") is Stage 14 (PVaL / ORCH-015) and is out of scope here except at the handoff.
2. Strategy (severity-first)
1. Risk-first ordering. The two carried-Open program items are scheduled earliest with the longest lead because they gate architecture and release: D3 −30 °C / 90-day cold endurance + cold-charge safety (PV-REL-001, PV-ENRG-001, PV-REL-002/003/004) and RF-in-container (PV-REL-005, PV-COMM-003). The 90-day endurance test is the schedule critical path (it takes ≥90 days plus setup).
2. Fault injection for safety. Every safety-critical control is proven by forcing the fault (PV-ENRG-, PV-TFC-, PV-SW- FI set, PV-SAFE-003/004). Robustness is not inferred from functional-correctness tests.
3. Independent adversarial verification for security. All security controls are pen-tested by a party independent of development (PV-CYBER-), scoped from the threat model (PV-ANAL-012).
4. Analysis predicts, test confirms. Safety-critical analyses (power budget, battery, RF, single-fault) are predicates; the load-bearing claim is confirmed by test (PRD-VER-003/004/005).
5. Objective acceptance. No subjective safety verdict; Evidence-Pending numeric targets carry their named closing test and are never invented (hallucination-policy).
6. Reconcile safety with HAL. Every PV-SAFE activity is a HAL hazard-record Verification Reference; no residual risk closes without it (PRD-VER-007; HAL-2600/2800/2900).
3. Verification levels
• Unit — code units, single components (PV-SW-014/015; PV-INSP-004/005; PV-HW component tests).
• Subsystem — battery/BMS, sensing chain, enclosure, radios (PV-ENRG-, PV-HW-, PV-MECH-, PV-PERF-).
• Integration — device↔cloud↔mobile interfaces (PV-COMM-, PV-CLOUD-, PV-SW-016).
• System — full device + cloud + mobile in representative use (PV-ENV-, PV-SAFE-003/004, PV-REL-).
• Production — EOL/line processes (PV-MFG-, PV-INSP production items).
4. Phases & sequencing
Phase 0 — Verification readiness (entry gate, PV-GOV-006)
Requirements frozen & quality-reviewed (PV-REQ-), coverage matrix complete (PV-PLAN-002), design outputs released (PV-INSP-009), UUT baseline controlled (PV-GOV-003), equipment calibrated (PV-EVID-003), protocols approved (PV-EVID-001), analyses complete (PV-ANAL-). No verification test starts before its readiness items.
Phase 1 — Analysis & inspection (early, low-sample)
PV-ANAL- (tolerance, thermal, power budget, battery, RF link, worst-case, reliability prediction, threat model, single-fault), PV-INSP- (drawing/PCB/code/label/supplier). These predict and de-risk the physical tests and set the Evidence-Pending values to be confirmed.
Phase 2 — Design Verification Test (DVT) — subsystem then system
Long-lead first | start PV-REL-001 (90-day cold endurance) and PV-REL-005 (RF-in-container) as early as the final cell/antenna are available — they gate release and dominate schedule.
Subsystem | PV-HW-, PV-ENRG-001/002 (charge/protection FI), PV-PERF- (accuracy/latency), PV-SW- (FI safety proofs + unit/coverage/timing), PV-MECH- and PV-ENV- (non-destructive first, destructive last).
Integration | PV-COMM-, PV-CLOUD-, PV-SW-016, PV-FUNC- end-to-end.
System | PV-SAFE-003/004 (single-fault + record-integrity-spine FI), PV-ENV-001…009, EMC (PV-EMC-), altitude, and the safety qualification (PV-TFC-002/003).
Independent/accredited | PV-CYBER- (independent pen-test), PV-EMC-/radio and PV-COMP-004 (UN 38.3), PV-COMP-007/008 at accredited houses.
Destructive-test ordering | UN 38.3, fire-enclosure containment (PV-TFC-002/003), and abuse tests run on dedicated samples after non-destructive characterisation.
Phase 3 — Process/Production Verification Test (PVT)
PV-MFG- (EOL calibration, incoming/AOI, poka-yoke, identity, config, latent-defect coverage, FAI/Cpk), PV-INSP production items, PV-REL-007 HASS screen definition, PV-FUNC-015 production regression.
Phase 4 — Verification summary & release recommendation (exit gate, PV-GOV-006 / PV-SUM-)
Coverage/pass-rate roll-up (PV-SUM-001), open-issue/residual status (PV-SUM-002), verification conclusion + go/no-go (PV-SUM-003), lessons-learned + handoff (PV-SUM-004). The two Open items must close by test (or the release is conditional/blocked); ISO 14971 residual/benefit-risk signed (RCR-058 / PV-COMP-012).
5. Sequencing dependencies (critical)
• Final cell datasheet (concept §7) gates PV-ENRG-001 charge limit, PV-ANAL-003/004, PV-REL-001/002/003.
• Antenna decision gates PV-REL-005 / PV-COMM-003 (integral vs external-antenna fallback).
• DO-VER-009 (cold-chamber + representative-shipper config) enables PV-REL-001 and PV-REL-005.
• WHO PQS category (Evidence-Pending) anchors detection/accuracy targets (PV-PERF-001, PV-SW-001, O-7).
• Test-house / pen-test vendor selection (Evidence-Pending) gates PV-EMC-, PV-COMP-004/006/007/008, PV-CYBER-.
• Non-destructive before destructive; subsystem before system; analysis before the test it predicts.
6. Resources, facilities & independence
Internal | climatic/humidity chambers, vibration/shock/drop, IP/leak rigs, bench electrical, battery cyclers, HALT/HASS, long-run harnesses, network emulators, cloud test tenants, device farm.
Accredited external | EMC/radio (PV-EMC-, PV-COMP-006/008), UN 38.3 (PV-COMP-004), product safety (PV-COMP-007), independent pen-test (PV-CYBER-), offgas lab (PV-TFC-003).
References | ISO/IEC 17025-traceable temperature/RH references, RF equipment (DO-VER-008; PV-EVID-003).
Independence (PV-GOV-002) | verifier ≠ author for safety-critical; security controls verified by an external party. Owners per module (Battery/Power, Electrical/EMC, Mechanical, Embedded SW, Cloud/Mobile, Product-Security, Metrology, Reliability, System-Safety, Quality/Regulatory).
7. Verification risks & mitigations
Risk | Impact | Mitigation
90-day endurance fails / marginal (D3, O-2) | Release blocker; architecture rework | Start earliest; analysis predicate PV-ANAL-003/004; reserve design RCR-006; parallel cold-capacity/fade tests
RF-in-container link cannot close (O-1) | Real-time-alert claim withheld on lanes | External-antenna fallback (RCR-031); consequence bounded by store-and-forward (RCR-030); early PV-REL-005
Final cell datasheet late | Blocks charge/endurance tests | Interim analysis on candidate chemistries; lock chemistry before Phase 2
Test-house / pen-test lead time | Schedule slip | Book accredited slots in Phase 0; scope pen-test from PV-ANAL-012 early
Cold-chamber / long-run capacity | Endurance critical path | Reserve chamber time; run endurance on a dedicated sample set
Detection/alarm targets unset (O-7) | Acceptance ambiguity | Confirm WHO PQS category early; anchor PV-SW-001/PV-USE-002 to it
8. Milestones & deliverables
• M1 Readiness gate passed (Phase 0) — plan, coverage matrix, released inputs, calibrated equipment.
• M2 Analysis/inspection complete (Phase 1) — Evidence-Pending values scoped; predicates done.
• M3 DVT complete (Phase 2) — incl. D3 endurance + RF-in-container closed or Open-status known; all FI/IV safety/security proofs executed.
• M4 PVT complete (Phase 3) — EOL/line processes verified.
• M5 Verification summary + release recommendation (Phase 4) — RCR-058 gate; handoff to Stage 14/16/18.
Deliverables | approved protocols (PV-EVID-001), test reports + raw data (PV-EVID-002), calibration certificates (PV-EVID-003), coverage matrix (PV-TRACE-), verification summary (PV-SUM-003).
9. Traceability & feeds
This plan executes the PVL and the coverage matrix; it feeds Validation (Stage 14, PVaL) (the VAL- targets), Release (Stage 16) (coverage/open-item evidence), and ORCH-018 / EM-ENG-012 (Stage 18) (the verification segment of the end-to-end matrix). Safety verification reconciles with HAL (Stage 10).
10. Review history
• 2026-07-19 v1.0 — Initial SCM Verification Plan (Stage 13, ORCH-014). 20 PVL modules, 195 activities; 58/58 RCR coverage; D3 and RF-in-container carried Open. All Planned/Evidence-Pending (pre-execution).
Verification Traceability & Coverage Matrix — Smart Cold Chain Monitor (SCM)
Stage 13 (ORCH-014) · Owner: EM-ENG-006 + EM-ENG-001 with EM-ENG-012. Maps Requirement / RCR → PV activity → method → acceptance basis → status, and identifies unverified requirements, controls without a PV, and safety controls lacking fault injection / security controls lacking independent verification. This is the Stage-13 segment of the end-to-end traceability spine handed to ORCH-018 / EM-ENG-012 (Stage 18). All PV status is Planned / Evidence-Pending (pre-execution — no results fabricated).
---
1. Coverage summary
Domain | Population | Covered by ≥1 PV | Gap
Product Requirements (PRD, Stage 05) | 220 | All applicable modules mapped to PV modules (see §3) | 0 (spot-exceptions noted §5)
System Requirements (SYS-001…190, Stage 06) | 190 | Mapped by category to PV modules (§3) | 0 applicable
Subsystem/SW (HWR/MER/FWR/IFR + SRS, Stage 07) | 488 | Mapped by discipline to PV modules (§3) | 0 applicable
Risk-Control Requirements (RCR-001…058, Stage 12) | 58 | 58/58 (§2) | 0
PRD-VER headline (PRD-VER-004/005) | 2 | PV-REL-001 / PV-REL-005 | 0
Verification activities authored (this PVL) | 195 | across 20 modules | —
FI/IV posture: every Stage-12 FI-flagged safety-critical control maps to a Test with fault injection; every IV-flagged security control maps to an independent pen-test/assessment (§2, FI/IV column).
---
2. RCR → PV coverage (all 58 controls — Stage-12 handshake)
Legend: FI = fault injection (safety-critical); IV = independent adversarial verification (security); T/A/I/D = Test/Analysis/Inspection/Demonstration. Status = Planned unless the RCR is carried Open (D3/RF). Acceptance basis is the PV activity's objective criterion (see the module entry).
RCR | Control (short) | PV activity(ies) | Method | FI/IV | Status
RCR-001 | Charge-inhibit <0 °C (HW interlock) | PV-ENRG-001; PV-HW-003; PV-COMP-004 (UN 38.3) | T | FI | Open (D3)
RCR-002 | Autonomous protection + secondary backstop | PV-ENRG-002; PV-HW-004 | T | FI | Planned
RCR-003 | Out-of-spec charge-source rejection | PV-ELEC-002; PV-EMC-006 | T | — | Evidence-Pending
RCR-004 | Cell retention (drop/vibe/transport) | PV-MECH-004; PV-ENV-009; PV-COMP-004 | T | FI-adj | Planned
RCR-005 | Self-heat bounded / cannot bypass inhibit | PV-TFC-001 | T | FI | Evidence-Pending
RCR-006 | Energy governor + reserve (90-day) | PV-REL-001 (=PRD-VER-004); PV-HW-007; PV-ANAL-003 | T/A | — | Open (D3)
RCR-007 | UN 38.3 / IATA-DGR transport | PV-COMP-004; PV-COMP-005; PV-INSP-007 | T/I | — | Evidence-Pending
RCR-008 | Fire-enclosure containment / offgas | PV-TFC-002; PV-TFC-003 | T | — | Planned/EP
RCR-009 | No accessible sharp/pinch/crush | PV-SAFE-002 | I | — | Planned
RCR-010 | Small-parts retention | PV-MECH-003 | T | — | Evidence-Pending
RCR-011 | Enclosure/seal IP + containment through life | PV-ENV-005; PV-MECH-001/002; PV-ANAL-001 | T/I | — | Evidence-Pending
RCR-012 | Fasteners resist vibration/thermal loosening | PV-MECH-001; PV-ENV-003 | T | — | Evidence-Pending
RCR-013 | SELV; no accessible hazardous voltage | PV-SAFE-001; PV-HW-001 | T | — | Planned
RCR-014 | EMC immunity (no false/missed) + emissions | PV-ELEC-004; PV-EMC-001; PV-COMP-008 | T | FI-adj | Planned
RCR-015 | Conformal coat + seal (record electronics) | PV-TFC-006; PV-ENV-003 | T | — | Evidence-Pending
RCR-016 | AOI + EOL latent-defect screening | PV-MFG-008; PV-REL-006/007 (HALT/HASS) | A/I | — | Evidence-Pending
RCR-017 | Operate + retain record −30…+60 (endurance) | PV-ENV-001; PV-REL-001 (=PRD-VER-004) | T | — | Open (D3)
RCR-018 | Tolerate condensation without record loss | PV-ENV-003; PV-TFC-006 | T | — | Evidence-Pending
RCR-019 | Two-sided detection, no undetected path | PV-SW-001; PV-FUNC-001; PV-ANAL-013 | T | FI | Evidence-Pending
RCR-020 | Probe-fault / implausible-read | PV-SW-008; PV-FUNC-003 | T | FI | Planned
RCR-021 | Verdict integrity (suppress false accept) | PV-SW-002; PV-CLOUD-003; PV-FUNC-002 | T | FI | Planned
RCR-022 | Honest gap / no clean pass on incomplete | PV-SW-002 / PV-SW-013; PV-DATA-004 | T | FI | Planned
RCR-023 | Safe state on fault + watchdog recovery | PV-SW-009; PV-FUNC-014 | T | FI | Planned
RCR-024 | Power-fail-safe storage (no torn/lost) | PV-SW-013; PV-REL-007 (flash); PV-PERF-008 | T | FI | Planned
RCR-025 | Trusted monotonic time / drift bound | PV-SW-005; PV-ANAL-014 | T/A | — | Evidence-Pending
RCR-026 | Calibration integrity / expiry blocks arming | PV-SW-004; PV-DATA-001; PV-MFG-001 | T | FI | Planned
RCR-027 | Availability self-check ("not monitoring") | PV-SW-009 | T | FI | Planned
RCR-028 | Excursion committed locally (offline-first) | PV-COMM-001 | T | — | Planned
RCR-029 | Alert-ahead, ack, escalate | PV-SW-006; PV-COMM-001; PV-USE-002 | T | FI | Planned
RCR-030 | Degrade to store-and-forward, no record loss | PV-COMM-001; PV-COMM-004; PV-PERF-009 | T | — | Planned
RCR-031 | RF closes in-container budget / ext-antenna | PV-REL-005 (=PRD-VER-005); PV-COMM-003; PV-ANAL-006 | T/A | — | Open (RF)
RCR-032 | Cloud never drops a safety event | PV-COMM-002 / PV-CLOUD-001; PV-PERF-012 | T | FI (cloud) | Planned
RCR-033 | Device↔trip↔owner↔responder binding | PV-COMM-008; PV-MFG-005 | T | — | Planned
RCR-034 | Capture-time authenticity/integrity | PV-CYBER-001; PV-SEC-001 | T | IV | Evidence-Pending
RCR-035 | Alert path not silently suppressible | PV-CYBER-002 | T | IV | Evidence-Pending
RCR-036 | Config/threshold changes auth'd/bounded/audited | PV-CYBER-003; PV-SEC-003 | T | IV | Evidence-Pending
RCR-037 | Signed, rollback-safe A/B OTA | PV-SW-010 + PV-CYBER-005; PV-FUNC-010 | T | FI+IV | Planned/EP
RCR-038 | Secure boot / debug lock / tamper-evident | PV-CYBER-006; PV-INSP-010 | T/I | IV | Evidence-Pending
RCR-039 | Cloud resists compromise (RBAC/SSO/DR) | PV-CYBER-009 | T | IV | Evidence-Pending
RCR-040 | Disposition surface shows qualifier | PV-DATA-005; PV-DATA-004 | I/D | — | Planned
RCR-041 | Estimated data flagged, not measured | PV-DATA-007 | I/T | — | Planned
RCR-042 | Export/API preserves attribution/time/gaps | PV-DATA-006 | T | — | Planned
RCR-043 | Legible language-independent verdict (cold) | PV-USE-001; PV-FUNC-002 | T | — | Evidence-Pending
RCR-044 | Verdict-override deliberate/recorded | PV-DATA-008 | T/D | — | Planned
RCR-045 | Pre-arm readiness gate | PV-SW-003; PV-FUNC-004 | T | FI | Planned
RCR-046 | Charge-inhibited/fault states legible | PV-USE-003 | T/D | — | Planned
RCR-047 | Alerting tuned to validated targets | PV-USE-002 (+ VAL-USE-002, Stage 14) | T | — | Planned
RCR-048 | Provisioning validates trip profile | PV-SW-003; PV-FUNC-011 | T | FI | Planned
RCR-049 | Recalibration interval enforced (pre-arm) | PV-SVC-001; PV-SW-004 | T | — | Planned
RCR-050 | Service restores seal/config/fw + BIST | PV-SVC-002; PV-FUNC-013 | T/D | — | Planned
RCR-051 | Debug/dev modes locked in field | PV-CYBER-006; PV-INSP-010; PV-FUNC-012 | T/I | IV | Evidence-Pending
RCR-052 | Safe EOL (battery/data/key erasure) | PV-SEC-005 (key erasure) + disposal-review; PV-INSP-006 | T/I | — | Planned
RCR-053 | EOL calibration coverage + birth record | PV-MFG-001 | T | — | Evidence-Pending
RCR-054 | Correct chemistry + protection parts (poka-yoke/AOI) | PV-MFG-002 / PV-MFG-003; PV-INSP-008 | T/I | — | Planned
RCR-055 | Unique cryptographic identity + serial | PV-MFG-005; PV-SEC-005 | T | — | Planned
RCR-056 | Correct regional/firmware config at EOL | PV-MFG-006; PV-COMM-006 | T | — | Planned
RCR-057 | Record-integrity chain (aggregate) | PV-SAFE-004 (+ PV-SW-001/002/008/013, PV-DATA-004/006) | T | FI | Planned
RCR-058 | Overall residual + benefit-risk gate | PV-SAFE-005 / PV-COMP-012; PV-SUM-003 | I/A | — | Planned
Result: 58/58 RCRs have ≥1 matching PV activity. No control is on paper only. Every FI-flagged safety-critical control maps to a fault-injection Test; every IV-flagged security control maps to an independent pen-test/assessment. The two carried-Open items are RCR-001/005/006/017 (D3 battery, via PV-ENRG-001 + PV-REL-001) and RCR-031 (RF-in-container, via PV-REL-005).
---
3. Requirement-category → PV-module map (coverage chains)
Requirement source | Primary PV modules
PRD-FUN / SYS functional (SYS-001…032) / SRS-DEV/CONN/CLOUD/MOB | PV-FUNC (500), PV-SW (1100)
PRD-PERF / SYS performance (SYS-033…050) / SRS-PERF | PV-PERF (600), PV-ANAL (400), PV-REL (1400)
PRD-ENV / SYS environmental (SYS-069…085) / MER | PV-ENV (700), PV-REL (1400), PV-ANAL (400)
PRD-ELEC / SYS power (SYS-086…103) / HWR | PV-HW (1200), PV-SAFE (900), PV-ANAL (400)
PRD-CONN / SYS connectivity (SYS-104…118) / SRS-CONN | PV-FUNC (500), PV-REL (1400 RF), PV-EMC (800)
PRD-CLOUD / MOB / SYS data (SYS-119…131) / SRS-CLOUD/MOB/DATA | PV-FUNC (500), PV-SW (1100), PV-SEC (1000)
PRD-SEC / SYS security (SYS-132…148) / SRS-SEC | PV-SEC (1000), PV-SW (1100), PV-ANAL (400)
PRD-SAFE / SYS safety (SYS-149…162) / SRS-SAFE | PV-SAFE (900), PV-SW (1100), PV-SEC (1000)
PRD-REL / SYS reliability (SYS-163…176) / SRS-REL | PV-REL (1400), PV-ANAL (400), PV-SW (1100)
PRD-REG / SYS regulatory (SYS-177…190) / SRS-REG | PV-COMP (1500), PV-EMC (800), PV-SAFE (900)
PRD-MECH / MER | PV-ENV (700), PV-INSP (300), PV-SAFE (900)
PRD-MFG / PRD-SVC / PRD-QUAL | PV-MFG (1300), PV-FUNC (500 service), PV-INSP (300)
PRD-VER (verification strategy) | all modules; PV-GOV/PLAN/REQ/TRACE
Design outputs (DO-, incl DO-VER) | PV-INSP (300), PV-TRACE (1600); DO-VER-009 enables PV-REL-001/005
---
4. PRD-VER headline requirements → PV
PRD-VER | Requirement | PV activity | Status
PRD-VER-001 | Every requirement verifiable w/ method | PV-REQ-001/004; PV-GOV-001 | Planned
PRD-VER-002 | Bidirectional traceability | PV-REQ-006; PV-TRACE-001…004 | Planned
PRD-VER-003 | Method allocation appropriate to risk | PV-GOV-001; PV-REQ-004 | Planned
PRD-VER-004 | D3 endurance/charge validated by test | PV-REL-001 (90-day) + PV-ENRG-001 (charge FI) + PV-REL-002/003/004 | Open (D3)
PRD-VER-005 | RF-in-container qualification | PV-REL-005 + PV-COMM-003 | Open (RF)
PRD-VER-006 | Env/regulatory qualification suite | PV-ENV-, PV-EMC-, PV-SAFE-, PV-COMP-004/007/008/009 | Planned/EP
PRD-VER-007 | Safety verification reconciled with HAL | PV-SAFE-001…005 (HAL verification references) | Planned
PRD-VER-008 | Validation of intended use/needs | Stage 14 (PVaL) — VAL-USE-001/002/004/007; VAL-SW-001/002; VAL-COMM- | Deferred to Stage 14
---
5. Gap analysis & findings
1. No RCR gap. All 58 RCRs map to ≥1 PV activity with correct FI/IV flags (§2). This satisfies the critical Stage-12→13 requirement.
2. Fault-injection completeness. Every safety-critical control the RCR register flagged FI has a Test with fault injection: PV-ENRG-001/002, PV-TFC-001, PV-SW-001/002/003/004/006/008/009/010/013, PV-MECH-004 (FI-adjacent), PV-ELEC-004 (FI-adjacent, interference-as-fault), PV-SAFE-003/004 (system-level FI).
3. Independent verification completeness. Every security control flagged IV has an independent pen-test/ assessment: PV-CYBER-001/002/003/005/006/009. Independence is governed by PV-GOV-002.
4. Two carried-Open items (release-gating), not gaps in coverage. The PV activities exist and are defined; their effectiveness values are Evidence-Pending and close by test: D3 (PV-REL-001 + PV-ENRG-001) and RF-in-container (PV-REL-005). These are correctly carried Open per Stage-12 §7.
5. Evidence-Pending numeric targets (not gaps). Several acceptance limits are carried Judgment/Evidence- Pending and close by their named test: ±0.5 °C (PV-PERF-001, anchored to WHO PQS O-7), 5-min latency (PV-PERF-004), IP67 (PV-ENV-005), storage sizing (PV-PERF-007), detection targets (PV-SW-001, O-7), alarm targets (PV-USE-002 + VAL Stage 14), calibration coverage (PV-MFG-001), drift bound (PV-SW-005), DR RTO/RPO (PV-CYBER-009), MTBF (PV-REL-009). None is invented.
6. Requirement spot-exceptions to confirm at execution. A minority of low-criticality requirements verified by Analysis/Inspection alone (e.g., mass SYS-084, fit SYS-083, marking durability) are acceptable per the severity-first method-allocation rule (PV-VER-003 / PV-REQ-004); each is confirmed in PV-REQ-003.
7. Validation handoff. Use-effectiveness (RCR-043/044/047 comprehension, detection false-alarm/false-reject targets) is verified functionally here (PV-USE-, PV-SW-001) and validated in Stage 14 (VAL-USE-, VAL-SW-) — recorded in the VAL- column of the RCR register.
---
6. Handoff
This matrix is the Stage-13 verification segment of the end-to-end traceability spine. It feeds ORCH-018 / EM-ENG-012 (Stage 18) for the complete bidirectional matrix, Validation (Stage 14, PVaL) for the VAL- targets, and Release (Stage 16) for the coverage/open-item evidence. All PV status is Planned/ Evidence-Pending until executed.

## 14 Validation
Validation
The plan to prove the product works for real people, in real conditions, for its real purpose.
↩ Contents
In this section:
• Product Validation Library (PVaL) — Smart Cold Chain Monitor (SCM)
• VAL-BIZ-800 (PVaL-800) — Business Validation
• VAL-CLIN-500 (PVaL-500) — Clinical / Operational Validation
• VAL-ECO-600 (PVaL-600) — Ecosystem Validation
• VAL-EVID-1100 (PVaL-1100) — Validation Evidence
• VAL-FIELD-900 (PVaL-900) — Field Validation
• VAL-GOV-000 (PVaL-000) — Validation Governance
• VAL-HF-400 (PVaL-400) — Human-Factors Validation
• VAL-PLAN-100 (PVaL-100) — Validation Planning
• VAL-REL-700 (PVaL-700) — Reliability Validation
• VAL-RISK-1000 (PVaL-1000) — Risk Validation
• VAL-SUM-1200 (PVaL-1200) — Validation Summary
• VAL-USE-300 (PVaL-300) — Intended-Use Validation
• VAL-USER-200 (PVaL-200) — User Validation
• Validation Coverage Matrix — Smart Cold Chain Monitor (SCM)
• Validation Plan — Smart Cold Chain Monitor (SCM)
Product Validation Library (PVaL) — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 14 — Validation (ORCH-015) · Owner/Lead: Principal Human-Factors / Design-Validation / Clinical-or-Operational + Product engineer · Contributors: Clinical/Operational; Product/Business; Quality/Regulatory (EM-ENG-006) · Integrator: Chief Systems (EM-ENG-001) · Traceability: ORCH-018 / EM-ENG-012 · Status: Authored playbook — every activity Planned / Evidence Pending (pre-execution; no product built, no users tested; no results fabricated).
This library is the SCM's validation playbook: it answers did we build the correct product? — does the SCM meet user needs and intended use in the real world, with representative users in representative contexts. It is entirely distinct from Verification (Stage 13, PVL), which asks did we build the product correctly?. Validation is outcome-focused and traces to User Needs / Intended Use and to residual-risk acceptability, never to design outputs. It follows 00-governance/validation-standard.md and the 13-module PVaL map, and is based on 00-concept/concept-and-decisions.md (D1–D4), the Stage-04 User Needs (esp. UN-VAL-001…009), the Stage-10 Hazard Analysis (HAL-1500 use-related, HAL-2800 individual residuals, HAL-2900 overall benefit-risk), the Stage-12 Risk-Control Requirements (VAL- targets), and the Stage-13 Verification Library.
---
The dominant validation outcome (the SCM's reason to exist)
Representative users — including the minimally trained, gloved, offline last-mile health worker — can
correctly detect, understand, and act on temperature excursions so that no temperature-compromised vaccine
is administered and no potent vaccine is needlessly discarded, in representative environments, with the
device's residual risks acceptable to those users.
Everything in this library serves that outcome. The two most load-bearing studies are VAL-USE-001 (read the verdict → correct disposition) and VAL-USE-004 (never trust a void/unmonitored trip) — the two convergence choke points of the harm pathway (HAL-2900 §5.2) — consolidated in the summative human-factors validation VAL-HF-001, and their residual acceptability validated in VAL-RISK-001/003.
---
Product in one line
A reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shippers to log temperature/humidity/shock, detect two-sided excursions, alert in time (connectivity permitting), give an offline accept/reject verdict at the door, and provide an auditable, tamper-evident cold-chain record for product-disposition decisions across a multi-market (WHO PQS / EU GDP / US FDA / air-transport), multi-trip fleet. It is an information/monitoring device supporting human decisions — not therapeutic, not temperature-controlling.
---
Module table (13 modules · 128 validation activities)
Module | Code | File | Activities | Focus
PVaL-000 | VAL-GOV | [VAL-GOV-000-governance.md](VAL-GOV-000-governance.md) | 6 | Philosophy, intended use/users, outcome-vs-implementation, acceptance authority, honesty, handoffs
PVaL-100 | VAL-PLAN | [VAL-PLAN-100-planning.md](VAL-PLAN-100-planning.md) | 8 | Strategy, participants/representative users, environments, facilities, phasing, criteria, deliverables, dependencies
PVaL-200 | VAL-USER | [VAL-USER-200-user-validation.md](VAL-USER-200-user-validation.md) | 15 | Every user class (last-mile → EPI → operators → sustainment → governance → enterprise) + accessibility
PVaL-300 | VAL-USE | [VAL-USE-300-intended-use-validation.md](VAL-USE-300-intended-use-validation.md) | 18 | Primary workflows + VAL-USE-001…009 = HA-USE-001…009 critical tasks + misuse/emergency/recovery/maintenance/installation
PVaL-400 | VAL-HF | [VAL-HF-400-human-factors-validation.md](VAL-HF-400-human-factors-validation.md) | 11 | Summative HF of the excursion-response critical task, legibility, comprehension, alarms, error avoidance, accessibility, training
PVaL-500 | VAL-CLIN | [VAL-CLIN-500-clinical-operational-validation.md](VAL-CLIN-500-clinical-operational-validation.md) | 10 | Operational effectiveness (both directions), decision support, wastage reduction, audit defensibility, benefit target
PVaL-600 | VAL-ECO | [VAL-ECO-600-ecosystem-validation.md](VAL-ECO-600-ecosystem-validation.md) | 10 | Cloud dashboard, mobile, fleet scale, offline floor, store-and-forward, sync/reconciliation, cross-tier consistency, remote mgmt
PVaL-700 | VAL-REL | [VAL-REL-700-reliability-validation.md](VAL-REL-700-reliability-validation.md) | 8 | Long-term in-use; D3 cold endurance + RF-in-container (Open); environmental/network/power/install/region diversity; calibration
PVaL-800 | VAL-BIZ | [VAL-BIZ-800-business-validation.md](VAL-BIZ-800-business-validation.md) | 9 | Satisfaction, NPS, retention, support burden, return/RMA, reuse TCO/cost-per-trip, tender/enterprise fit
PVaL-900 | VAL-FIELD | [VAL-FIELD-900-field-validation.md](VAL-FIELD-900-field-validation.md) | 10 | Pilot → beta → production trial on LMIC/GDP lanes; installer/service/support feedback; both-direction safety signals
PVaL-1000 | VAL-RISK | [VAL-RISK-1000-risk-validation.md](VAL-RISK-1000-risk-validation.md) | 12 | Residual-risk acceptability (HA-RES), overall benefit-risk (HAL-2900), foreseeable misuse, hazard communication, security IV
PVaL-1100 | VAL-EVID | [VAL-EVID-1100-validation-evidence.md](VAL-EVID-1100-validation-evidence.md) | 6 | Pre-registered protocols, observation/video, datasets, usability engineering file, field reports, signed acceptance
PVaL-1200 | VAL-SUM | [VAL-SUM-1200-validation-summary.md](VAL-SUM-1200-validation-summary.md) | 5 | Coverage, user/clinical/business acceptance, residual concerns, go-to-market recommendation
Total |  | 13 modules | 128
Companion documents: [Validation-Plan.md](Validation-Plan.md) (strategy, participants, environments, phasing/sequencing, success criteria, deliverables, dependencies/gates) · [Validation-Coverage-Matrix.md](Validation-Coverage-Matrix.md) (User Need / Intended-Use → VAL activity → method → success criterion → status; gaps: user needs without validation, safety-critical workflows not validated, use errors not validated).
---
ID scheme & the HA-USE ↔ VAL-USE correspondence
Validation activities use VAL-<CODE>-NNN. To honour the Stage-10/12 references, the twelve use-related critical tasks preserve the exact IDs the hazard records cite: VAL-USE-001…009 map one-to-one to HA-USE-001…009 (HAL-1500), and are the "Validation Reference" that closes each use residual. The upstream also cites VAL-COMM-001/002/008 (hosted here as VAL-ECO-004/005/006), VAL-SW-001 (hosted as the operational detection-effectiveness activity VAL-CLIN-002), and VAL-USE-005/008 — all resolved to real activities so the traceability spine closes. Each activity carries the 14-field register entry (Objective [user outcome] · Intended Use/User · Method · Participants/Representative Users · Environment/Context · Success/Acceptance Criteria · User Need Traceability · Risk/Use-Error Relationship · Evidence Required · Owner · Status · Notes).
---
Applicability Register
All 13 modules are Applicable — the SCM is a globally distributed, regulated, reusable, connected, safety- significant IoT device with device + cloud + mobile tiers, operated at its most safety-critical point by its least-trained user, so every module carries real, outcome-based, product-specific validation.
Module | Status | Note
All 13 (VAL-GOV … VAL-SUM) | Applicable | Every module carries real, product-specific, outcome-based activities.
VAL-CLIN (500) | Applicable — operational | Not a clinical trial (non-therapeutic device); validates operational effectiveness on vaccine-disposition outcomes. Benefit claim provisional until WHO PQS accuracy target confirmed (D4).
VAL-REL (700) | Applicable — 2 Open items | D3 90-day cold endurance (PRD-VER-004) and RF-in-container (PRD-VER-005) are release-gating and gate reliability/field validation.
VAL-RISK (1000) | Applicable — gates release | Validates the HAL-2800/2900 residual conclusions; the seven HAL-2900 benefit-risk conditions must close before an unconditional release.
Borderline within modules (confirm, not excluded): WHO PQS category/accuracy target (VAL-CLIN-010, D4); US FDA classification (regulatory stage); all numeric outcome targets (use-error rate, false-reject rate, NPS, return rate, TCO) are Evidence-Pending with named closing protocols.
---
Coverage summary (see Validation-Coverage-Matrix.md)
• Every use error (HA-USE-001…012) has a validation reference (VAL-USE-001…009; VAL-USE-007 also closes HA-USE-012; HA-USE-010/011 are process-controlled with reconditioning/expiry-gate verification and covered by VAL-USE-017 / VAL-REL-008).
• Every top residual (HA-RES-001…022) has a VAL-RISK (or VAL-USE/VAL-HF/VAL-REL/VAL-ECO) validation reference; the overall benefit-risk (HAL-2900) is validated by VAL-RISK-003 against its seven conditions.
• Every safety-critical workflow (disposition read, alert response, provisioning/arming, override, probe placement, charge-inhibit interpretation) has a summative activity.
• Every user class (SH-01…SH-15) has a validation activity; the beneficiary (SH-11) is represented.
• Named coverage gaps / open items (carried honestly): D3 cold endurance and RF-in-container (Open, release-gating); WHO PQS accuracy/benefit target (Open, D4); all quantitative outcome targets (Evidence-Pending until set in protocol).
Evidence & honesty
Per evidence-standard.md / hallucination-policy.md: every objective is a user outcome; every acceptance criterion is outcome-based and, where a numeric target is not yet set, Evidence-Pending with a named closing protocol; summative endpoints are pre-registered; no result, use-error rate, satisfaction score, return rate, or field figure is asserted — none has been measured. Every activity's Status is Planned / Evidence Pending. Standards are cited by designation (revisions to confirm).
Chain position & feeds
Flow: … Design Outputs (09) → Hazard Analysis (10) → ERFKL (11) → Risk-Control Requirements (12) → Verification (13, PVL) → Validation (14, this library) → Manufacturing/Transfer (15) → Release (16) → Field/PMS (17) → Traceability (18). Validation runs after Verification and before/at Design Transfer, traces back to User Needs / Intended Use and residual-risk acceptability, and feeds Release (Stage 16) (coverage + user/residual acceptance + go-to-market recommendation), Field (Stage 17) (the both-direction safety-signal monitoring and the HAL-2800 disclosure/monitoring obligations), and ORCH-018 / EM-ENG-012 (Stage 18) (the validation segment of the end-to-end traceability spine, incl. the HA-USE ↔ VAL-USE closures).
Convening: ORCH-015 (Validation). Lead: HF/Usability. Contributors: Clinical/Operational; Product/Business; EM-ENG-006 (residual/benefit-risk, evidence). Integrator: EM-ENG-001 (release).
VAL-BIZ-800 (PVaL-800) — Business Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-800, activity code VAL-BIZ. Owner: Product/Business, with Clinical/Operational and Finance stakeholders. This module validates the commercial and adoption outcomes that determine whether the safety benefit actually reaches beneficiaries at scale: customer satisfaction, NPS, retention, support burden, warranty/return rate, reuse TCO/cost-per-trip, and market/tender acceptance. Business validation is real but secondary to safety (severity-first): a commercially attractive device that fails the excursion critical task is not shippable. Status: Planned / Evidence Pending — no satisfaction, NPS, return-rate, or TCO figure has been measured; none is asserted.
---
1. Purpose
Confirm that the SCM is adopted, retained, supportable, and economically viable so its public-health benefit scales — because a device that programmes will not buy, procurement cannot tender, or finance cannot justify never protects a beneficiary. The dominant business unknown is the reuse economics (return rate, turnaround, cost-per-trip), on which the entire reusable (D1) business case rests and which only field data can confirm.
2. Scope
Customer satisfaction and NPS by user class; adoption and retention; support burden and warranty/return rate; reuse TCO, cost-per-trip, and utilization; PQS-eligibility and pooled-tender qualification; enterprise/OEM commercial fit; and the false-reject/wastage cost balance (the accepted trade also has a commercial face). Validated with real customers via pilots, surveys, and interviews.
3. Applicable standards & references
• WHO PQS / GAVI-UNICEF procurement — pooled-tender eligibility and TCO evidence expectations.
Upstream | UN-BIZ-001…008, UN-VAL-009 (reuse economics), UN-STK-008/009/015; conflict C6 (reuse TCO vs reverse-logistics), C1 (cost vs cold endurance); HA-RES-019 (wastage trade, commercial face).
Evidence | all financial and satisfaction figures are Judgment/Assumption ("Evidence-Pending") per the UN library — measured in pilot, never asserted.
4. Governance & approach
• Safety-first, business-second. Business acceptance never overrides a failed safety/HF validation; it is weighed after the dominant-outcome conclusion (PVaL-1200).
• Reuse economics by field measurement. Return rate and cost-per-trip are measured in the pilot/production trial (VAL-FIELD), not modelled and asserted — the swing variable in reuse TCO (conflict C6).
• Both directions have a cost. False-accept (safety) and false-reject (wastage) both carry commercial and public-health cost; business validation includes the wastage-cost balance.
• Honest pre-registration. Satisfaction/NPS/retention endpoints and targets are pre-registered and Evidence-Pending; no benchmark is asserted.
5. Validation categories
Customer satisfaction · NPS · adoption & retention · support burden · warranty/return rate · reuse TCO/ cost-per-trip · market/tender acceptance · enterprise/OEM commercial fit · wastage-cost balance.
6. Validation Register (VAL-BIZ activities)
VAL-BIZ-001 — Customer satisfaction by user class
Objective (user outcome) | Each user class is satisfied enough with the device in real use to keep using and recommend it.
Intended Use / User | All customer-facing classes (SH-01…SH-10, SH-13…SH-15).
Validation Method | Survey (importance–satisfaction) + interview post-pilot.
Participants / Representative Users | Representative users across classes after real use.
Environment / Context of Use | Post-pilot / in-deployment.
Success / Acceptance Criteria | Satisfaction meets a pre-registered threshold across classes; no class systematically dissatisfied; targets Evidence-Pending.
User Need Traceability | UN-BIZ; UN-VAL-002; STN success measures.
Risk / Use-Error Relationship | Dissatisfaction → workaround/abandonment → unmonitored trips (HA-USE-004).
Evidence Required | Survey/interview data; importance–satisfaction map.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Ties the recommended Stage-02 §6 importance–satisfaction survey.
VAL-BIZ-002 — Net Promoter Score / willingness to recommend
Objective (user outcome) | Users and buyers would recommend the SCM — a signal it delivers enough value to spread.
Intended Use / User | Program managers, QA, distributors, procurement (SH-01/SH-02/SH-03/SH-08).
Validation Method | NPS survey post-deployment.
Participants / Representative Users | Representative buyers/decision-makers.
Environment / Context of Use | Post-deployment.
Success / Acceptance Criteria | NPS meets a pre-registered target; detractor themes root-caused; target Evidence-Pending.
User Need Traceability | UN-BIZ; UN-VAL-002.
Risk / Use-Error Relationship | Low NPS predicts poor scale-up of the safety benefit.
Evidence Required | NPS data; detractor analysis.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | No NPS figure asserted pre-measurement.
VAL-BIZ-003 — Adoption & retention (sustained use, not pilot novelty)
Objective (user outcome) | Customers keep using the device past the pilot novelty period, integrated into SOPs.
Intended Use / User | Programmes, distributors, 3PLs (SH-01/SH-03/SH-04).
Validation Method | Longitudinal field study (retention/usage over time).
Participants / Representative Users | Representative deploying organizations.
Environment / Context of Use | Sustained deployment beyond pilot.
Success / Acceptance Criteria | Sustained usage/retention meets a pre-registered target; SOP integration observed; no post-pilot abandonment cluster.
User Need Traceability | UN-BIZ; UN-STK-004 (SOP inclusion); UN-VAL-002.
Risk / Use-Error Relationship | Abandonment → return to unmonitored distribution.
Evidence Required | Usage/retention time series; SOP-integration evidence.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | 3PL/program SOP inclusion is the durable-adoption signal.
VAL-BIZ-004 — Support burden & serviceability cost
Objective (user outcome) | The device is supportable at acceptable cost — support tickets, RMA rate, and field-support effort are within a viable envelope.
Intended Use / User | Support ops, service, ecosystem partners (SH-13/SH-14).
Validation Method | Support-data field study.
Participants / Representative Users | Support/service operations during pilot/trial.
Environment / Context of Use | Live support during deployment.
Success / Acceptance Criteria | Support/ticket/RMA rates within a pre-registered envelope; no dominant recurring support driver unaddressed; targets Evidence-Pending.
User Need Traceability | UN-BIZ; UN-SVC; UN-STK-013.
Risk / Use-Error Relationship | High support burden → workaround/disable behaviors.
Evidence Required | Support/ticket/RMA data; driver analysis.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Feeds Stage-17 field-support planning.
VAL-BIZ-005 — Warranty & return (RMA) rate
Objective (user outcome) | Device failures returned under warranty stay within a viable rate over the reuse life.
Intended Use / User | Finance, service, procurement (SH-10/SH-13/SH-08).
Validation Method | Field warranty/RMA tracking.
Participants / Representative Users | Deployed fleet over time.
Environment / Context of Use | In-deployment over reuse cycles.
Success / Acceptance Criteria | Warranty/RMA rate within a pre-registered envelope; failure modes root-caused; target Evidence-Pending.
User Need Traceability | UN-BIZ; UN-SVC; UN-VAL-009.
Risk / Use-Error Relationship | Distinct from reverse-logistics return rate (VAL-BIZ-006); a reliability signal (ties VAL-REL).
Evidence Required | Warranty/RMA data; failure-mode analysis.
Owner | Product/Business + Reliability.
Status | Planned / Evidence Pending.
Notes | Warranty return ≠ reverse-logistics return; both tracked.
VAL-BIZ-006 — Reuse economics: return rate, turnaround & cost-per-trip ★ (dominant business unknown)
Objective (user outcome) | The reusable business case holds — devices come back, turn around, and cost-per-trip is low enough to justify the reusable model over single-use.
Intended Use / User | Finance, procurement, service ops, executive (SH-10/SH-08/SH-13/SH-15).
Validation Method | Reuse/return pilot with cost-per-trip measurement (ties VAL-USE-017 / VAL-FIELD).
Participants / Representative Users | Representative programmes/lanes across regions (C6 region-specific).
Environment / Context of Use | Real return loops (LMIC last-mile and GDP lanes).
Success / Acceptance Criteria | Return rate, turnaround, and cost-per-trip meet a pre-registered viability threshold; all figures Evidence-Pending and measured, not modelled; region-specific results reported.
User Need Traceability | UN-VAL-009; UN-BIZ; UN-SVC; UN-STK-008/009; D1; conflict C6.
Risk / Use-Error Relationship | No direct hazard; the economics that fund the safety mission at scale.
Evidence Required | Return-pilot report; return-rate/turnaround/cost-per-trip by region.
Owner | Product/Business + Finance.
Status | Planned / Evidence Pending.
Notes | Return rate is the swing variable in reuse TCO (conflict C6); THE business validation.
VAL-BIZ-007 — Market / pooled-tender acceptance (PQS eligibility path)
Objective (user outcome) | The device qualifies into institutional pooled tenders (GAVI/UNICEF) — the channel that scales the benefit to LMIC volume.
Intended Use / User | Procurement (GAVI/UNICEF), programmes (SH-08/SH-01).
Validation Method | Tender-qualification evidence review + procurement interview.
Participants / Representative Users | Representative institutional procurement.
Environment / Context of Use | Tender qualification / framework award.
Success / Acceptance Criteria | PQS-eligibility path and TCO/supply evidence judged sufficient to qualify (contingent on PQS listing, Evidence-Pending, D4).
User Need Traceability | UN-STK-008; UN-VAL-009; UN-REG (PQS); UN-BIZ.
Risk / Use-Error Relationship | No direct hazard; gates whether the benefit reaches LMIC scale.
Evidence Required | Tender-qualification assessment; PQS-path status.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Shares evidence with VAL-USER-012; PQS listing has long lead times.
VAL-BIZ-008 — Enterprise / OEM commercial fit & channel viability
Objective (user outcome) | Enterprise customers and OEM/reseller partners find the commercial and integration terms viable enough to standardize and resell.
Intended Use / User | Executive/IT (SH-15/SH-09); ecosystem partners (SH-14).
Validation Method | Commercial pilot + partner interview.
Participants / Representative Users | Representative enterprise buyers and channel partners.
Environment / Context of Use | Enterprise procurement and partner enablement.
Success / Acceptance Criteria | Enterprise/partner commercial fit judged viable; channel enablement sufficient; no blocking commercial gap.
User Need Traceability | UN-STK-013/015; UN-ECO; UN-BIZ.
Risk / Use-Error Relationship | No direct hazard; channel that scales the benefit into fragmented markets.
Evidence Required | Commercial-pilot notes; partner interviews.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Ties VAL-USER-014 (partner integration) and VAL-ECO-010 (enterprise integration).
VAL-BIZ-009 — Wastage-cost balance (the accepted trade, commercial face)
Objective (user outcome) | The conservative safety bias's wastage/false-reject does not impose an unacceptable commercial or stock-out cost, and its reduction versus the counterfactual is a value story.
Intended Use / User | Programmes, procurement, finance (SH-01/SH-08/SH-10).
Validation Method | Field cost analysis (ties VAL-CLIN-004).
Participants / Representative Users | Representative programmes/distributors.
Environment / Context of Use | Real lanes with wastage tracking.
Success / Acceptance Criteria | Net value positive (loss reduction ≫ wastage cost) vs counterfactual; false-reject cost within a pre-registered bound; Evidence-Pending.
User Need Traceability | UN-BIZ; UN-VAL-009; HA-RES-019.
Risk / Use-Error Relationship | Commercial face of the accepted public-health trade HA-RES-019.
Evidence Required | Wastage/loss-reduction cost analysis; net-value model with field inputs.
Owner | Product/Business + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Balances the safety bias's cost against its benefit in commercial terms.
7. Traceability notes
VAL-BIZ validates the adoption/economic outcomes that scale the safety benefit (UN-BIZ, UN-VAL-009). VAL-BIZ-006 (reuse economics) is the dominant business unknown and shares field evidence with VAL-USE-017 and VAL-FIELD; VAL-BIZ-007 gates the LMIC tender channel; VAL-BIZ-009 is the commercial face of HA-RES-019. All figures are Evidence-Pending and measured in the field, never asserted. Feeds PVaL-900 and PVaL-1200.
8. Acceptance / summary notes
Business validation is real but subordinate to safety. Its load-bearing item is the reuse economics (return rate / cost-per-trip), measured in the pilot. All activities are Planned / Evidence Pending; no satisfaction, NPS, return-rate, or TCO number is asserted before measurement.
VAL-CLIN-500 (PVaL-500) — Clinical / Operational Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-500, activity code VAL-CLIN. Owner: Clinical/Operational contributor, with HF/Usability and EM-ENG-006. Applicability: APPLICABLE (operational). The SCM is not a therapeutic device, so this module is operational-effectiveness validation, not a clinical trial: it confirms the device actually protects vaccine-potency decisions in real distribution — that it improves the correctness of disposition decisions, reduces both compromised-dose administration and needless wastage, provides effective decision support, and improves cold-chain outcomes at programme scale. Status: Planned / Evidence Pending — no operational study executed; no effectiveness figure asserted.
---
1. Purpose
Answer the outcome question the device exists for at the level of real distribution outcomes, not individual tasks: does deploying the SCM lead to better, more correct, more defensible vaccine-disposition decisions than the counterfactual (no monitoring, or a dumb single-use indicator) — protecting potency (no compromised dose administered) while avoiding needless destruction of good product (no potent dose wasted)? This module validates the device's public-health effectiveness and its role as decision support for human/organizational disposition decisions.
2. Scope
Operational effectiveness across representative LMIC/GDP distribution lanes and programmes: correctness of disposition decisions with vs without the SCM; two-sided detection effectiveness against real product- stability profiles; decision-support quality (does the verdict/record lead users to the right call); wastage/false-reject reduction; workflow improvement in receiving and dispatch; audit/defensibility of disposition; and programme-scale cold-chain outcome signals. The beneficiary outcome is represented via disposition correctness and potency-protection proxies — the device does not touch a patient.
3. Applicable standards & references
• WHO PQS E006 / WHO-UNICEF EVM (Effective Vaccine Management) — the operational effectiveness frame, cold-chain performance indicators, and wastage metrics.
• EU GDP / 21 CFR Part 11 / ALCOA+ — defensibility of the disposition decision and its record.
• ISO 14971 — the disposition-correctness outcome is the severity anchor (STN-020).
Upstream | UN-VAL-002/005/006, UN-VIS-001, UN-SAFE-001/002; HAL-2900 §6 (benefit frame, both directions); RCR-referenced VAL-SW-001 (detection effectiveness) and VAL-USE-002 (false-reject/wastage targets).
4. Governance & approach
• Counterfactual comparison. Effectiveness is validated against the realistic counterfactual (no monitor / dumb indicator), not against perfection — the benefit claim in HAL-2900 §6 is exactly this comparison.
• Both harm directions are endpoints. Disposition correctness is measured on both axes: compromised-dose administration avoided and good-dose wastage avoided (the accepted-trade residual HA-RES-019 is a first- class metric, not a footnote).
• Decision support, not automation. The SCM supports a human decision; effectiveness is validated as the quality of the supported decision, including the honest-gap and honest-reject surfaces.
• Represented beneficiary, honest proxies. Potency protection is validated via disposition correctness and cold-chain-integrity proxies; no clinical potency assay is claimed unless a partner study provides it (Evidence-Pending).
• Provisional benefit. Per HAL-2900 §5.6, the benefit claim is provisional until the WHO PQS accuracy target is fixed and met (D4) — effectiveness is validated against that target once confirmed.
5. Validation categories
Disposition-decision correctness (effectiveness) · two-sided detection effectiveness · decision support · wastage/false-reject reduction · workflow improvement · audit/defensibility · programme-scale outcome signals.
6. Validation Register (VAL-CLIN activities)
VAL-CLIN-001 — Operational effectiveness: correct disposition decisions vs the counterfactual ★
Objective (user outcome) | Deploying the SCM measurably increases the rate of correct vaccine- disposition decisions (accept good / reject compromised) versus no monitoring or a dumb indicator, in real distribution.
Intended Use / User | Programmes, distributors, QA (SH-01/SH-03/SH-02); protects beneficiary (SH-11).
Validation Method | Field study / operational comparison across representative lanes.
Participants / Representative Users | Representative programmes/distributors running SCM vs counterfactual lanes.
Environment / Context of Use | Real LMIC/GDP distribution lanes over a representative period.
Success / Acceptance Criteria | Disposition-decision correctness (both directions) is higher with the SCM than the counterfactual, with no increase in compromised-dose administration; effect-size target Evidence- Pending (set with programme partners).
User Need Traceability | UN-VIS-001; UN-VAL-002/005; UN-SAFE-001; STN-020.
Risk / Use-Error Relationship | Validates the device's core benefit and the HAL-2900 §6 benefit claim (both directions).
Evidence Required | Field-study protocol/report; disposition-correctness data with/without SCM; adverse- disposition log.
Owner | Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | The operational-effectiveness anchor; feeds the benefit-risk conclusion (PVaL-1000/1200).
VAL-CLIN-002 — Two-sided detection effectiveness against real product-stability profiles
Objective (user outcome) | Real freeze and heat excursions on real vaccine profiles are correctly caught and lead to the correct disposition — freeze detected as reliably as heat.
Intended Use / User | Programmes/QA relying on the verdict (SH-01/SH-02); beneficiary.
Validation Method | Operational validation against reference excursion scenarios (ties VAL-SW-001).
Participants / Representative Users | Programme/QA decision-makers; reference product-stability library.
Environment / Context of Use | Scripted and real freeze/heat excursions across profiles/lanes.
Success / Acceptance Criteria | No missed excursion type; freeze and heat detection lead to correct disposition; no systematic false-accept; accuracy anchored to the WHO PQS target (Evidence-Pending, D4).
User Need Traceability | UN-SAFE-002; UN-FUN-004/005; UN-VAL-005.
Risk / Use-Error Relationship | HA-SW-001 (missed excursion); this is the operational face of VAL-SW-001, validating HA-RES-001 effectiveness in real use.
Evidence Required | Detection-effectiveness report; per-profile freeze/heat outcomes.
Owner | Clinical/Operational + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Referenced upstream as VAL-SW-001; operational effectiveness of two-sided detection.
VAL-CLIN-003 — Decision support: the verdict/record leads users to the right call
Objective (user outcome) | The device's verdict, honest-gap qualifier, and record actively lead users to the correct disposition — including on incomplete/gapped/borderline trips — rather than to over-trust or confusion.
Intended Use / User | Disposition decision-makers (SH-01/SH-05/SH-02).
Validation Method | Simulated-use decision-support study + field.
Participants / Representative Users | Representative decision-makers.
Environment / Context of Use | Receipt disposition incl. gapped/borderline records.
Success / Acceptance Criteria | Supported decisions are more correct and better calibrated than unsupported ones; the honest-gap surface prevents false-reassurance; no automation-bias-driven wrong acceptance.
User Need Traceability | UN-UX-011; UN-SAFE-003; UN-VAL-002; UN-JTBD-002.
Risk / Use-Error Relationship | HA-DATA-004/005 (false reassurance/misleading viz); validates HA-RES-002 in real decisions.
Evidence Required | Decision-support study report; decision-correctness on gapped/borderline cases.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Validates the "supported human decision" choke point (HAL-2900 §5.2) operationally.
VAL-CLIN-004 — Wastage / false-reject reduction (the accepted-trade validated)
Objective (user outcome) | The SCM reduces needless destruction of good product versus the counterfactual, and its conservative safety bias does not drive an unacceptable false-reject/stock-out rate.
Intended Use / User | Programmes, procurement, finance (SH-01/SH-08/SH-10); beneficiary (avoids stock-out).
Validation Method | Field study (wastage/false-reject measurement) — ties VAL-USE-002 targets.
Participants / Representative Users | Representative programmes/distributors.
Environment / Context of Use | Real lanes with real disposition pressure.
Success / Acceptance Criteria | Wastage of good product is reduced vs counterfactual; false-reject rate within the validated two-sided target (Evidence-Pending); a rising false-reject rate is reportable as a first-class metric.
User Need Traceability | UN-VAL-005; UN-SAFE-004; HA-RES-019.
Risk / Use-Error Relationship | Validates the accepted public-health trade HA-RES-019 and HAL-2900 both-directions balance.
Evidence Required | Wastage/false-reject field data; counterfactual comparison.
Owner | Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Wastage is measured as a first-class outcome, balancing the false-accept residuals.
VAL-CLIN-005 — Workflow improvement at receiving & dispatch
Objective (user outcome) | The SCM makes receiving/disposition and dispatch faster and more reliable than the incumbent process, so it is adopted rather than worked around.
Intended Use / User | Receiving/dispatch staff, QA (SH-05/SH-02/SH-03).
Validation Method | Time-and-motion field study + interview.
Participants / Representative Users | Representative receiving/dispatch operations.
Environment / Context of Use | Real receiving docks and pack-out lines.
Success / Acceptance Criteria | Disposition/dispatch time and error rate improve vs incumbent; no covert workaround; adoption sustained over the study window.
User Need Traceability | UN-STK-001/005; UN-UX-005/006; UN-JTBD-002/004.
Risk / Use-Error Relationship | Friction → skipped/void trips (HA-USE-004); adoption gate.
Evidence Required | Time-and-motion data; workaround observations.
Owner | Clinical/Operational + Product.
Status | Planned / Evidence Pending.
Notes | Workflow fit is what turns a validated device into an adopted one.
VAL-CLIN-006 — Audit / defensibility of the disposition decision (audit simulation)
Objective (user outcome) | A disposition decision made on the SCM record can be defended in a real GDP/ Part 11 audit — the record is complete, attributable, and unambiguous.
Intended Use / User | QA/RA, regulators/auditors (SH-02/SH-07).
Validation Method | Audit-simulation study with real auditors (shared with VAL-USER-011).
Participants / Representative Users | Representative GDP/Part 11 auditors and QA reviewers.
Environment / Context of Use | Simulated audit/inspection.
Success / Acceptance Criteria | Auditors judge the record admissible and the disposition defensible; no ambiguity that would invalidate a decision or listing.
User Need Traceability | UN-VAL-006; UN-STK-007; UN-JTBD-007/008; UN-REG.
Risk / Use-Error Relationship | HA-MPE-004 / HA-DATA-006; validates HA-RES-020/018 defensibility.
Evidence Required | Audit-simulation report; auditor findings; admissibility judgments.
Owner | EM-ENG-006 + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Defensibility is part of operational effectiveness — an indefensible correct call still fails.
VAL-CLIN-007 — Programme-scale cold-chain outcome signals
Objective (user outcome) | At programme scale, deploying the SCM is associated with improved cold-chain integrity outcomes (fewer undetected excursions reaching point-of-care; better loss visibility).
Intended Use / User | EPI programmes, executive/supply-chain leadership (SH-01/SH-15).
Validation Method | Programme field study (observational, multi-site).
Participants / Representative Users | Representative national/regional programmes.
Environment / Context of Use | Programme-scale deployment over a representative period.
Success / Acceptance Criteria | Programme-level cold-chain integrity indicators improve or become visible where previously blind; no adverse programme-scale effect; effect Evidence-Pending.
User Need Traceability | UN-STK-005/015; UN-VIS; UN-VAL-002.
Risk / Use-Error Relationship | Validates the aggregate benefit (HAL-2900 §5.1/§6) at programme scale.
Evidence Required | Programme-scale study report; cold-chain KPI deltas.
Owner | Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Observational; feeds ISO 14971 §9 post-market benefit-risk revalidation (Stage 17).
VAL-CLIN-008 — Attribution effectiveness: where/when a break occurred is correctly localized
Objective (user outcome) | Users can correctly attribute where and when a cold-chain break happened, so corrective action targets the right lane/leg.
Intended Use / User | Distributors, programmes, QA (SH-03/SH-01/SH-02).
Validation Method | Field study + record-review.
Participants / Representative Users | Representative logistics/QA analysts.
Environment / Context of Use | Multi-leg lanes with induced/real breaks.
Success / Acceptance Criteria | Break location/time correctly attributed within useful bounds; time/ location provenance disclosed and understood.
User Need Traceability | UN-JTBD-008; UN-FUN-016; UN-VAL-002.
Risk / Use-Error Relationship | HA-DATA-002/003 (attribution/time); validates HA-RES-016/022.
Evidence Required | Attribution-accuracy report; provenance comprehension.
Owner | Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Attribution turns detection into corrective action; provenance honesty is validated too.
VAL-CLIN-009 — Effectiveness holds for the low-training last-mile decision (operational)
Objective (user outcome) | The operational effectiveness gains hold specifically at the low-training last-mile point of use, not only at well-resourced central sites.
Intended Use / User | Last-mile programmes/clinics (SH-01/SH-05); beneficiary.
Validation Method | Field study at last-mile sites (subset of VAL-CLIN-001).
Participants / Representative Users | Representative last-mile health workers/clinics.
Environment / Context of Use | Rural health posts / mobile teams.
Success / Acceptance Criteria | Disposition correctness at last-mile sites matches or approaches central- site performance; no degradation concentrated at the last mile.
User Need Traceability | UN-STK-006; UN-VAL-003; UN-SAFE-005.
Risk / Use-Error Relationship | HA-CTX-005 (vulnerable user) at scale; validates HA-RES-014 in the field.
Evidence Required | Last-mile field-study report; site-stratified disposition correctness.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Ties the summative HF result to real last-mile operational outcomes.
VAL-CLIN-010 — Benefit provisional on WHO PQS accuracy target (confirmation gate)
Objective (user outcome) | The device's effectiveness is validated against the recognized cold-chain- monitor accuracy/performance target it must meet for its benefit claim to hold.
Intended Use / User | Programmes/procurement/regulators relying on the benefit claim.
Validation Method | Effectiveness-against-target confirmation (gated on D4 PQS category/target).
Participants / Representative Users | EM-ENG-006, Clinical/Operational, regulatory.
Environment / Context of Use | Operational effectiveness reconciled to the confirmed PQS target.
Success / Acceptance Criteria | Once the WHO PQS category/target is confirmed (D4, Evidence-Pending), the device's demonstrated effectiveness meets it; benefit claim upgraded from provisional to demonstrated.
User Need Traceability | UN-REG (PQS); UN-VAL-005; HAL-2900 §5.6/condition 6.
Risk / Use-Error Relationship | Gates the benefit side of the overall benefit-risk conclusion.
Evidence Required | Confirmed PQS target; effectiveness-vs-target reconciliation.
Owner | EM-ENG-006 + Clinical/Operational.
Status | Planned / Evidence Pending (PQS category/target is Open, D4).
Notes | HAL-2900 condition 6 — the benefit claim is provisional until this closes.
7. Traceability notes
VAL-CLIN validates the operational effectiveness that HAL-2900 §6 asserts as the device's benefit — both directions of the disposition. VAL-CLIN-001/002/003 are the effectiveness/decision-support anchors (VAL-CLIN-002 is referenced upstream as VAL-SW-001); VAL-CLIN-004 validates the accepted-trade residual HA-RES-019; VAL-CLIN-006 provides the defensibility outcome (with VAL-USER-011); VAL-CLIN-010 gates the benefit claim on the D4 PQS target. All trace to UN-VIS/UN-VAL/UN-SAFE and feed PVaL-1000/1200.
8. Acceptance / summary notes
Operational effectiveness is validated against the realistic counterfactual, on both harm directions, with the low-training last mile included, and with the benefit claim explicitly provisional until the WHO PQS accuracy target is confirmed and met (D4). All activities are Planned / Evidence Pending; no effectiveness figure, wastage rate, or benefit magnitude is asserted.
VAL-ECO-600 (PVaL-600) — Ecosystem Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-600, activity code VAL-ECO. Owner: Product, with HF/Usability, Clinical/Operational, and EM-ENG-006. The SCM is a device + cloud dashboard + mobile/BLE app + multi-device fleet, operated offline-first with store-and-forward sync. This module validates that the ecosystem works as a system for real users — the cloud dashboard, the mobile app, fleet management at scale, multi-device operations, offline operation, synchronization/reconciliation, and remote management — so that no user outcome falls through a gap between tiers. Status: Planned / Evidence Pending — no ecosystem study executed.
---
1. Purpose
Validate the SCM as an integrated ecosystem, not three separate products. The device holds the safety truth (offline-first, D2), the cloud provides fleet visibility, alerting, and the cross-check, and the mobile app provisions and disposes at the edge. The user outcomes — timely alerts, fleet oversight, correct sync, consistent verdict/record across tiers, remote management — only hold if the tiers cooperate. This module validates that cooperation with representative fleet/IT/program users, including the hard cases: offline periods, RF-hostile in-container operation, and multi-device fleets at scale.
2. Scope
Cloud dashboard (fleet view, alerting, cross-check, export); mobile/BLE app (provisioning, disposition, readiness); multi-device fleet operations; offline operation and store-and-forward buffering; synchronization and end-to-end reconciliation; remote fleet management (config, OTA campaigns, monitoring); enterprise integration (SSO/RBAC/API); and cross-tier consistency (device verdict = cloud recompute = record). Validated across the RF-hostile in-container and intermittent-connectivity contexts (UN-ENV-006/007).
3. Applicable standards & references
• EU GDP / 21 CFR Part 11 / ALCOA+ — record integrity across tiers; enterprise data governance.
• WHO PQS E006 — fleet/programme operating context. ISO 14971 — cross-tier residuals (delivery, cloud, sync).
Upstream | UN-ECO-001…009, UN-FUN-009/010/015/018, UN-JTBD-006/009, UN-ENV-006/007; HAL-1100/1200 (connectivity/cloud), HA-COMM-001/002/003, HA-CLOUD-001…008; RCR-referenced VAL-COMM-001/002/003/008.
4. Governance & approach
• System-level user outcomes. Validation endpoints are user outcomes across tiers (e.g. "the responder gets the alert and acts"), never single-tier features.
• Offline-first is validated, not assumed. The local safety floor (detect/commit/verdict/edge-read with zero connectivity, HAL-2900 §5.4) is explicitly validated; the cloud is validated as an overlay whose loss degrades to a logger-with-verdict, not to unsafe.
• RF-in-container honesty (RF Open, PRD-VER-005). Connectivity-dependent outcomes are validated against real RF-hostile in-container conditions; where the link cannot close, the alerting claim defaults to best-effort and the on-receipt verdict is the compensating outcome.
• Reconciliation over silence. Sync/delivery validation confirms losses surface as detectable backlog/gaps, never as silent absence (RCR-032/035, SYS-176).
5. Validation categories
Cloud dashboard · mobile app · fleet scale · multi-device · offline operation · synchronization/reconciliation · remote management · enterprise integration · cross-tier consistency.
6. Validation Register (VAL-ECO activities)
VAL-ECO-001 — Cloud dashboard: fleet watchers see only what needs action
Objective (user outcome) | Fleet/program users monitor many shipments and surface exactly the action- required items, without interpreting raw data or missing a real excursion.
Intended Use / User | Distributor coordinators, programmes, IT/fleet (SH-03/SH-01/SH-09).
Validation Method | Summative usability test (dashboard) + field.
Participants / Representative Users | Representative fleet coordinators/programme staff.
Environment / Context of Use | Operations center; many concurrent shipments.
Success / Acceptance Criteria | Users triage action-required vs informational correctly, find shipment status at a glance, and do not miss genuine excursions; usability targets Evidence-Pending.
User Need Traceability | UN-ECO-001; UN-JTBD-006; UN-FUN-018; UN-UX-006; UN-STK-002.
Risk / Use-Error Relationship | HA-USE-002 (fatigue) at fleet scale; ties HA-CLOUD-001.
Evidence Required | Dashboard usability report; triage accuracy; miss incidence.
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Shares evidence with VAL-USER-005 and VAL-HF-005.
VAL-ECO-002 — Mobile/BLE app: provision, disposition & readiness in the field
Objective (user outcome) | Field users complete provisioning, disposition, and readiness checks on the mobile/BLE app reliably, including on constrained devices and links.
Intended Use / User | Provisioners, technicians, program staff (SH-02/SH-06/SH-01).
Validation Method | Summative usability test (mobile) — shared with VAL-HF-011.
Participants / Representative Users | Representative app users across device types.
Environment / Context of Use | Field provisioning; variable devices/connectivity.
Success / Acceptance Criteria | App tasks complete without critical error on representative devices; BLE provisioning reliable; profile selection error-proofed.
User Need Traceability | UN-ECO-002; UN-FUN-011; UN-UX-005/007.
Risk / Use-Error Relationship | HA-USE-003 (profile via app); ties HA-RES-014.
Evidence Required | Mobile usability report; task/error data by device.
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | BLE reliability in the field is a distinct outcome from app UX.
VAL-ECO-003 — Offline operation: local safety floor with zero connectivity
Objective (user outcome) | With no connectivity for a whole trip, the device still logs, detects, commits, and gives a correct on-receipt verdict — users are never left without a disposition basis.
Intended Use / User | Last-mile programmes/clinics, distributors (SH-01/SH-05/SH-03).
Validation Method | Field study / simulated offline trips.
Participants / Representative Users | Representative last-mile and transit users.
Environment / Context of Use | Zero/blackout connectivity lanes; remote posts.
Success / Acceptance Criteria | On-receipt verdict correct with zero connectivity; no data loss; users interpret the offline-derived verdict correctly; degrades to logger-with-verdict, never unsafe.
User Need Traceability | UN-ECO-006; UN-ENV-007; UN-FUN-009; UN-FLOW-005.
Risk / Use-Error Relationship | Validates the local safety floor (HAL-2900 §5.4); ties HA-COMM-001/003.
Evidence Required | Offline-trip report; verdict correctness offline; data-completeness.
Owner | Clinical/Operational + Product.
Status | Planned / Evidence Pending.
Notes | The property that makes systemic connectivity dependence acceptable — validated directly.
VAL-ECO-004 — Store-and-forward buffering & opportunistic upload
Objective (user outcome) | Data buffered while offline uploads correctly and completely when a link appears, so the fleet record is eventually complete.
Intended Use / User | Distributors, IT/fleet, programmes (SH-03/SH-09/SH-01).
Validation Method | Field study (intermittent connectivity) — referenced upstream as VAL-COMM-001.
Participants / Representative Users | Representative transit/logistics operations.
Environment / Context of Use | Intermittent-connectivity lanes; in-container RF attenuation.
Success / Acceptance Criteria | Buffered data uploads completely on link availability; no loss; backlog visible; behavior holds under RF-hostile in-container conditions (RF Open, PRD-VER-005).
User Need Traceability | UN-ECO-006; UN-FUN-009/010; UN-ENV-006/007; UN-VAL-008.
Risk / Use-Error Relationship | HA-COMM-001 (latency); validates HA-RES-007 with users; RF Open gate.
Evidence Required | Store-and-forward field report; upload-completeness; backlog visibility.
Owner | Product + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Referenced upstream as VAL-COMM-001; gated on the RF-in-container Open item.
VAL-ECO-005 — Alert delivery & escalation end-to-end (device → cloud → responder)
Objective (user outcome) | A genuine excursion alert reaches a responder who can act, and escalates if unacknowledged — end-to-end across tiers.
Intended Use / User | Responders, distributors, programmes (SH-03/SH-01).
Validation Method | Field study + simulated delivery-path scenarios — referenced as VAL-COMM-002.
Participants / Representative Users | Representative responders and escalation chains.
Environment / Context of Use | Real lanes with connectivity variation; escalation policies configured.
Success / Acceptance Criteria | Alerts delivered ahead of bulk, ack-required, and escalate on no-ack; no silent non-delivery (surfaces as backlog); responder acts within window where connectivity allows.
User Need Traceability | UN-ECO-003; UN-FUN-006/007; UN-JTBD-009; UN-SAFE-004.
Risk / Use-Error Relationship | HA-COMM-002 (missing alert); validates HA-RES-008 with users.
Evidence Required | Delivery/escalation field report; ack/escalation data; non-delivery backlog.
Owner | Product + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Referenced upstream as VAL-COMM-002; the human twin of alarm fatigue (VAL-USE-002).
VAL-ECO-006 — Synchronization & end-to-end reconciliation (no silent gap between tiers)
Objective (user outcome) | The device record, cloud record, and export always reconcile; any loss between tiers surfaces as a detectable gap, never a silent absence.
Intended Use / User | IT/fleet, QA/RA, auditors (SH-09/SH-02/SH-07).
Validation Method | Simulated-use + field reconciliation scenarios — referenced as VAL-COMM-008.
Participants / Representative Users | Representative fleet/QA users.
Environment / Context of Use | Multi-tier operation with induced losses.
Success / Acceptance Criteria | Records reconcile end-to-end; induced losses appear as detectable backlog/gaps; no silent divergence; users can trust cross-tier consistency.
User Need Traceability | UN-ECO-007; UN-FUN-014/015; UN-UX-011.
Risk / Use-Error Relationship | HA-DATA-002/006, HA-CLOUD-004 (mis-attribution/loss); validates HA-RES-016/018.
Evidence Required | Reconciliation report; loss-visibility data; cross-tier consistency checks.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Referenced upstream as VAL-COMM-008; reconciliation over silence.
VAL-ECO-007 — Cross-tier consistency: device verdict = cloud recompute = record
Objective (user outcome) | Users always see a consistent verdict and record across device, dashboard, and export — the cloud cross-check agrees with the device, and disagreements are surfaced, not hidden.
Intended Use / User | Disposition decision-makers, QA, auditors (SH-01/SH-02/SH-07).
Validation Method | Simulated-use consistency scenarios + field.
Participants / Representative Users | Representative decision-makers/QA.
Environment / Context of Use | Device + dashboard + export side-by-side.
Success / Acceptance Criteria | Device verdict, cloud recompute, and record agree; any divergence is surfaced and adjudicated toward "not a clean accept," never toward a false clean pass.
User Need Traceability | UN-UX-011; UN-FUN-008/014; UN-SAFE-003.
Risk / Use-Error Relationship | HA-CLOUD-001/002 (rules/notify), HA-SW-002; validates HA-RES-002/017 cross-tier.
Evidence Required | Consistency-scenario report; divergence-handling observations.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | The cloud cross-check must never mask on-device safety logic (HA-CIR-003).
VAL-ECO-008 — Multi-device fleet operation at scale
Objective (user outcome) | Fleet users manage many devices at once — status, readiness, config, attribution — without a device or shipment falling through the cracks.
Intended Use / User | IT/fleet admins, distributors, service (SH-09/SH-03/SH-06).
Validation Method | Scale field study / simulated fleet.
Participants / Representative Users | Representative fleet operators at realistic device counts.
Environment / Context of Use | Large multi-device fleet across lanes.
Success / Acceptance Criteria | Fleet operations scale without loss of visibility or mis-attribution; readiness/status accurate across the fleet.
User Need Traceability | UN-ECO-004; UN-FUN-017/018; UN-STK-009/015.
Risk / Use-Error Relationship | HA-DATA-002 (mis-attribution) at scale; validates HA-RES-016.
Evidence Required | Fleet-scale study report; attribution accuracy at scale.
Owner | Product.
Status | Planned / Evidence Pending.
Notes | Ties fleet ROI/visibility outcomes (VAL-BIZ / VAL-USER-013).
VAL-ECO-009 — Remote management: config change & OTA campaign at fleet scale
Objective (user outcome) | Fleet admins push config changes and firmware campaigns safely across the fleet — staged, monitored, reversible — without blinding or bricking devices.
Intended Use / User | IT/fleet admins (SH-09); service (SH-06).
Validation Method | Simulated/field OTA & config campaign — ties VAL-USE-014.
Participants / Representative Users | Representative fleet admins.
Environment / Context of Use | Fleet OTA/config campaign with staged rollout.
Success / Acceptance Criteria | Campaigns roll out staged with halt-on-regression; no bricked devices; any unmonitored update window understood; config changes authenticated/audited.
User Need Traceability | UN-ECO-008; UN-FLOW-010; UN-SEC.
Risk / Use-Error Relationship | HA-CYBER-005/HA-SW-010 (OTA), HA-CLOUD-002 (config); validates HA-RES-012/017 operationally (fleet common-cause containment, HAL-2900 §5.3).
Evidence Required | Campaign report; rollout/rollback observations; config-audit review.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | OTA/cloud is a fleet common cause — its operational containment (canary) is validated here.
VAL-ECO-010 — Enterprise integration: SSO/RBAC/API & data governance in a real customer stack
Objective (user outcome) | An enterprise customer integrates the platform into their stack with SSO/RBAC, retention, and API integration meeting their governance — so it is allowed on the network and used.
Intended Use / User | IT/security admins, ecosystem partners (SH-09/SH-14).
Validation Method | Integration pilot in a representative customer environment.
Participants / Representative Users | Representative enterprise IT and partner integrators.
Environment / Context of Use | Real/representative enterprise stack.
Success / Acceptance Criteria | Integration completed to the customer's security/governance bar; SSO/RBAC/ retention/API function; no blocking gap.
User Need Traceability | UN-ECO-005; UN-STK-003/013; UN-SEC.
Risk / Use-Error Relationship | HA-CYBER-009/HA-CLOUD-008 (access/compromise); ties HA-RES-013; conflict C9.
Evidence Required | Integration-pilot report; security/governance sign-off.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Enterprise/partner adoption gate; shares evidence with VAL-USER-009/014.
7. Traceability notes
VAL-ECO validates the ecosystem outcomes that the connectivity/cloud hazards (HA-COMM/HA-CLOUD) and their residuals (HA-RES-007/008/013/016/017) turn on. VAL-ECO-004/005/006 are referenced upstream as VAL-COMM-001/002/008; VAL-ECO-003 validates the local safety floor; VAL-ECO-007 the cross-tier consistency choke point; VAL-ECO-009 the fleet common-cause containment. All trace to UN-ECO/UN-FUN/UN-ENV and feed PVaL-700 (reliability), PVaL-900 (field), and PVaL-1000 (residual).
8. Acceptance / summary notes
The ecosystem is validated as a system with the local safety floor proven directly, connectivity-dependent outcomes validated honestly against RF-hostile conditions (RF Open, PRD-VER-005), and cross-tier reconciliation proven to surface losses rather than hide them. All activities are Planned / Evidence Pending; no sync, delivery, or fleet-scale result is asserted.
VAL-EVID-1100 (PVaL-1100) — Validation Evidence
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-1100, activity code VAL-EVID. Owner: EM-ENG-006 + HF/Usability. This module defines the evidence each validation activity must produce and how it is captured, controlled, and retained — the observations, videos, transcripts, survey data, task/ error/time records, usability reports, and acceptance records — kept distinct from the playbook. The playbook (VAL-GOV…VAL-SUM) defines what must be validated and with whom; this module governs the records of what was observed. Status: Planned / Evidence Pending — no evidence exists yet; this defines its form, control, and honesty, not its content.
---
1. Purpose
Ensure every validation conclusion is backed by a real, controlled, retrievable, honestly-labelled record — so that "validated" is auditable, and so no result enters a module (or the summary) without an evidence record behind it. This is the SCM's usability engineering file plus the field/operational evidence base.
2. Scope
Evidence types: pre-registered protocols; user observations and video; interview/survey instruments and data; task-completion, use-error (root-caused), time-study, and error-rate records; usability reports; audit- simulation findings; field-study and pilot reports; residual-risk determinations and benefit-risk records; signed acceptance records; and the IEC 62366-1 usability engineering file. Covers evidence provenance, integrity, retention, and access.
3. Applicable standards & references
• IEC 62366-1 (usability engineering file); FDA HF (summative report content); ISO 13485 §4.2.5 (records control); 21 CFR Part 11 / ALCOA+ (data integrity of the evidence itself); ISO 14971 (residual-risk/benefit-risk records).
Upstream/companion | every VAL- activity's Evidence Required field; evidence-standard.md, hallucination-policy.md.
4. Governance & approach
• Pre-registration. Summative critical-task protocols (endpoints, criteria, analysis) are registered before execution; post-hoc criteria changes are recorded, never silent.
• Provenance & integrity. Evidence is attributable, dated, and ALCOA+; the evidence base is itself data-integrity-controlled.
• No result without a record. A result may appear in a module or the summary only when its evidence record exists here and is referenced.
• Honesty. Pre-execution, all records are placeholders marked Evidence-Pending; no fabricated data.
5. Validation categories
Protocols · observations/video · interviews/surveys · task/error/time data · usability reports · audit findings · field/pilot reports · residual/benefit-risk records · acceptance records · usability engineering file.
6. Validation Register (VAL-EVID activities)
VAL-EVID-001 — Pre-registered protocols for all summative & residual studies
Objective (user outcome) | Every reader can trust that summative endpoints and acceptance criteria were fixed before data, not fitted after.
Intended Use / User | Regulators/auditors, QA/RA (SH-07/SH-02).
Validation Method | Protocol pre-registration & control.
Participants / Representative Users | HF/Usability lead, EM-ENG-006.
Environment / Context of Use | Design-control.
Success / Acceptance Criteria | Each summative (VAL-USE/VAL-HF) and residual (VAL-RISK) study has a pre-registered protocol with endpoints/criteria; deviations documented.
User Need Traceability | UN-VAL-001; VAL-GOV-005.
Risk / Use-Error Relationship | Prevents false-reassurance of validated safety.
Evidence Required | Registered protocols; deviation log.
Owner | HF/Usability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | The honesty backbone of the evidence base.
VAL-EVID-002 — User observation & video records (critical tasks)
Objective (user outcome) | Critical-task performance is captured objectively (video/observation) so use errors are root-caused, not merely counted.
Intended Use / User | HF analysts; auditors.
Validation Method | Structured observation + video capture.
Participants / Representative Users | Summative-study participants (consented).
Environment / Context of Use | Usability lab / field.
Success / Acceptance Criteria | Each critical-task session has an analyzable record; use errors root- caused; consent/privacy controlled.
User Need Traceability | UN-VAL-003; IEC 62366-1.
Risk / Use-Error Relationship | Evidence for HA-RES-014 residual determination.
Evidence Required | Observation logs; video; root-cause analyses.
Owner | HF/Usability.
Status | Planned / Evidence Pending.
Notes | Root-cause is what makes a residual judgeable.
VAL-EVID-003 — Interview, survey & task/error/time datasets
Objective (user outcome) | Quantitative and qualitative outcome data (satisfaction, NPS, task-completion, error/time) are captured with valid instruments.
Intended Use / User | Product/Business, HF, Clinical/Operational.
Validation Method | Instrument design + data capture.
Participants / Representative Users | Study/pilot participants.
Environment / Context of Use | Studies, pilots, deployments.
Success / Acceptance Criteria | Validated instruments; complete, attributable datasets; no fabricated values.
User Need Traceability | UN-VAL-002; UN-BIZ.
Risk / Use-Error Relationship | Feeds VAL-BIZ/VAL-CLIN conclusions.
Evidence Required | Instruments; datasets; analysis outputs.
Owner | Product/Business + HF/Usability.
Status | Planned / Evidence Pending.
Notes | All figures Evidence-Pending until measured.
VAL-EVID-004 — Usability reports & the IEC 62366-1 usability engineering file
Objective (user outcome) | The summative human-factors conclusions are consolidated into the regulatory usability engineering file that programmes/regulators rely on.
Intended Use / User | Regulators/auditors, QA/RA, procurement (SH-07/SH-02/SH-08).
Validation Method | Report authoring + UEF assembly.
Participants / Representative Users | HF/Usability lead, EM-ENG-006.
Environment / Context of Use | Design-control documentation.
Success / Acceptance Criteria | Summative reports + UEF complete per IEC 62366-1/FDA HF; residual use-error acceptability documented.
User Need Traceability | UN-VAL-003/006; UN-REG.
Risk / Use-Error Relationship | The evidentiary home of the HA-USE residual closures.
Evidence Required | Summative reports; usability engineering file.
Owner | HF/Usability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | A WHO PQS / regulatory listing input.
VAL-EVID-005 — Field / pilot / audit-simulation reports
Objective (user outcome) | Real-world validation is documented so the release and PMS decisions rest on actual field evidence.
Intended Use / User | Release decision (Stage 16); PMS (Stage 17); auditors.
Validation Method | Field/pilot/audit report authoring.
Participants / Representative Users | Field/clinical/QA teams.
Environment / Context of Use | Deployments and audit simulations.
Success / Acceptance Criteria | Each field/pilot/audit activity has a complete report incl. both-direction safety signals.
User Need Traceability | UN-VAL-002/006.
Risk / Use-Error Relationship | Feeds VAL-RISK residual/benefit-risk validation.
Evidence Required | Field/pilot/audit reports; safety-signal logs.
Owner | Clinical/Operational + Product.
Status | Planned / Evidence Pending.
Notes | Bridges to Stage-17 PMS.
VAL-EVID-006 — Residual-risk / benefit-risk & signed acceptance records
Objective (user outcome) | Every residual acceptance and the overall benefit-risk conclusion is a signed, authority-appropriate, retrievable record.
Intended Use / User | EM-ENG-001/006; regulators.
Validation Method | Determination memos + sign-off records.
Participants / Representative Users | EM-ENG-006, EM-ENG-001, discipline owners.
Environment / Context of Use | Release governance.
Success / Acceptance Criteria | Each VAL-RISK conclusion and the overall benefit-risk carries a signed record by the HAL-2200 §7 authority; retained per records control.
User Need Traceability | UN-VAL-005; VAL-GOV-004.
Risk / Use-Error Relationship | The acceptance evidence for HAL-2800/2900.
Evidence Required | Signed residual/benefit-risk records; retention register.
Owner | EM-ENG-006 + EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | Closes the loop from evidence to accepted residual.
7. Traceability notes
Every VAL- activity's Evidence Required field points here; VAL-EVID-004/006 are the regulatory homes of the usability engineering file and the residual/benefit-risk acceptance records. No result enters any module or PVaL-1200 without a record in this module.
8. Acceptance / summary notes
Pre-execution, the evidence base is empty by design — its structure, control, and honesty rules are defined so that when studies run, their records are pre-registered, ALCOA+, root-caused, and signed. All activities are Planned / Evidence Pending.
VAL-FIELD-900 (PVaL-900) — Field Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-900, activity code VAL-FIELD. Owner: Product + Clinical/Operational, with HF/Usability, Reliability, and EM-ENG-006. This module validates the SCM in the real world at increasing scale — pilot deployment, beta program, and production trial in real LMIC/GDP distribution lanes — and captures installer, service, and support feedback from actual operations. Field validation is where the summative/lab conclusions meet reality: the last-mile disposition, the alerting-under-real-connectivity, the reuse return loop, and the reliability-in-use all get their ecological confirmation here. Status: Planned / Evidence Pending — no field deployment has occurred; no field result is asserted.
---
1. Purpose
Confirm, with real users in real distribution, that the SCM delivers the dominant outcome and its supporting outcomes at deployment scale — progressively de-risking from a small controlled pilot, through a wider beta, to a production trial on representative commercial and LMIC lanes. Field validation is the setting where the two carried-Open unknowns (D3 cold-endurance, RF-in-container) show their true behavior, where alarm fatigue and override pressure are real, and where the reuse economics are actually measured.
2. Scope
Staged real-world deployment: (1) controlled pilot (few lanes, close observation); (2) beta program (wider, representative user classes and regions); (3) production trial (near-release scale on real LMIC/GDP lanes). Across all stages: last-mile disposition outcomes, alerting/escalation under real connectivity, in-container RF behavior, reuse return loop and turnaround, reliability-in-use, and installer/service/support feedback. Feeds the release decision (Stage 16) and the post-market surveillance plan (Stage 17).
3. Applicable standards & references
• ISO 13485 §7.3.7 (validation under actual use conditions); FDA HF (residual-risk in real use); WHO PQS E006 / WHO-UNICEF EVM (field performance); EU GDP (commercial-lane practice); ISO 14971 §9 (production/post-production information).
Upstream | UN-VAL-002/007/008/009; HAL-2800/2900 (residuals, disclosure/monitoring obligations); PV-REL-001/005 (D3/RF); all summative/operational conclusions this module confirms in the field.
4. Governance & approach
• Staged, gated de-risking. Each stage gates the next: the pilot must show no unrecovered critical use error before beta; beta before production trial. Safety-critical field signals (missed excursion, wrong disposition, thermal event) halt/roll back per governance.
• Real conditions, not demos. Field validation uses real trips, real users, real connectivity, real pressure — the on-receipt verdict and alerting are validated as users actually experience them.
• D3/RF confirmed in the field. The endurance and RF-in-container behaviors are observed under real deep- cold and in-container conditions; a field contradiction of the lab result is a release-blocking finding.
• Both harm directions monitored as first-class field metrics. Missed-excursion signals AND false-reject/ wastage/stock-out rates are tracked from the first pilot (feeds ISO 14971 §9 loop, HAL-2900 condition 7).
• Feedback loops. Installer, service, and support feedback is captured structurally, not anecdotally, and fed to design/CAPA and the Stage-17 PMS plan.
5. Validation categories
Pilot deployment · beta program · production trial · LMIC last-mile lanes · GDP commercial lanes · installer feedback · service/reverse-logistics feedback · support feedback · field safety-signal monitoring.
6. Validation Register (VAL-FIELD activities)
VAL-FIELD-001 — Controlled pilot: dominant outcome holds in real distribution ★
Objective (user outcome) | On a small set of closely observed real lanes, users correctly detect, understand, and act on excursions — no compromised dose released, no needless wastage — confirming the summative/HF conclusions in the field.
Intended Use / User | Real last-mile, program, distributor users on pilot lanes (SH-05/SH-01/SH-03).
Validation Method | Pilot deployment (field study, close observation).
Participants / Representative Users | Representative users on ≥2–3 real lanes incl. last-mile.
Environment / Context of Use | Real trips end-to-end incl. cold/offline/in-container segments.
Success / Acceptance Criteria | No unrecovered critical use error; disposition correctness in the field consistent with the summative studies; any safety signal captured and root-caused; gates the beta.
User Need Traceability | UN-VAL-002/003; UN-VIS-001; UN-SAFE-001.
Risk / Use-Error Relationship | Field confirmation of HA-RES-014 and HAL-2900 condition 1.
Evidence Required | Pilot protocol/report; field disposition-correctness; safety-signal log.
Owner | Clinical/Operational + Product.
Status | Planned / Evidence Pending.
Notes | The first real-world confirmation of the device's reason to exist; gates scale-up.
VAL-FIELD-002 — Beta program: representative user classes & regions at wider scale
Objective (user outcome) | Across a wider, representative set of users and regions, outcomes hold without a class- or region-specific failure emerging.
Intended Use / User | All primary field user classes across regions.
Validation Method | Beta field program.
Participants / Representative Users | Representative organizations across LMIC/GDP regions.
Environment / Context of Use | Broader real deployment.
Success / Acceptance Criteria | Outcomes sustained across classes/regions; no disproportionate failure/ return cluster; issues fed to design/CAPA; gates the production trial.
User Need Traceability | UN-VAL-002; UN-ENV; UN-VIS.
Risk / Use-Error Relationship | Aggregate/field manifestation of the residual set (HAL-2800).
Evidence Required | Beta report; class/region-stratified outcomes; issue/CAPA log.
Owner | Product + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Widens representativeness before near-release scale.
VAL-FIELD-003 — Production trial on real LMIC/GDP lanes (near-release scale)
Objective (user outcome) | At near-release scale on real commercial and LMIC lanes, the device delivers its outcomes reliably and is ready to release.
Intended Use / User | Full user chain at scale.
Validation Method | Production trial (field validation at scale).
Participants / Representative Users | Representative production-scale deployment.
Environment / Context of Use | Real LMIC last-mile and GDP commercial lanes.
Success / Acceptance Criteria | Outcomes, reliability-in-use, and reuse loop hold at scale; open field issues within release tolerance; feeds the release recommendation.
User Need Traceability | UN-VAL-002/007/008/009; UN-VIS.
Risk / Use-Error Relationship | Final field confirmation of the overall residual (HAL-2900) before release.
Evidence Required | Production-trial report; scale outcome/reliability/return data.
Owner | Product + Clinical/Operational + Reliability.
Status | Planned / Evidence Pending.
Notes | The last field gate before Stage-16 release.
VAL-FIELD-004 — Last-mile LMIC lane: disposition & offline verdict in the field
Objective (user outcome) | Real last-mile health workers on LMIC lanes reach correct dispositions using the offline on-device verdict.
Intended Use / User | Last-mile health workers, programmes (SH-05/SH-01); beneficiary.
Validation Method | Field study at LMIC last-mile sites (subset across stages).
Participants / Representative Users | Representative last-mile workers.
Environment / Context of Use | Rural health posts / mobile teams; offline, low-light, cold.
Success / Acceptance Criteria | Correct offline dispositions at last-mile sites; performance matches or approaches central sites (ties VAL-CLIN-009).
User Need Traceability | UN-STK-006; UN-ENV-007; UN-VAL-003.
Risk / Use-Error Relationship | HA-RES-014 in the hardest real context; HAL-2900 §5.4 local floor.
Evidence Required | Last-mile field report; site-stratified disposition data.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | The device's most consequential real-world use, validated where it matters most.
VAL-FIELD-005 — GDP commercial lane: alerting & escalation under real connectivity
Objective (user outcome) | On real GDP commercial lanes, alerts reach responders and escalate under real connectivity, and the on-receipt verdict compensates where the link fails.
Intended Use / User | Distributors, 3PLs, QA (SH-03/SH-04/SH-02).
Validation Method | Field study on GDP lanes.
Participants / Representative Users | Representative commercial distribution operations.
Environment / Context of Use | Real commercial lanes; variable/in-container connectivity.
Success / Acceptance Criteria | Alert delivery/escalation characterized per lane; responders act within window where connectivity allows; on-receipt verdict catches excursions where alerts could not; RF behavior consistent with PRD-VER-005.
User Need Traceability | UN-VAL-008; UN-ENV-006; UN-FUN-006/007.
Risk / Use-Error Relationship | HA-COMM-001/002/003; field validation of HA-RES-007/008/009.
Evidence Required | GDP-lane report; per-lane delivery/latency/escalation data.
Owner | Product + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Where the RF-in-container Open item meets reality; shapes the real-time-alert claim.
VAL-FIELD-006 — Reuse return loop & cost-per-trip measured in the field
Objective (user outcome) | Devices actually come back, turn around, and redeploy — the reuse loop and its economics are confirmed by real return data.
Intended Use / User | Service/reverse-logistics, finance, procurement (SH-13/SH-10/SH-08).
Validation Method | Reuse/return field pilot (ties VAL-BIZ-006 / VAL-USE-017).
Participants / Representative Users | Representative return loops across regions.
Environment / Context of Use | Real return logistics (LMIC last-mile + GDP).
Success / Acceptance Criteria | Return rate, turnaround, and cost-per-trip measured and within viability thresholds; region-specific (conflict C6); all figures Evidence-Pending.
User Need Traceability | UN-VAL-009; UN-SVC; UN-BIZ; D1.
Risk / Use-Error Relationship | No direct hazard; the economics that scale the safety mission.
Evidence Required | Return-loop field data; return-rate/turnaround/cost-per-trip by region.
Owner | Product/Business + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | The field source of the reuse-economics validation (VAL-BIZ-006).
VAL-FIELD-007 — Installer / OEM-integrator field feedback
Objective (user outcome) | Installers in the field achieve secure fit and correct placement across real shipper/container/OEM variants, and their feedback improves guidance.
Intended Use / User | Installers/OEM integrators (SH-12).
Validation Method | Field feedback capture + installation observation.
Participants / Representative Users | Representative field installers.
Environment / Context of Use | Real installations across variants.
Success / Acceptance Criteria | Fit/placement outcomes acceptable across variants; feedback structured and fed to design/guidance; no install-type placement-error cluster.
User Need Traceability | UN-STK-011; UN-FLOW-003; UN-FUN-013.
Risk / Use-Error Relationship | HA-USE-005/006 across real installs; ties VAL-REL-006.
Evidence Required | Installer feedback log; field placement-error data.
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Fit issues block the OEM-embed channel.
VAL-FIELD-008 — Service / reverse-logistics field feedback
Objective (user outcome) | Service/reconditioning operations in the field can clean, recalibrate, recharge, and re-provision at throughput, and their feedback improves the loop.
Intended Use / User | Service/reverse-logistics, technicians (SH-13/SH-06).
Validation Method | Field feedback capture + depot observation.
Participants / Representative Users | Representative depot/reconditioning operations.
Environment / Context of Use | Real service depots.
Success / Acceptance Criteria | Reconditioning throughput and quality acceptable; cleaning/recal/recharge workflows usable; feedback fed to design/CAPA.
User Need Traceability | UN-STK-010/012; UN-SVC; UN-FLOW-007/008/009.
Risk / Use-Error Relationship | HA-USE-009/010/011 in the field; ties VAL-REL-005/008.
Evidence Required | Service feedback log; reconditioning throughput/quality data.
Owner | Product + Reliability.
Status | Planned / Evidence Pending.
Notes | Turnaround sets fleet utilization; feeds reuse economics.
VAL-FIELD-009 — Support feedback & issue-to-CAPA loop
Objective (user outcome) | Support experiences in the field are captured, root-caused, and drive design/ process improvement — customers get their issues resolved.
Intended Use / User | Support ops, all customer classes.
Validation Method | Support feedback analysis (field).
Participants / Representative Users | Support operations and customers during deployment.
Environment / Context of Use | Live support across stages.
Success / Acceptance Criteria | Support issues captured, themed, and fed to CAPA; recurring drivers addressed; support burden within envelope (ties VAL-BIZ-004).
User Need Traceability | UN-BIZ; UN-SVC; UN-VAL-002.
Risk / Use-Error Relationship | Support signals can surface latent use errors/hazards (feeds Stage-17).
Evidence Required | Support-feedback analysis; CAPA linkage.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Establishes the field feedback→CAPA loop before release.
VAL-FIELD-010 — Field safety-signal monitoring (both harm directions) → PMS handoff
Objective (user outcome) | From the first pilot, field safety signals in both directions — missed excursion / wrong disposition AND false-reject / wastage / stock-out — are monitored as first-class metrics.
Intended Use / User | Programmes, QA/RA, EM-ENG-006; protects beneficiary.
Validation Method | Field safety-signal monitoring (feeds ISO 14971 §9 / Stage-17 PMS).
Participants / Representative Users | Field operations and safety/quality reviewers.
Environment / Context of Use | All field stages.
Success / Acceptance Criteria | Both-direction signals collected, reviewed, and fed to the residual/ benefit-risk revalidation and the PMS plan; disclosure/monitoring obligations from HAL-2800 operationalized.
User Need Traceability | UN-VAL-002/005; HAL-2900 condition 7.
Risk / Use-Error Relationship | Operationalizes the HAL-2800 disclosure/monitoring obligations and keeps the benefit-risk conclusion valid (ISO 14971 §9).
Evidence Required | Safety-signal monitoring plan; both-direction metric definitions; PMS handoff.
Owner | EM-ENG-006 + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | The bridge from validation to post-market surveillance (Stage 17).
7. Traceability notes
VAL-FIELD stages the real-world confirmation of every prior module's conclusions: the dominant outcome (VAL-FIELD-001/004), alerting/RF (VAL-FIELD-005, ties VAL-REL-004/PRD-VER-005), reuse economics (VAL-FIELD-006, ties VAL-BIZ-006), and the installer/service/support feedback loops. VAL-FIELD-010 operationalizes the HAL-2800 disclosure/monitoring obligations and hands to Stage-17. All trace to UN-VAL and feed the release decision (Stage 16) and PMS (Stage 17).
8. Acceptance / summary notes
Field validation de-risks in stages, halts on safety signals, confirms the D3/RF behaviors in the real world, and monitors both harm directions from the first pilot. All activities are Planned / Evidence Pending; no field outcome, delivery rate, return rate, or safety signal is asserted before deployment.
VAL-GOV-000 (PVaL-000) — Validation Governance
Product Validation Library (PVaL) — Smart Cold Chain Monitor (SCM) · Stage 14 (ORCH-015) · Module PVaL-000, activity code VAL-GOV. Owner: Principal Human-Factors / Design-Validation / Clinical-or-Operational + Product engineer, with Quality/Regulatory (EM-ENG-006) and Chief Systems (EM-ENG-001). This module states the SCM's validation philosophy, the intended use and users the validation must confirm, the outcome-first principles, the acceptance and approval authority, and the honesty posture. It follows 00-governance/validation-standard.md and the 13-module PVaL map, and is based on 00-concept/concept-and-decisions.md (D1–D4), the Stage-04 User Needs, the Stage-10 Hazard Analysis (HAL-1500 use-related, HAL-2800/2900 residual), the Stage-12 Risk-Control Requirements, and the Stage-13 Verification Library. Status: authored playbook — every activity is Planned / Evidence Pending; no product has been built and no users have been tested. No results are recorded or fabricated.
---
1. Purpose
Verification (Stage 13) asks did we build the product correctly? Validation asks the different, harder question: did we build the correct product? — does the SCM meet the user needs and intended use of real people in the real vaccine cold chain. This module governs that question for the SCM. The governing outcome the whole library exists to confirm is stated once, here, and every downstream module traces to it:
The dominant validation outcome (the SCM's reason to exist): representative users — including the
minimally trained, gloved, offline last-mile health worker — can correctly detect, understand, and act on
temperature excursions so that no temperature-compromised vaccine is administered and no potent vaccine is
needlessly discarded, in representative environments, with the residual risks of the device acceptable to
those users.
Validation is outcome-based. Every objective in this library describes what a user must be able to achieve in the world, not what the device implements. Where verification proves a control works (e.g. "the cold display meets a contrast/response-time budget at −30 °C", PV-USE-001), validation proves the human outcome the control exists for (e.g. "a low-training nurse, gloved and cold, reads the verdict correctly and cannot be led to the wrong disposition", VAL-USE-001).
2. Scope
All SCM SKUs (reusable, multi-trip) across the full multi-market distribution chain and lifecycle — from manufacturer pack-out through distributors, 3PLs, air/ground transport, national immunization programs, and the last-mile point of use, and back through return, reconditioning, recalibration, and re-provisioning (D1 reusable). The device + cloud dashboard + mobile/BLE app are all in scope (the ecosystem is validated as a system, PVaL-600). Validation covers primary and secondary workflows, foreseeable misuse, emergency and recovery situations, maintenance/installation, human-factors critical tasks, operational effectiveness, and residual-risk acceptability.
Out of scope (and where it lives instead): the engineering test evidence that a design output meets a design input — that is verification (Stage 13, PVL); the executed validation records themselves (observations, videos, transcripts, survey data, usability reports) — those are validation evidence, referenced from PVaL-1100 and each entry's Evidence Required field, kept distinct from this playbook.
3. Applicable standards & references
• ISO 13485:2016 §7.3.7 — design and development validation (representative product, representative use, representative users; performed before design transfer / release).
• FDA Human Factors guidance (Applying Human Factors and Usability Engineering to Medical Devices) — formative + summative human-factors validation of critical tasks; the SCM adopts this posture because its most safety-critical interaction is performed by its least-trained user.
• IEC 62366-1 — usability engineering: use specification, hazard-related use scenarios, and summative usability evaluation of critical tasks; the twelve HA-USE critical tasks (HAL-1500) become the twelve VAL-USE summative activities.
• ISO 14971 §7.4 / §8 — validation of residual-risk acceptability and the overall benefit-risk conclusion with representative users (PVaL-1000 validates the HAL-2800 individual residuals and the HAL-2900 overall benefit-risk conditions).
• WHO PQS E006 / WHO-UNICEF EVM — the public-health benefit frame, the operating context of use, and the cold-chain-monitor performance the benefit claim is judged against (exact PQS category/revision Evidence-Pending, D4).
• EU GDP / 21 CFR Part 11 / ALCOA+ — the record-defensibility outcome validated by audit simulation.
Upstream | Stage-04 User Needs (UN-VAL-001…009 seed this library; UN-SAFE, UN-UX, UN-FUN, UN-JTBD, UN-FLOW, UN-ENV, UN-STK, UN-PER supply the outcomes and users); Stage-10 HAL-1500 (HA-USE-001…012 use errors), HAL-2800 (HA-RES individual residuals), HAL-2900 (overall benefit-risk conditions 1–7); Stage-12 RCR register (VAL- targets, esp. VAL-USE-001…012); Stage-13 PVL (verification/validation distinction).
4. Governance & approach
• Outcome-first, user-centered. Every activity is authored from the user's job and lived context (UN-JTBD, UN-FLOW, UN-ENV), not from a requirement or design output. If an objective could be satisfied by inspecting the design rather than observing a user, it belongs in verification, not here.
• Representative on all three axes (ISO 13485 §7.3.7). Validation uses representative product (production-equivalent DVT/PVT units and the release-candidate cloud/app), representative users (the actual user classes SH-01…SH-15, with the low-training last-mile user deliberately over-sampled), and representative environments (deep cold, low light, gloved, offline, in-container, LMIC/GDP lanes) — never a lab convenience sample of engineers in an office.
• Severity-first (CLAUDE.md rule 3). Validation effort and acceptance rigor follow Action Priority, not volume. The excursion-response critical task and the residual-risk acceptability of the S4/S5 hazards dominate; convenience/UX-satisfaction validations are secondary.
• The warnings-challenge rule carries into validation (HAL-2200 §6.4). For the low-training last-mile user, training/labels/IFU are never the primary control — so a validation that merely confirms "users who read the manual succeed" is insufficient for a critical task. Summative usability must confirm the outcome for the untrained/decay-of-training user, because that is who will use it.
• Formative then summative. Formative studies (early, iterative, diagnostic) de-risk the design; the summative studies (final, on representative product, per IEC 62366-1 / FDA HF) are the ones that carry acceptance. This library defines both and marks which activities are summative-critical.
• Evidence honesty (hallucination-policy.md). Every acceptance criterion is stated as an outcome target to be met; where the target value is not yet set it is Evidence-Pending and named for closure in PRD/protocol. No use-error rate, satisfaction score, NPS, return rate, or field result is asserted — none has been measured. Every activity's Status is Planned / Evidence Pending.
5. Validation categories (governance-level)
The library is organized into the 13 PVaL modules. Governance recognizes four acceptance tiers that determine authority and rigor:
1. Summative human-factors critical tasks (the excursion detect→understand→act chain) — highest rigor; IEC 62366-1 summative; acceptance by HF lead + EM-ENG-006 + EM-ENG-001 with benefit-risk.
2. Residual-risk / benefit-risk validation (PVaL-1000) — confirms the HAL-2800/2900 conclusions hold with real users; acceptance by EM-ENG-006 + EM-ENG-001.
3. Operational / clinical effectiveness (PVaL-500) and ecosystem/reliability/field (600/700/900) — confirms the device does its public-health job in real distribution; acceptance by Clinical/Operational + Product + EM-ENG-001.
4. Business acceptance (PVaL-800) — satisfaction/NPS/retention/TCO; acceptance by Product/Business.
6. Governance Register (VAL-GOV activities)
VAL-GOV-001 — Validation philosophy & the dominant-outcome charter are approved
Objective (user outcome) | Every stakeholder is assured the SCM will be judged by whether real users achieve the dominant outcome (correct excursion detect→understand→act; no compromised dose given, no potent dose wasted), not by a spec sheet.
Intended Use / User | All user classes; the program and regulators who rely on the validation conclusion.
Validation Method | Governance review & sign-off (not a user study).
Participants / Representative Users | HF/Usability lead, Clinical/Operational, Product/Business, EM-ENG-006, EM-ENG-001.
Environment / Context of Use | Design-control review.
Success / Acceptance Criteria | Approved validation charter stating the dominant outcome, the three representativeness axes, severity-first rigor, and the warnings-challenge posture; signed before any protocol executes.
User Need Traceability | UN-VAL-001/002; UN-VIS-001; UN-SAFE-001 (STN-020 severity anchor).
Risk / Use-Error Relationship | Governs validation of the whole HA-USE cluster (HAL-1500) and HAL-2900.
Evidence Required | Approved charter; module map; applicability register.
Owner | HF/Usability (lead) + EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | This charter is the parent of every VAL- activity; downstream modules cite it.
VAL-GOV-002 — Intended use & intended users for validation are confirmed
Objective (user outcome) | The user classes and use environments the validation must represent are agreed, complete, and traceable to the intended-use statement, so no real user or context is validated by proxy when it should be represented directly.
Intended Use / User | The full user set SH-01…SH-15, with the last-mile low-training user (SH-05) and the represented beneficiary (SH-11) explicitly flagged.
Validation Method | Use-specification review (IEC 62366-1 use specification).
Participants / Representative Users | HF/Usability lead, Clinical/Operational, Product; reviewed by EM-ENG-006.
Environment / Context of Use | Documented use environments UN-ENV-001…009 (deep cold, low light, gloved, offline, in-container, humid, LMIC last-mile, depot).
Success / Acceptance Criteria | Approved use specification mapping every user class to at least one validation activity and every safety-critical workflow to a summative activity; gaps recorded, none silent.
User Need Traceability | Intended-Use Statement (UN index §Intended Use); UN-STK-001…015; UN-PER-001…011; UN-ENV-001…009; UN-VAL-002/003.
Risk / Use-Error Relationship | HA-CTX-005 (vulnerable user); anchors HA-USE critical-task coverage.
Evidence Required | Approved use specification; participant-representativeness plan (feeds VAL-PLAN-003).
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | The last-mile user is deliberately over-represented because they perform the most critical task.
VAL-GOV-003 — Validation ≠ verification boundary is enforced
Objective (user outcome) | Stakeholders can trust that "validated" means a real user achieved the outcome, never that a design output merely passed a bench test.
Intended Use / User | Program, regulators, QA/RA relying on the distinction for release/listing.
Validation Method | Register review against the outcome-vs-implementation test.
Participants / Representative Users | HF/Usability lead, EM-ENG-006.
Environment / Context of Use | Design-control review.
Success / Acceptance Criteria | Every activity's Objective is outcome-worded; any implementation-worded item is moved to PVL (Stage 13) or reworded; the coverage matrix cross-references the paired PV- activity.
User Need Traceability | UN-VAL-001; validation-standard §"Validation ≠ verification".
Risk / Use-Error Relationship | Prevents a false-reassurance that residuals are validated when only verified (ties HA-USE-004 posture).
Evidence Required | Reviewed register; PV↔VAL cross-reference table.
Owner | HF/Usability (lead) + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | e.g. PV-USE-001 (cold-display bench legibility) is paired with VAL-USE-001 (user reaches correct disposition) — both required, neither substitutes for the other.
VAL-GOV-004 — Acceptance authority & benefit-risk sign-off model is defined
Objective (user outcome) | Each validation conclusion is accepted by an authority proportionate to its consequence, so no S4/S5 residual is accepted by the wrong hand.
Intended Use / User | All; the release decision (Stage 16).
Validation Method | Governance definition & approval.
Participants / Representative Users | EM-ENG-001, EM-ENG-006, HF/Usability lead, Clinical/Operational, Product/Business.
Environment / Context of Use | Design-control / release-governance review.
Success / Acceptance Criteria | Approved authority matrix: summative critical tasks & residual-risk validation → EM-ENG-006 + EM-ENG-001 + documented benefit-risk; operational/ecosystem → Clinical/Operational + Product + EM-ENG-001; business → Product/Business; the S5 aircraft-fire and security residuals carry the most stringent sign-off (per HAL-2200 §7).
User Need Traceability | UN-VAL-005; UN-STK-007 (auditor); UN-STK-014 (beneficiary, represented).
Risk / Use-Error Relationship | Governs acceptance of HA-RES-001…022 and HAL-2900 conditions 1–7.
Evidence Required | Approved acceptance-authority matrix; benefit-risk record template.
Owner | EM-ENG-001 + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Mirrors HAL-2200 §7 acceptance authority so hazard closure and validation acceptance agree.
VAL-GOV-005 — Evidence-honesty & no-fabrication rule for validation results
Objective (user outcome) | Every reader can trust that a stated validation result is a real, dated observation of a real user — never an assumed or invented number.
Intended Use / User | All; regulators/auditors especially.
Validation Method | Evidence-governance rule + review gate.
Participants / Representative Users | EM-ENG-006, HF/Usability lead.
Environment / Context of Use | Design-control review; evidence intake (PVaL-1100).
Success / Acceptance Criteria | All acceptance criteria are outcome-based and pre-registered; all targets not yet set are labelled Evidence-Pending with a named protocol; no result appears in any module until its evidence record exists in PVaL-1100; Status of every activity is Planned/Evidence-Pending pre-execution.
User Need Traceability | UN-VAL-001; evidence-standard.md, hallucination-policy.md.
Risk / Use-Error Relationship | Prevents a false-reassurance of validated safety (the meta-form of HA-USE-004).
Evidence Required | Evidence-review checklist; pre-registration of summative protocols.
Owner | EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Pre-registration of the summative critical-task protocols (endpoints, criteria) before execution.
VAL-GOV-006 — Validation feeds release, field, and traceability correctly
Objective (user outcome) | The release decision-maker and the post-market team receive a complete, honest validation conclusion — coverage, user acceptance, residual concerns, and the go/no-go recommendation.
Intended Use / User | Chief Systems (release, Stage 16); Field/PMS team (Stage 17); traceability (ORCH-018 / EM-ENG-012).
Validation Method | Governance handoff definition.
Participants / Representative Users | EM-ENG-001, EM-ENG-006, Product, Field/PMS lead.
Environment / Context of Use | Stage-gate governance.
Success / Acceptance Criteria | PVaL-1200 summary + coverage matrix are release-stage deliverables; the disclosure/monitoring obligations from residual validation feed the Stage-17 PMS plan; every VAL- ID resolves in the ORCH-018 matrix.
User Need Traceability | UN-VAL-002/005/006; UN index §Traceability.
Risk / Use-Error Relationship | Carries HAL-2900 conditions 1–7 into the release gate.
Evidence Required | Handoff map to Stages 16/17/18; open-conditions register.
Owner | EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | Feeds Release Stage 16, Field Stage 17, and ORCH-018 / EM-ENG-012.
7. Traceability notes
This module is the parent of the library. It binds the dominant outcome (VAL-GOV-001) to the intended-use users/environments (VAL-GOV-002), enforces the outcome-vs-implementation boundary (VAL-GOV-003), sets acceptance authority (VAL-GOV-004) aligned to HAL-2200 §7, fixes the honesty posture (VAL-GOV-005), and wires the outputs to release/field/traceability (VAL-GOV-006). Backward it traces to UN-VAL, UN-VIS, UN-SAFE, HAL-1500/2800/2900, and the RCR VAL- targets; forward to PVaL-100…1200, the Coverage Matrix, Stage 16, Stage 17, and ORCH-018.
8. Acceptance / summary notes
No validation has executed. The governance posture is: outcome-first, representative on all three axes, severity-first, warnings-challenge-aware, formative-then-summative, and evidence-backed/honest. The single non-negotiable is that the dominant-outcome critical tasks (VAL-USE-001/002/004/007, VAL-HF summative) and the residual-risk validation (PVaL-1000) close acceptably before an unconditional benefit-risk conclusion (HAL-2900) can be signed for release.
VAL-HF-400 (PVaL-400) — Human-Factors Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-400, activity code VAL-HF. Owner: HF/Usability lead. This module holds the human-factors validation that underpins the intended-use critical tasks: the summative human-factors validation of the excursion-response critical task, the cold/low-light/gloved display legibility outcome, alarm comprehension vs alarm fatigue, correct excursion response, use-error avoidance, accessibility, and training/decay. Where VAL-USE frames each critical task as a workflow, VAL-HF frames the underlying human-perception/cognition/action outcome and consolidates the IEC 62366-1 summative usability engineering evidence. Status: Planned / Evidence Pending — no HF study executed.
---
1. Purpose
Validate that the SCM's human interface produces the right human outcome for the user as they really are — minimally trained, gloved, cold, in poor light, offline, possibly low-literacy or reading a non-primary language, deciding once and acting. This module is the home of the summative human-factors validation of the excursion-response critical task (detect → understand → act) demanded by the brief, and of the perceptual/ cognitive outcomes (legibility, comprehension, alarm salience, error avoidance) that the VAL-USE workflows depend on. Per the warnings-challenge rule, training and labels are validated only as supplements; the design must carry the outcome.
2. Scope
The on-device HMI (OLED verdict/status, button, status LED), the mobile/BLE provisioning & disposition app, and the cloud alert presentation — validated for: at-a-glance verdict comprehension; cold/low-light/gloved legibility; language-independent (icon+colour+word) understanding; safe-cold vs damaging-frozen distinction; alarm salience/prioritization and fatigue resistance; correct excursion response; feedback of action success; error avoidance (mis-arm, wrong profile, accidental override); accessibility (literacy/language/vision/age); and the effect (and decay) of training. Consolidates the IEC 62366-1 usability engineering file.
3. Applicable standards & references
• IEC 62366-1 — usability engineering process and summative evaluation of critical tasks; use-error root-cause and residual acceptability; the usability engineering file.
• FDA Human Factors guidance — critical-task summative validation, ≥15 users/group, simulated-use realism.
• ISO 14971 — use-related residual risk. IEC 60601-1-8 (alarm principles, referenced) — alarm distinctiveness/prioritization concepts applied to the excursion alert.
Upstream | UN-UX-001…012, UN-SAFE-004/005, UN-VAL-003/004; HAL-1500 (HA-USE cluster); DO-HF-001…010 (HMI design outputs); PV-USE-001/002/008/009 (paired verification).
4. Governance & approach
• Summative is the acceptance gate. Formative HF studies iterate the design; the summative studies (final design, representative product, representative users/environments) carry acceptance and feed the usability engineering file and the benefit-risk conclusion.
• Perception + cognition + action, decomposed. For the excursion critical task, HF validates each link — can the user perceive the state (legibility), comprehend it (icon/colour/word, safe-cold vs frozen), and act correctly (disposition/response) — so a failure is diagnosable, not just a pass/fail.
• Fatigue and automation-bias as design outcomes. Alarm fatigue (over-muting) and automation bias (over-trust of a clean record / distrust of a reject) are validated as design properties, not user failings.
• Training-decay honesty. Because real users' training decays, HF validates the untrained/decayed user; training benefit is measured but never credited as the primary control.
5. Validation categories
Summative critical-task HF · display legibility (cold/low-light/gloved) · comprehension & language-independence · alarms & fatigue · excursion-response action · error avoidance · accessibility · training & decay · mobile/ app HF.
6. Validation Register (VAL-HF activities)
VAL-HF-001 — Summative human-factors validation of the excursion-response critical task ★ (detect → understand → act)
Objective (user outcome) | Representative users perceive the excursion/verdict state, understand it correctly, and take the correct disposition/response action — the complete safety-critical chain — with no unrecovered critical use error.
Intended Use / User | Last-mile health worker & responders (SH-05/SH-03/SH-01); protects beneficiary.
Validation Method | Summative usability validation (IEC 62366-1 / FDA HF) — the module's anchor study; consolidates VAL-USE-001/002/004/007 into the critical-task HF conclusion.
Participants / Representative Users | ≥15 representative users per distinct user group; low-training, gloved, low-literacy/non-primary-language, low-vision included.
Environment / Context of Use | Simulated use in cold/low-light/gloved/offline with realistic alert load and time pressure.
Success / Acceptance Criteria | Across scripted clean/freeze/heat/gapped/borderline scenarios, users complete detect→understand→act correctly; every critical use error is root-caused and judged acceptable per IEC 62366-1; no residual critical use error reaches harm without recovery; targets pre-registered, Evidence-Pending.
User Need Traceability | UN-VAL-003; UN-SAFE-001/004/005; UN-UX-001/002/003; UN-FUN-008.
Risk / Use-Error Relationship | The consolidated HF closure of HA-USE-001/002/004/007 → HA-RES-014; validates HAL-2900 conditions 1 (supported human decision).
Evidence Required | Summative usability report; usability engineering file; per-link (perceive/comprehend/ act) and per-scenario outcome data; use-error root-cause analysis.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | The brief's required summative HF validation for the critical excursion-response task; the single most important HF activity in the program.
VAL-HF-002 — Verdict/state legibility outcome at −30 °C, low light & glare, gloved
Objective (user outcome) | Users can actually read the verdict/state on the OLED/LED when the device is cold-soaked, in dim or glaring light, with gloves — the perceptual precondition of the critical task.
Intended Use / User | Cold-store operator & last-mile user (SH-05).
Validation Method | Summative usability test (legibility outcome) — paired with PV-USE-001.
Participants / Representative Users | ≥15 representative users incl. low-vision; gloved.
Environment / Context of Use | −30 °C cold-soaked device; dim/bright/direct-sun lighting rigs.
Success / Acceptance Criteria | Users correctly read the verdict/state within a glance-time budget under each condition; no condition produces a systematic misread; legibility thresholds Evidence-Pending (set with PV-USE-001).
User Need Traceability | UN-UX-001/002; UN-ENV-002/008; UN-VAL-004.
Risk / Use-Error Relationship | HA-USE-001, HA-ENV-002; the perceptual link of HA-RES-014.
Evidence Required | Cold/low-light legibility study report; read-accuracy & time by condition.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Verification (PV-USE-001) proves the display meets a physical budget; this proves users read it.
VAL-HF-003 — Language-independent comprehension (icon + colour + word) without training or a manual
Objective (user outcome) | Users understand the verdict/state from the icon+colour+word encoding without reading text, a manual, or training, across literacy levels and languages.
Intended Use / User | Last-mile & operational users across locales (SH-05/SH-01/SH-02).
Validation Method | Summative comprehension test (multi-locale) — paired with PV-USE-008.
Participants / Representative Users | Representative users across ≥3 locales incl. low literacy.
Environment / Context of Use | Field/simulated across markets.
Success / Acceptance Criteria | Comprehension achieved via the encoding without text reliance; no icon is systematically misunderstood in any locale; comprehension threshold Evidence-Pending.
User Need Traceability | UN-UX-003; UN-SAFE-005; UN-VAL-003; HA-IUE-010.
Risk / Use-Error Relationship | HA-USE-008; validates the language-independent control (RCR-043) with its intended users.
Evidence Required | Multi-locale comprehension report; per-icon/per-locale results.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | The cognitive link of the critical task; the warnings-challenge design bet.
VAL-HF-004 — Distinguish safe-cold from damaging-frozen (and heat) states
Objective (user outcome) | Users do not confuse a safe cold reading with a damaging freeze, so a freeze-sensitive vaccine is correctly protected without needless rejection of a safely-cold one.
Intended Use / User | Last-mile & program users (SH-05/SH-01).
Validation Method | Summative comprehension/decision test.
Participants / Representative Users | Representative low-training users.
Environment / Context of Use | Receipt disposition; freeze vs safe-cold scenarios.
Success / Acceptance Criteria | Users correctly distinguish safe-cold, frozen, and overheated states and disposition accordingly; no systematic freeze/cold confusion.
User Need Traceability | UN-UX-008; UN-SAFE-002; UN-FUN-004.
Risk / Use-Error Relationship | Two-sided detection comprehension; ties HA-SW-001, HA-RES-019 (avoid false-reject of safe-cold).
Evidence Required | Comprehension/decision report; freeze/cold/heat confusion rates.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Guards both harm directions — freeze damage and needless wastage.
VAL-HF-005 — Alarm salience, prioritization & fatigue resistance
Objective (user outcome) | A genuine excursion alert is salient and distinguishable from low-value ones, and users do not habituate into muting real alerts.
Intended Use / User | Responders/coordinators (SH-03/SH-01/SH-05).
Validation Method | Summative usability test under representative alert load — paired with PV-USE-002.
Participants / Representative Users | ≥15 representative responders.
Environment / Context of Use | Realistic multi-shipment alert load over a sustained session.
Success / Acceptance Criteria | Genuine excursion alerts are noticed and acted on; muting behavior does not defeat a real excursion; consolidation/escalation prevents single-responder dependence; false-alarm tolerance Evidence-Pending.
User Need Traceability | UN-SAFE-004; UN-UX-006; UN-FUN-007.
Risk / Use-Error Relationship | HA-USE-002 (alarm fatigue); validates HA-RES-014/019.
Evidence Required | Alarm-load study report; notice/act/mute rates; escalation observations.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Alarm fatigue validated as a design property (tuning + consolidation + escalation).
VAL-HF-006 — Correct excursion-response action (the "act" link)
Objective (user outcome) | Once an excursion is understood, users take the correct action (segregate/ re-ice/quarantine/reject) rather than a wrong or no action.
Intended Use / User | Responders & clinic staff (SH-03/SH-05/SH-01).
Validation Method | Summative usability test (action outcome).
Participants / Representative Users | Representative responders/clinic staff.
Environment / Context of Use | In-trip and receipt excursion scenarios.
Success / Acceptance Criteria | Correct response action taken within the actionable window; no systematic wrong/omitted action; wrong-action rate acceptable, Evidence-Pending.
User Need Traceability | UN-JTBD-003; UN-FLOW-006; UN-SAFE-001.
Risk / Use-Error Relationship | Completes the detect→understand→act chain; ties HA-USE-002/007.
Evidence Required | Action-outcome report; response correctness/timing.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Distinct from comprehension — a user can understand and still act wrongly under pressure.
VAL-HF-007 — Use-error avoidance: mis-arm, wrong profile, accidental override
Objective (user outcome) | The design prevents the high-consequence slips — arming a void trip, setting a wrong profile, accidentally overriding a reject.
Intended Use / User | Provisioners & disposition decision-makers (SH-02/SH-05/SH-01).
Validation Method | Summative usability test (error-provocation scenarios).
Participants / Representative Users | ≥15 representative provisioners/decision-makers.
Environment / Context of Use | Provisioning and disposition under time/production pressure.
Success / Acceptance Criteria | Mis-arm/void, wrong-profile, and accidental-override slips are prevented or caught (pre-arm gate, profile validation, deliberate-recorded override); no unrecovered critical slip.
User Need Traceability | UN-SAFE-003/006; UN-UX-007/009.
Risk / Use-Error Relationship | HA-USE-003/004/007; validates HA-RES-014.
Evidence Required | Error-provocation study report; slip incidence/recovery.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Validates that design prevention (not warnings) carries these errors.
VAL-HF-008 — Feedback of action success (arm/place/charge/update confirmed)
Objective (user outcome) | Users get clear, unambiguous feedback that an action succeeded, so they do not proceed on a false belief (e.g. armed when not).
Intended Use / User | Operational users & technicians (SH-02/SH-05/SH-06).
Validation Method | Usability test (feedback comprehension).
Participants / Representative Users | Representative operational users.
Environment / Context of Use | Arming, placement, charging, firmware update.
Success / Acceptance Criteria | Users correctly recognize success vs failure of each action; no false-success misinterpretation.
User Need Traceability | UN-UX-009/012; UN-FUN-012.
Risk / Use-Error Relationship | HA-USE-004 (believed armed); ties HA-RES-014.
Evidence Required | Feedback-comprehension report.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | False-success belief is a root cause of the void-trip hazard.
VAL-HF-009 — Accessibility HF: low vision, colour-vision deficiency, aging, gloved dexterity
Objective (user outcome) | Users with low vision, colour-vision deficiency, age-related decline, or heavy gloves still perceive, comprehend, and operate the core tasks.
Intended Use / User | The accessibility tail of last-mile/operational users.
Validation Method | Summative accessibility sub-study (shared with VAL-USER-015).
Participants / Representative Users | Quota-sampled low-vision/CVD/older/gloved users.
Environment / Context of Use | Representative cold/low-light/gloved conditions.
Success / Acceptance Criteria | These users achieve the critical-task outcomes; colour is never the sole channel (icon+word redundancy works); no accessibility group with concentrated critical use error.
User Need Traceability | UN-UX-003; UN-SAFE-005; HA-IUE-010.
Risk / Use-Error Relationship | HA-USE-001/008; validates the redundancy of the encoding.
Evidence Required | Accessibility sub-study report; per-group outcome data.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Colour-only coding would fail CVD users — validated redundancy is the safeguard.
VAL-HF-010 — Training effect & decay; IFU/labelling comprehension (supplement only)
Objective (user outcome) | Training and the IFU/labelling help but the design outcome holds when training has decayed and the manual is unavailable.
Intended Use / User | All operational users; especially untrained last-mile.
Validation Method | Comparative usability study (trained vs untrained vs decayed) + IFU comprehension.
Participants / Representative Users | Representative users at fresh-trained, untrained, and time-decayed states.
Environment / Context of Use | Simulated real use without a manual to hand.
Success / Acceptance Criteria | Critical-task outcomes hold for untrained/decayed users; training measured as a supplement, never required for the critical task; IFU/labels comprehensible where used.
User Need Traceability | UN-SAFE-005; UN-VAL-003; warnings-challenge (HAL-2200 §6.4).
Risk / Use-Error Relationship | Confirms training/IFU are supplements (HAL-1500 governing rule), not the primary control.
Evidence Required | Comparative study report; trained/untrained/decayed outcome deltas; IFU comprehension.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Directly tests the warnings-challenge premise for the SCM's least-trained user.
VAL-HF-011 — Mobile/BLE provisioning & disposition app HF
Objective (user outcome) | Users complete provisioning, profile selection, and any app-side disposition on the mobile/BLE app without error, including on constrained/older devices.
Intended Use / User | Provisioners, technicians, program staff (SH-02/SH-06/SH-01).
Validation Method | Summative usability test (mobile app).
Participants / Representative Users | Representative app users across device types.
Environment / Context of Use | Field provisioning; variable devices/connectivity.
Success / Acceptance Criteria | App tasks completed without critical error; profile selection error-proofed (ties VAL-USE-003); works on representative low-end devices.
User Need Traceability | UN-UX-005/007; UN-FUN-011; UN-ECO (mobile).
Risk / Use-Error Relationship | HA-USE-003 (profile via app); ties HA-RES-014.
Evidence Required | Mobile-app usability report; task-completion/error data.
Owner | HF/Usability (lead) + Product.
Status | Planned / Evidence Pending.
Notes | Shares provisioning outcomes with VAL-USE-003; feeds VAL-ECO mobile validation.
7. Traceability notes
VAL-HF-001 is the summative human-factors validation of the excursion-response critical task required by the brief; it consolidates the VAL-USE critical tasks into the IEC 62366-1 usability engineering conclusion for HA-RES-014 and HAL-2900 condition 1. VAL-HF-002…011 validate the perceptual, cognitive, action, error- avoidance, accessibility, training, and mobile outcomes those tasks rest on, each paired with a PV-USE verification and traced to UN-UX / UN-SAFE. All feed the usability engineering file (PVaL-1100) and the benefit-risk validation (PVaL-1000).
8. Acceptance / summary notes
Human-factors validation is the safety spine of this library. Acceptance requires that the untrained, gloved, cold, low-light, possibly low-literacy last-mile user achieves the excursion detect→understand→act outcome on representative product — with training and labels proven to be supplements only. All activities are Planned / Evidence Pending; no HF outcome or use-error rate is asserted.
VAL-PLAN-100 (PVaL-100) — Validation Planning
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-100, activity code VAL-PLAN. Owner: HF/Usability lead + Product, with EM-ENG-006 and EM-ENG-001. This module defines the validation strategy — objectives, the representative participants, users, environments and facilities, the study phasing and sequencing, the success-criteria philosophy, the deliverables, and the dependencies and open items — so the validation program is planned before any protocol runs. Status: Planned / Evidence Pending.
---
1. Purpose
Turn the governance charter (VAL-GOV) into an executable plan: who we validate with, where, in what order, against what outcome targets, and what must be true first (which verification must pass, which representative product and DVT/PVT units must exist). The plan is deliberately built around the dominant outcome — the excursion detect→understand→act critical task with the last-mile user — with everything else sequenced to protect and inform that summative study.
2. Scope
The whole SCM validation program: formative and summative human-factors studies, intended-use workflow validation, operational/clinical effectiveness, ecosystem validation, long-term reliability-in-use, business validation, field pilots/beta/production trials, and residual-risk/benefit-risk validation. Covers the device, cloud dashboard, and mobile/BLE app as one system, across multi-market LMIC/GDP lanes.
3. Applicable standards & references
• ISO 13485 §7.3.7; FDA Human Factors guidance (formative→summative, critical-task focus, ≥15 representative users per distinct user group for summative); IEC 62366-1 (use scenarios, summative evaluation); ISO 14971 §7.4/§8 (residual/benefit-risk with users); WHO PQS E006 / WHO-UNICEF EVM (operational context and benefit frame).
Upstream | UN-VAL-001…009 (validation needs), UN-ENV-001…009 (environments), UN-STK/UN-PER (users), HAL-1500 (critical tasks), HAL-2800/2900 (residuals/conditions), Stage-13 PVL (verification dependencies).
4. Governance & approach
• Formative → summative gate. No summative critical-task study runs until (a) formative studies have stabilized the design and (b) the paired verification (e.g. PV-USE-001 cold-display legibility) has passed — a summative study on an unverified display would waste representative users and confound the result.
• Representative-product gate. Summative and field validation run on production-equivalent DVT/PVT hardware and the release-candidate cloud/app, not prototypes; formative may use earlier builds.
• Two carried-Open dependencies gate field/reliability validation. The D3 cold-endurance (PRD-VER-004) and RF-in-container (PRD-VER-005) verification items are release-gating and their outcomes shape the field/reliability plan (e.g. an unclosed RF budget defaults the real-time-alert claim to best-effort).
• Both harm directions are first-class. The plan validates both "no compromised dose administered" and "no potent dose needlessly wasted" (the accepted-trade residual HA-RES-019) — false-reject/wastage is a planned endpoint, not an afterthought.
• Outcome targets, honestly pending. Success criteria are outcome-based; numeric targets (use-error rate, time-to-decision, false-reject rate, NPS, return rate) are Evidence-Pending and set in the protocol/PRD, not invented here.
5. Validation categories (planning topics)
Strategy · objectives · participants/representative users · facilities/environments · phasing & sequencing · success-criteria philosophy · deliverables · timeline/dependencies · risk-based prioritization.
6. Validation Register (VAL-PLAN activities)
VAL-PLAN-001 — Validation strategy & objective hierarchy
Objective (user outcome) | The program validates, in priority order, that users achieve the dominant outcome, then that residual risks are acceptable, then that the device is operationally effective, ecosystem- complete, reliable, and commercially adopted.
Intended Use / User | All user classes; the release decision.
Validation Method | Strategy definition & review.
Participants / Representative Users | HF/Usability lead, Clinical/Operational, Product, EM-ENG-006, EM-ENG-001.
Environment / Context of Use | Program planning.
Success / Acceptance Criteria | Approved strategy that ranks objectives by Action Priority, names the summative-critical set (VAL-USE-001/002/004/007; VAL-HF summative), and sequences everything else to inform or follow them.
User Need Traceability | UN-VAL-001/002/003/005; UN-VIS-001; UN-SAFE-001.
Risk / Use-Error Relationship | Prioritizes the HA-USE cluster and the HAL-2900 conditions.
Evidence Required | Approved validation strategy; objective-priority list.
Owner | HF/Usability (lead) + Product.
Status | Planned / Evidence Pending.
Notes | Severity-first: the excursion critical task dominates the budget.
VAL-PLAN-002 — Representative-user recruitment & sampling plan
Objective (user outcome) | Each summative conclusion rests on enough of the right users — especially the minimally trained last-mile worker — that it generalizes to real deployment.
Intended Use / User | SH-01…SH-15, with SH-05 last-mile over-sampled; SH-11 beneficiary represented.
Validation Method | Recruitment/sampling plan (FDA HF ≥15 per distinct user group for summative).
Participants / Representative Users | HF/Usability lead, field/program partners, recruiting.
Environment / Context of Use | LMIC programs, clinics, cold stores, depots, 3PL/distributor operations.
Success / Acceptance Criteria | Approved plan specifying ≥15 representative users per distinct user group for each summative study, quotas covering literacy/language/gloved/vision diversity, and no engineer/ convenience proxies for critical-task users; inclusion of decayed-training users.
User Need Traceability | UN-VAL-002/003; UN-STK-006; UN-PER-007; HA-CTX-005.
Risk / Use-Error Relationship | Under-representing the low-training user would under-estimate P1 of the HA-USE errors — the central validity risk.
Evidence Required | Recruitment plan; user-group definitions; representativeness rationale.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Sample sizes are the FDA HF floor; final counts set per protocol.
VAL-PLAN-003 — Representative environments & context-of-use plan
Objective (user outcome) | Users are validated where they actually work — deep cold, low light, gloved, offline, in-container, humid, dusty last-mile, depot — not in a comfortable lab.
Intended Use / User | All field users; especially receipt/point-of-use and cold-store/dispatch.
Validation Method | Environment/context plan (simulated-use chambers + real field lanes).
Participants / Representative Users | HF/Usability lead, Reliability, field partners.
Environment / Context of Use | Cold chambers to −30 °C, low-light/glare rigs, gloved protocols, RF-hostile in-container mockups, LMIC lanes (UN-ENV-001…009).
Success / Acceptance Criteria | Approved plan mapping every summative critical task to at least one representative environment; cold/low-light/gloved/offline conditions mandatory for the disposition and provisioning tasks.
User Need Traceability | UN-ENV-001…009; UN-UX-001/002; UN-VAL-004.
Risk / Use-Error Relationship | HA-USE-001 (cold/low-light misread), HA-ENV-002.
Evidence Required | Environment plan; chamber/rig specs; lane list.
Owner | HF/Usability (lead) + Reliability.
Status | Planned / Evidence Pending.
Notes | Simulated-use cold/low-light is the minimum; field lanes add ecological validity.
VAL-PLAN-004 — Facilities, equipment & simulated-use apparatus
Objective (user outcome) | The studies can be run credibly — cold chambers, glove kits, low-light rigs, reference thermal sources for scripted excursions, RF-attenuating shipper mockups, and a recording setup.
Intended Use / User | Study operators; supports all summative users.
Validation Method | Facilities/equipment plan.
Participants / Representative Users | HF/Usability lead, test lab, Reliability.
Environment / Context of Use | Usability lab + environmental chambers + field kits.
Success / Acceptance Criteria | Approved facilities list with calibrated reference sources for scripted excursion scenarios, glove/lighting standardization, and validated video/observation capture.
User Need Traceability | UN-VAL-003/004; UN-ENV.
Risk / Use-Error Relationship | Ensures scripted excursion scenarios are realistic and repeatable.
Evidence Required | Facilities/equipment plan; calibration certs for reference sources.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Scripted excursions must be product-relevant (freeze and heat) for VAL-USE-001/CLIN studies.
VAL-PLAN-005 — Study phasing & sequencing (formative → summative → field)
Objective (user outcome) | Findings flow in the right order so the summative studies are clean and the field pilots deploy a validated design.
Intended Use / User | Program governance.
Validation Method | Phasing/sequencing plan.
Participants / Representative Users | HF/Usability lead, Product, EM-ENG-001.
Environment / Context of Use | Program timeline; DVT/PVT gates.
Success / Acceptance Criteria | Approved sequence: (1) formative HF; (2) verification dependencies pass (PV-USE-, PV-SW-, PV-COMM-); (3) summative HF critical tasks; (4) operational/ecosystem/reliability; (5) field pilot → beta → production trial; residual/benefit-risk validation runs alongside (3)–(5).
User Need Traceability | UN-VAL-001/007/008/009.
Risk / Use-Error Relationship | Prevents summative studies on unverified controls (HA-USE cluster).
Evidence Required | Sequenced plan with gate criteria.
Owner | HF/Usability (lead) + EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | D3/RF Open items (PRD-VER-004/005) gate phase (4)–(5).
VAL-PLAN-006 — Success-criteria philosophy (outcome-based, both harm directions)
Objective (user outcome) | Every study is judged by whether users achieved the real-world outcome, with both "no compromised dose" and "no wasted dose" as explicit endpoints.
Intended Use / User | All; regulators reviewing acceptance criteria.
Validation Method | Criteria-framework definition.
Participants / Representative Users | HF/Usability lead, Clinical/Operational, EM-ENG-006.
Environment / Context of Use | Protocol design.
Success / Acceptance Criteria | Approved framework: outcome endpoints (correct disposition, timely action, correct provisioning), critical-task use-error analysis (root-caused, not just counted), and dual false- accept/false-reject endpoints; all numeric targets pre-registered and Evidence-Pending until set in PRD.
User Need Traceability | UN-VAL-001/005; UN-SAFE-001/002/003/004; HA-RES-019 (wastage trade).
Risk / Use-Error Relationship | Ties acceptance to HA-USE residual acceptability, both directions.
Evidence Required | Criteria framework; pre-registration template.
Owner | HF/Usability (lead) + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Per IEC 62366-1, residual use errors are analyzed for acceptability, not required to be zero.
VAL-PLAN-007 — Deliverables & usability engineering file plan
Objective (user outcome) | The validation produces the artifacts programs, regulators, and procurement need to adopt — a usability engineering file, study reports, and the summary/coverage/benefit-risk record.
Intended Use / User | Regulators/auditors (SH-07), procurement (SH-08), QA/RA (SH-02).
Validation Method | Deliverables plan.
Participants / Representative Users | HF/Usability lead, EM-ENG-006, Product.
Environment / Context of Use | Design-control documentation.
Success / Acceptance Criteria | Approved deliverables list: IEC 62366-1 usability engineering file, per-study protocols/reports (PVaL-1100), coverage matrix, benefit-risk validation record, PVaL-1200 summary.
User Need Traceability | UN-VAL-006; UN-STK-007/008; UN-REG.
Risk / Use-Error Relationship | The UEF is the evidentiary home of the HA-USE residual closures.
Evidence Required | Deliverables register; UEF outline.
Owner | EM-ENG-006 + HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | UEF is a WHO PQS / regulatory listing input.
VAL-PLAN-008 — Dependencies, timeline & open-item register
Objective (user outcome) | No summative or field validation is blocked by a missing prerequisite; the two carried-Open engineering unknowns are tracked as validation gates.
Intended Use / User | Program governance; release decision.
Validation Method | Dependency/timeline plan + open-item register.
Participants / Representative Users | EM-ENG-001, HF/Usability lead, Reliability, Product.
Environment / Context of Use | Program planning.
Success / Acceptance Criteria | Approved register listing each validation activity's prerequisites (representative product, passed verification, recruited users), the D3/RF Open items as explicit gates, and a realistic sequence; no activity marked ready whose prerequisites are unmet.
User Need Traceability | UN-VAL-007/008; UN-VAL-009.
Risk / Use-Error Relationship | Carries the HAL-2900 conditions (esp. D3/RF) into the validation timeline.
Evidence Required | Dependency map; open-item register; timeline.
Owner | EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | PRD-VER-004 (D3) and PRD-VER-005 (RF) gate reliability/field validation and the alert claim.
7. Traceability notes
VAL-PLAN operationalizes VAL-GOV. It maps users (UN-STK/UN-PER), environments (UN-ENV), and outcomes (UN-VAL/UN-SAFE) into a phased, resourced program, and carries the Stage-13 verification dependencies and the D3/RF Open items forward as gates. Each downstream module's activities inherit the sampling (VAL-PLAN-002), environment (VAL-PLAN-003), and criteria (VAL-PLAN-006) posture defined here.
8. Acceptance / summary notes
The plan is authored, not executed. Its load-bearing commitments: over-sample the last-mile user; validate in cold/low-light/gloved/offline/in-container conditions; run formative before summative and only summative on verified, representative product; treat wastage/false-reject as a first-class endpoint; and gate field/ reliability validation on the D3 and RF Open items.
VAL-REL-700 (PVaL-700) — Reliability Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-700, activity code VAL-REL. Owner: Product + Reliability, with Clinical/Operational and EM-ENG-006. This module validates that the SCM keeps delivering its user outcomes over long-term real use, across the diversity of real environments, networks, power, and regions it will meet — not just on day one in a lab. It is the in-use complement to the Stage-13 reliability verification (PV-REL, incl. the D3 90-day cold-endurance and RF-in-container Open items): verification proves the design meets a reliability requirement; this module proves real users keep getting the outcome across the fleet, over the reuse life. Status: Planned / Evidence Pending — no long-term use data exists; the D3 and RF Open items gate this module.
---
1. Purpose
Confirm that the dominant outcome (correct excursion detect→understand→act; complete, defensible record) is sustained by real deployed devices over their multi-trip, multi-year reusable life, across the environmental, network, power, installation, and regional diversity of global vaccine distribution — so the device does not quietly stop protecting the beneficiary after a few cold cycles, a firmware update, or a humid lane.
2. Scope
Long-term real-use reliability of the fielded fleet: cold-endurance-in-use (the D3 90-day/−30 °C outcome as users experience it); battery capacity fade and charge-safety over reuse cycles; environmental diversity (deep cold, tropical humidity/condensation, shock/vibration, dust, altitude, UV); network diversity (LTE-M/ Wi-Fi/BLE, RF-hostile in-container, intermittent/blackout); power diversity (mains-scarce depots, cold-charge inhibit); installation diversity (shipper/container types, OEM embeds); regional diversity (LMIC lanes, GDP lanes). Feeds the field program (PVaL-900) and the residual-risk revalidation (PVaL-1000).
3. Applicable standards & references
• IEC 60068 (environmental), IEC 60529 (ingress), CISPR/IEC 61000 + FCC/CE (EMC/radio), UN 38.3/IATA (transport) — the verification basis these in-use validations reference.
• WHO PQS E006 / WHO-UNICEF EVM — field reliability expectations. ISO 14971 §9 — production/post- production information keeps the residual/benefit-risk conclusion valid.
Upstream | UN-ENV-001…009, UN-PERF (endurance/accuracy), UN-VAL-007/008/009; PV-REL-001 (PRD-VER-004 D3) / PV-REL-005 (PRD-VER-005 RF); HA-ENRG-006, HA-SW-004 (calibration drift), HA-RES-004/015.
4. Governance & approach
• In-use, over life, across diversity. Endpoints are the user outcome sustained (verdict/record still correct), measured across reuse cycles and the diversity matrix — not a single lab MTBF.
• D3/RF Open items are gates. Reliability-in-use validation cannot conclude until the D3 cold-endurance (PRD-VER-004) and RF-in-container (PRD-VER-005) verification close; their field manifestation is validated here and both are release-gating (HAL-2900 conditions 3).
• Calibration drift is a fleet common cause. Per HAL-2900 §5.3, systematic calibration drift over reuse is the information chain's dominant common cause — its in-use manifestation and the recal-interval control are validated across the fleet with standing monitoring.
• Honest degradation. Any loss of function over life must surface as an honest gap / not-ready state, never a silent wrong verdict — validated in use.
5. Validation categories
Long-term use · cold endurance in use (D3) · battery fade/charge-safety over cycles · environmental diversity · network diversity · power diversity · installation diversity · regional diversity · calibration stability.
6. Validation Register (VAL-REL activities)
VAL-REL-001 — Cold endurance in use: 90-day / −30 °C outcome as users experience it ★ (D3, gated)
Objective (user outcome) | Deployed devices keep logging, detecting, and giving a correct verdict for the full trip duration in deep cold — users are not left with a dead-in-cold device and a lost record.
Intended Use / User | Distributors, programmes, procurement (SH-03/SH-01/SH-08).
Validation Method | Field/extended reliability study following PV-REL-001 (PRD-VER-004).
Participants / Representative Users | Representative deep-cold lanes/operations.
Environment / Context of Use | −30 °C / long-duration real and simulated trips.
Success / Acceptance Criteria | Endurance target met in use (no dead-in-cold data loss over the trip); any energy shortfall surfaces as an honest low-battery/gap, never a silent wrong verdict; target Evidence- Pending and gated on PRD-VER-004.
User Need Traceability | UN-VAL-007; UN-ENV-001; UN-PERF (endurance); D3.
Risk / Use-Error Relationship | HA-ENRG-006 / HA-SW-013 (silent gap / dead-in-cold); validates HA-RES-004 in use; HAL-2900 condition 3.
Evidence Required | Endurance field/reliability report; energy-margin-in-use data; gap-honesty evidence.
Owner | Reliability + Product.
Status | Planned / Evidence Pending (Open — D3, PRD-VER-004).
Notes | The top program risk closed by test; release-gating.
VAL-REL-002 — Battery capacity fade & charge safety over reuse cycles
Objective (user outcome) | Over many reuse cycles, devices retain enough capacity for a trip and never charge unsafely in the cold — users keep getting a trip-ready device without a safety event.
Intended Use / User | Service/technicians, distributors (SH-06/SH-03).
Validation Method | Cycle-life reliability study + field aging.
Participants / Representative Users | Representative reuse operations over cycles.
Environment / Context of Use | Repeated cold/thermal cycling; depot recharge.
Success / Acceptance Criteria | Capacity fade stays within a usable-trip margin over the rated cycle life; charge-inhibit-below-0 °C holds every cycle; no cold-charge safety event; margins Evidence-Pending.
User Need Traceability | UN-VAL-007; UN-SAFE-007; UN-SVC; D3.
Risk / Use-Error Relationship | HA-ENRG-001 (cold-charge), HA-USE-009 (inhibit misread); validates HA-RES-005 over life.
Evidence Required | Cycle-life data; charge-inhibit verification over cycles; fade curve.
Owner | Reliability.
Status | Planned / Evidence Pending.
Notes | Ties charge-inhibit legibility (VAL-USE-009) — the human side of the same control.
VAL-REL-003 — Environmental diversity: humidity/condensation, shock/vibration, dust, altitude, UV in use
Objective (user outcome) | The device keeps measuring accurately and giving correct verdicts through the full environmental range of real lanes over life.
Intended Use / User | Programmes, distributors, 3PLs (SH-01/SH-03/SH-04).
Validation Method | Field study across environmental-diverse lanes + reliability soak.
Participants / Representative Users | Representative operations in humid/rugged/high-altitude lanes.
Environment / Context of Use | Tropical humidity, freeze-thaw condensation, transport shock/vibration, dust, altitude, UV.
Success / Acceptance Criteria | No measurement error or false verdict induced by environmental exposure over life; ingress/condensation controls hold; no environment-specific failure cluster.
User Need Traceability | UN-ENV-003/004/005; UN-PERF (accuracy).
Risk / Use-Error Relationship | HA-ENV-003/004, HA-TFC-005/006 (condensation/ingress); validates HA-RES-021 in use.
Evidence Required | Environmental field report; accuracy-over-exposure data.
Owner | Reliability.
Status | Planned / Evidence Pending.
Notes | Condensation on the probe and USB-C sealing are named concerns (flag materials).
VAL-REL-004 — Network diversity: LTE-M / Wi-Fi / BLE across regions & in-container (RF, gated)
Objective (user outcome) | Across the network diversity of real regions and the RF-hostile in-container context, the device still delivers its data and alerts (best-effort) and always the on-receipt verdict.
Intended Use / User | Distributors, IT/fleet, programmes (SH-03/SH-09/SH-01).
Validation Method | Field study across regional networks following PV-REL-005 (PRD-VER-005).
Participants / Representative Users | Representative multi-region lanes/carriers.
Environment / Context of Use | Diverse carriers/bands; in-container RF attenuation.
Success / Acceptance Criteria | Data eventually delivered across networks; alert delivery characterized per lane; where the in-container link cannot close, the on-receipt verdict compensates; RF margin Evidence- Pending and gated on PRD-VER-005.
User Need Traceability | UN-ENV-006; UN-VAL-008; UN-FUN-010.
Risk / Use-Error Relationship | HA-COMM-003 (RF attenuation); validates HA-RES-009 in use; HAL-2900 condition 3.
Evidence Required | Multi-region delivery report; per-lane latency/delivery-success; RF-in-container data.
Owner | Reliability + Product.
Status | Planned / Evidence Pending (Open — RF, PRD-VER-005).
Notes | The single largest connectivity uncertainty; release-gating; shapes the alert claim.
VAL-REL-005 — Power diversity: mains-scarce depots & cold-charge readiness timing
Objective (user outcome) | In depots with scarce/unreliable mains and cold-returned devices, technicians still get trip-ready devices in a workable time.
Intended Use / User | Service/technicians, depots (SH-06/SH-13).
Validation Method | Field study at representative depots.
Participants / Representative Users | Representative depot operations.
Environment / Context of Use | Mains-scarce depots; cold-returned/charge-inhibited devices.
Success / Acceptance Criteria | Readiness achievable within workable turnaround despite power scarcity and warm-then-charge timing; no fleet-availability collapse.
User Need Traceability | UN-ENV-007; UN-SVC; UN-FLOW-009; UN-VAL-009.
Risk / Use-Error Relationship | HA-USE-009 (charge-inhibit/availability); ties HA-RES-014 availability path.
Evidence Required | Depot power/readiness field report; turnaround-under-scarcity data.
Owner | Reliability + Product.
Status | Planned / Evidence Pending.
Notes | Cold-charge inhibit affects readiness timing (D3) — a real depot constraint.
VAL-REL-006 — Installation diversity: shipper/container/OEM-embed variants
Objective (user outcome) | Across the variety of shipper types, container geometries, and OEM embeds, the device fits, mounts, and measures the product zone reliably over life.
Intended Use / User | Installers/OEM integrators, operators (SH-12/SH-05).
Validation Method | Field study across installation variants.
Participants / Representative Users | Representative installers across shipper/container types.
Environment / Context of Use | Diverse insulated shippers/containers; OEM embeds.
Success / Acceptance Criteria | Secure fit and representative product-zone measurement across variants; no install-type-specific placement-error cluster over life.
User Need Traceability | UN-STK-011; UN-PHY; UN-FUN-003/013.
Risk / Use-Error Relationship | HA-USE-005/006 (placement) across installs; validates HA-RES-003/014.
Evidence Required | Installation-diversity report; per-variant fit/placement data.
Owner | Reliability + Product.
Status | Planned / Evidence Pending.
Notes | Fit failures block the OEM-embed channel; ties VAL-USER-010.
VAL-REL-007 — Regional diversity: LMIC & GDP lanes end-to-end over time
Objective (user outcome) | In the full diversity of real LMIC and GDP regions, the device sustains its outcomes over a representative deployment period.
Intended Use / User | Programmes, procurement, distributors (SH-01/SH-08/SH-03).
Validation Method | Multi-region field reliability study (ties PVaL-900 pilots).
Participants / Representative Users | Representative programmes across regions.
Environment / Context of Use | LMIC last-mile and GDP commercial lanes.
Success / Acceptance Criteria | Sustained outcome delivery across regions; no region with disproportionate failure/return; region-stratified reliability acceptable, Evidence-Pending.
User Need Traceability | UN-ENV; UN-VAL-002/009; UN-VIS.
Risk / Use-Error Relationship | Aggregate reliability of the benefit across deployment diversity (HAL-2900 §5.1).
Evidence Required | Multi-region reliability report; region-stratified outcome/return data.
Owner | Reliability + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Shares deployments with the field pilots (VAL-FIELD).
VAL-REL-008 — Calibration stability over reuse life (fleet common cause)
Objective (user outcome) | Devices stay in-calibration over reuse life, or are blocked from arming when out — so a drifted device never silently biases a verdict in the field.
Intended Use / User | Programmes/QA/technicians (SH-01/SH-02/SH-06); beneficiary.
Validation Method | Fleet drift study + recal-interval-in-use validation.
Participants / Representative Users | Representative fleet over reuse cycles.
Environment / Context of Use | Multi-trip cold/thermal cycling; depot recal.
Success / Acceptance Criteria | Calibration drift stays within tolerance or the pre-arm expiry gate blocks arming; no fleet-wide systematic bias escapes; recal interval effective in use.
User Need Traceability | UN-VAL-005; UN-SVC (calibration); UN-SAFE-001.
Risk / Use-Error Relationship | HA-SW-004 / HA-USE-011 (drift/reuse-without-recal); validates HA-RES-015 (the dominant common cause, HAL-2900 §5.3).
Evidence Required | Fleet drift data; expiry-gate-in-use evidence; recal-escape monitoring.
Owner | Reliability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Largest information-chain common cause; needs standing post-market monitoring (Stage 17).
7. Traceability notes
VAL-REL validates the in-use, over-life, across-diversity persistence of the outcomes that Stage-13 PV-REL verifies at design level. VAL-REL-001 (D3) and VAL-REL-004 (RF) carry the two release-gating Open items into in-use validation (HAL-2900 condition 3); VAL-REL-008 validates the dominant fleet common cause (HA-RES-015). All trace to UN-ENV/UN-PERF/UN-VAL and feed PVaL-900 (field) and PVaL-1000 (residual).
8. Acceptance / summary notes
Reliability-in-use cannot conclude until the D3 cold-endurance and RF-in-container verifications close and the fleet has accumulated representative use across the diversity matrix. Degradation over life must surface honestly, never as a silent wrong verdict. All activities are Planned / Evidence Pending; no field reliability, capacity-fade, or drift figure is asserted.
VAL-RISK-1000 (PVaL-1000) — Risk Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-1000, activity code VAL-RISK. Owner: EM-ENG-006 (Quality/Regulatory) + EM-ENG-001 (Chief Systems), with HF/Usability and Clinical/ Operational. This module validates the residual-risk conclusions of the Hazard Analysis in real use — that the HAL-2800 individual residuals and the HAL-2900 overall residual/benefit-risk hold with representative users in representative contexts, that foreseeable misuse is acceptably controlled, and that hazard communication (warnings, disclosures, IFU) achieves its risk-reduction with the users it targets. Each activity here is the "Validation Reference" that closes the corresponding HAL residual and confirms the seven HAL-2900 benefit-risk conditions. Status: Planned / Evidence Pending — no residual has been validated as accepted; the HAL residual verdicts are projected, not demonstrated, until these close.
---
1. Purpose
ISO 14971 §7.4/§8 require that residual risks — individual and overall — be judged acceptable and that the judgement be validated where it depends on real-world use (esp. use-related and benefit-risk residuals). The HAL residual verdicts (HAL-2800) and the overall benefit-risk conclusion (HAL-2900) are explicitly projected, Evidence-Pending, conditional. This module is where those conditions are validated with representative users: that the risk controls remain effective in real use, that the residuals are acceptable to the users who bear them, that the conservative safety bias's wastage trade is accepted, and that hazard communication works for the low-training user (not just a literate reviewer).
2. Scope
Validation of: the individual residual acceptability of the top hazards (the S4/S5 records in HAL-2800, esp. the use-related cluster HA-RES-014 and the residuals whose acceptability turns on user behavior); the overall residual and benefit-risk (HAL-2900) with both harm directions; foreseeable-misuse acceptability; hazard-communication effectiveness (warnings/disclosures/IFU) for representative users; and confirmation of the seven HAL-2900 release conditions. The physical S5 (aircraft-fire) residual is validated by verification + UN 38.3 (not a user study) but its disclosure obligation is validated here.
3. Applicable standards & references
• ISO 14971 §7.4 / §8 / §9 — individual and overall residual-risk acceptability, benefit-risk, and the production/post-production loop that keeps the conclusion valid.
• IEC 62366-1 — use-related residual-risk acceptability from summative usability.
• HAL-2200 §6/§7 — acceptability regions, decision rules, acceptance authority; HAL-2800/2900 — the residual verdicts and the seven benefit-risk conditions this module validates.
Upstream | every HA-RES record; the VAL-USE/VAL-HF summative studies; VAL-CLIN effectiveness; VAL-FIELD signals; RCR register.
4. Governance & approach
• Validate the residual, not just the control. A control can pass verification yet leave a residual unacceptable to real users (e.g. an override barrier that is defeated under stock-out pressure). This module validates the residual as users experience it.
• Acceptance authority per HAL-2200 §7. S4/S5 residual acceptance requires EM-ENG-006 + EM-ENG-001 + documented benefit-risk; the S5 aircraft-fire and security residuals carry the most stringent sign-off.
• Both harm directions in the benefit-risk. The overall benefit-risk validation weighs false-accept (compromised dose) and false-reject (wastage/stock-out) together, as HAL-2900 §6 requires.
• Projected until closed. Every residual stays Controlled/Open-pending-validation until its named VAL-RISK/VAL-USE/VAL-HF activity reports; nothing is accepted on projection.
• Hazard communication tested with the target user. Warnings/disclosures/IFU are validated for the low-training last-mile user, per the warnings-challenge rule — comprehension by a literate reviewer is insufficient.
5. Validation categories
Individual residual acceptability · use-related residual cluster · overall residual & benefit-risk · both- directions balance · foreseeable-misuse acceptability · hazard communication / warnings · disclosure obligations · security-residual (IV) confirmation · benefit-target confirmation · post-market loop.
6. Validation Register (VAL-RISK activities)
VAL-RISK-001 — Use-related residual cluster acceptable in real use ★ (HA-RES-014)
Objective (user outcome) | The residual risk of the use-error cluster — misread verdict, muted alert, mis-armed/void trip, override, probe placement — is acceptable to representative users: no unrecovered critical use error reaches harm, and residuals are ALARP.
Intended Use / User | Last-mile & operational users (SH-05/SH-01/SH-02/SH-03); beneficiary.
Validation Method | Residual-risk determination from the VAL-USE/VAL-HF summative studies + field signals.
Participants / Representative Users | The summative-study populations (≥15/group, low-training included).
Environment / Context of Use | Representative cold/low-light/gloved/offline + field.
Success / Acceptance Criteria | Residual use-error acceptability confirmed per IEC 62366-1 across HA-USE-001/002/004/005/007; residual region A (ALARP) with benefit-risk sign-off; no residual remains region R; placement remainder (HA-USE-005) ALARP via HF+IFU.
User Need Traceability | UN-SAFE-001…006; UN-VAL-003/005.
Risk / Use-Error Relationship | The Validation Reference for HA-RES-014; closes the use residuals in HAL-2800; the human twin of HAL-2900 condition 1.
Evidence Required | Residual-risk determination memo; summative use-error acceptability analysis; field signal review.
Owner | EM-ENG-006 + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Rolls up VAL-USE-001/002/004/005/007 and VAL-HF-001 into the residual verdict.
VAL-RISK-002 — Honest-verdict / no-false-reassurance choke point holds with users (HA-RES-002)
Objective (user outcome) | Users never dispose a spoiled/void/incomplete trip as a clean accept — the first pathway choke point holds in real decisions.
Intended Use / User | Disposition decision-makers (SH-01/SH-05/SH-02).
Validation Method | Residual determination from VAL-USE-004 + VAL-CLIN-003 + field.
Participants / Representative Users | Summative + field decision-makers.
Environment / Context of Use | Receipt disposition of gapped/void/incomplete trips.
Success / Acceptance Criteria | No false-clean-accept over a gap/void in summative or field; honest-gap qualifier comprehended; residual region A with benefit-risk.
User Need Traceability | UN-SAFE-003; UN-VAL-003; UN-UX-011.
Risk / Use-Error Relationship | Validation Reference for HA-RES-002; HAL-2900 condition 1 (choke point).
Evidence Required | Residual memo; VAL-USE-004/VAL-CLIN-003 evidence; field gap-disposition data.
Owner | EM-ENG-006 + HF/Usability.
Status | Planned / Evidence Pending.
Notes | One of the two most load-bearing residual validations (with VAL-RISK-001).
VAL-RISK-003 — Overall residual risk & benefit-risk acceptable (both directions) ★ (HA-RES / HAL-2900)
Objective (user outcome) | The whole-device residual risk is acceptable relative to the public-health benefit — protecting potency AND availability — with representative users, per ISO 14971 §8.
Intended Use / User | All; the release decision; beneficiary (represented).
Validation Method | Overall benefit-risk determination integrating all validation evidence.
Participants / Representative Users | EM-ENG-001, EM-ENG-006, HF/Usability, Clinical/Operational; user evidence from all summative/field studies.
Environment / Context of Use | Design-control / release governance, based on field evidence.
Success / Acceptance Criteria | Overall residual judged acceptable vs benefit on demonstrated (not merely projected) evidence; both harm directions weighed; the seven HAL-2900 conditions confirmed closed; documented benefit-risk signed.
User Need Traceability | UN-VIS-001; UN-VAL-002/005; HAL-2900 §6/§7.
Risk / Use-Error Relationship | Validates the HAL-2900 overall benefit-risk conclusion; converts it from conditional to demonstrated.
Evidence Required | Overall benefit-risk record; conditions-1–7 closure evidence; sign-off.
Owner | EM-ENG-001 + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | The capstone risk validation; feeds the Stage-16 release recommendation.
VAL-RISK-004 — Accepted public-health trade: false-reject / wastage acceptable (HA-RES-019)
Objective (user outcome) | The conservative safety bias's wastage/false-reject residual is acceptable — reduced by validated two-sided targets, balanced against the false-accept residuals, and monitored.
Intended Use / User | Programmes, procurement, finance (SH-01/SH-08/SH-10); beneficiary (stock-out).
Validation Method | Residual determination from VAL-USE-002 targets + VAL-CLIN-004 field wastage.
Participants / Representative Users | Summative responders + field programmes.
Environment / Context of Use | Real disposition under scarcity/pressure.
Success / Acceptance Criteria | False-reject/wastage within validated two-sided targets; net public-health benefit positive; residual accepted as a monitored trade (region A) with benefit-risk.
User Need Traceability | UN-VAL-005; UN-SAFE-004; HA-RES-019.
Risk / Use-Error Relationship | Validation Reference for HA-RES-019; HAL-2900 both-directions balance.
Evidence Required | False-reject/wastage data; two-sided target reconciliation; benefit-risk balance.
Owner | EM-ENG-006 + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Wastage is a first-class monitored residual, not a free trade.
VAL-RISK-005 — Foreseeable-misuse residual acceptable (HA-RES-014 twin / HA-USE-012)
Objective (user outcome) | The residual risk from foreseeable non-malicious circumvention (force a pass, mute, edit thresholds, override) is acceptable — casual bypass prevented, deliberate bypass recorded/ detectable.
Intended Use / User | Operational users under pressure; QA/regulator (record integrity).
Validation Method | Residual determination from VAL-USE-016 misuse study + audit review.
Participants / Representative Users | Misuse-scenario participants + audit reviewers.
Environment / Context of Use | Simulated production/stock-out pressure.
Success / Acceptance Criteria | Casual bypass prevented; deliberate bypass deliberate/recorded/ detectable; residual region A with benefit-risk; residual programme-policy dependency documented.
User Need Traceability | UN-SAFE-001/003/006; HA-CTX-003.
Risk / Use-Error Relationship | Validation Reference for the HA-USE-012 residual; ties HA-CYBER-002/003.
Evidence Required | Misuse-residual memo; bypass-resistance evidence; audit-trail review.
Owner | EM-ENG-006 + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Residual retains an out-of-device accountability-policy dependency (flagged in benefit-risk).
VAL-RISK-006 — Hazard communication / warnings effective for the low-training user (warnings-challenge)
Objective (user outcome) | Warnings, disclosures, and IFU achieve their intended risk reduction with the users they target — comprehended by the low-training last-mile user, not just a literate reviewer.
Intended Use / User | Last-mile & operational users (SH-05/SH-01).
Validation Method | Comprehension/summative validation of warnings & IFU (ties VAL-HF-010).
Participants / Representative Users | Representative low-training/low-literacy/non-primary-language users.
Environment / Context of Use | Real use without a manual to hand.
Success / Acceptance Criteria | Where a warning/disclosure/IFU is credited with risk reduction, its target users comprehend and act on it; where they do not, the residual reverts to design (warnings-challenge escalation), not to the warning.
User Need Traceability | UN-SAFE-005; UN-VAL-003; HAL-2200 §6.4.
Risk / Use-Error Relationship | Validates the information-for-safety controls credited in HAL-2800; confirms training/IFU are supplements only.
Evidence Required | Warnings/IFU comprehension report; per-user-group results.
Owner | HF/Usability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | A warning credited with risk reduction that its target user cannot use is not a valid control.
VAL-RISK-007 — Disclosure obligations validated & user-facing (record limits, connectivity, transport)
Objective (user outcome) | The mandatory disclosures — detection limits, connectivity-bound alerting, honest-gap, in-container link limits, transport (UN 38.3/IATA) — are present, understood, and honest to the users/parties who rely on them.
Intended Use / User | Programmes, distributors, 3PLs, regulators (SH-01/SH-03/SH-04/SH-07).
Validation Method | Disclosure-comprehension review + audit walkthrough.
Participants / Representative Users | Representative users/auditors for each disclosure.
Environment / Context of Use | Point of use, transport, audit.
Success / Acceptance Criteria | Each HAL-2800 disclosure obligation is present, comprehensible, and honest; no over-claim (e.g. real-time alerting where the link cannot close).
User Need Traceability | UN-VAL-006; UN-REG; UN-SAFE-009.
Risk / Use-Error Relationship | Operationalizes the HAL-2800 disclosure obligations; ties HA-RES-007/009 (connectivity), HA-RES-006 (transport).
Evidence Required | Disclosure inventory; comprehension/audit evidence.
Owner | EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Honest disclosure is a control; over-claiming is a hazard (false reassurance).
VAL-RISK-008 — Physical S5 (aircraft-hold fire) disclosure & monitoring obligations validated (HA-RES-006)
Objective (user outcome) | The rare-catastrophic battery-fire residual carries its mandatory transport disclosure and post-market monitoring — the parties who bear it are informed and it is watched.
Intended Use / User | 3PLs/air-cargo, regulators, EM-ENG-001/006 (SH-04/SH-07).
Validation Method | Disclosure/monitoring-obligation validation (residual itself closed by PV-TFC-002 + UN 38.3, not a user study).
Participants / Representative Users | Transport/regulatory stakeholders.
Environment / Context of Use | Air/ground transport; post-market thermal-event monitoring.
Success / Acceptance Criteria | UN 38.3/IATA-DGR documentation present and accepted; post-market thermal- event monitoring established; disclosure validated; residual region-A-High accepted per HAL-2200 §7.
User Need Traceability | UN-SAFE-008/009; UN-REG (transport).
Risk / Use-Error Relationship | Validation Reference for the HA-RES-006 disclosure/monitoring; HAL-2900 condition 2 (stand-alone benefit-risk line).
Evidence Required | UN 38.3/IATA documentation; thermal-event monitoring plan; disclosure evidence.
Owner | EM-ENG-006 + EM-ENG-001.
Status | Planned / Evidence Pending.
Notes | The residual can never be broadly acceptable; carries the most stringent sign-off.
VAL-RISK-009 — Security residuals confirmed by independent adversarial verification (HA-RES-010/011/012/013)
Objective (user outcome) | Users and programmes can trust that the security controls hold against a real adversary — telemetry authenticity, non-suppressible alarms, safe OTA, cloud resilience.
Intended Use / User | IT/security admins, programmes, regulators (SH-09/SH-01/SH-07).
Validation Method | Confirmation that the independent pen-tests (PV-CYBER-001/002/005/006/009) closed; operational trust validation.
Participants / Representative Users | Independent security assessors; IT/security admins.
Environment / Context of Use | Representative deployment/threat model.
Success / Acceptance Criteria | All security residuals close on independent adversarial verification; operational trust confirmed with admins; residuals region A with benefit-risk + security IV sign-off.
User Need Traceability | UN-SEC; UN-STK-003; HAL-2900 condition 5.
Risk / Use-Error Relationship | Validates HA-RES-010/011/012/013 acceptability (require IV).
Evidence Required | Independent pen-test closure; operational-trust evidence.
Owner | EM-ENG-006 + IT/security.
Status | Planned / Evidence Pending.
Notes | Security residuals cannot be accepted on projection — independent verification required.
VAL-RISK-010 — Common-cause containment validated in use (calibration & OTA/cloud) (HA-RES-015/012)
Objective (user outcome) | The dominant fleet common causes — systematic calibration bias and OTA/cloud rollout — do not defeat the safety chain fleet-wide in real operation.
Intended Use / User | Programmes/QA/IT (SH-01/SH-02/SH-09); beneficiary.
Validation Method | In-use common-cause validation (ties VAL-REL-008, VAL-ECO-009) + monitoring.
Participants / Representative Users | Fleet over reuse cycles + OTA campaigns.
Environment / Context of Use | Fleet operations across cycles/campaigns.
Success / Acceptance Criteria | No fleet-wide systematic calibration bias escapes (expiry gate + recal effective); OTA/config staged/reversible with no fleet-blinding; standing post-market monitoring established.
User Need Traceability | UN-VAL-005; UN-SEC; HAL-2900 condition 4/§5.3.
Risk / Use-Error Relationship | Validates HA-RES-015 (dominant common cause) and HA-RES-012 containment in use.
Evidence Required | Fleet drift/recal monitoring; OTA campaign-health evidence.
Owner | EM-ENG-006 + Reliability + Product.
Status | Planned / Evidence Pending.
Notes | These common causes rely on standing operational control, not one-time tests (HAL-2900 §5.3).
VAL-RISK-011 — Carried-Open unknowns confirmed acceptable (D3 endurance, RF-in-container)
Objective (user outcome) | The two release-gating unknowns close acceptably in use — devices survive the cold trip and deliver/verdict despite in-container RF — so their residuals become demonstrated-acceptable.
Intended Use / User | Distributors, programmes, procurement (SH-03/SH-01/SH-08).
Validation Method | Residual confirmation from VAL-REL-001/004 + PV-REL-001/005 (PRD-VER-004/005).
Participants / Representative Users | Deep-cold and multi-region field operations.
Environment / Context of Use | −30 °C long trips; in-container RF-hostile lanes.
Success / Acceptance Criteria | D3 endurance and RF behavior close acceptably; where RF cannot close, the alert claim is restricted and the on-receipt verdict compensates; residuals HA-RES-004/009 demonstrated.
User Need Traceability | UN-VAL-007/008; UN-ENV-001/006; D3.
Risk / Use-Error Relationship | Validation Reference for HA-RES-004/009; HAL-2900 condition 3 (the two carried-Open items).
Evidence Required | D3/RF in-use closure evidence; restricted-claim decision if RF cannot close.
Owner | EM-ENG-001 + Reliability.
Status | Planned / Evidence Pending (Open — PRD-VER-004/005).
Notes | These gate the release recommendation; a restricted claim is the fallback if RF cannot close.
VAL-RISK-012 — Post-market benefit-risk loop established (ISO 14971 §9)
Objective (user outcome) | The benefit-risk conclusion stays valid after release — field signals in both directions feed back to re-affirm or revise it.
Intended Use / User | EM-ENG-006/001; programmes; regulators.
Validation Method | Post-market surveillance loop definition (ties VAL-FIELD-010 / Stage-17).
Participants / Representative Users | Safety/quality reviewers; field operations.
Environment / Context of Use | Ongoing post-release.
Success / Acceptance Criteria | Both-direction field metrics, thermal events, security incidents feed the ISO 14971 §9 loop; the disclosure/monitoring obligations from HAL-2800 are operational; conclusion re-review cadence set.
User Need Traceability | UN-VAL-002/005; HAL-2900 condition 7.
Risk / Use-Error Relationship | Keeps the overall benefit-risk (HAL-2900) valid over life.
Evidence Required | PMS plan; §9 feedback loop definition; re-review cadence.
Owner | EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Hands the disclosure/monitoring obligations to Stage-17 PMS.
7. Traceability notes
Each VAL-RISK activity is the "Validation Reference" that closes a HAL residual: VAL-RISK-001→HA-RES-014, -002→HA-RES-002, -003→HAL-2900 overall, -004→HA-RES-019, -005→HA-USE-012 residual, -006/007→information-for- safety controls, -008→HA-RES-006, -009→HA-RES-010/011/012/013, -010→HA-RES-015/012, -011→HA-RES-004/009, -012→ISO 14971 §9. Together they confirm the seven HAL-2900 benefit-risk conditions and convert the projected overall residual to demonstrated. They consume the VAL-USE/VAL-HF/VAL-CLIN/VAL-ECO/VAL-REL/VAL-FIELD evidence.
8. Acceptance / summary notes
No residual is accepted on projection. Each S4/S5 residual stays Controlled/Open-pending-validation until its VAL-RISK reference reports; the overall benefit-risk (VAL-RISK-003) is signed only when the seven HAL-2900 conditions close — including the D3/RF Open items (VAL-RISK-011), the security IV (VAL-RISK-009), and the WHO PQS benefit target (VAL-CLIN-010). All activities are Planned / Evidence Pending.
VAL-SUM-1200 (PVaL-1200) — Validation Summary
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-1200, activity code VAL-SUM. Owner: EM-ENG-001 + EM-ENG-006, with HF/Usability, Clinical/Operational, Product/Business. This module rolls up validation coverage, user acceptance, operational/clinical acceptance, business acceptance, and residual concerns into a single go-to-market recommendation for the release decision (Stage 16). Because no validation has executed, the summary is a pre-execution frame: it states what the conclusion will be judged on and carries the honest verdict that the go-to-market recommendation is Not Yet Determined — validation not executed. No acceptance or recommendation is asserted. Status: Planned / Evidence Pending.
---
1. Purpose
Give the release decision-maker one honest, integrated view: did we build the correct product for real users in the real cold chain? The summary aggregates the library's conclusions across the dominant outcome, the users, the workflows, the ecosystem, reliability-in-use, business, and — decisively — the residual-risk/ benefit-risk validation, and states the go-to-market recommendation with its conditions.
2. Scope
Whole-library roll-up: coverage completeness (every user need / safety-critical workflow validated), user acceptance (summative HF), operational/clinical acceptance (effectiveness), ecosystem/reliability acceptance, business acceptance (adoption/economics), residual concerns (the HAL-2800/2900 conditions), and the resulting recommendation. Feeds Stage 16 (Release), Stage 17 (Field/PMS), and ORCH-018 / EM-ENG-012.
3. Applicable standards & references
• ISO 13485 §7.3.7 (validation complete before transfer); ISO 14971 §8 (overall benefit-risk); IEC 62366-1 (summative conclusion); WHO PQS E006 (benefit target).
Upstream | all PVaL modules; HAL-2900 (the seven benefit-risk conditions); the Coverage Matrix.
4. Governance & approach
• Outcome-first roll-up. The dominant-outcome critical tasks and the residual/benefit-risk validation dominate the recommendation; business acceptance is weighed but cannot override a failed safety conclusion.
• Conditions, not vibes. The recommendation is tied to the closure of the seven HAL-2900 conditions and the coverage gaps in the matrix — each named, each with a status.
• Honest verdict. Pre-execution, the recommendation is Not Yet Determined; the summary never asserts an acceptance that has not been evidenced.
5. Validation categories
Coverage · user acceptance · clinical/operational acceptance · business acceptance · residual concerns · go-to-market recommendation.
6. Validation Register (VAL-SUM activities)
VAL-SUM-001 — Validation coverage roll-up
Objective (user outcome) | Confirm every user need and every safety-critical workflow has a validation activity and evidence, with gaps named.
Validation Method | Coverage-matrix roll-up (from Validation-Coverage-Matrix.md).
Success / Acceptance Criteria | No user need without ≥1 validation; no safety-critical workflow without a summative activity; every use error (HA-USE) has a validation reference; gaps explicitly listed.
User Need Traceability | all UN modules; UN-VAL-001.
Risk / Use-Error Relationship | Confirms HA-USE / HA-RES validation coverage.
Evidence Required | Completed coverage matrix; gap list.
Owner | EM-ENG-006 + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Coverage is structurally complete in the playbook; execution pending.
VAL-SUM-002 — User & human-factors acceptance roll-up
Objective (user outcome) | State whether representative users — especially the low-training last mile — achieved the dominant detect→understand→act outcome.
Validation Method | Summative HF/use conclusion roll-up (VAL-USE/VAL-HF/VAL-USER).
Success / Acceptance Criteria | All summative critical tasks passed with acceptable residual use errors; otherwise recommendation blocked.
User Need Traceability | UN-VAL-003; UN-SAFE.
Risk / Use-Error Relationship | HA-RES-014; HAL-2900 condition 1.
Evidence Required | Summative conclusion; usability engineering file reference.
Owner | HF/Usability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | The safety spine of the recommendation.
VAL-SUM-003 — Clinical/operational & ecosystem/reliability acceptance roll-up
Objective (user outcome) | State whether the device is operationally effective (both directions), ecosystem-complete, and reliable-in-use across diversity.
Validation Method | VAL-CLIN/VAL-ECO/VAL-REL conclusion roll-up.
Success / Acceptance Criteria | Effectiveness demonstrated vs counterfactual; ecosystem/offline/sync outcomes hold; reliability-in-use sustained; D3/RF closed or claim restricted.
User Need Traceability | UN-VAL-002/005/007/008.
Risk / Use-Error Relationship | HA-RES effectiveness/reliability residuals; HAL-2900 conditions 3/4.
Evidence Required | Operational/ecosystem/reliability conclusions.
Owner | Clinical/Operational + Product + Reliability.
Status | Planned / Evidence Pending.
Notes | Carries the two carried-Open items (D3/RF) into the recommendation.
VAL-SUM-004 — Business acceptance & residual-concerns roll-up
Objective (user outcome) | State adoption/economic acceptance and consolidate the residual concerns (HAL-2800/2900) and their acceptance status.
Validation Method | VAL-BIZ + VAL-RISK roll-up.
Success / Acceptance Criteria | Reuse economics viable (measured); residuals accepted per authority; the seven HAL-2900 conditions each with a closure status; residual programme-policy dependencies flagged.
User Need Traceability | UN-VAL-009; UN-BIZ; UN-VAL-005.
Risk / Use-Error Relationship | Consolidates the residual/benefit-risk validation (VAL-RISK-003).
Evidence Required | Business conclusion; residual-concerns register; conditions-1–7 status.
Owner | Product/Business + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Residual concerns include out-of-device policy dependencies (override, bypass).
VAL-SUM-005 — Go-to-market recommendation
Objective (user outcome) | A single, honest release recommendation for the decision-maker, tied to the closure of the seven HAL-2900 conditions and any restricted-claim conditions.
Validation Method | Integrated benefit-risk recommendation (VAL-RISK-003 + all roll-ups).
Success / Acceptance Criteria | Recommendation stated with conditions; if any of the seven HAL-2900 conditions is unclosed, recommendation is withhold/restrict per that condition (e.g. no air-transport claim without UN 38.3; no real-time-alert claim on lanes where RF cannot close).
User Need Traceability | UN-VIS-001; UN-VAL; HAL-2900 §6/§7.
Risk / Use-Error Relationship | The whole-device benefit-risk conclusion for release.
Evidence Required | Signed go-to-market recommendation with conditions.
Owner | EM-ENG-001 + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Pre-execution verdict: Not Yet Determined — validation not executed. Consistent with the HAL "Conditionally Ready — release pending V&V" posture (HAL-2900 §7).
7. Traceability notes
VAL-SUM consumes every module and the Coverage Matrix, and confirms the seven HAL-2900 benefit-risk conditions. VAL-SUM-005 is the release-decision input (Stage 16); the residual-concerns and disclosure/monitoring obligations feed Stage-17 PMS; all VAL- IDs resolve in the ORCH-018 / EM-ENG-012 matrix.
8. Acceptance / summary notes
No acceptance and no go-to-market recommendation is asserted — validation has not executed. The summary frames the decision: the recommendation will be positive only when the dominant-outcome critical tasks (VAL-USE-001/002/004/007, VAL-HF-001), the residual/benefit-risk validation (VAL-RISK-003), and the seven HAL-2900 conditions (incl. D3/RF, security IV, and the WHO PQS benefit target) all close acceptably. Until then, the honest verdict is Not Yet Determined / Conditionally Ready — pending validation.
VAL-USE-300 (PVaL-300) — Intended-Use Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-300, activity code VAL-USE. Owner: HF/Usability lead. This module validates that the intended-use workflows succeed with representative users in representative conditions — the primary chain (place-in-container → monitor → read status → respond to alert/excursion → make the disposition decision → return/re-provision), plus configuration, misuse, emergency/exceptional situations, recovery, maintenance, and installation. It carries the twelve IEC 62366-1 critical-task summative studies that close the HAL-1500 use hazards: VAL-USE-001…009 map one-to-one to HA-USE-001…009 (the exact Validation References cited in the hazard records), so hazard closure and validation acceptance resolve to the same activities. Status: Planned / Evidence Pending — no workflow has been validated with users.
---
1. Purpose
Confirm that the SCM does its job as a whole workflow, not just as isolated features — that a real user can carry the device through its intended trip and reach correct outcomes at every decision point, and that the foreseeable ways use goes wrong (misuse, exceptions, faults) are handled acceptably. The heart of the module is the dominant outcome: users correctly detect, understand, and act on temperature excursions so no compromised dose is administered and no potent dose is wasted. The five headline HA-USE hazards (misread verdict, muted alert, void-trip-believed-monitored, override, charge-inhibit) each become a summative critical-task study here.
2. Scope
All intended-use scenarios of the reusable SCM across the multi-market chain:
Primary workflows | provision/configure → arm → place-in-container/confirm placement → transit monitoring → in-trip alert response → receipt disposition read (accept/quarantine/reject) → return → reconditioning/recharge/recalibrate → re-provision.
Configuration | trip-profile selection, threshold setting, probe/channel selection.
Exceptional/emergency | mass/campaign excursion, stock-out pressure, connectivity blackout, device fault in transit.
Recovery | power loss / dead-in-cold, offline buffering, resync.
Maintenance/installation | recharge-to-ready, firmware update, new-fleet onboarding.
Misuse | foreseeable non-malicious circumvention (the use twin of HA-USE-012).
3. Applicable standards & references
• IEC 62366-1 — hazard-related use scenarios and summative usability evaluation of critical tasks; each VAL-USE-001…009 is a summative critical-task study.
• FDA HF guidance — critical-task validation, ≥15 representative users per group, root-cause of residual use errors. ISO 14971 — residual use-error acceptability. WHO PQS E006 — disposition context.
Upstream | HAL-1500 HA-USE-001…012 (the use hazards and their VAL-USE references); UN-FLOW-001…011, UN-JTBD-001…012, UN-FUN-001…018, UN-SAFE-001…010, UN-UX-001…012; Stage-13 PV-USE-001…012 (paired verification).
4. Governance & approach
• Critical tasks are summative and warnings-challenge-aware. VAL-USE-001…009 use minimally trained / training-decayed users in cold/low-light/gloved/offline conditions — the design must produce the outcome for the user as they really are.
• Paired with verification, never substituting for it. Each VAL-USE-00N runs after its paired PV-USE-00N passes (e.g. VAL-USE-001 after PV-USE-001 cold-display legibility) — validation confirms the human outcome the verified control exists for.
• Both harm directions. Scenarios include clean, freeze, heat, gapped/void, and borderline trips so both false-accept (compromised dose) and false-reject (wastage/stock-out) outcomes are measured.
• Root-cause, not just count. Residual use errors are analyzed for cause and acceptability per IEC 62366-1; a critical use error that reaches harm without recovery fails the study.
5. Validation categories
Primary disposition & alert critical tasks · provisioning/arming/placement critical tasks · override & state- interpretation critical tasks · end-to-end workflow · exceptional/emergency · recovery · maintenance · installation · foreseeable misuse.
6. Validation Register (VAL-USE activities)
VAL-USE-001 — Read the verdict & reach the correct disposition (gloved / cold / low-light / low-literacy) ★ dominant critical task — HA-USE-001
Objective (user outcome) | The last-mile user reads the accept/reject/quarantine verdict correctly and makes the correct disposition — and cannot be led to the wrong one — under real conditions.
Intended Use / User | Last-mile health worker (SH-05); protects beneficiary (SH-11).
Validation Method | Summative usability test (IEC 62366-1 critical task).
Participants / Representative Users | ≥15 minimally trained health workers per sub-group; low-literacy, non-primary-language, gloved, training-decayed included.
Environment / Context of Use | Cold-soaked device, low light/glare, gloved, offline, time pressure; simulated rural post (field subset in PVaL-900).
Success / Acceptance Criteria | Correct disposition on clean/freeze/heat/gapped scenarios via the icon+colour+word encoding; no unrecovered critical use error (accept-a-reject / reject-an-accept); residual use errors root-caused and acceptable; use-error-rate target Evidence-Pending.
User Need Traceability | UN-SAFE-001/005; UN-UX-001/002/003; UN-FUN-008; UN-FLOW-005; UN-JTBD-002; UN-VAL-003.
Risk / Use-Error Relationship | HA-USE-001 (Validation Reference); ties HA-ENV-002, HA-SW-007, HA-DATA-005; validates HA-RES-014 & HAL-2900 condition 1.
Evidence Required | Summative protocol/report; video; use-error root-cause; disposition-correctness by scenario/sub-group.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | The headline critical task; runs after PV-USE-001 (cold-display legibility) passes.
VAL-USE-002 — Respond to a genuine excursion alert within the actionable window (no alarm fatigue) — HA-USE-002
Objective (user outcome) | A responder acts on a real excursion alert in time — even amid many prior low-value alerts — so the product is rescued/segregated before harm.
Intended Use / User | Distributor coordinator / responder (SH-03); clinic staff (SH-05).
Validation Method | Summative usability test under representative alert load + simulated-use.
Participants / Representative Users | ≥15 representative responders per group.
Environment / Context of Use | Operations center / clinic with realistic background alert volume; time-limited rescue window.
Success / Acceptance Criteria | Responders act on the genuine excursion within the actionable window, do not habitually mute, and unacknowledged alerts escalate to another responder; false-reject/wastage measured as the counter-endpoint; false-alarm-tolerance target Evidence-Pending.
User Need Traceability | UN-SAFE-004; UN-FUN-006/007; UN-JTBD-003/009; UN-FLOW-006; UN-UX-006.
Risk / Use-Error Relationship | HA-USE-002 (Validation Reference); ties HA-COMM-001/002, HA-SW-006; validates HA-RES-014 & HA-RES-019 (both directions).
Evidence Required | Alert-load protocol/report; time-to-act; mute-rate; escalation observations.
Owner | HF/Usability + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Runs after PV-USE-002 (alarm-management tuning/escalation) passes; alarm fatigue is a design problem, validated as such.
VAL-USE-003 — Configure the correct trip profile & thresholds (wrong profile caught) — HA-USE-003
Objective (user outcome) | The provisioner selects/confirms the correct product-stability profile and a wrong/implausible profile is caught before arming.
Intended Use / User | Dispatch/QA provisioner (SH-02); cold-store operator (SH-05).
Validation Method | Summative usability test (provisioning workflow).
Participants / Representative Users | ≥15 representative provisioners.
Environment / Context of Use | Pack-out / cold store; mobile/BLE provisioning app + device.
Success / Acceptance Criteria | Correct profile selectable and confirmable; a mismatched/implausible profile is flagged by validation and the pre-arm gate blocks arming; no silent wrong-profile arming.
User Need Traceability | UN-SAFE-006; UN-FUN-011; UN-JTBD-004/005; UN-FLOW-002; UN-UX-005/007.
Risk / Use-Error Relationship | HA-USE-003 (Validation Reference); ties HA-SW-003; validates HA-RES-014.
Evidence Required | Provisioning protocol/report; wrong-profile catch rate; task-completion.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Runs after PV-USE-003 (profile validation + pre-arm gate) passes.
VAL-USE-004 — Never dispatch an un-armed device believed armed; never read a void trip as clean ★ false-reassurance critical task — HA-USE-004
Objective (user outcome) | Users cannot dispatch believing a device is armed when it is not, and never read a void/unmonitored trip as a clean pass.
Intended Use / User | Dispatch/QA (SH-02) and last-mile user (SH-05).
Validation Method | Summative usability test (arming/dispatch + receipt of a void trip).
Participants / Representative Users | ≥15 representative dispatch and last-mile users.
Environment / Context of Use | Pack-out/dispatch (arming) and receipt (reading a void/gapped trip); cold/gloved/offline.
Success / Acceptance Criteria | No dispatch of an un-armed/void device believed armed; a void/unmonitored trip is never dispositioned as a clean accept (the honest-gap surface is understood); critical false- reassurance error rate acceptable, Evidence-Pending.
User Need Traceability | UN-SAFE-003/006; UN-FUN-012; UN-FLOW-002; UN-UX-009/012; UN-JTBD-001.
Risk / Use-Error Relationship | HA-USE-004 (Validation Reference); ties HA-DATA-004, HA-SW-013; validates HA-RES-002/014 & HAL-2900 condition 1 (a pathway choke point).
Evidence Required | Arming/void protocol/report; false-reassurance error incidence; comprehension of the honest-gap qualifier.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | One of the two most load-bearing studies in the whole program (with VAL-USE-001).
VAL-USE-005 — Place the probe in the product zone (unrepresentative placement caught) — HA-USE-005
Objective (user outcome) | Loading staff place the probe in the product zone so readings reflect the doses, and an unrepresentative placement (air/wall) is caught.
Intended Use / User | Loader / cold-store operator (SH-05); installer (SH-12).
Validation Method | Summative usability test (placement task).
Participants / Representative Users | ≥15 representative loaders/installers, gloved.
Environment / Context of Use | Shipper loading; cold store; before sealing.
Success / Acceptance Criteria | Correct product-zone placement achieved with the provided guidance; an air/wall placement is flagged by plausibility / caught before sealing; placement-error rate acceptable, Evidence-Pending.
User Need Traceability | UN-SAFE-005; UN-FUN-003/013; UN-FLOW-003; UN-UX-010.
Risk / Use-Error Relationship | HA-USE-005 (Validation Reference); ties HA-SW-008, HA-DATA-001; validates the HA-RES-003/014 placement remainder (the plausibility-uncatchable case).
Evidence Required | Placement protocol/report; placement-error incidence; catch-rate.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Placement is the plausibility-uncatchable remainder — HF/guidance is load-bearing here.
VAL-USE-006 — Fit/select the correct probe & channel (wrong one caught) — HA-USE-006
Objective (user outcome) | The correct probe type/channel is fitted/selected and a wrong or out-of-range one is caught before arming.
Intended Use / User | Provisioner / technician (SH-02/SH-06).
Validation Method | Usability test (subset of the provisioning summative study).
Participants / Representative Users | Representative provisioners/technicians (multi-probe configs).
Environment / Context of Use | Provisioning; multi-probe/channel fleet.
Success / Acceptance Criteria | Correct probe/channel fitted/selected and confirmed; a wrong/out-of-range/ mis-mapped probe is caught by identity/plausibility + pre-arm gate; no wrong-channel arming.
User Need Traceability | UN-SAFE-005; UN-FUN-003; UN-FLOW-002; UN-UX-007.
Risk / Use-Error Relationship | HA-USE-006 (Validation Reference); ties HA-DATA-001, HA-SW-008.
Evidence Required | Probe/channel task records; wrong-probe catch rate.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Runs with VAL-USE-003 provisioning study; after PV-USE-006.
VAL-USE-007 — A reject cannot be accidentally overridden; a deliberate override is understood — HA-USE-007 / HA-USE-012
Objective (user outcome) | A user cannot accidentally override a reject/quarantine verdict, and a deliberate override is consciously understood as accepting a compromised dose (and is recorded).
Intended Use / User | Disposition decision-maker under pressure (SH-01/SH-05).
Validation Method | Summative usability test (override workflow).
Participants / Representative Users | ≥15 representative disposition decision-makers; include stock-out/ production-pressure scenarios.
Environment / Context of Use | Receipt/disposition under simulated scarcity/pressure.
Success / Acceptance Criteria | No accidental override occurs; a deliberate override requires a confirmed, recorded action and users articulate that they are accepting a rejected lot; override rate under pressure measured, Evidence-Pending; legitimate policy-bounded override remains achievable (not a lockout).
User Need Traceability | UN-SAFE-001; UN-UX-011; UN-JTBD-002; UN-VAL-003/005.
Risk / Use-Error Relationship | HA-USE-007 & HA-USE-012 (Validation Reference); ties HA-DATA-004/008, HA-SW-002; validates HA-RES-014 & HAL-2900 condition 1; residual carries a programme-policy dependency.
Evidence Required | Override protocol/report; accidental-override incidence; comprehension data; audit of recorded overrides.
Owner | HF/Usability (lead) + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Residual depends on out-of-device disposition/accountability policy — flagged in benefit-risk.
VAL-USE-008 — Comprehension across representative locales, languages & literacy levels — HA-USE-008
Objective (user outcome) | Users across the deployment locales comprehend the verdict/state and configure correctly without depending on reading UI text.
Intended Use / User | Last-mile & provisioning users across markets (SH-05/SH-01/SH-02).
Validation Method | Summative usability test across locales (multi-site).
Participants / Representative Users | Representative users in ≥3 locales/languages incl. low literacy.
Environment / Context of Use | Multi-market field/simulated conditions.
Success / Acceptance Criteria | Verdict/state comprehension and correct configuration achieved via icon+colour+word across locales; no locale/literacy group with concentrated critical use error; icon-comprehension threshold Evidence-Pending.
User Need Traceability | UN-UX-003; UN-SAFE-005; UN-VAL-003; UN-REG (localisation).
Risk / Use-Error Relationship | HA-USE-008 (Validation Reference); feeds HA-USE-001/003.
Evidence Required | Multi-locale protocol/report; per-locale comprehension & disposition rates.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Runs after PV-USE-008 (symbol/label comprehension + localisation completeness).
VAL-USE-009 — Correctly interpret a cold charge-inhibited device (warm-then-charge, not discard) — HA-USE-009
Objective (user outcome) | A technician correctly identifies a safely charge-inhibited (cold) device and warms-then-charges it rather than discarding it or sending an unmonitored substitute.
Intended Use / User | Depot/service technician (SH-06); dispatch (SH-02).
Validation Method | Summative usability test (state-interpretation task).
Participants / Representative Users | ≥15 representative technicians.
Environment / Context of Use | Depot recharge/triage; cold-returned devices.
Success / Acceptance Criteria | Technicians distinguish charge-inhibited-with-cause from a dead/faulty device and take the correct action; no serviceable device wrongly retired; misidentification rate acceptable, Evidence-Pending.
User Need Traceability | UN-STK-010; UN-FUN-017; UN-FLOW-009; UN-SAFE-007/010.
Risk / Use-Error Relationship | HA-USE-009 (Validation Reference); ties HA-ENRG-001, HA-USE-004, HA-SVC-007; validates the HA-RES-014 availability/stock-out path.
Evidence Required | State-interpretation protocol/report; correct-action rate.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Runs after PV-USE-009 (charge-inhibit-with-cause vs dead legibility).
VAL-USE-010 — End-to-end primary workflow: place → monitor → read → respond → disposition → return
Objective (user outcome) | A shipment completes the full intended trip with correct outcomes at every decision point when handled by the real chain of users, with no user action required in transit.
Intended Use / User | The full user chain (SH-02 → SH-03/SH-04 → SH-01/SH-05 → SH-06/SH-13).
Validation Method | Simulated-use end-to-end scenario + field pilot leg (ties PVaL-900).
Participants / Representative Users | Representative users at each stage of a scripted trip.
Environment / Context of Use | A full scripted lane (pack-out → transit → receipt → return) incl. cold/ offline segments.
Success / Acceptance Criteria | Every stage's outcome is correct (armed, monitored, alerted where connectivity allows, correctly dispositioned, returned/re-provisioned); no hand-off breaks the chain; no in-transit user action required.
User Need Traceability | UN-JTBD-001…011 (workflow spine); UN-FLOW-001…011; UN-VIS.
Risk / Use-Error Relationship | Integrates the HA-USE cluster; validates the convergence pathway (HAL-2900 §5.2) end-to-end with real hand-offs.
Evidence Required | End-to-end scenario report; per-stage outcome log.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | The integrative study that proves the whole is coherent, not just each node.
VAL-USE-011 — In-trip excursion response & product rescue (re-ice / segregate)
Objective (user outcome) | When a mid-transit excursion alert is received (connectivity permitting), the responder rescues or segregates the product within the actionable window.
Intended Use / User | Distributor/3PL responder (SH-03/SH-04); program (SH-01).
Validation Method | Simulated-use scenario + field study (connectivity-dependent).
Participants / Representative Users | Representative responders on real/simulated lanes.
Environment / Context of Use | In-transit excursion with variable connectivity; re-ice/segregate options.
Success / Acceptance Criteria | Responders take effective rescue/segregation action within the window where connectivity allows; where connectivity is absent, the on-receipt verdict correctly catches the excursion (the local safety floor).
User Need Traceability | UN-JTBD-003; UN-FLOW-006; UN-FUN-006/007; UN-ENV-006/007; UN-SAFE-004.
Risk / Use-Error Relationship | HA-COMM-001/003 (latency/RF); validates HA-RES-007/009 acceptability with users; HAL-2900 §5.4 (connectivity dependence).
Evidence Required | Excursion-response scenario report; rescue-action effectiveness; latency observations.
Owner | Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | The alerting benefit is conditional on connectivity (RF Open, PRD-VER-005) — validated honestly.
VAL-USE-012 — Recovery from device fault / power loss / offline resync
Objective (user outcome) | After a power loss, dead-in-cold, or connectivity blackout, users recover the trip and its record without a silent gap or a false clean pass.
Intended Use / User | Field users and technicians (SH-05/SH-06/SH-03).
Validation Method | Simulated-use fault/recovery scenarios.
Participants / Representative Users | Representative field users and technicians.
Environment / Context of Use | Induced power loss / offline periods / resync at receipt or depot.
Success / Acceptance Criteria | Any monitoring gap surfaces as an honest gap (never a clean pass); buffered data resyncs correctly; users interpret the recovered state correctly.
User Need Traceability | UN-SAFE-003/010; UN-FUN-009; UN-FLOW-004; UN-UX-012.
Risk / Use-Error Relationship | HA-SW-013 / HA-ENRG-006 (silent gap); validates HA-RES-004/018 acceptability with users.
Evidence Required | Recovery-scenario report; gap-comprehension; resync correctness.
Owner | HF/Usability + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Ties the D3 energy-margin Open item (PRD-VER-004) — dead-in-cold must surface honestly.
VAL-USE-013 — Exceptional / emergency situation: campaign-scale excursion under stock-out pressure
Objective (user outcome) | Under a mass/campaign excursion and stock-out pressure, users still make correct, defensible dispositions — neither releasing compromised lots nor needlessly destroying good ones.
Intended Use / User | Program managers and clinic staff (SH-01/SH-05) in a campaign.
Validation Method | Simulated-use high-stakes scenario + interview.
Participants / Representative Users | Representative program/clinic decision-makers.
Environment / Context of Use | Simulated campaign with scarcity and multiple concurrent excursions.
Success / Acceptance Criteria | Correct dispositions maintained under pressure; overrides remain deliberate/recorded; false-reject/wastage bounded; decision defensibility preserved.
User Need Traceability | UN-SAFE-001; UN-JTBD-002; UN-VAL-005; HA-RES-019 (wastage trade).
Risk / Use-Error Relationship | HA-USE-007/012 under pressure; validates the accepted-trade residual HA-RES-019 and HAL-2900 both-directions balance.
Evidence Required | Emergency-scenario report; disposition correctness under pressure; wastage measure.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Validates the S5-scale campaign case (mass false-reject) as well as false-accept.
VAL-USE-014 — Maintenance: recharge-to-ready & safe firmware update
Objective (user outcome) | A technician recharges a returned device and knows when it is trip-ready, and updates firmware knowing it succeeded — without bricking or an unnoticed unmonitored window.
Intended Use / User | Cold-chain technician / IT admin (SH-06/SH-09).
Validation Method | Simulated-use (maintenance workflow).
Participants / Representative Users | Representative technicians and fleet admins.
Environment / Context of Use | Service depot; fleet OTA campaign.
Success / Acceptance Criteria | Trip-ready state is correctly recognized; firmware update completes with clear success/failure feedback and no bricking; users understand any unmonitored update window.
User Need Traceability | UN-FLOW-009/010; UN-FUN-017; UN-JTBD-010; UN-UX-009.
Risk / Use-Error Relationship | HA-CYBER-005 / HA-SW-010 (OTA); validates HA-RES-012 acceptability from the user/operator side.
Evidence Required | Maintenance-workflow report; readiness/OTA outcome log.
Owner | HF/Usability + Product.
Status | Planned / Evidence Pending.
Notes | OTA safety is verification/security-anchored; this validates the human operator outcome.
VAL-USE-015 — Installation / new-fleet onboarding without specialist effort
Objective (user outcome) | A customer onboards a new device fleet and integrates it without specialist effort, ready to run trips.
Intended Use / User | IT admin / installer / program (SH-09/SH-12/SH-01).
Validation Method | Simulated-use / out-of-box onboarding study.
Participants / Representative Users | Representative admins/installers at first-use.
Environment / Context of Use | Enterprise/program onboarding; OEM embedding.
Success / Acceptance Criteria | Fleet onboarded and first trip provisioned without vendor hand-holding; onboarding-time acceptable, Evidence-Pending.
User Need Traceability | UN-FLOW-001; UN-STK-003/011; UN-ECO.
Risk / Use-Error Relationship | Misconfiguration at onboarding → downstream void/mis-armed trips (HA-USE-003/004).
Evidence Required | Onboarding-study report; task-completion; time.
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Onboarding friction is an enterprise/OEM adoption gate.
VAL-USE-016 — Foreseeable misuse: non-malicious circumvention resists casual bypass
Objective (user outcome) | A user under pressure who tries to make a trip "pass" (mute alerts, edit thresholds, override) cannot casually or accidentally do so, and any bypass is deliberate, recorded, and detectable.
Intended Use / User | Any operational user under production/stock-out pressure.
Validation Method | Simulated-use misuse/abuse scenario.
Participants / Representative Users | Representative operational users instructed to try to force a pass.
Environment / Context of Use | Simulated pressure to bypass a control.
Success / Acceptance Criteria | Casual bypass is prevented; alarm suppression self-escalates; config changes are authenticated/audited; override is deliberate/recorded; users cannot silently defeat the safety chain.
User Need Traceability | UN-SAFE-001/003/006; HA-CTX-003 (foreseeable misuse).
Risk / Use-Error Relationship | HA-USE-012 (deliberate bypass); ties HA-CYBER-002/003; validates HA-RES-014 non-malicious twin.
Evidence Required | Misuse-scenario report; bypass-resistance observations; audit-trail review.
Owner | HF/Usability + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Non-malicious twin of the cyber controls; residual depends on accountability policy.
VAL-USE-017 — Return & re-provision workflow closes the reusable loop
Objective (user outcome) | Field users return devices through a simple defined workflow and depots re-provision them for the next trip, keeping the reusable fleet moving.
Intended Use / User | Field users, service/reverse-logistics, technicians (SH-05/SH-13/SH-06).
Validation Method | Simulated-use + field pilot (return-loop) — ties VAL-BIZ return-rate pilot.
Participants / Representative Users | Representative field users and depot staff.
Environment / Context of Use | Last-mile return; depot reconditioning/re-provisioning.
Success / Acceptance Criteria | Return workflow completed by field users; re-provisioning restores a trip-ready device; return-rate target Evidence-Pending (measured in pilot).
User Need Traceability | UN-FLOW-007/008/009; UN-JTBD-010/011; UN-SVC; UN-VAL-009.
Risk / Use-Error Relationship | HA-USE-010/011 (cleaning/recal on reuse); ties reuse-economics viability.
Evidence Required | Return-loop workflow report; return-rate & turnaround data (pilot).
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Return rate is the swing variable in reuse TCO (conflict C6) — measured, not assumed.
VAL-USE-018 — Secondary workflow: fleet monitoring & audit-record export on demand
Objective (user outcome) | Fleet watchers see only what needs action and produce an audit-ready export on demand, so oversight and defence of dispositions are effortless.
Intended Use / User | Distributor/IT/program fleet users (SH-03/SH-09/SH-01); auditor (SH-07).
Validation Method | Simulated-use (dashboard + export) + audit walkthrough.
Participants / Representative Users | Representative fleet users and auditors.
Environment / Context of Use | Operations center; audit/inspection.
Success / Acceptance Criteria | Users surface action-required items and export a complete, attributable, shareable record accepted in the audit walkthrough.
User Need Traceability | UN-JTBD-006/007/008; UN-FUN-015/018; UN-FLOW; UN-VAL-006.
Risk / Use-Error Relationship | HA-DATA-006 (record loss); ties HA-RES-018/020.
Evidence Required | Dashboard/export task records; audit walkthrough notes.
Owner | Product + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Shares evidence with VAL-CLIN audit simulation and VAL-ECO dashboard validation.
7. Traceability notes
VAL-USE-001…009 are the exact Validation References cited in HA-USE-001…009 (HAL-1500) and are the summative critical-task closures for the use-related residual cluster HA-RES-014; VAL-USE-007 also closes HA-USE-012. VAL-USE-010…018 cover the end-to-end workflow, in-trip response, recovery, emergency, maintenance, installation, misuse, and secondary workflows, tracing to UN-JTBD/UN-FLOW/UN-FUN/UN-SAFE. Each critical task is paired with its PV-USE-00N verification and runs only after it passes.
8. Acceptance / summary notes
The intended-use workflows are validated as a coherent whole, centered on the dominant excursion detect→understand→act outcome. VAL-USE-001 and VAL-USE-004 are the two most load-bearing studies (the pathway choke points, HAL-2900 §5.2/condition 1). All activities are Planned / Evidence Pending; no workflow outcome is asserted, and every residual use-error acceptability judgment awaits the summative evidence.
VAL-USER-200 (PVaL-200) — User Validation
Product Validation Library (PVaL) — SCM · Stage 14 (ORCH-015) · Module PVaL-200, activity code VAL-USER. Owner: HF/Usability lead, with Clinical/Operational and Product. This module validates that each intended user class can do the job the SCM asks of them, in their real context — the primary/critical last-mile and program users, the operational users (cold-store, dispatch, distributor, 3PL), the sustainment users (technician, service/reverse-logistics, installer), and the governance/enterprise users (QA/RA, regulator, IT admin, procurement, finance, executive). It confirms fit for the person, not the feature. The single most important activity is VAL-USER-001 — the minimally trained, offline last-mile health worker. Status: Planned / Evidence Pending — no users have been tested.
---
1. Purpose
Validate that the SCM meets the needs of every user class it will pass through, with rigor proportional to consequence. A reusable device passes through a dozen hands on one trip (UN-100); each hand has a job, and the device fails commercially or in safety if any critical hand cannot do it. The module deliberately centers the last-mile low-training user (SH-05 / UN-PER-007) — the least-trained operator performing the most safety-critical task — and confirms the others (program, operational, sustainment, governance) can do their jobs without specialist effort.
2. Scope
The fifteen stakeholder classes (SH-01…SH-15) reduced to distinct validation user groups by the job they do: (1) last-mile health worker (vulnerable/low-training); (2) EPI/program manager; (3) cold-store/warehouse operator; (4) manufacturer QA/RA (dispatch); (5) distributor coordinator; (6) 3PL/air-cargo operator; (7) cold-chain/calibration technician; (8) service/reverse-logistics operator; (9) IT/security admin; (10) installer/OEM integrator; (11) regulator/auditor; (12) procurement (GAVI/UNICEF); (13) finance; (14) executive/supply-chain leadership; (15) ecosystem partner (OEM/reseller/SI). The beneficiary (SH-11) is represented, not tested. Accessibility, literacy, language, and training-decay diversity are validated within each relevant group.
3. Applicable standards & references
• IEC 62366-1 / FDA HF guidance — user profiles, distinct user groups, summative evaluation per group.
• ISO 13485 §7.3.7 — representative users. WHO PQS E006 — the program/operator context.
Upstream | UN-STK-001…015 (anchor need per class), UN-PER-001…011 (personas), UN-VAL-002/003/004, HA-CTX-005 (vulnerable user), HAL-1500 (the critical-task use errors these users can commit).
4. Governance & approach
• Rigor follows consequence. The last-mile disposition user (VAL-USER-001) and provisioning users get summative studies; governance/commercial users get interviews/surveys/task walkthroughs sufficient to confirm adoption fit.
• Test the untrained user as they will really be. Per the warnings-challenge rule, the last-mile study uses minimally trained / training-decayed users; a result that only holds for freshly trained users does not validate the critical task.
• Accessibility & inclusivity in-group. Literacy levels, non-primary UI languages, glove use, low vision / color-vision deficiency, and age are quotas within the relevant groups, not a separate afterthought.
• No convenience proxies for critical users. Engineers/employees may not stand in for the last-mile or clinic users.
5. Validation categories
Primary/critical users · secondary operational users · sustainment users · governance & enterprise users · accessibility & inclusivity · represented non-users (beneficiary).
6. Validation Register (VAL-USER activities)
VAL-USER-001 — Last-mile health worker (minimally trained, gloved, offline) reaches the correct disposition ★ dominant user
Objective (user outcome) | A minimally trained last-mile nurse/health worker, gloved and without a phone, network, or manual, reads the device at receipt and correctly decides fridge / quarantine / return — and is never led to accept a rejected lot or reject an accepted one.
Intended Use / User | Clinic/last-mile health worker (SH-05, vulnerable/low-training); protects the beneficiary (SH-11).
Validation Method | Summative usability test (IEC 62366-1) — the module's critical study (shared root with VAL-USE-001 / VAL-HF-001).
Participants / Representative Users | ≥15 representative low-training health workers per distinct sub-group; include low-literacy, non-primary-language, gloved, and training-decayed participants.
Environment / Context of Use | Simulated rural health post — low light and glare, cold-soaked device, gloves, offline, time pressure; a subset at real health posts (field, PVaL-900).
Success / Acceptance Criteria | Users reach the correct accept/reject/quarantine disposition on scripted clean, freeze, heat, and gapped/void trips; no critical use error (accepting a reject / rejecting an accept) is unrecovered; residual use errors root-caused and judged acceptable per IEC 62366-1; use-error rate target Evidence-Pending (set in protocol).
User Need Traceability | UN-STK-006; UN-PER-007; UN-VAL-003; UN-SAFE-001/005; UN-UX-001/002/003; UN-FUN-008; UN-FLOW-005; UN-JTBD-002.
Risk / Use-Error Relationship | HA-USE-001 (misread verdict), HA-USE-004 (void trip read as clean), HA-USE-008 (language); HA-RES-014; validates HAL-2900 condition 1 (supported human decision).
Evidence Required | Summative protocol & report; video observations; use-error root-cause analysis; correct-disposition rates by scenario/user sub-group.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | THE validation of the device's reason to exist; gates the benefit-risk conclusion.
VAL-USER-002 — EPI / immunization program manager judges consignments correctly & defensibly
Objective (user outcome) | A program manager applies the SCM's PQS-consistent accept/reject verdict to arriving consignments and makes correct, defensible, program-wide disposition decisions.
Intended Use / User | EPI/immunization program manager & MoH (SH-01).
Validation Method | Simulated-use test + interview.
Participants / Representative Users | Representative EPI/program managers across ≥2 programs/countries.
Environment / Context of Use | Program receiving office / national store; dashboard + device verdict.
Success / Acceptance Criteria | Managers correctly disposition scripted consignments (incl. partial/ gapped records), can defend each decision, and find the verdict PQS-consistent and program-usable; no systematic mis-disposition.
User Need Traceability | UN-STK-005; UN-PER-006; UN-JTBD-002/007; UN-REG (PQS).
Risk / Use-Error Relationship | HA-USE-007 (override), HA-DATA-004; validates HA-RES-002 acceptability at program scale.
Evidence Required | Simulated-use records; interview transcripts; disposition-correctness log.
Owner | Clinical/Operational + HF/Usability.
Status | Planned / Evidence Pending.
Notes | Program-level decision fit; complements the individual last-mile study.
VAL-USER-003 — Cold-store / warehouse operator provisions, arms & places gloved and cold
Objective (user outcome) | A cold-store operator, gloved and in a freezer, provisions, arms, confirms placement, and seals a shipper quickly and without a void/mis-armed trip.
Intended Use / User | Warehouse/cold-store operators (SH-05).
Validation Method | Simulated-use test + time-and-motion.
Participants / Representative Users | ≥15 representative cold-store operators, gloved.
Environment / Context of Use | Cold store (−20 °C) / freezer staging; gloved; high throughput.
Success / Acceptance Criteria | Operators complete provision→arm→place→confirm without a mis-arm/void, the compliant path is the fastest path, and cold/gloved operation succeeds; provisioning-time target Evidence- Pending.
User Need Traceability | UN-PER-005; UN-VAL-004; UN-FLOW-002/003; UN-UX-005; UN-FUN-012/013.
Risk / Use-Error Relationship | HA-USE-003/004/005 (profile/arm/placement); validates HA-RES-014.
Evidence Required | Time-and-motion data; mis-arm/void incidence; observation records.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | Speed is a covert adoption gate (conflict C4); compliant-fastest is the design bet.
VAL-USER-004 — Manufacturer QA/RA releases with a defensible record without slowing dispatch
Objective (user outcome) | QA/RA staff release each shipment with a record they can defend to any inspector, produced without adding time that slows high-throughput pack-out.
Intended Use / User | Manufacturer QA/RA (SH-02).
Validation Method | Simulated-use + interview at pack-out.
Participants / Representative Users | Representative QA/RA and pack-out staff.
Environment / Context of Use | Manufacturer pack-out line; high throughput.
Success / Acceptance Criteria | Release completed within pack-out takt; record judged defensible; no workaround/skip behavior induced by friction.
User Need Traceability | UN-STK-001; UN-PER-001; UN-JTBD-004/007; UN-UX-005.
Risk / Use-Error Relationship | Friction → skipped/void trips (HA-USE-004); ties HA-RES-014.
Evidence Required | Cycle-time data; interview transcripts; defensibility judgments.
Owner | Product + HF/Usability.
Status | Planned / Evidence Pending.
Notes | QA manager is the economic approver; adoption depends on not slowing the line.
VAL-USER-005 — Distributor coordinator sees status & acts only on action-required alerts
Objective (user outcome) | A coordinator watching hundreds of concurrent shipments sees each shipment's status and acts on the alerts that matter — without interpreting raw data and without alarm fatigue.
Intended Use / User | Distributor/wholesaler coordinator (SH-03).
Validation Method | Simulated-use (dashboard) + field study under real alert load.
Participants / Representative Users | Representative distributor coordinators.
Environment / Context of Use | Distributor operations center; multi-shipment dashboard.
Success / Acceptance Criteria | Coordinators correctly triage action-required vs informational alerts, act on genuine excursions within the actionable window under realistic load, and do not habitually mute; false-alarm tolerance target Evidence-Pending.
User Need Traceability | UN-STK-002; UN-PER-002; UN-JTBD-006/009; UN-UX-006; UN-SAFE-004.
Risk / Use-Error Relationship | HA-USE-002 (alarm fatigue); validates HA-RES-014 / HA-RES-019 (both directions).
Evidence Required | Triage-accuracy data; mute-rate; time-to-act; field observations.
Owner | HF/Usability + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | Daily active user; alert fatigue is the failure mode to disprove.
VAL-USER-006 — 3PL / air-cargo operator: device never disrupts handling or transport
Objective (user outcome) | 3PL handlers handle and fly shippers with the device inside without extra steps, delays, or dangerous-goods hold-ups.
Intended Use / User | 3PL/logistics/air-cargo operations (SH-04).
Validation Method | Field study / operational walkthrough + interview.
Participants / Representative Users | Representative 3PL/air-cargo operations staff.
Environment / Context of Use | Sorting/handling hubs; air/ground transport legs.
Success / Acceptance Criteria | Device is handling-neutral (no handler action required), accepted for air/ground transport with correct Li-battery documentation, and never the reason for a hold.
User Need Traceability | UN-STK-004; UN-PER-004; UN-FLOW-004; UN-SAFE-009.
Risk / Use-Error Relationship | Ties HA-TFC-002 transport safety (validated via UN 38.3/IATA compliance); channel-exclusion risk.
Evidence Required | Operational walkthrough notes; hold-incidence; transport-acceptance evidence.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | 3PL SOP inclusion multiplies pull-through; transport safety is verification-anchored.
VAL-USER-007 — Cold-chain / calibration technician turns a returned device around fast & verifiably
Objective (user outcome) | A technician quickly confirms a returned device is charged, in-calibration, and firmware-current and re-provisions it, without hand-checking each unit, keeping the fleet moving.
Intended Use / User | Cold-chain/calibration technician / field service (SH-06).
Validation Method | Simulated-use (depot workflow) + time study.
Participants / Representative Users | Representative service/calibration technicians.
Environment / Context of Use | Service depot; batch turnaround.
Success / Acceptance Criteria | Technicians correctly identify charge/calibration/firmware readiness (including a cold charge-inhibited device — warm-then-charge, not discard) and re-provision within a target turnaround; no serviceable device wrongly retired.
User Need Traceability | UN-STK-010; UN-PER-011; UN-JTBD-010; UN-FLOW-009; UN-FUN-017.
Risk / Use-Error Relationship | HA-USE-009 (charge-inhibit mistaken for dead), HA-USE-011 (reuse without recal); validates HA-RES-014 availability path.
Evidence Required | Turnaround-time data; readiness-identification accuracy; misidentification incidence.
Owner | HF/Usability + Product.
Status | Planned / Evidence Pending.
Notes | Turnaround sets utilization and reuse margin; ties charge-inhibit state legibility (VAL-USE-011).
VAL-USER-008 — Service / reverse-logistics operator recovers, cleans & closes the loop safely
Objective (user outcome) | Reverse-logistics staff recover devices, clean them between trips without damage, and dispose of end-of-life units/batteries safely and compliantly.
Intended Use / User | Service/reverse-logistics operations (SH-13).
Validation Method | Simulated-use (reconditioning) + procedure walkthrough.
Participants / Representative Users | Representative reverse-logistics/reconditioning staff.
Environment / Context of Use | Depot reconditioning line; return logistics.
Success / Acceptance Criteria | Cleaning/decontamination and EOL/battery disposal completed per procedure without device damage or missed cleaning; return workflow usable at throughput.
User Need Traceability | UN-STK-012; UN-JTBD-011/012; UN-FLOW-007/008/011; UN-SVC.
Risk / Use-Error Relationship | HA-USE-010 (missed cleaning); ties HA-SVC / Li-disposal (HA-SVC-007).
Evidence Required | Reconditioning walkthrough; cleaning-completion audit; EOL/disposal compliance.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | Return rate is the swing variable in reuse TCO (conflict C6) — ties VAL-BIZ.
VAL-USER-009 — IT / security administrator can validate, govern & integrate at fleet scale
Objective (user outcome) | An enterprise IT/security admin can satisfy computerized-system validation, configure SSO/RBAC and retention, and integrate the platform securely — clearing the bar to allow it on the network.
Intended Use / User | IT/security administrators (SH-09).
Validation Method | Simulated-use (admin console) + structured evaluation.
Participants / Representative Users | Representative enterprise IT/security admins.
Environment / Context of Use | Enterprise onboarding; ongoing fleet administration.
Success / Acceptance Criteria | Admins complete onboarding, RBAC/SSO, retention config, and integration without vendor hand-holding; the platform passes the admins' validation/security checklist.
User Need Traceability | UN-STK-003; UN-PER-003; UN-SEC; UN-ECO-005.
Risk / Use-Error Relationship | Uncontrolled access / unsigned updates → record/disposition corruption (ties HA-CYBER, HA-RES-013); conflict C9.
Evidence Required | Admin task-completion records; security/CSV checklist results.
Owner | Product + EM-ENG-006.
Status | Planned / Evidence Pending.
Notes | A siloed device stalls at procurement (conflict C9); enterprise adoption gate.
VAL-USER-010 — Installer / OEM integrator fits, mounts & guides correct placement
Objective (user outcome) | An installer/OEM integrator mounts the device securely inside insulated shippers/containers and places the probe so readings reflect the product zone.
Intended Use / User | Installers / OEM integrators (SH-12).
Validation Method | Simulated-use (installation) + interview.
Participants / Representative Users | Representative installers/OEM integrators.
Environment / Context of Use | Shipper/container installation; OEM embedding.
Success / Acceptance Criteria | Installers achieve secure fit and correct probe placement following the provided guidance; placement errors caught before sealing.
User Need Traceability | UN-STK-011; UN-FLOW-003; UN-FUN-013; UN-UX-010; UN-PHY.
Risk / Use-Error Relationship | HA-USE-005/006 (probe placement/channel); validates HA-RES-003/014 placement remainder.
Evidence Required | Installation observation; placement-error incidence; fit assessment.
Owner | HF/Usability + Product.
Status | Planned / Evidence Pending.
Notes | Poor fit blocks OEM embedding (a key channel); flag mechanical/materials neighbours.
VAL-USER-011 — Regulator / auditor finds the record the easiest artifact in the audit
Objective (user outcome) | An auditor verifies data integrity, calibration currency, and excursion handling from records that are complete, attributable, and unambiguous — without friction.
Intended Use / User | Regulators / QA / auditors (SH-07).
Validation Method | Audit-simulation study with real auditors (shared with VAL-CLIN audit sim).
Participants / Representative Users | Representative GDP/Part 11 auditors and QA reviewers.
Environment / Context of Use | Simulated audit / inspection; export + dashboard review.
Success / Acceptance Criteria | Auditors reconstruct the trip, verify ALCOA+ attributes, and judge the record admissible; no ambiguity that would invalidate a disposition.
User Need Traceability | UN-STK-007; UN-PER-009; UN-VAL-006; UN-JTBD-007/008; UN-REG; UN-SEC.
Risk / Use-Error Relationship | HA-DATA-006 / HA-MPE-004 (record integrity); validates HA-RES-018/020.
Evidence Required | Audit-simulation report; auditor findings; admissibility judgments.
Owner | EM-ENG-006 + Clinical/Operational.
Status | Planned / Evidence Pending.
Notes | "Make the record the easiest artifact in the audit" — a differentiator (UN-PER-009).
VAL-USER-012 — Procurement (GAVI/UNICEF) can qualify the device into a pooled tender
Objective (user outcome) | Institutional procurement confirms the device is PQS-eligible and carries the reuse-TCO and supply-assurance evidence needed to place it into a pooled tender.
Intended Use / User | Procurement / GAVI / UNICEF (SH-08).
Validation Method | Interview + evidence-review (documentary validation).
Participants / Representative Users | Representative institutional procurement officers.
Environment / Context of Use | Tender qualification / framework award.
Success / Acceptance Criteria | Procurement judges the PQS eligibility path and TCO/supply evidence sufficient to qualify (contingent on PQS listing, Evidence-Pending, D4).
User Need Traceability | UN-STK-008; UN-PER-008; UN-VAL-009; UN-REG (PQS); UN-BIZ.
Risk / Use-Error Relationship | No direct hazard; gates whether the safety benefit reaches LMIC scale.
Evidence Required | Procurement interview; evidence-completeness assessment; PQS-path status.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Largest aggregated LMIC channel; PQS listing has long lead times — plan early.
VAL-USER-013 — Finance & executive get the fleet ROI / cost-per-trip evidence they need
Objective (user outcome) | Finance can justify capital outlay against loss reduction, and executives see network-level loss reduction, compliance exposure, and cost-per-trip to fund and standardize the program.
Intended Use / User | Finance/asset accounting (SH-10); executive/supply-chain leadership (SH-15).
Validation Method | Interview + business-case evidence review (ties VAL-BIZ).
Participants / Representative Users | Representative finance and supply-chain leaders.
Environment / Context of Use | Business-case approval; multi-year contract decisions.
Success / Acceptance Criteria | Decision-makers judge the fleet reporting and cost-per-trip evidence sufficient to fund/standardize; return-rate/TCO figures come from the pilot, Evidence-Pending.
User Need Traceability | UN-STK-009/015; UN-PER-010; UN-VAL-009; UN-BIZ; UN-ECO fleet KPIs.
Risk / Use-Error Relationship | No direct hazard; funds deployment of the safety benefit at scale.
Evidence Required | Interviews; business-case validation notes.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Executive cares about the network, not one device; distinct from QA needs.
VAL-USER-014 — Ecosystem partner (OEM/reseller/SI) can embed, integrate & support in-region
Objective (user outcome) | A partner can build in, resell, integrate, and support the device in their market using the form factor, integration points, provisioning tooling, and enablement provided.
Intended Use / User | Ecosystem partners — OEMs/resellers/SIs (SH-14).
Validation Method | Partner pilot / integration study + interview.
Participants / Representative Users | Representative OEM/reseller/SI partners.
Environment / Context of Use | Partner onboarding, integration, regional support.
Success / Acceptance Criteria | Partners complete an integration and support workflow with the provided tooling; enablement judged sufficient for in-region support.
User Need Traceability | UN-STK-013; UN-ECO API/partner needs; UN-PHY embeddable form.
Risk / Use-Error Relationship | Open integration vs security lockdown (conflict C9); indirect data-integrity.
Evidence Required | Integration-study report; partner interviews.
Owner | Product/Business.
Status | Planned / Evidence Pending.
Notes | Designed-in distribution channel into fragmented markets.
VAL-USER-015 — Accessibility & inclusivity across low-literacy, non-primary-language, low-vision & gloved users
Objective (user outcome) | Users at the margins — low literacy, non-primary/absent-translation UI language, low vision or color-vision deficiency, aging, gloved — still reach the correct disposition and operate the core tasks.
Intended Use / User | The inclusive tail of SH-05/SH-01 last-mile and operational users.
Validation Method | Summative usability test (accessibility-focused sub-study of VAL-USER-001/USE-001).
Participants / Representative Users | Quota-sampled low-literacy, non-primary-language, low-vision/ color-vision-deficient, and older users; gloved.
Environment / Context of Use | Representative low-light/cold/gloved field conditions across locales.
Success / Acceptance Criteria | These users reach correct dispositions via the icon+colour+word encoding without relying on reading text or a manual; no critical use error concentrated in an accessibility group.
User Need Traceability | UN-UX-003; UN-VAL-003; UN-SAFE-005; HA-IUE-010.
Risk / Use-Error Relationship | HA-USE-001/008 (misread/language); validates the language-independent design control (RCR-043) with the users it exists for.
Evidence Required | Accessibility sub-study report; per-group correct-disposition rates.
Owner | HF/Usability (lead).
Status | Planned / Evidence Pending.
Notes | The warnings-challenge rule demands the design work for these users, not just literate ones.
7. Traceability notes
Each activity anchors to a stakeholder need (UN-STK), a persona (UN-PER), and the jobs/flows/UX/functional needs that class exercises. The critical users (VAL-USER-001/003/010) share summative roots with VAL-USE and VAL-HF; the governance/commercial users (VAL-USER-011/012/013/014) feed VAL-CLIN and VAL-BIZ. Every safety- relevant activity ties to the HA-USE errors that user can commit and the HA-RES residual it helps validate as acceptable.
8. Acceptance / summary notes
The user set is completely represented — no class is silently dropped and the beneficiary is explicitly represented not tested. The dominant activity is VAL-USER-001 (last-mile disposition). All activities are Planned / Evidence Pending; no user acceptance is asserted.
Validation Coverage Matrix — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 14 — Validation (ORCH-015) · Owner: EM-ENG-006 + HF/Usability · Status: Planned / Evidence Pending (pre-execution). Companion to index.md and [Validation-Plan.md](Validation-Plan.md). This matrix maps User Need / Intended Use → validation activity → method → outcome-based success criterion → status, and explicitly identifies coverage gaps: user needs without validation, safety-critical workflows not validated, and use errors not validated.
Status legend: P/EP = Planned / Evidence Pending · Open = release-gating engineering unknown carried open (D3/RF/D4). All statuses are pre-execution; no result is asserted.
---
1. Dominant-outcome coverage (the SCM's reason to exist)
Outcome | VAL activity | Method | Success criterion (outcome) | Status
Read verdict → correct disposition (gloved/cold/low-light/low-literacy) | VAL-USE-001 / VAL-USER-001 / VAL-HF-001/002/003 | Summative usability | Correct disposition on clean/freeze/heat/gapped; no unrecovered critical use error | P/EP
Act on a genuine excursion alert in time (no alarm fatigue) | VAL-USE-002 / VAL-HF-005/006 | Summative under alert load | Acts within actionable window; no habitual muting; escalates on no-ack | P/EP
Never trust a void/unmonitored trip as clean (false reassurance) | VAL-USE-004 / VAL-HF-008 | Summative | No false-clean-accept over gap/void; honest-gap comprehended | P/EP
No accidental override of a reject | VAL-USE-007 / VAL-HF-007 | Summative under pressure | No accidental override; deliberate override recorded & understood | P/EP
Residual of the use-error cluster acceptable | VAL-RISK-001 | Residual determination | ALARP, region A, benefit-risk signed | P/EP
Overall benefit-risk acceptable (both directions) | VAL-RISK-003 | Benefit-risk determination | HAL-2900 seven conditions closed on demonstrated evidence | P/EP
---
2. User-need module → validation coverage
User-need module | Representative needs | VAL activity(ies) | Method | Status
UN-000 Vision (VIS) | UN-VIS-001 (severity anchor) | VAL-CLIN-001/007, VAL-RISK-003, VAL-SUM-005 | Field/benefit-risk | P/EP
UN-100 Stakeholders (STK) | UN-STK-001…015 | VAL-USER-001…014 (per class) | Summative/field/interview | P/EP
UN-200 Personas (PER) | UN-PER-001…011 | VAL-USER-001…013; VAL-HF | Summative/interview | P/EP
UN-300 JTBD | UN-JTBD-001…012 | VAL-USE-010/011/018; VAL-CLIN-008 | End-to-end/field | P/EP
UN-400 Environments (ENV) | UN-ENV-001…009 | VAL-REL-001…007; VAL-HF-002; VAL-ECO-003/004 | Field/reliability | P/EP (D3/RF Open)
UN-500 Workflow (FLOW) | UN-FLOW-001…011 | VAL-USE-003/004/005/010/014/015/017 | Summative/simulated | P/EP
UN-600 Functional (FUN) | UN-FUN-001…018 | VAL-USE, VAL-CLIN-002/008, VAL-ECO-006/007 | Operational/simulated | P/EP
UN-700 UX | UN-UX-001…012 | VAL-HF-001…011; VAL-USE-001 | Summative usability | P/EP
UN-800 Performance (PERF) | endurance/accuracy/latency | VAL-REL-001/004; VAL-CLIN-002; VAL-ECO-004/005 | Field/reliability | P/EP (D3/RF Open)
UN-900 Safety (SAFE) | UN-SAFE-001…010 | VAL-USE-001…009; VAL-HF-001; VAL-RISK-001…008 | Summative/residual | P/EP
UN-1000 Security & Privacy (SEC) | access/OTA/integrity | VAL-USER-009; VAL-ECO-009/010; VAL-RISK-009 | IV/integration | P/EP
UN-1100 Physical (PHY) | fit/mounting/display | VAL-USER-010; VAL-REL-006; VAL-HF-002 | Field/summative | P/EP
UN-1200 Service (SVC) | turnaround/cleaning/recal | VAL-USER-007/008; VAL-USE-017; VAL-REL-005/008; VAL-FIELD-008 | Field/simulated | P/EP
UN-1300 Regulatory (REG) | GDP/Part 11/PQS | VAL-USER-011; VAL-CLIN-006/010; VAL-RISK-007 | Audit sim/confirmation | P/EP (PQS Open)
UN-1400 Quality (QUAL) | testable acceptance | VAL-GOV-005; VAL-EVID-001 | Governance | P/EP
UN-1500 Ecosystem (ECO) | cloud/mobile/fleet/API | VAL-ECO-001…010 | Summative/field | P/EP
UN-1600 Business (BIZ) | TCO/adoption/support | VAL-BIZ-001…009; VAL-USER-012/013/014 | Survey/pilot | P/EP
UN-1700 Validation (VAL) | UN-VAL-001…009 | see §3 (direct map) | — | P/EP
---
3. UN-VAL (validation needs) → validation activity (direct map)
User need | What it demands | VAL activity | Method | Success criterion | Status
UN-VAL-001 | Testable acceptance basis for every need | VAL-GOV-003/005; VAL-EVID-001 | Governance/pre-registration | Every need has an outcome-based criterion; pre-registered | P/EP
UN-VAL-002 | Real users confirm fit in context | VAL-USER-; VAL-CLIN-001; VAL-FIELD-001/002/003 | Field/pilot | Users confirm fit in real context | P/EP
UN-VAL-003 | Novice offline user reaches right disposition | VAL-USE-001; VAL-HF-001; VAL-USER-001 | Summative | Correct disposition, no unrecovered critical error | P/EP
UN-VAL-004 | Gloved/cold + dispatch tasks succeed | VAL-USER-003; VAL-USE-003; VAL-HF-002 | Summative/time-motion | Cold/gloved provisioning succeeds; no void | P/EP
UN-VAL-005 | Two-sided detection & correct disposition validated | VAL-CLIN-002; VAL-USE-001; VAL-RISK-004 | Operational/summative | Freeze=heat; no false accept; false-reject bounded | P/EP
UN-VAL-006 | Audit defensibility with real auditors | VAL-USER-011; VAL-CLIN-006 | Audit simulation | Record admissible; disposition defensible | P/EP
UN-VAL-007 | Cold-endurance closed by test (90d/−30°C) | VAL-REL-001; VAL-RISK-011 | Reliability/field | Endurance met in use; honest gap on shortfall | Open (D3)
UN-VAL-008 | Offline resilience & alert latency vs real connectivity | VAL-ECO-003/004/005; VAL-REL-004; VAL-FIELD-005 | Field | Offline verdict correct; latency characterized | Open (RF)
UN-VAL-009 | Reuse economics (return rate/cost-per-trip) by pilot | VAL-BIZ-006; VAL-USE-017; VAL-FIELD-006 | Reuse pilot | Return/turnaround/cost-per-trip within viability | P/EP
No UN-VAL need is without a validation activity.
---
4. Use-error (HA-USE) → validation reference (safety-critical closure)
Use error (HAL-1500) | Validation reference | Method | Success criterion | Status
HA-USE-001 Misread verdict | VAL-USE-001 / VAL-HF-001/002/003 | Summative | Correct disposition; no misread-driven wrong action | P/EP
HA-USE-002 Ignored/muted alert (fatigue) | VAL-USE-002 / VAL-HF-005/006 | Summative | Acts within window; no habitual mute | P/EP
HA-USE-003 Wrong trip profile | VAL-USE-003 / VAL-HF-011 | Summative | Correct profile; wrong one caught | P/EP
HA-USE-004 Void trip believed monitored | VAL-USE-004 / VAL-HF-008 | Summative | No false-clean-accept over void | P/EP
HA-USE-005 Probe not in product zone | VAL-USE-005 / VAL-USER-010 | Summative | Correct placement; unrepresentative caught | P/EP
HA-USE-006 Wrong probe/channel | VAL-USE-006 | Usability (subset) | Correct probe/channel; wrong caught | P/EP
HA-USE-007 Override a reject | VAL-USE-007 / VAL-HF-007 | Summative | No accidental override; deliberate understood | P/EP
HA-USE-008 Confusing labelling/language | VAL-USE-008 / VAL-HF-003; VAL-USER-015 | Multi-locale summative | Comprehension across locales | P/EP
HA-USE-009 Charge-inhibit mistaken for dead | VAL-USE-009 / VAL-USER-007 | Summative | Warm-then-charge, not discard | P/EP
HA-USE-010 Missed cleaning (reuse) | VAL-USE-017 / VAL-FIELD-008 (+ PV-USE-010) | Process/field | Cleaning completed; post-service BIST | P/EP
HA-USE-011 Reuse without recalibration | VAL-REL-008 (+ PV-USE-011 expiry gate) | Reliability/gate | Out-of-cal cannot arm; no drift escape | P/EP
HA-USE-012 Deliberate bypass | VAL-USE-007 / VAL-USE-016 / VAL-RISK-005 | Misuse scenario | Casual bypass prevented; deliberate recorded | P/EP
Every HA-USE use error has a validation reference. HA-USE-010/011 are design/process-controlled (their HAL records list validation N/A for the disposition-critical-task sense); they are covered by reconditioning/ expiry-gate verification plus the reuse-loop (VAL-USE-017) and reliability-in-use (VAL-REL-008) validations.
---
5. Residual risk (HA-RES) → validation reference (residual acceptability)
Residual (HAL-2800) | Validation reference | Success criterion | Status
HA-RES-001 Missed excursion | VAL-CLIN-002 (VAL-SW-001) | No missed excursion type in use | P/EP
HA-RES-002 False accept over gap | VAL-USE-004 / VAL-RISK-002 | No false-clean-accept (choke point) | P/EP
HA-RES-003 Probe fault as valid | VAL-USE-005/006 | Unrepresentative/wrong probe caught | P/EP
HA-RES-004 Silent gap / dead-in-cold | VAL-USE-012 / VAL-REL-001 / VAL-RISK-011 | Honest gap; endurance met | Open (D3)
HA-RES-005 Cold-charge plating | VAL-REL-002 (+ PV-ENRG-001) | Charge-inhibit holds over cycles | P/EP
HA-RES-006 Aircraft-hold fire (S5) | VAL-RISK-008 (+ PV-TFC-002/UN 38.3) | Disclosure + monitoring validated | P/EP
HA-RES-007 Delayed alert (latency) | VAL-USE-011 / VAL-ECO-004 / VAL-FIELD-005 | Acts within window where link allows | Open (RF)
HA-RES-008 Missing alert | VAL-ECO-005 (VAL-COMM-002) | No silent non-delivery; escalates | P/EP
HA-RES-009 In-container RF | VAL-REL-004 / VAL-RISK-011 | Delivery/verdict acceptable; claim restricted if not | Open (RF)
HA-RES-010…013 Security | VAL-RISK-009 (IV pen-tests) | Controls hold vs adversary (IV) | P/EP
HA-RES-014 Use-error cluster | VAL-RISK-001 (VAL-USE/VAL-HF summative) | ALARP, region A | P/EP
HA-RES-015 Calibration common cause | VAL-REL-008 / VAL-RISK-010 | No fleet-wide bias escape | P/EP
HA-RES-016 Mis-attribution/routing | VAL-ECO-006 (VAL-COMM-008); VAL-CLIN-008 | Correct attribution; reconcile | P/EP
HA-RES-017 Cloud notify/rules | VAL-ECO-007 | Cross-tier consistency | P/EP
HA-RES-018 Record/DB loss | VAL-ECO-006 / VAL-USER-011 | Loss surfaces as gap; admissible | P/EP
HA-RES-019 False-reject/wastage | VAL-CLIN-004 / VAL-RISK-004 | Wastage bounded; net benefit positive | P/EP
HA-RES-020 GDP/data-integrity | VAL-CLIN-006 / VAL-USER-011 | Record integrity in audit | P/EP
HA-RES-021 Condensation/ingress | VAL-REL-003 | No env-induced false verdict | P/EP
HA-RES-022 Time/units/estimate | VAL-CLIN-008 | Provenance disclosed & understood | P/EP
HAL-2900 Overall benefit-risk | VAL-RISK-003 (+ VAL-SUM-005) | Seven conditions closed | P/EP
Every top residual has a validation reference.
---
6. Safety-critical workflow → summative coverage
Safety-critical workflow | Summative activity | Status
Disposition read (accept/reject/quarantine) | VAL-USE-001 / VAL-HF-001 | P/EP
Excursion alert response | VAL-USE-002 / VAL-HF-005/006 | P/EP
Provisioning / trip-profile config | VAL-USE-003 / VAL-HF-011 | P/EP
Arming / avoid void-believed-monitored | VAL-USE-004 / VAL-HF-008 | P/EP
Probe placement / channel | VAL-USE-005/006 | P/EP
Verdict override | VAL-USE-007 / VAL-HF-007 | P/EP
Charge-inhibit state interpretation | VAL-USE-009 / VAL-USER-007 | P/EP
Cross-locale comprehension | VAL-USE-008 / VAL-HF-003 | P/EP
No safety-critical workflow is without a summative activity.
---
7. Coverage gaps & open items (named, not hidden)
Per the honesty posture, gaps and open items are stated explicitly rather than silently closed:
1. User needs without validation: None. Every UN module and every UN-VAL need maps to ≥1 activity.
2. Safety-critical workflows not validated: None. Every HA-USE critical task has a summative activity (VAL-USE-001…009); HA-USE-010/011 are design/process-controlled with reuse-loop/reliability validation.
3. Use errors not validated: None. All HA-USE-001…012 have a validation reference (§4).
4. Residuals not validated: None. All HA-RES-001…022 + the overall benefit-risk have a reference (§5).
5. Release-gating engineering unknowns carried Open (not validation gaps, but gates):
D3 cold endurance (90d/−30 °C) | PRD-VER-004 → VAL-REL-001 / VAL-RISK-011 — Open.
RF-in-container link margin | PRD-VER-005 → VAL-REL-004 / VAL-ECO-004/005 / VAL-FIELD-005 / VAL-RISK-011 — Open; if it cannot close, the real-time-alert claim is restricted and the on-receipt verdict is the compensating benefit.
WHO PQS accuracy/benefit target (D4) | VAL-CLIN-010 — Open; the benefit claim is provisional until confirmed and met.
6. Residuals with an out-of-device dependency (validated but not fully closable by device design): HA-USE-007 (clinical override) and HA-USE-012 (deliberate bypass) retain a programme disposition/ accountability-policy dependency — flagged in VAL-RISK-005 and the benefit-risk record.
7. All quantitative outcome targets (use-error rate, time-to-decision, false-reject rate, latency, NPS, return rate, cost-per-trip) are Evidence-Pending with a named closing protocol; none is asserted.
---
8. Coverage summary
• 128 validation activities across 13 modules; outcome-based, traced to User Needs / Intended Use and to residual-risk acceptability.
Structural coverage is complete | no user need, safety-critical workflow, use error, or top residual is without a validation reference.
• Execution coverage is zero — all Planned / Evidence Pending; the go-to-market recommendation is Not Yet Determined / Conditionally Ready — pending V&V (VAL-SUM-005), gated on the dominant-outcome summative studies, the residual/benefit-risk validation, the D3/RF Open items, the security IV, and the WHO PQS benefit target.
Validation Plan — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage: 14 — Validation (ORCH-015) · Owner: HF/Usability lead + Product · Approvers: EM-ENG-006 (Quality/Regulatory), EM-ENG-001 (Chief Systems) · Status: Planned / Evidence Pending (pre-execution). This plan operationalizes the PVaL library (index.md, VAL-GOV…VAL-SUM) into an executable program. It is the companion to [Validation-Coverage-Matrix.md](Validation-Coverage-Matrix.md).
---
1. Objective & dominant outcome
Validate that the SCM meets user needs and intended use in the real world, with the dominant outcome first: representative users — including the low-training, gloved, offline last-mile health worker — correctly detect, understand, and act on temperature excursions so no compromised dose is administered and no potent dose is wasted, with residual risks acceptable to those users. Everything else (operational effectiveness, ecosystem, reliability, business) is validated in support of and after that outcome.
2. Strategy (severity-first, outcome-based)
Priority order (Action-Priority-driven):
1. Summative human-factors critical tasks — VAL-USE-001/002/004/007 + VAL-HF-001 (the excursion detect→understand→act chain and the two pathway choke points).
2. Residual-risk / benefit-risk validation — VAL-RISK-001/002/003 + the HAL-2900 seven conditions.
3. Operational/clinical effectiveness — VAL-CLIN (both harm directions).
4. Ecosystem & reliability-in-use — VAL-ECO, VAL-REL (incl. the D3/RF Open items).
5. Field pilots → beta → production trial — VAL-FIELD (staged de-risking).
6. Business acceptance — VAL-BIZ (reuse economics is the key unknown).
Validation is outcome-based; it never substitutes for verification (Stage 13) and each critical task runs only after its paired PV-USE verification passes.
3. Participants & representative users
Per VAL-PLAN-002. Distinct user groups (from SH-01…SH-15): last-mile health worker (over-sampled; low-literacy/non-primary-language/gloved/training-decayed quotas), EPI/program manager, cold-store operator, QA/RA (dispatch), distributor coordinator, 3PL operator, cold-chain/calibration technician, service/reverse- logistics operator, IT/security admin, installer/OEM integrator, regulator/auditor, procurement, finance, executive, ecosystem partner. The beneficiary (SH-11) is represented, not tested.
Sample size | ≥15 representative users per distinct user group for each summative study (FDA HF floor); final counts per protocol.
• No convenience proxies for critical-task (last-mile/provisioning) users.
• Accessibility & inclusivity are quotas within the relevant groups (VAL-USER-015 / VAL-HF-009).
4. Representative environments & context of use
Per VAL-PLAN-003. Cold chambers to −30 °C; low-light/glare/direct-sun rigs; gloved protocols; RF-attenuating in-container shipper mockups; humid/condensation and shock/vibration exposure; real LMIC last-mile and GDP commercial lanes; service depots (mains-scarce). Mandatory conditions for the disposition and provisioning critical tasks: cold, low-light, gloved, offline.
5. Facilities, equipment & apparatus
Per VAL-PLAN-004. Usability lab + environmental chambers; calibrated reference thermal sources for scripted freeze/heat excursion scenarios; standardized glove/lighting kits; validated video/observation capture; release-candidate cloud/app; production-equivalent DVT/PVT devices for summative and field studies.
6. Phasing, sequencing & gates
Per VAL-PLAN-005/008.
1. Formative HF (early builds) — iterate the design; not acceptance-bearing.
2. Verification dependencies pass — PV-USE-001/002/003/004/005/006/008/009, PV-SW-, PV-COMM-, and the safety/security verifications; no summative study on an unverified control.
3. Summative HF critical tasks — on production-equivalent product with representative users/environments.
4. Operational/ecosystem/reliability — incl. the D3 cold-endurance (PRD-VER-004) and RF-in-container (PRD-VER-005) Open items (release-gating).
5. Field: pilot → beta → production trial — staged, gated on the prior stage's safety signals.
6. Residual/benefit-risk validation (VAL-RISK) — runs alongside 3–5 and concludes with VAL-RISK-003.
Gates: representative-product gate (summative/field on DVT/PVT + release-candidate cloud/app); verification gate (paired PV passed); safety-signal gate (a critical use error / missed excursion / thermal event halts scale-up); the two carried-Open items (D3/RF) gate reliability/field conclusion and the release recommendation.
7. Success-criteria philosophy
Per VAL-PLAN-006. Outcome-based endpoints (correct disposition, timely action, correct provisioning, effective rescue); critical-task use errors root-caused and judged for acceptability (IEC 62366-1), not required to be zero; both harm directions are endpoints (false-accept AND false-reject/wastage). All numeric targets (use-error rate, time-to-decision, false-reject rate, latency, NPS, return rate, cost-per-trip) are pre-registered and Evidence-Pending — set in PRD/protocol, never asserted here.
8. Deliverables
Per VAL-PLAN-007. IEC 62366-1 usability engineering file; per-study pre-registered protocols and reports (PVaL-1100); the Validation-Coverage-Matrix; the residual-risk/benefit-risk validation record; and the PVaL-1200 summary + go-to-market recommendation. These feed Stage 16 (Release) and Stage 17 (Field/PMS).
9. Dependencies, open items & timeline gates
Verification (Stage 13) | each summative critical task depends on its paired PV-USE passing.
Carried-Open, release-gating | D3 cold endurance (PRD-VER-004 → VAL-REL-001/VAL-RISK-011) and RF-in-container (PRD-VER-005 → VAL-REL-004/VAL-RISK-011); if RF cannot close, the real-time-alert claim is restricted and the on-receipt verdict is the compensating benefit.
Benefit target (D4) | WHO PQS accuracy/category confirmation gates the benefit claim (VAL-CLIN-010).
Security IV | independent adversarial verification (PV-CYBER-) gates the security residuals (VAL-RISK-009).
Representative product | production-equivalent DVT/PVT + release-candidate cloud/app for summative/field.
10. Roles & acceptance authority
Per VAL-GOV-004 (aligned to HAL-2200 §7): summative critical tasks & residual-risk → EM-ENG-006 + EM-ENG-001 + documented benefit-risk (S5 aircraft-fire and security residuals carry the most stringent sign-off); operational/ecosystem/reliability → Clinical/Operational + Product + EM-ENG-001; business → Product/Business; release recommendation → EM-ENG-001 + EM-ENG-006.
11. Honesty posture
No study has executed. All Status = Planned / Evidence Pending. No use-error rate, effectiveness figure, satisfaction score, return rate, or field signal is asserted. Summative endpoints are pre-registered; residuals are accepted only on demonstrated (not projected) evidence; the go-to-market recommendation is Not Yet Determined — Conditionally Ready pending V&V (consistent with HAL-2900 §7).

## 15 Manufacturing & Transfer
Manufacturing & Transfer
How the design becomes a product that can be built the same way, correctly, every time.
↩ Contents
In this section:
• Manufacturing & Design Transfer — Index — Smart Cold Chain Monitor (SCM)
• MDT-100 — Manufacturing Process Flow — Smart Cold Chain Monitor (SCM)
• MDT-1000 — Design Transfer Package & Manufacturing Readiness — Smart Cold Chain Monitor (SCM)
• MDT-200 — Process FMEA (PFMEA) — Smart Cold Chain Monitor (SCM)
• MDT-300 — Control Plan — Smart Cold Chain Monitor (SCM)
• MDT-400 — Critical & Special Characteristics — Smart Cold Chain Monitor (SCM)
• MDT-500 — Work Instructions Summary — Smart Cold Chain Monitor (SCM)
• MDT-600 — Production Test Specification — Smart Cold Chain Monitor (SCM)
• MDT-700 — Supplier Controls & Incoming Inspection — Smart Cold Chain Monitor (SCM)
• MDT-800 — Calibration & Measurement-System Analysis (MSA) Plan — Smart Cold Chain Monitor (SCM)
• MDT-900 — Serialization, UDI & Production Traceability — Smart Cold Chain Monitor (SCM)
Manufacturing & Design Transfer — Index — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 15 — Manufacturing & Design Transfer · Owner: Manufacturing/Quality Engineer with Chief Systems (EM-ENG-001) and Quality/Regulatory (EM-ENG-006). This stage authors the SCM's production process baseline and design-transfer package: the process flow, Process FMEA, Control Plan, Critical/Special Characteristics, Work Instructions, Production Test Specification, Supplier Controls, Calibration/MSA plan, Serialization/UDI/traceability, and the design-transfer package + manufacturing readiness assessment. Its central job is to preserve the safety- and measurement-critical controls into every production unit and to state honestly whether the product is ready to transfer to volume.
Position in the chain: consumes Stage 09 (Design Outputs, esp. DOL-1400 manufacturing / DOL-1500 inspection-test / DOL-1600 packaging-labeling / DOL-1700 service / DOL-2000 transfer), Stage 10 (HAL residuals + Risk-Management Summary), Stage 12 (Risk-Control Requirements to preserve), Stage 13 (PV-MFG production verification), and Stage 14 (validation status). Feeds Stage 16 (Production Release), Stage 17 (Field Monitoring), and ORCH-018 / EM-ENG-012 (Stage 18) — the end-to-end lifecycle traceability spine.
---
1. Module files
Module | File | Scope
MDT-100 | [MDT-100-process-flow.md](MDT-100-process-flow.md) | Process flow (PROC-###/STEP-###): procurement → SMT/reflow → coat → ICT/FCT → battery → seal/IP → USB-C → FW flash → identity → calibration → final/safety test → serialization/UDI → packaging → ship; safety-control preservation map
MDT-200 | [MDT-200-pfmea.md](MDT-200-pfmea.md) | Process FMEA — 28 process failure modes (PFM-###) across SMT/reflow, battery, seal/IP, USB-C, calibration, provisioning, FW flash, labeling; S/O/D + severity-first Action Priority
MDT-300 | [MDT-300-control-plan.md](MDT-300-control-plan.md) | Control Plan (CP-###) — CC/SC → process control → measurement → reaction; preserves charge-inhibit, protection, calibration, seal/IP, crypto identity, signed FW
MDT-400 | [MDT-400-critical-special-characteristics.md](MDT-400-critical-special-characteristics.md) | Critical (CC-01…08) & Special (SC-01…09) Characteristics from DO-/RCR-/PFM-
MDT-500 | [MDT-500-work-instructions-summary.md](MDT-500-work-instructions-summary.md) | Work-instruction register + mandatory content in safety-critical WIs + operator qualification/doc control
MDT-600 | [MDT-600-production-test-specification.md](MDT-600-production-test-specification.md) | Production Test Spec (PT-###) — the acceptance spine; ties to PV-MFG (PVL-1300)
MDT-700 | [MDT-700-supplier-controls-incoming.md](MDT-700-supplier-controls-incoming.md) | Supplier controls (SUP-###) + incoming inspection (INC-###); enhanced control on cell/BMS/SE/fuse
MDT-800 | [MDT-800-calibration-msa-plan.md](MDT-800-calibration-msa-plan.md) | Product calibration integrity (CAL-###) + Measurement-System Analysis (MSA-###)
MDT-900 | [MDT-900-serialization-udi-traceability.md](MDT-900-serialization-udi-traceability.md) | Serialization/UDI/identity binding + Device History Record traceability spine
MDT-1000 | [MDT-1000-design-transfer-package-readiness.md](MDT-1000-design-transfer-package-readiness.md) | Design-transfer package (DO-XFER-), four reconciliations, readiness assessment, open-items, open-risk disposition
2. The preservation principle (why this stage matters)
The SCM's entire safety case rests on a small set of controls holding across the whole fleet (HAL-2900): the S5 battery-fire spine (charge-inhibit, protection, correct chemistry/parts, retention), the measurement integrity (calibration), and the record integrity/authenticity (seal/IP, signed FW, crypto identity). This stage's job is to guarantee those controls are built into and proven on every unit — as Critical Characteristics (MDT-400) held by zero-defect Control-Plan lines (MDT-300), realized by safety-critical Work Instructions (MDT-500), and proven by the acceptance-spine production tests (MDT-600) whose detection systems are themselves MSA-qualified and master-challenged (MDT-800). The two process controls that are operational, not one-time facts — calibration and config/OTA — carry standing field monitoring into Stage 17.
3. Critical & Special Characteristics at a glance
CC (8) | charge-inhibit <0 °C (CC-01), protection+backstop (CC-02), correct chemistry/protection parts (CC-03), cell retention/connection (CC-04), temperature-sensor calibration accuracy (CC-05), signed FW + region config (CC-06), unique crypto identity + debug-lock (CC-07), regulatory/UDI/Li-transport marking (CC-08).
SC (9) | seal/IP integrity (SC-01), coat coverage (SC-02), final functional/system (SC-03), debug-lock verify (SC-04), PCBA safety-net/coverage (SC-05), serial/identity match (SC-06), correct-config/line- clearance (SC-07), reflow CPP (SC-08), torque CPP (SC-09).
4. Evidence & honesty statement
No invented numbers | no reflow profile, torque, weld parameter, leak rate, calibration accuracy value, capability index (Cpk), or gauge-R&R % is asserted — all are Evidence Pending against the released design outputs (DO-MFG-003/006, DO-TST-006/011/014) and the MSA (MDT-800), or anchored to the WHO PQS accuracy target (O-7).
• Severity-first Action Priority, not raw RPN, throughout the PFMEA (CLAUDE.md rule 3).
Honest readiness | manufacturing readiness is NOT READY for volume — the process design is transfer-mature but the AML freeze (RA-07), numeric limits/MSA (RA-09), regulatory/UDI (RA-10), pilot build, and HAL open-risk reconciliation on executed V&V (RA-05) are open. This is carried explicitly into folder 16.
5. Feeds
Stage 16 (Production Release) — the readiness assessment, four reconciliations, and open-risk disposition are the release-gate inputs. Stage 17 (Field Monitoring) — the two process common causes (calibration drift, config/OTA) and the Device History Record are the standing-monitoring + recall/CAPA spine. Stage 18 (ORCH-018 / EM-ENG-012) — the manufacturing/transfer segment of the end-to-end traceability spine (PROC-/CP-/STEP- ↔ DO-/RCR-/HA-/PV-).
MDT-100 — Manufacturing Process Flow — Smart Cold Chain Monitor (SCM)
Stage 15 (ORCH-018 chain) · Module MDT-100 · Owner: Manufacturing Engineer with Quality (EM-ENG-006) and Chief Systems (EM-ENG-001). This module defines the end-to-end production process flow for the SCM — from component procurement to finished-goods ship — as the process baseline from which the Process FMEA (MDT-200), Control Plan (MDT-300), Critical/Special Characteristics (MDT-400), Work Instructions (MDT-500), and Production Test Spec (MDT-600) are derived. Each production stage carries a PROC-### ID; each operation carries a STEP-### ID. Safety-critical steps (those that implement or verify a risk control) are flagged ★ and reconciled to the HAL (Stage 10) and the Risk-Control Register (Stage 12).
Evidence posture (per CLAUDE.md rule 1 + hallucination-policy.md): no process window, temperature, torque, leak-rate, or capability number is asserted as fact. Reflow profile, seal torque, leak limit, and calibration points are Evidence Pending against released design outputs (DO-MFG-003 reflow, DO-MFG-006 torque, DO-TST-011 leak, DO-TST-006 cal) — this flow states what is controlled and why, not invented values. The two program-top unknowns (D3 −30 °C/90-day battery + charge safety; RF-in-container) and the AML freeze (chemistry/antenna, RA-07) are carried Open and gate the line's ability to reach volume.
---
1. Purpose & scope
Define the manufacturing process the contract manufacturer (CM) executes to build every SCM SKU (SCM-100 / -NA / -EU / -ROW / -XA per DO-CFG-002), and to establish where the safety- and measurement- critical controls are built in (charge-inhibit hardware, seal/IP integrity, cryptographic identity, signed firmware, sensor calibration) and where they are proven (production tests). Scope covers PCBA (SMT), battery handling and connection, enclosure seal/IP assembly, USB-C and mechanical final assembly, firmware flashing and identity provisioning, sensor calibration, final/safety test, serialization/UDI labeling, packaging, and ship. Excludes design verification methods (Stage 13) and validation (Stage 14).
2. High-level flow (value stream)
PROC-010 Procurement & supplier release │ PROC-020 Incoming inspection (enhanced for cell / BMS-AFE / secure element / fuse) ★ │ PROC-030 SMT / PCBA assembly ── paste → SPI → place → reflow → AOI → X-ray ★ (reflow, cell/BMS placement) │ PROC-040 Conformal coat + cure (keep-outs: sensors / antenna / contacts / probe) ★ (condensation control) │ PROC-050 ICT / FCT (in-circuit + functional; safety-net verification) ★ (DO-TST-004) │ PROC-060 Battery handling, connection / spot-weld, pack build + retention ★ (cell abuse prevention) │ PROC-070 Enclosure seal / IP assembly (gasket seat, closure sequence, torque) ★ (seal integrity) │ PROC-080 USB-C + mechanical final assembly (probe gland, HMI window, mount) │ PROC-090 Firmware flashing (signed image) + config ★ (signed FW) │ PROC-100 Cryptographic identity / key provisioning + debug-lock ★ (crypto identity) │ PROC-110 Sensor calibration (internal T / probe T / RH) against traceable reference ★ (calibration accuracy) │ PROC-120 Final functional & system test (BIST, sensors, HMI, arm-and-capture) ★ (DO-TST-005) │ PROC-130 Battery-protection / charge-inhibit fault-injection safety test ★ (DO-TST-009, D3) │ PROC-140 Seal-integrity / leak (IP) production test ★ (DO-TST-011) │ PROC-150 Serialization / UDI marking + label application + label inspection ★ (UDI/Li-transport) │ PROC-160 Packaging, IFU/quick-start kitting, Li-transport marking, line clearance │ PROC-170 Finished-goods QA gate, hold/DHR review, release to ship │ → Ship (+ Design Transfer readiness, MDT-1000; Release gate, folder 16)
Rework/repair (PROC-900) and non-conforming material handling (PROC-910) are off-line loops available from several stages, with mandatory re-test on return to flow (see §5).
3. Stage-by-stage process definition
PROC-010 — Procurement & supplier release
• STEP-011 Release POs only against the frozen AML (DO-CFG-004 / DO-MFG-001 mBOM); block any part not on the approved-manufacturer list. Critical parts (cell CMP-006/007, BMS/AFE CMP-008, secure element CMP-016, protection fuse CMP-010) are single-source-controlled with PCN/PCN-hold notification.
• STEP-012 Require lot traceability and Certificate of Conformance (CoC) for critical parts; require UN 38.3 test summary + IATA classification evidence for cells/packs (DO-LBL-005).
Trace | DO-CFG-004 (AML), DO-MFG-001 (mBOM), SUP-controls (MDT-700). Open: AML cannot freeze until chemistry (CMP-006) and antenna (CMP-028/029) are selected — RA-07.
PROC-020 — Incoming inspection ★
• STEP-021 Standard AML parts: sampling per DO-TST-001 / ANSI-ASQ Z1.4 plan; dimensional/electrical attributes per released datasheet.
• STEP-022 ★ Safety-critical parts (cell, BMS/AFE, secure element, fuse): enhanced acceptance — 100 % CoC + lot-traceability check, sampled electrical characterization, cell chemistry / protection-rating verification (poka-yoke against wrong chemistry — RCR-054), counterfeit screening on the secure element. Non-conforming lots quarantined (PROC-910).
Trace | DO-TST-001, PV-MFG-002 (incoming chemistry/protection), RCR-054, HA-MFG-002/003.
PROC-030 — SMT / PCBA assembly ★
• STEP-031 Solder-paste print; SPI (solder-paste inspection) — volume/offset within limits.
• STEP-032 Pick-and-place per DO-MFG-002 (populate/no-populate per SKU); orientation poka-yoke for polarized/protection parts.
• STEP-033 ★ Reflow to the DO-MFG-003 profile — a Critical Process Parameter (CPP); profile must give reliable fine-pitch/QFN and cell/BMS-area joints that survive −30↔+60 °C cycling. Profile values Evidence Pending (DO-MFG-003).
• STEP-034 Post-reflow AOI to IPC-A-610 target class; defect classes per DO-TST-002.
• STEP-035 ★ X-ray of BGA/QFN and the cell/BMS/protection area — solder-defect escape on a safety net (charge-inhibit, protection) must be caught before it is buried by coat/assembly. AOI/ICT coverage on safety-critical nets is a named control (RCR-016, PV-MFG-008).
Trace | DO-MFG-002/003, DO-TST-002, PV-MFG-003/008, RCR-016/054.
PROC-040 — Conformal coat + cure ★
• STEP-041 Cleanliness check before coat (ionic/residue within limit).
• STEP-042 ★ Apply conformal coat per DO-MFG-005 with keep-outs on sensors, antenna, connectors, and the probe interface; cure. Coat + seal are the condensation/ingress barrier for the record electronics across cold↔warm cycling (RCR-015/018).
Trace | DO-MFG-005, DO-PCB-010, RCR-015/018, HA-RES-021.
PROC-050 — ICT / FCT (board level) ★
• STEP-051 ★ In-circuit / functional test of the bare PCBA via test points (DO-PCB-011): power rails, MCU boot, sensor buses, radios, charge path, and every safety net — protection nets, charge-inhibit net, probe-presence. This is the first gate that the safety-relevant circuitry is populated and functional before it is sealed inside the enclosure (DO-TST-004, Critical).
• STEP-052 Record coverage; fail-code any board that misses a safety-net check; route fails to PROC-900.
Trace | DO-TST-004, DO-PCB-011, PV-MFG-008, RCR-016.
PROC-060 — Battery handling, connection / spot-weld, pack build + retention ★
• STEP-061 ★ ESD-safe, no-crush/no-puncture cell handling per DO-MFG-007; controlled shipping/handling state-of-charge; line-safety controls (fire-safe storage, quarantine for damaged cells).
• STEP-062 ★ Cell-to-PCBA / pack connection — spot-weld or connector per the released pack design (DO-EE-004); weld/joint parameters are a CPP; no reflow/solder directly on the cell can bypass the pack build rules. Weld parameters Evidence Pending (DO-EE-004).
• STEP-063 ★ Cell retention (no crush/puncture through drop/vibration/transport — RCR-004); verify retention fixture engagement (poka-yoke).
Trace | DO-MFG-007, DO-EE-004, DO-MECH-012 (retention), RCR-002/004, HA-RES-005/006. Open: low-temp chemistry not frozen (D3, RA-07).
PROC-070 — Enclosure seal / IP assembly ★
• STEP-071 ★ Seat the seal/gasket in the groove per DO-MFG-004; verify groove-fill and correct seal part (poka-yoke for the -XA aperture variant).
• STEP-072 ★ Closure sequence + fastening/adhesive to the DO-MFG-006 torque spec (a CPP) — even seal-compression without over-stress; seal the probe gland and USB-C aperture per spec. Torque values Evidence Pending (DO-MFG-006).
Trace | DO-MFG-004/006, DO-MECH-005 (seal), RCR-011/015, HA-RES-021. Seal integrity is proven at PROC-140 (DO-TST-011).
PROC-080 — USB-C + mechanical final assembly
• STEP-081 Install/verify USB-C connector retention and aperture seal (DO-MECH-008); HMI/OLED window fit; external-probe and mount interfaces.
• STEP-082 -XA: install/verify external-antenna connector and its IP-sealed gland.
Trace | DO-MECH-008, DO-CFG-003 (options).
PROC-090 — Firmware flashing (signed image) + config ★
• STEP-091 ★ Flash the released, signed factory image (hash/version = DO-FW-019) via the controlled provisioning tool (DO-SW-001); load the correct regional/firmware config (band plan, market profile) per DO-CFG-006. An unsigned/mismatched image or wrong-region config is rejected — the station will not pass an incorrectly-loaded unit (RCR-037/056, PV-MFG-006/007).
Trace | DO-MFG-008, DO-FW-019/020, DO-CFG-005/006, RCR-037/056, PV-MFG-006/007.
PROC-100 — Cryptographic identity / key provisioning + debug-lock ★
• STEP-101 ★ Provision a unique per-device cryptographic identity + keys into the secure element and bind it to the serial (DO-CFG-010); duplicates prevented/detected (RCR-055, PV-MFG-005).
• STEP-102 ★ Set the production debug lock (readout-protection / lifecycle fuse) so keys/records cannot be read from a fielded unit (RCR-038, DO-SEC-008); read-back verified (DO-TST-012).
• STEP-103 The station refuses to pass an incompletely-provisioned unit; every unit auditable per serial.
Trace | DO-MFG-008, DO-CFG-010, DO-SEC-007/008, DO-TST-008/012, RCR-038/055.
PROC-110 — Sensor calibration ★
• STEP-111 ★ Calibrate/verify internal-T, external-probe-T, and RH channels against a NIST/traceable reference across the operating range (DO-MFG-010 / DO-TST-006); post-cal error ≤ the released accuracy budget (±0.5 °C internal/probe target); write coefficients with metadata; read-back-verify.
• STEP-112 ★ Bind the calibration/birth record to the serial and enforce cal-limits; a miscalibrated unit cannot pass EOL (RCR-053, PV-MFG-001) — the primary defence against the fleet-wide calibration common-cause (HA-RES-015, S5-worst). Reference chain / cal points Evidence Pending (DO-TST-006).
Trace | DO-MFG-010, DO-TST-006, DO-FW-009, RCR-026/049/053, HA-RES-015, PV-MFG-001.
PROC-120 — Final functional & system test ★
• STEP-121 ★ Assembled-unit boot + BIST; read all sensors within limits; exercise HMI/verdict; arm and capture a sample record; confirm radios attach on the correct region band plan (DO-TST-005/007). Last full-unit acceptance of the capture→detect→warn→decide→prove chain (Critical). Fail-code recorded.
Trace | DO-TST-005/007, DO-FW-012 (BIST), RCR-019/020/021/043.
PROC-130 — Battery-protection / charge-inhibit fault-injection safety test ★
• STEP-131 ★ Inject/simulate over-voltage, over-current, and cold-cell conditions; confirm autonomous protection trips and charging is inhibited below 0 °C (or the qualified datasheet limit) with the inhibited state reported; confirm the non-resettable secondary backstop coordinates. No unit ships with a defeated safety control (DO-TST-009, Critical; the production proof of the D3 charge-safety controls). 100 % vs. sampled plan Evidence Pending.
Trace | DO-TST-009, DO-EE-005/007/010, RCR-001/002, HA-RES-005/006, PV-SAFE / PV-ENRG-001.
PROC-140 — Seal-integrity / leak (IP) production test ★
• STEP-141 ★ Each sealed unit passes a leak test (e.g., pressure-decay) correlated to the target IP rating (DO-TST-011); a leaking unit — which can corrupt the evidentiary record via ingress/condensation — is rejected or reworked (PROC-900) and re-tested. Leak method + limit + IP correlation Evidence Pending.
Trace | DO-TST-011, DO-MECH-005, DO-MFG-004, RCR-011/015/018, HA-RES-021.
PROC-150 — Serialization / UDI marking + label application + inspection ★
• STEP-151 ★ Apply serial + machine-readable identity (barcode/QR, DO-LBL-003), regulatory-markings label (DO-LBL-002), and UDI content (DO-LBL-004) — reconciled with the regulatory determination (DOL-1800); scan-verify grade + uniqueness.
• STEP-152 ★ Apply the lithium-battery transport marking (UN 38.3 / IATA, DO-LBL-005) to device/pack.
• STEP-153 Label inspection per DO-LBL-012: correct content, legibility, placement, adhesion (survives −30↔+60 °C / condensation / cleaning), correct language variant. Mismatch → reject.
Trace | DO-MFG-009, DO-LBL-002/003/004/005/012, DO-CFG-010, RCR-055, PV-MFG-005. Open: exact UDI regime + markings set follow DOL-1800 (RA-10).
PROC-160 — Packaging, kitting, Li-transport, line clearance
• STEP-161 Kit device + probe + accessories + IFU/quick-start (correct-language) into primary pack per DO-LBL-006/010/011; protective cushioning per DO-LBL-007.
• STEP-162 ★ Master-carton / palletization with the correct DG/lithium shipping configuration and outer marking (DO-LBL-008); storage/transport SoC controls per DO-LBL-009.
• STEP-163 ★ Line clearance between SKUs / regional variants (DO-MFG-011) — prevents mixed-config builds (wrong band plan / wrong markings shipped).
Trace | DO-LBL-006/007/008/009/010/011, DO-MFG-011, RCR-056.
PROC-170 — Finished-goods QA gate & release to ship
• STEP-171 Device History Record (DHR) review: all critical-step records present (calibration/birth, protection FI, leak, provisioning/lock, label); hold any unit with a missing/failed critical record.
• STEP-172 Disposition holds (MRB); release conforming finished goods to ship.
Trace | ISO 13485 §7.5; feeds MDT-1000 (transfer readiness) and folder 16 (production release).
4. Safety-control preservation map (design control → process step)
The whole point of this flow is that the safety and measurement controls established in design are preserved into every production unit. The load-bearing map:
Safety/measurement control | RCR | Built at | Proven at | Design output
Charge-inhibit < 0 °C (HW interlock) | RCR-001 | PROC-030/060 | PROC-130 ★ | DO-EE-007 / DO-TST-009
Autonomous protection + secondary backstop | RCR-002 | PROC-030/060 | PROC-130 ★ | DO-EE-005/010 / DO-TST-009
Cell retention (no crush/puncture) | RCR-004 | PROC-060 ★ | PROC-060 fixture + drop (design) | DO-MECH-012
Seal / IP integrity | RCR-011/015/018 | PROC-070 ★ | PROC-140 ★ | DO-MECH-005 / DO-TST-011
Sensor calibration accuracy | RCR-026/053 | PROC-110 ★ | PROC-110 read-back + PV-MFG-001 | DO-TST-006
Correct chemistry / protection parts | RCR-054 | PROC-010/020 | PROC-020 ★ / PROC-030 X-ray | DO-TST-001 / PV-MFG-002
Signed firmware only | RCR-037 | PROC-090 ★ | PROC-090 hash + DO-TST-008 | DO-FW-019
Unique crypto identity | RCR-055 | PROC-100 ★ | PROC-100 + PV-MFG-005 | DO-CFG-010
Debug-lock (no key/record readout) | RCR-038 | PROC-100 ★ | DO-TST-012 | DO-SEC-008
Correct regional config | RCR-056 | PROC-090 | PROC-120 / PV-MFG-006 | DO-CFG-006
No latent SPF on safety nets | RCR-016 | PROC-030 | PROC-050/030 (AOI/ICT/X-ray) | PV-MFG-008
5. Rework / non-conforming loops (mandatory re-test)
• PROC-900 Rework/repair (per DO-MFG-014): allowed rework limits; prohibited rework on cell and seal; ESD/handling controls; mandatory re-test after rework — a unit that returns to flow re-runs the affected critical test (leak PROC-140, calibration PROC-110, protection PROC-130, functional PROC-120). No rework may compromise seal, cell, or coat.
• PROC-910 Non-conforming material (MRB): quarantine, disposition (use-as-is only via approved deviation, rework, scrap); critical-part non-conformance feeds Supplier CAPA (MDT-700) and, if fielded implications exist, the closed loop (folder 17).
6. Line qualification & readiness gates (feed MDT-1000)
The flow is not production-ready until: (a) fixtures/tooling qualified (DO-MFG-012/013), (b) reflow/torque/ weld/leak/cal process windows established and process-validated (IQ/OQ/PQ), (c) golden units established for each critical test station (DO-TST-013), (d) MSA/gauge-R&R complete on the calibration and leak measurements (MDT-800), and (e) a pilot build passes with deviations dispositioned (MDT-1000, DO-XFER-009). Until the AML freezes (RA-07) and D3/RF unknowns close (PRD-VER-004/005), the line can build pilot/engineering units for V&V but not release to volume.
7. Traceability
Backward | DO-MFG- (DOL-1400), DO-TST- (DOL-1500), DO-LBL- (DOL-1600), DO-CFG- (DOL-100), the RCR register (Stage 12), and the HAL residuals (HAL-2800).
Forward | MDT-200 (PFMEA — each PROC/STEP is a PFMEA function), MDT-300 (Control Plan — each ★ step is a controlled characteristic), MDT-400 (CSC), MDT-600 (Production Test Spec ties to PV-MFG), MDT-1000 (transfer readiness), folder 16 (release), and ORCH-018 / EM-ENG-012 (Stage 18 lifecycle spine).
MDT-1000 — Design Transfer Package & Manufacturing Readiness — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-1000 · Owner: Chief Systems Engineer (EM-ENG-001) + Manufacturing + Quality/Regulatory (EM-ENG-006). The design-transfer package and manufacturing-readiness assessment: the honest ledger of what design outputs are released, whether the BOM/drawing/software/manufacturing baselines reconcile, whether suppliers/tooling/fixtures/test-line are ready, and whether the pilot build has proven the process — the inputs to the production-release decision (folder 16). Realizes DOL-2000 (DO-XFER-001…010) and operates per ISO 13485:2016 §7.3.8 (design & development transfer) and the FDA QMSR (21 CFR 820, incorporating ISO 13485:2016).
Evidence honesty: this module is Evidence Pending by nature — the transfer baseline is not frozen. The honest top blockers are RA-07 (low-temp cell chemistry + antenna → AML freeze), RA-05 (HAL open-risk reconciliation / V&V), and RA-10 (regulatory classification / UDI / WHO PQS target). No output is claimed Released that is not; no readiness is asserted without its evidence.
---
1. Design-transfer package contents (DO-XFER-001…010)
DO-XFER | Content | This-stage module | Status
001 Released-output index | Snapshot of the Design Output Register at the transfer baseline | (references DOL register) | Draft — revs TBD until freeze
002 Open-items list | Every Evidence-Pending/Draft output with owner + closure plan | §5 below | Draft (mirrors top gaps)
003 Config audit (FCA/PCA) | Functional + physical configuration audit | §3 | Not started (pre-freeze)
004 BOM/drawing/SW/mfg reconciliation | eBOM=mBOM=AML=drawings=SW versions | §2 | Draft
005 Production-test readiness | Test line defined/fixtured/golden/limits/MSA | §4.3 | Draft (limits/MSA EP)
006 Supplier/tooling/fixture readiness | Suppliers qualified; tooling/fixtures capable | §4.2 | Draft (RA-07 blocked)
007 Packaging/labeling/service/training readiness | Downstream launch-ready | §4.4 | Draft
008 ★ Open-risk disposition & HAL reconciliation | Every HAL control → released output; residual accepted | §6 | Preliminary (RA-05)
009 Pilot-build baseline & deviation log | Pilot output revs + deviation disposition | §4.5 | Not started
010 ★ Production-release baseline & approvals | Frozen baseline + RA-01…12 + signatures | folder 16 | Not started
2. The four reconciliations (DO-XFER-004)
Before transfer, four baselines must reconcile with zero unresolved deltas:
1. BOM reconciliation — eBOM (DO-EE-020) = PCBA BOM (DO-PCB-013) = mechanical BOM (DO-MECH-019) = mBOM (DO-MFG-001) = AML (DO-CFG-004). Blocked: AML cannot freeze until chemistry (CMP-006) + antenna (CMP-028/029) are selected — RA-07.
2. Drawing reconciliation — released drawings match models; CTQ/GD&T (DO-VER-002, DO-MECH-018) match the dimensional inspection plan (DO-TST-003).
3. Software reconciliation — released firmware/cloud/mobile versions + hashes reconcile to the compatibility matrix (DO-CFG-005) and the compliance-critical register (DO-REG-011). Version numbers are the release baseline (folder 16, REL-200) — not yet frozen.
4. Manufacturing reconciliation — WIs (DOL-1400), test spec (DOL-1500), Control Plan (MDT-300), and PFMEA (MDT-200) all reference the same process baseline (MDT-100) and the same released outputs.
3. Configuration audit (DO-XFER-003)
FCA (functional configuration audit) | each requirement (PRD/SYS/SSR/HWR/SWR) verified against a design output and its verification (Stage 13) — cannot complete until V&V executes (RA-05).
PCA (physical configuration audit) | each design output/document present at its released revision; discrepancies dispositioned via ECO (DO-GOV-004). Every Critical output (DO-MFG-004 seal, DO-MFG-007 battery handling, DO-MFG-008 programming, DO-TST-004/005/006/009/011, DO-CFG-004/005, DO-XFER-008/010) must be Released for PCA to pass — presently Draft.
4. Manufacturing-readiness assessment
4.1 BOM / mBOM / AML readiness
mBOM (DO-MFG-001) drafted; AML freeze blocked by RA-07 (chemistry + antenna). Critical parts identified and tiered (MDT-700); T1 supplier qualification cannot complete until selection. Not ready.
4.2 Supplier / tooling / fixture readiness (DO-XFER-006)
Suppliers | T1 (cell, BMS/AFE, fuse, secure element) qualification Open (SUP-001…003, RA-07); T2/T3 qualifiable. Second-source status for critical parts recorded once selected. Not ready (T1).
Tooling | injection-mould (DO-MFG-013) + fixtures (DO-MFG-012) designed; final resin grade EP; tool qualification + first-article pending. Partially ready.
Fixtures | assembly/seal/press fixtures + test fixtures (DO-TST-013) to be built, tried out, and qualified; poka-yoke for orientation/SKU. Partially ready.
4.3 Production-test readiness (DO-XFER-005)
All production tests defined (MDT-600 PT-001…011) and mapped to PV-MFG (PVL-1300); but golden-unit correlation, numeric limits + guard-bands (DO-TST-014), leak method/limit (DO-TST-011), calibration reference chain + accuracy target (DO-TST-006 / WHO PQS O-7), and MSA (MDT-800, MSA-001/002) are Evidence Pending; protection-FI 100 %-vs-sample reconciles with HAL. Not ready (limits/MSA/leak/cal EP).
4.4 Packaging / labeling / service / training readiness (DO-XFER-007)
Packaging/protective design (DOL-1600) drafted; regulatory/UDI/Li-transport markings gated by DOL-1800 (RA-10); IFU/quick-start + translations drafted; service manual/FRU/cleaning/recal/disposal (DOL-1700) drafted with cleaning-agent/recal-interval EP; field/depot staff training material to be released + delivered. Partially ready.
4.5 Pilot build (DO-XFER-009)
Pilot build has not run. When it does: it records the exact output revisions used, proves the process on representative equipment, re-scores the PFMEA O/D estimates with real data, runs MSA/capability, and logs + dispositions every transfer deviation (accept/ECO/blocker) before volume. Pilot units are also the V&V build for the safety/measurement verifications (Stage 13/14). Not started.
5. Open-items list (DO-XFER-002) — honest ledger
Item | Blocks | Closes on | Owner | Ref
Low-temp cell chemistry not selected | AML freeze, T1 supplier qual, battery WIs, D3 endurance | Selection + datasheet + PRD-VER-004 | Battery/Power | RA-07 / O-2
Antenna branch (integral vs external) undecided | AML freeze, RF test, -XA readiness | PRD-VER-005 (RF-in-container) | RF/Antenna + Connectivity | RA-07 / O-1
Numeric test limits + MSA (Cpk/gauge-R&R) | Production-test readiness | DO-TST-014 + MSA (MDT-800) | Quality/Metrology | RA-09
Leak method + limit + IP correlation | Seal test readiness | DO-TST-011 | Quality/Mech | —
Calibration reference chain + accuracy target | Calibration readiness | DO-TST-006 + WHO PQS target (D4) | Metrology/Regulatory | O-7
Regulatory classification / UDI / markings | Labeling readiness, release | DOL-1800 (D4) | EM-ENG-006 | RA-10
HAL open-risk reconciliation / V&V execution | Release safety gate | Stages 13–14 + DO-XFER-008 | Reliability/Safety | RA-05
Cleaning agents/cycle + recal interval | Service readiness | material-compat test + drift data | Service/Metrology | —
6. Open-risk disposition & HAL reconciliation (DO-XFER-008) — the safety gate
Per the Risk-Management Summary (Conditionally Ready — release pending V&V) and HAL-2900, transfer to volume production requires that every HAL risk control maps to a Released implementing output with completed risk-control verification, and the residual risk is dispositioned + accepted by authority (ISO 14971). At this stage:
• Every DO-SAFE control has a defined implementing output and a Control-Plan line that preserves it into production (MDT-300/400) — good.
• But the verifications are Evidence-Pending (Stage 13/14 not executed), so the disposition is Preliminary: the safety controls are designed and process-preserved, not demonstrated. RA-05 is not satisfied.
• The two carried-Open unknowns (D3 cold endurance PRD-VER-004; RF-in-container PRD-VER-005) and the S5 aircraft-fire qualification (UN 38.3 + PV-TFC-002) remain conditions.
This disposition is completed in folder 16 (release gate) once V&V returns.
7. Readiness conclusion (feeds folder 16)
Manufacturing readiness: NOT READY for volume production; process design is transfer-mature. The
process flow, PFMEA, Control Plan, Critical/Special Characteristics, Work Instructions, Production Test
Spec, Supplier Controls, Calibration/MSA plan, and Serialization/UDI/traceability are authored and
internally consistent, and every safety/measurement control is preserved into production via a
zero-defect Control-Plan line. The line can build pilot/engineering units for V&V. It is not ready to
release to volume until: (RA-07) the cell chemistry + antenna are selected and the AML freezes;
(RA-09) numeric limits + MSA close; the leak/calibration methods close; (RA-10) the regulatory
classification/UDI/markings resolve (DOL-1800); the pilot build runs with deviations dispositioned; and
(RA-05) the HAL open-risk reconciliation completes on executed V&V. These are the explicit blockers
carried into the production-release readiness assessment (folder 16).
8. Traceability
Backward | DOL-2000 (DO-XFER-), DOL-100 (DO-CFG-), the entire DOL register, the HAL (Stage 10, open-risk disposition + Risk-Management Summary), the RCR register (Stage 12), Stages 13–14 (V&V status), and all of MDT-100…900.
Forward | folder 16 (production release) — this readiness assessment + the four reconciliations + open-risk disposition are the direct inputs to REL-100 (release readiness) and the REL-400 release decision; ORCH-018 / EM-ENG-012 (Stage 18) — the transfer segment of the lifecycle spine.
MDT-200 — Process FMEA (PFMEA) — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-200 · Owner: Manufacturing/Quality Engineer with System-Safety and Battery/Power. Process Failure Mode & Effects Analysis for the SCM production process (MDT-100). Deferred here from Stage 11 (the ERFKL index defers FAL-1100 PFMEA to Stage 15). Each record keys to a PROC-###/STEP-### process function, states failure mode → effect → cause, current prevention/detection controls, S/O/D, and Action Priority (severity-dominant, NOT raw RPN) per CLAUDE.md rule 3 / AIAG-VDA / ISO 14971, then the recommended action and target AP. Every safety-significant process failure mode reconciles to a HAL residual (HA-RES-###) and a Risk-Control Requirement (RCR-###), and drives a Control-Plan line (MDT-300) and a Critical/Special Characteristic (MDT-400).
Evidence honesty: O and D are pre-production estimates with a stated basis (Judgment / handbook-analogue / Evidence-Pending until a named study — MSA, capability, pilot). No capability index (Cpk), escape rate, or process-defect rate is asserted as fact; those close with MSA (MDT-800) and the pilot build (MDT-1000).
---
1. Scoring convention (aligned to Stage-11 ERFKL)
• Severity (S) 1–10, outcome-fixed (same scale as the FMEA family): S5 (9–10) catastrophic (battery fire / aircraft); S4 (7–8) critical (compromised dose administered as good / thermal event / fleet-blinding); S3 (5–6) serious (wrong disposition / lost record / burn); S2 (3–4) minor; S1 (1–2) negligible. Process severity is the end-effect on the fielded device, not the factory scrap cost.
• Occurrence (O) / Detection (D) 1–10 with a stated basis. D=10 = the defect is undetectable by any current process control and escapes to the field.
Action Priority (severity-dominant) | an unmitigated safety end-effect (S≥7) is High unless both O≤3 and D≤4 (then Medium); S3 loss-of-function follows S/O/D bands; S1–S2 nuisance is Low unless detection is poor. AP shown is the residual direction after the current controls; the recommended action drives it to the target.
2. PFMEA register
Legend: S/O/D current · AP current Action Priority → AP\ target after recommended action. ★ = safety-critical (preserves a risk control into production).
---
PROC-020 — Incoming inspection
PFM-001 ★ — Wrong cell chemistry / wrong-rated protection part accepted
Process function | STEP-022 verify critical-part identity/rating.
Failure mode | A cell of incorrect chemistry (not the qualified low-temp grade) or a wrong-rated/ counterfeit protection IC/fuse passes incoming.
Effect (field) | Charge-safety / protection envelope wrong → cold-charge plating or protection failure → thermal runaway (S5 aircraft-fire path); or endurance shortfall. HA-RES-005/006.
Cause | Mixed lots, mislabeled reels, supplier substitution, counterfeit.
Prevention | Single-source AML with PCN hold (DO-CFG-004); supplier qualification (MDT-700); CoC + lot traceability.
Detection | Enhanced incoming — chemistry/rating verification + counterfeit screen (DO-TST-001, PV-MFG-002).
S/O/D | S=9 · O=3 (Judgment) · D=3 → AP: High (S5 dominates).
RCR/HA | RCR-054 / HA-RES-005/006 / HA-MFG-002/003.
Recommended action | 100 % CoC verification + sampled electrical characterization on cell/BMS/fuse; poka-yoke reel verification; supplier lot-level acceptance. AP\: High (S5 can never leave High; O/D driven to floor). CC in MDT-400.
PFM-002 — Out-of-spec sensor / secure-element lot accepted
Failure mode | T/RH sensor or secure element outside datasheet accepted.
Effect | Calibration cannot bring T-channel into ±0.5 °C → biased verdict (S4); or provisioning fails.
Cause | Lot variation; wrong part.
Prevention | AML; sampled electrical check.
Detection | Incoming sampling; caught later at calibration (PROC-110) / provisioning (PROC-100).
S/O/D | S=7 · O=3 · D=4 → AP: High.
RCR/HA | RCR-026/053 / HA-RES-015.
Recommended action | Sampled characterization on the T-sensor lot; correlate to cal yield. AP\: Medium.
---
PROC-030 — SMT / reflow
PFM-003 ★ — Cold/insufficient reflow on a safety-net joint (charge-inhibit / protection / probe-presence)
Process function | STEP-033 reflow to DO-MFG-003 profile.
Failure mode | Marginal solder joint on a charge-inhibit, protection, or probe-presence net (head-in- pillow, cold joint, insufficient wetting) that passes AOI but opens later under thermal cycling.
Effect | A safety net silently fails in the field after −30↔+60 °C cycling → defeated charge-inhibit or protection (S5 path) or probe-fault read as valid temp (S4). HA-RES-005/006/003.
Cause | Reflow profile drift, oxidized pad, paste volume, warpage.
Prevention | Profiled reflow (CPP, DO-MFG-003); SPI (STEP-031); N2 as specified.
Detection | AOI (STEP-034) + X-ray on cell/BMS/protection area (STEP-035); ICT on the net (PROC-050); cross-section on periodic sample.
S/O/D | S=9 · O=4 · D=4 → AP: High.
RCR/HA | RCR-016 / HA-RES-005/006 / HA-ELEC-008.
Recommended action | Reflow profiling + periodic thermocouple re-profile; AOI/ICT coverage target on safety nets (RCR-016, PV-MFG-008) with cross-section audit; HASS to precipitate latent joints. AP\: High (S5) with O/D at floor. SC in MDT-400.
PFM-004 — Component mis-orientation / wrong value (non-safety)
Failure mode | Polarized/valued part placed reversed or wrong value.
Effect | Board functional failure → caught at ICT/FCT (scrap/rework, not field escape).
Cause | Feeder setup, program error.
Prevention | Placement poka-yoke; program verification; first-article.
Detection | AOI + ICT/FCT (PROC-050); 100 % functional.
S/O/D | S=3 · O=3 · D=3 → AP: Low.
Recommended action | Standard SMT controls; first-article per setup.
PFM-005 ★ — Mis-orientation / wrong value on a protection component (escapes screening)
Failure mode | Reversed/wrong protection IC, fuse, or interlock component that still "works" at FCT but has the wrong trip behavior.
Effect | Protection trips late/never → thermal event (S5). HA-RES-006.
Cause | Feeder mix, wrong reel, program error on the safety part.
Prevention | Poka-yoke feeders + AOI on protection parts (PV-MFG-003).
Detection | AOI orientation + protection fault-injection at PROC-130 (proves actual trip behavior).
S/O/D | S=9 · O=2 · D=3 → AP: High.
RCR/HA | RCR-054/002 / HA-RES-006.
Recommended action | 100 % protection fault-injection (PROC-130) as the backstop; challenge-unit verification of AOI. AP\: High (S5). CC in MDT-400.
---
PROC-040 — Conformal coat
PFM-006 ★ — Coat coverage gap / cure defect over record electronics
Failure mode | Voids, thin coat, or under-cure over the sensing/record electronics; or coat encroaches a sensor/antenna keep-out.
Effect | Condensation/ingress reaches electronics across cold↔warm cycling → intermittent wet-fault reads / record-integrity fault (S3→S4); or degraded sensor/RF. HA-RES-021.
Cause | Spray parameter drift, masking error, cure time/temp.
Prevention | Coat WI + keep-out fixtures (DO-MFG-005); cleanliness before coat.
Detection | UV-trace coat inspection; cure verification; downstream leak/insulation (PROC-140) partially.
S/O/D | S=6 · O=4 · D=5 → AP: Medium.
RCR/HA | RCR-015/018 / HA-RES-021.
Recommended action | UV inspection + periodic cross-section; cleanliness SPC. AP\: Low/Medium.
---
PROC-060 — Battery handling, connection / spot-weld, retention
PFM-007 ★ — Cell mechanical damage in handling (crush / puncture / dent)
Failure mode | Cell dented, punctured, or internally damaged during handling/installation.
Effect | Latent internal short → delayed thermal runaway/fire (S5), possibly after shipment. HA- RES-006.
Cause | Drop, wrong fixture, sharp tooling, over-insertion force, ESD.
Prevention | DO-MFG-007 handling WI (ESD-safe, no-crush tooling, force-limited fixtures); operator certification; fire-safe cell storage.
Detection | Visual/handling audit; OCV/impedance check post-connection; hard to detect a latent internal short (D high).
S/O/D | S=9 · O=3 · D=6 → AP: High.
RCR/HA | RCR-004/002 / HA-RES-006.
Recommended action | Poka-yoke insertion fixtures + force monitoring; 100 % post-handling OCV/impedance + visual; quarantine any dropped cell; HASS/aging to precipitate latents on a sample. AP\: High (S5 latent). CC in MDT-400.
PFM-008 ★ — Bad cell connection / spot-weld (weak, over-welded, or missing retention)
Failure mode | Weak/over-penetrating spot-weld or loose connector; or cell retention not fully engaged.
Effect | Intermittent power loss (silent gap S4) or, if retention fails, cell movement → puncture → thermal event (S5). HA-RES-004/006.
Cause | Weld parameter drift, electrode wear, fixture misload.
Prevention | Weld-parameter CPP + electrode maintenance (DO-EE-004); retention poka-yoke fixture (DO-MECH-012).
Detection | Weld pull-test on sample; connection resistance check; retention engagement verification; vibration screen on sample.
S/O/D | S=9 · O=3 · D=5 → AP: High.
RCR/HA | RCR-004 / HA-RES-004/006.
Recommended action | Weld monitoring (energy/displacement) + periodic pull-test; 100 % connection- resistance + retention verification. AP\: High (S5). SC in MDT-400.
---
PROC-070 — Enclosure seal / IP assembly
PFM-009 ★ — Seal mis-seated / pinched / wrong part → IP breach
Process function | STEP-071/072 seat seal + closure torque.
Failure mode | Gasket not fully in groove, pinched, twisted, contaminated, or wrong seal part fitted (esp. -XA aperture variant); or under/over-torque.
Effect | Ingress/condensation over reuse life → record-electronics fault / corrupted evidentiary record (S3→S4). HA-RES-021.
Cause | Manual seating error, debris, wrong seal kit, torque tool drift.
Prevention | Seal-seat WI + groove-fill check (DO-MFG-004); seal poka-yoke; torque tool calibrated to DO-MFG-006 CPP.
Detection | 100 % production leak test (PROC-140, DO-TST-011) — the primary detection; visual groove- fill; torque audit.
S/O/D | S=6 (S4 worst if record corrupted) · O=4 · D=3 → AP: Medium (High if record-integrity effect).
RCR/HA | RCR-011/015/018 / HA-RES-021.
Recommended action | 100 % leak test with limit correlated to IP; torque monitoring; seal-kit poka-yoke. AP\: Medium. SC in MDT-400.
PFM-010 — Over-torque / cross-thread → cracked boss or stressed seal
Failure mode | Fastener over-torqued/cross-threaded; boss cracked; seal over-compressed.
Effect | Latent seal/mechanical failure → IP loss over reuse (S3).
Cause | Torque tool drift, operator technique, no torque sequence.
Prevention | Torque spec + sequence (DO-MFG-006); calibrated drivers.
Detection | Torque audit; leak test (PROC-140).
S/O/D | S=5 · O=3 · D=4 → AP: Medium.
Recommended action | Torque monitoring + sequence poka-yoke.
---
PROC-080 — USB-C + mechanical final assembly
PFM-011 — USB-C aperture / probe gland seal not applied
Failure mode | Connector aperture or probe gland seal omitted/damaged.
Effect | Localized ingress path (S3, record-integrity). HA-RES-021.
Cause | Missed step, damaged part.
Prevention | Assembly WI (DO-MECH-008); poka-yoke.
Detection | Leak test (PROC-140); visual.
S/O/D | S=5 · O=3 · D=3 → AP: Medium.
Recommended action | Include gland/aperture in leak-path acceptance.
PFM-012 — USB-C connector poor retention / solder → charge-path fault
Failure mode | Weak connector mechanical/solder retention.
Effect | Intermittent charging → dead-on-arrival / loss of monitoring (S4 availability). HA-RES-004.
Cause | Solder/retention process.
Prevention | Connector WI; AOI/X-ray of connector.
Detection | FCT charge-path check (PROC-120); connector retention pull-test on sample.
S/O/D | S=7 · O=3 · D=4 → AP: High.
Recommended action | 100 % charge-path FCT; periodic pull-test. AP\: Medium.
---
PROC-090 — Firmware flashing + config
PFM-013 ★ — Wrong / unsigned / mismatched firmware image loaded
Failure mode | An incorrect, unsigned, or version-mismatched image is flashed.
Effect | Detection/verdict logic wrong or unverifiable across the fleet → compromised dose administered (S4, fleet-scale S5-worst). HA-RES-012/001.
Cause | Wrong image on the station, script error, tampered image.
Prevention | Controlled provisioning tool loads only the released signed image (DO-SW-001/DO-FW-019); signature enforced by secure boot; image-hash allow-list on the station.
Detection | Hash/version read-back (DO-TST-008) + secure-boot rejection; PV-MFG-007.
S/O/D | S=8 · O=2 · D=2 → AP: High (S4).
RCR/HA | RCR-037/038 / HA-RES-012.
Recommended action | Station allow-list + mandatory read-back; two-person control on image release to the line. AP\: Medium. CC in MDT-400.
PFM-014 ★ — Wrong regional / firmware config provisioned (band plan / market profile)
Failure mode | Unit flashed with the wrong region config (e.g., -EU band plan on an -NA unit).
Effect | Radio non-compliant / cannot attach → no alert egress in that market (S4 availability) + regulatory non-compliance. HA-RES / HA-MFG-006.
Cause | Line mix, wrong config bundle, no line clearance.
Prevention | Config bundle bound to SKU/traveler; line clearance between variants (PROC-160).
Detection | EOL config verification (PROC-120, PV-MFG-006) — wrong region caught.
S/O/D | S=7 · O=3 · D=3 → AP: High.
RCR/HA | RCR-056 / HA-MFG-006 / HA-COMM-006.
Recommended action | 100 % EOL config verify + traveler/SKU poka-yoke. AP\: Medium. SC in MDT-400.
---
PROC-100 — Cryptographic identity / key provisioning + debug-lock
PFM-015 ★ — Duplicate / non-unique cryptographic identity provisioned
Failure mode | Two units receive the same identity/key, or a key-generation/entropy fault produces a predictable key.
Effect | Device↔trip binding ambiguous → mis-attribution / spoofable authenticity fleet-wide (S4). HA-RES-010/016.
Cause | Provisioning-script bug, HSM/counter fault, re-used key material.
Prevention | Key-management with uniqueness enforcement (DO-SEC-007); HSM-backed generation; registry binding (DO-CFG-010).
Detection | Uniqueness check at provisioning + registry (PV-MFG-005); read-back verify.
S/O/D | S=7 · O=2 · D=3 → AP: High.
RCR/HA | RCR-055/033 / HA-RES-010/016.
Recommended action | Registry-enforced uniqueness (reject duplicate) + entropy self-test on the station. AP\: Medium. CC in MDT-400.
PFM-016 ★ — Debug lock not set / readout-protection not fused
Failure mode | Unit ships with the debug/EOL port open (RDP/lifecycle fuse not set).
Effect | Keys/records readable from a fielded unit → falsified telemetry / suppressed alarm (S4). HA-RES-010/011.
Cause | Skipped step, fuse-write failure, station error.
Prevention | Provisioning sequence sets and verifies the lock (DO-SEC-008); station refuses to pass an unlocked unit.
Detection | DO-TST-012 post-lock read-attempt refused (per serial).
S/O/D | S=7 · O=2 · D=2 → AP: High.
RCR/HA | RCR-038 / HA-RES-010/011.
Recommended action | 100 % post-lock verification (unauthenticated read refused); hard station gate. AP\: Medium. CC in MDT-400.
PFM-017 ★ — Incompletely-provisioned unit passes
Failure mode | A unit missing identity, keys, birth record, or lock passes to the next stage.
Effect | Unattributable or unsafe unit in the field (S4).
Cause | Station override, partial write, missing gate.
Prevention | Station refuses incomplete provisioning; DHR completeness gate (PROC-170).
Detection | Provisioning read-back + DHR review.
S/O/D | S=7 · O=2 · D=3 → AP: High.
RCR/HA | RCR-055/038/053.
Recommended action | Hard interlock — no pass without complete, verified provisioning record. AP\: Medium.
---
PROC-110 — Sensor calibration
PFM-018 ★ — Miscalibration / uncorrected bias written (single unit)
Failure mode | Wrong coefficients written, or a unit calibrated outside the accuracy budget passes.
Effect | T-channel biased → wrong accept/reject verdict on that unit (S4). HA-RES-015.
Cause | Reference drift, contact/thermal-lag error, coefficient math bug, wrong cal points.
Prevention | Traceable reference (ISO/IEC 17025); cal WI (DO-MFG-010); cal-limit enforcement (DO-TST-006).
Detection | Post-cal read-back verification against reference; birth-record capture; a miscalibrated unit fails EOL (RCR-053, PV-MFG-001).
S/O/D | S=8 · O=3 · D=3 → AP: High.
RCR/HA | RCR-026/053 / HA-RES-015.
Recommended action | Guard-banded cal-accept limit (MDT-800 MSA); independent verify point; auto-reject out-of-tolerance; reference-standard recal schedule. AP\: Medium. CC in MDT-400.
PFM-019 ★★ — Systematic fleet-wide calibration bias (reference/station error) — COMMON CAUSE
Failure mode | A drifted/incorrect reference standard or a station math error biases every unit calibrated on that station/period in the same direction.
Effect | Fleet-wide detection+verdict bias → many spoiled doses administered as good, or mass false-reject (S5-worst common cause). HA-RES-015, HA-MPE-002.
Cause | Reference standard out of cal, station firmware/coefficient regression, uncontrolled cal-station update.
Prevention | Reference-standard calibration chain + interval (ISO/IEC 17025); cal-station software version control (DO-SW-002/008); change-control on any cal-station update.
Detection | Golden-unit drift monitoring per station (DO-TST-013); cross-station correlation; periodic reference re-verification; SPC on cal residual distributions.
S/O/D | S=9 · O=2 · D=4 → AP: High.
RCR/HA | RCR-053/026/049 / HA-RES-015 / FA-CCF-002.
Recommended action | Independent reference cross-check + golden-unit trending + SPC + staged cal-station changes with re-qualification; standing field calibration-drift monitoring (folder 17, FM-200) because this common cause is a process control, not a one-time fact. AP\: High (S5) with O/D at floor. CC + named PMS metric.
---
PROC-120 — Final functional & system test
PFM-020 ★ — Escaped functional defect (unit passes with a broken safety/verdict feature)
Failure mode | A unit with a marginal sensor path, verdict-display fault, or alarm defect passes final test (test-escape / limits too loose / golden-unit drift).
Effect | Field device misreads/misdisplays a verdict or fails to alarm (S4). HA-RES-001/002/014.
Cause | Loose limits, un-guard-banded pass/fail, drifted test station, incomplete coverage.
Prevention | Guard-banded limits (DO-TST-014); golden units (DO-TST-013); coverage review.
Detection | Final system test itself + golden-unit re-verification; MSA on measured parameters.
S/O/D | S=8 · O=3 · D=4 → AP: High.
RCR/HA | RCR-019/020/021/043 / HA-RES-001/002.
Recommended action | Guard-band all safety/measurement limits; MSA (MDT-800); scheduled golden-unit checks; zero-defect acceptance on safety subtests. AP\: Medium.
---
PROC-130 — Protection / charge-inhibit fault-injection
PFM-021 ★ — Charge-inhibit / protection FI test not run or wrongly passed
Failure mode | The safety fault-injection is skipped, a sampled plan misses a bad unit, or a false-pass occurs (test rig fault).
Effect | A unit with a defeated charge-inhibit or non-functional protection ships → cold-charge/thermal event (S5). HA-RES-005/006.
Cause | Sampling instead of 100 %, rig calibration fault, station bypass.
Prevention | DO-TST-009 as a hard gate; rig self-check with a known-bad reference; no override.
Detection | Rig self-verification each shift; DHR completeness (PROC-170).
S/O/D | S=9 · O=2 · D=3 → AP: High.
RCR/HA | RCR-001/002 / HA-RES-005/006.
Recommended action | 100 % (not sampled) safety FI unless a sampling plan is justified and HAL- reconciled; daily known-bad-unit rig challenge; hard station interlock. AP\: High (S5). CC in MDT-400. 100 %-vs-sample decision Evidence Pending — reconcile with HAL.
---
PROC-140 — Seal-integrity / leak test
PFM-022 ★ — Leak test escape (leaking unit passes)
Failure mode | A unit with a marginal seal passes the leak test (limit too loose, poor IP correlation, rig leak, temperature-uncompensated pressure decay).
Effect | Ingress over reuse life → record-electronics fault / corrupted record (S3→S4). HA-RES-021.
Cause | Un-correlated limit, rig drift, fixture seal leak, ambient temp/pressure not compensated.
Prevention | Leak method + limit correlated to IP (DO-TST-011); rig calibration; master leak standard.
Detection | Daily master-leak (known-leak + known-good) rig challenge; MSA on the leak measurement.
S/O/D | S=6 (S4 worst) · O=3 · D=4 → AP: Medium (High on record-integrity effect).
RCR/HA | RCR-011/015/018 / HA-RES-021.
Recommended action | Establish leak-limit↔IP correlation + guard-band; daily master-leak challenge; MSA (MDT-800). AP\: Medium. SC in MDT-400. Method/limit Evidence Pending.
---
PROC-150 — Serialization / UDI / labeling
PFM-023 ★ — Wrong / missing regulatory or lithium-transport marking
Failure mode | Wrong-market regulatory markings, missing UDI, or missing/incorrect UN 38.3/IATA lithium marking applied.
Effect | Non-compliant shipment / un-recallable device / unsafe transport declaration (S4 regulatory/transport). HA-RES / DOL-1800.
Cause | Wrong artwork variant, line mix, no reconciliation with DOL-1800.
Prevention | Artwork bound to SKU/traveler; DOL-1800 reconciliation; line clearance.
Detection | Label inspection (DO-LBL-012) — content/legibility/placement/language; scan-verify.
S/O/D | S=7 · O=3 · D=3 → AP: High.
RCR/HA | RCR-055 / DOL-1800 / DO-LBL-002/004/005.
Recommended action | Per-SKU artwork poka-yoke + 100 % label content/scan inspection; freeze on DOL-1800 determination (RA-10). AP\: Medium. CC in MDT-400.
PFM-024 ★ — Duplicate / unreadable serial or identity-label mismatch
Failure mode | Duplicate serial, poor barcode/QR grade, or label serial ≠ provisioned identity.
Effect | Record cannot be attributed to the physical unit → traceability/recall failure (S4 evidentiary). HA-RES-016.
Cause | Printer fault, counter reset, mis-pairing of label to unit.
Prevention | Serial from the provisioning record (DO-CFG-010); print-and-verify.
Detection | Scan-verify grade + match to provisioned identity; registry uniqueness.
S/O/D | S=7 · O=2 · D=3 → AP: High.
RCR/HA | RCR-055/033 / HA-RES-016.
Recommended action | Print-verify-bind (label read-back matched to provisioned identity before pass). AP\: Medium. SC in MDT-400.
PFM-025 — Label lift/illegibility over cold/condensation/cleaning
Failure mode | Label adhesive/material fails at −30↔+60 °C or after inter-trip cleaning.
Effect | Loss of identity/markings over reuse life (S3 traceability).
Cause | Wrong label material (Evidence Pending), surface prep.
Prevention | Qualified label material/adhesive (DO-LBL-001).
Detection | Adhesion/legibility test on sample; reuse-cycle aging (folder 17 feedback).
S/O/D | S=5 · O=3 · D=4 → AP: Medium.
Recommended action | Qualify label material for cold/condensation/cleaning; monitor field legibility.
---
PROC-160 — Packaging / line clearance
PFM-026 ★ — Mixed-config / mixed-SKU build (line-clearance failure)
Failure mode | Units of different regional/config variants intermixed on the line or in a carton.
Effect | Wrong-region/wrong-marking device shipped (S4 availability + compliance). HA-MFG-006.
Cause | No line clearance between variant runs.
Prevention | Line clearance (DO-MFG-011, STEP-163); traveler control; single-variant runs.
Detection | EOL config verify (PROC-120); packaging content check; carton scan reconciliation.
S/O/D | S=7 · O=3 · D=3 → AP: High.
RCR/HA | RCR-056 / HA-MFG-006.
Recommended action | Enforced line-clearance record + carton-level scan reconciliation. AP\: Medium.
PFM-027 — Wrong-language IFU / missing accessory in pack
Failure mode | Wrong-language quick-start/IFU or missing probe/accessory kitted.
Effect | Use-error risk (misplacement / mis-arming) at point of use (S3/S4 use-mediated). HA-RES-014.
Cause | Kitting error, wrong artwork variant.
Prevention | Kitting WI + per-market BOM; poka-yoke.
Detection | Packaging content inspection (DO-LBL-012).
S/O/D | S=6 · O=3 · D=3 → AP: Medium.
Recommended action | Per-market kitting poka-yoke + content checklist.
---
PROC-900 — Rework / repair
PFM-028 ★ — Rework compromises seal / cell / coat without re-test
Failure mode | A reworked unit returns to flow without re-running the affected safety test, or rework damages the cell/seal/coat.
Effect | A unit with a defeated seal/cell/coat control ships (S3→S5 depending on control). HA-RES-006/ 021.
Cause | Missing re-test rule, prohibited rework performed, ESD/handling damage.
Prevention | DO-MFG-014 rework limits (prohibited on cell/seal); ESD controls.
Detection | Mandatory re-test on return to flow (leak/cal/protection/functional); DHR flag.
S/O/D | S=9 (worst) · O=3 · D=4 → AP: High.
RCR/HA | RCR-002/004/011 / HA-RES-006/021.
Recommended action | Hard rule — reworked unit re-runs every affected critical test; prohibited-rework poka-yoke; DHR rework flag reviewed at PROC-170. AP\: High (S5). CC in MDT-400.
---
3. Action-Priority summary
AP (residual, current controls) | Count | IDs
High | 18 | PFM-001,002,003,005,007,008,012,013,014,015,016,017,018,019,020,021,023,024,026,028 (S≥7; several S5)
Medium | 8 | PFM-006,009,010,011,022,025,027
Low | 1 | PFM-004
Every S5 process failure mode (PFM-001,003,005,007,008,019,021,028) remains High by rule and can only have its O/D driven to the floor — never removed from High — consistent with the severity-first policy and the HAL S5 acceptance rules (HAL-2200 §6.1). These become Critical Characteristics (CC) in MDT-400 and 100 %/error-proofed lines in the Control Plan (MDT-300).
4. Traceability
Backward | MDT-100 (PROC/STEP), the HAL residuals (HA-RES-###, HAL-2800), the RCR register (Stage 12), the ERFKL FMEA family (FA-HWFMEA/SWFMEA — the design-side failure modes these process modes could induce), and the manufacturing design outputs (DO-MFG-, DO-TST-).
Forward | MDT-300 (Control Plan — each PFM drives a control line), MDT-400 (CC/SC list), MDT-600 (Production Test Spec — the detection tests), MDT-800 (MSA for the detection measurements), MDT-1000 (pilot build proves O/D), folder 16 (open-defect / deviation input to release), folder 17 (the two process common causes — calibration PFM-019 and OTA/config — carry standing field monitoring), and ORCH-018 / EM-ENG-012.
Living | O/D estimates are pre-production; they are re-scored after MSA (MDT-800) and the pilot build, and again from field data (folder 17) via the closed loop.
MDT-300 — Control Plan — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-300 · Owner: Quality Engineer (EM-ENG-006) with Manufacturing and System-Safety. The production Control Plan links each Critical / Special Characteristic (MDT-400) to its process step (MDT-100 PROC/STEP), the process control that produces it, the measurement/detection method and sample/frequency, and the reaction plan when the characteristic is out of control. Every safety- and measurement-critical control that the design and risk analysis rely on (temperature-sensor calibration accuracy, seal/IP integrity, battery protection/charge-inhibit, cryptographic identity, signed firmware) is carried here as a controlled line so it is preserved into every production unit. Each line has a CP-### ID. Control-Plan phase: Pre-launch / Pilot (production phase inherits it once the pilot and MSA close).
Evidence honesty: specification limits, guard-bands, leak-rate, torque, reflow profile, and capability targets are Evidence Pending against the released design outputs and the MSA (MDT-800); this plan states what is controlled, how it is measured, and the reaction — not invented numeric limits. Safety/measurement- critical characteristics carry zero-defect acceptance (100 % / error-proofed), not sampling, unless a sampling plan is justified and HAL-reconciled.
---
1. Legend & conventions
Class | CC = Critical Characteristic (safety, S≥7 / S5 process modes); SC = Special/Significant Characteristic (function/compliance, high AP); KC = Key (quality). CC/SC defined in MDT-400.
Control method | the process parameter or poka-yoke that makes the characteristic right (prevention).
Measurement / detection | the gauge/test that proves it (detection), with sample size & frequency.
Reaction plan | what the line does on a non-conformance (contain, quarantine, notify, disposition).
Sample | "100 %" = every unit; "master challenge" = daily known-good/known-bad rig verification; "SPC" = statistical process control on the parameter; "FAI" = first-article inspection.
2. Control Plan register
---
Battery safety — charge-inhibit & protection (D3; the S5 spine)
CP-001 ★ CC — Charge-inhibit below 0 °C operates (HW interlock independent of MCU/FW)
Char. (MDT-400) | CC-01 · Step: PROC-130 / STEP-131 (built PROC-030/060).
Spec/limit | Charge current = 0 A below the qualified cell lower-charge limit, with FW halted/comms lost; inhibited state + cause reported. Exact limit (0 °C vs datasheet) Evidence Pending (DO-EE-003/RCR-001).
Control method | MCU-independent hardware interlock on cell-coupled temperature (DO-EE-007); FW defense-in-depth (DO-FW-005).
Measurement/detection | Fault-injection cold-charge test (DO-TST-009), 100 %; rig master challenge each shift with a known-bad reference.
Reaction plan | Fail → quarantine unit (PROC-910); if rig master-challenge fails, stop the line, hold all units since last good challenge, re-verify rig, re-test held units.
Trace | RCR-001 / HA-RES-005/006 / PFM-021 / PV-ENRG-001 / PV-SAFE.
CP-002 ★ CC — Autonomous protection + non-resettable secondary backstop trips
Char. | CC-02 · Step: PROC-130 / STEP-131.
Spec/limit | Injected OV/OC/OT/short → primary protection trips and the independent secondary backstop (fuse/OVP) coordinates; single-fault-safe.
Control method | Correct protection IC + rated fuse placed/oriented (PROC-030, poka-yoke + AOI); pack design (DO-EE-005/010).
Measurement/detection | Protection fault-injection (DO-TST-009), 100 %; AOI on protection parts + X-ray (PROC-030); daily rig master challenge.
Reaction plan | Fail → quarantine; investigate placement/part (Supplier CAPA if incoming); stop-line on rig-challenge failure.
Trace | RCR-002 / HA-RES-006 / PFM-005,021 / PV-SAFE.
CP-003 ★ CC — Correct cell chemistry + correct-rated/oriented protection parts
Char. | CC-03 · Step: PROC-020 / STEP-022 (+ PROC-030 AOI/X-ray).
Spec/limit | Only the qualified low-temp cell chemistry and datasheet-rated protection parts enter the build; no counterfeit secure element.
Control method | Frozen single-source AML + PCN hold (DO-CFG-004); reel/lot poka-yoke; supplier qualification (MDT-700).
Measurement/detection | Enhanced incoming — 100 % CoC + lot check, sampled chemistry/rating characterization + counterfeit screen (DO-TST-001, PV-MFG-002); AOI/X-ray of protection area.
Reaction plan | Non-conforming lot → quarantine + supplier CAPA + PCN review; block AML deviation.
Trace | RCR-054 / HA-RES-005/006 / PFM-001 / PV-MFG-002.
CP-004 ★ CC — Cell retention (no crush/puncture) + sound cell connection/weld
Char. | CC-04 · Step: PROC-060 / STEP-061/062/063.
Spec/limit | Cell undamaged (OCV/impedance in band, no dent/puncture); weld/connection within parameter; retention fully engaged.
Control method | No-crush handling WI + force-limited insertion fixtures (DO-MFG-007); weld-parameter CPP + electrode maintenance; retention poka-yoke (DO-MECH-012).
Measurement/detection | 100 % post-handling OCV/impedance + visual; weld pull-test on sample + weld energy/displacement monitoring; 100 % retention verification.
Reaction plan | Any dropped/damaged cell → quarantine (fire-safe); weld out-of-parameter → hold lot, pull-test, re-qualify electrodes.
Trace | RCR-004/002 / HA-RES-004/006 / PFM-007,008.
---
Sensor calibration accuracy (the measurement-integrity spine)
CP-005 ★ CC — Temperature-sensor calibration accuracy (internal + probe) within budget
Char. | CC-05 · Step: PROC-110 / STEP-111/112.
Spec/limit | Post-cal error ≤ the released accuracy budget (±0.5 °C internal/probe target) across the operating range vs. a NIST/traceable reference; coefficients + metadata written; read-back verified. Cal points / reference chain Evidence Pending (DO-TST-006).
Control method | Traceable reference bath (ISO/IEC 17025); cal WI (DO-MFG-010); cal-limit enforcement in firmware (DO-FW-009); guard-banded accept limit (MDT-800).
Measurement/detection | 100 % calibrate + independent verify point read-back; birth record bound to serial; miscalibrated unit auto-fails EOL (RCR-053); golden-unit drift check per station (DO-TST-013); SPC on cal-residual distribution.
Reaction plan | Out-of-tolerance → auto-reject/re-cal; golden-unit drift or reference-out-of-cal → stop station, quarantine all units since last good reference verification, re-verify reference, re-cal held units (this is the fleet common-cause line, PFM-019).
Trace | RCR-026/049/053 / HA-RES-015 / PFM-018,019 / PV-MFG-001.
---
Seal / IP integrity (record-integrity barrier)
CP-006 ★ SC — Seal / IP integrity of the sealed enclosure
Char. | SC-01 · Step: PROC-140 / STEP-141 (built PROC-070/080).
Spec/limit | Each unit passes the production leak test correlated to the target IP rating; leak-rate ≤ limit. Leak method + limit + IP correlation Evidence Pending (DO-TST-011).
Control method | Seal-seat + groove-fill WI + seal poka-yoke (DO-MFG-004); closure torque CPP (DO-MFG-006); probe-gland/USB-C aperture seal (DO-MECH-008).
Measurement/detection | 100 % leak test (e.g., pressure-decay, temp-compensated); daily master-leak challenge (known-leak + known-good); torque audit; MSA on leak measurement (MDT-800).
Reaction plan | Fail → rework per DO-MFG-014 (re-seal) + re-test; rig master-leak failure → stop, hold, re-verify rig, re-test.
Trace | RCR-011/015/018 / HA-RES-021 / PFM-009,010,022 / DO-TST-011.
CP-007 SC — Conformal-coat coverage over record electronics
Char. | SC-02 · Step: PROC-040 / STEP-042.
Spec/limit | Coat coverage per DO-MFG-005 with keep-outs honored; cure complete; cleanliness before coat within limit.
Control method | Coat WI + keep-out masking fixtures; cure profile; pre-coat cleanliness.
Measurement/detection | UV-trace coat inspection (100 %); cure verification; cleanliness SPC; periodic cross-section.
Reaction plan | Coverage gap → re-coat within rework limits; keep-out encroachment → quarantine/scrap (sensor/antenna affected).
Trace | RCR-015/018 / HA-RES-021 / PFM-006.
---
Firmware & identity (record authenticity / correctness)
CP-008 ★ CC — Signed, correct firmware image + correct regional config loaded
Char. | CC-06 · Step: PROC-090 / STEP-091.
Spec/limit | Programmed image hash/version = released signed image (DO-FW-019); region config = SKU; unsigned/mismatched rejected.
Control method | Controlled provisioning tool + station image allow-list (DO-SW-001); secure-boot signature enforcement; config bundle bound to traveler/SKU.
Measurement/detection | 100 % hash/version read-back (DO-TST-008, PV-MFG-007); EOL config verify (PROC-120, PV-MFG-006); two-person control on line image release.
Reaction plan | Mismatch → re-flash from controlled source; repeated mismatch → hold line, audit station image set.
Trace | RCR-037/056 / HA-RES-012 / PFM-013,014 / PV-MFG-006/007.
CP-009 ★ CC — Unique cryptographic identity + keys provisioned; debug-lock set
Char. | CC-07 · Step: PROC-100 / STEP-101/102/103.
Spec/limit | Unique per-device identity/keys bound to serial (no duplicates); debug/EOL port locked (RDP/lifecycle fuse); station refuses incomplete provisioning.
Control method | HSM-backed key generation + registry uniqueness enforcement (DO-SEC-007); provisioning sequence sets + verifies lock (DO-SEC-008).
Measurement/detection | 100 % uniqueness check + read-back (PV-MFG-005); 100 % post-lock read-attempt refused (DO-TST-012); entropy self-test on station.
Reaction plan | Duplicate/predictable key → reject + investigate station/HSM; unlocked unit → do not pass; registry duplicate → quarantine both units.
Trace | RCR-055/038/033 / HA-RES-010/011/016 / PFM-015,016,017 / PV-MFG-005.
---
Final acceptance & assembly quality
CP-010 ★ SC — Final functional & system acceptance (BIST, sensors, HMI/verdict, arm-and-capture, radios)
Char. | SC-03 · Step: PROC-120 / STEP-121.
Spec/limit | Boot + BIST pass; all sensors within limits; HMI/verdict correct; sample record captured; radios attach on the correct band plan; every safety subtest pass/fail. Guard-banded limits (DO-TST-014).
Control method | Released test program under version control (DO-TST-013/DO-SW-008); golden units.
Measurement/detection | 100 % final system test (DO-TST-005/007); scheduled golden-unit re-verification; MSA on measured parameters.
Reaction plan | Fail-code → rework/re-test or scrap; golden-unit drift → stop station, re-correlate.
Trace | RCR-019/020/021/043 / HA-RES-001/002/014 / PFM-020 / DO-TST-005.
CP-011 ★ SC — Debug-lock verification (post-provisioning readout refused)
Char. | SC-04 · Step: PROC-100 / STEP-102 → DO-TST-012.
Spec/limit | Unauthenticated read attempt on the locked port is refused, recorded per serial.
Control method | Lifecycle-fuse set at provisioning.
Measurement/detection | 100 % post-lock read-attempt test (DO-TST-012).
Reaction plan | Read succeeds → quarantine, do not ship, investigate fuse-write.
Trace | RCR-038 / HA-RES-010/011 / PFM-016.
CP-012 SC — PCBA safety-net integrity (ICT/FCT + AOI/X-ray coverage, no latent SPF)
Char. | SC-05 · Step: PROC-050 / STEP-051 (+ PROC-030 AOI/X-ray).
Spec/limit | Every safety net (protection, charge-inhibit, probe-presence) populated and functional; AOI/ICT coverage on safety-critical nets ≥ target; HASS applied. Coverage target Evidence Pending (DfT).
Control method | DfT test points (DO-PCB-011); SPI/AOI/X-ray program; HASS.
Measurement/detection | 100 % ICT/FCT with coverage report (PV-MFG-008); AOI/X-ray; HASS record.
Reaction plan | Miss/fail → rework/re-test; coverage below target → DfT/program change (ECO).
Trace | RCR-016 / HA-RES-005/006 / PFM-003 / PV-MFG-008.
---
Labeling, transport, configuration integrity
CP-013 ★ CC — Regulatory / UDI / lithium-transport marking correct
Char. | CC-08 · Step: PROC-150 / STEP-151/152/153.
Spec/limit | Correct-market regulatory markings, UDI content, and UN 38.3/IATA lithium marking per DOL-1800 + certificates; no unsupported mark. Markings set / UDI regime Evidence Pending (RA-10).
Control method | Artwork bound to SKU/traveler; DOL-1800 reconciliation; line clearance.
Measurement/detection | 100 % label inspection (DO-LBL-012) — content/legibility/placement/language; scan-verify.
Reaction plan | Wrong/missing marking → hold + relabel + re-inspect; systemic → hold SKU run.
Trace | RCR-055 / DOL-1800 / DO-LBL-002/004/005 / PFM-023.
CP-014 SC — Serial/identity label matches provisioned identity; readable & durable
Char. | SC-06 · Step: PROC-150 / STEP-151.
Spec/limit | Serial + machine-readable identity = provisioned identity; barcode/QR ≥ min grade; unique; durable across reuse/cleaning/cold. Label material Evidence Pending (DO-LBL-001).
Control method | Serial from provisioning record (DO-CFG-010); print-and-verify; qualified label material.
Measurement/detection | 100 % scan-verify grade + match to provisioned identity; registry uniqueness; sampled adhesion/legibility.
Reaction plan | Mismatch/low-grade → reprint + re-bind; duplicate → quarantine.
Trace | RCR-055/033 / HA-RES-016 / PFM-024,025.
CP-015 ★ SC — Correct-config build; no mixed-SKU (line clearance)
Char. | SC-07 · Step: PROC-160 / STEP-163 (+ PROC-120 config verify).
Spec/limit | No intermixed regional/config variants; carton contents single-variant; correct-language IFU/accessories kitted.
Control method | Line clearance between variant runs (DO-MFG-011); traveler control; per-market kitting poka-yoke.
Measurement/detection | EOL config verify (PV-MFG-006); packaging content inspection; carton scan reconciliation.
Reaction plan | Mix detected → quarantine carton/run, re-verify each unit.
Trace | RCR-056 / HA-MFG-006 / PFM-026,027.
---
Process parameters (CPP) & first-article
CP-016 SC — Reflow profile (CPP)
Char. | SC-08 · Step: PROC-030 / STEP-033.
Spec/limit | Reflow to DO-MFG-003 profile window (fine-pitch/QFN + cell/BMS-area reliability across −30↔+60 °C). Profile Evidence Pending.
Control method | Profiled reflow; N2 as specified; paste control (SPI).
Measurement/detection | Periodic thermocouple re-profile; SPI 100 %; AOI; cross-section audit.
Reaction plan | Profile drift → re-profile, hold suspect boards, cross-section.
Trace | DO-MFG-003 / RCR-016 / PFM-003.
CP-017 SC — Closure torque / fastening (CPP)
Char. | SC-09 · Step: PROC-070 / STEP-072.
Spec/limit | Torque values + sequence per DO-MFG-006 (even seal compression, no over-stress). Values Evidence Pending.
Control method | Calibrated torque drivers; sequence poka-yoke.
Measurement/detection | Torque monitoring/audit; leak test correlation (CP-006).
Reaction plan | Out-of-torque → re-work per sequence, re-leak-test.
Trace | DO-MFG-006 / RCR-011 / PFM-010.
CP-018 KC — Critical-to-quality dimensions & first-article
Char. | KC-01 · Step: PROC-070/080 (+ incoming enclosure).
Spec/limit | CTQ dimensions (seal groove, probe/USB-C apertures, HMI window) per released drawings (DO-VER-002); capability meets Cpk target. Cpk target Evidence Pending.
Control method | Qualified moulds/tooling (DO-MFG-013); process capability.
Measurement/detection | FAI (DO-TST-003 / PV-MFG-009); CMM/gauge; SPC on key dims; capability study.
Reaction plan | Out-of-capability → tool/process correction (ECO); FAI fail → hold ramp.
Trace | DO-VER-002 / RCR-011 / PV-MFG-009.
---
3. Control-Plan summary
Class | Count | Sample posture
CC (Critical) | 8 (CP-001,002,003,004,005,008,009,013) | 100 % / error-proofed + daily rig master challenge; stop-line reaction
SC (Special) | 9 (CP-006,007,010,011,012,014,015,016,017) | 100 % where safety/measurement; SPC + master challenge on CPPs
KC (Key) | 1 (CP-018) | FAI + SPC + capability
Design rules preserved: every CC line carries zero-defect acceptance with a reaction plan that stops the line and holds units on a detection-system (rig/reference/golden-unit) failure — because for the S5 and fleet-common-cause characteristics (charge-inhibit, protection, calibration), a drifted detector is itself a fleet hazard. Sampling-vs-100 % on the safety fault-injection (CP-001/002) is Evidence Pending — to be reconciled with HAL Stage 10; the default posture is 100 %.
4. Traceability
Backward | MDT-400 (CC/SC definitions), MDT-200 (PFM-### that each line detects), MDT-100 (PROC/STEP), the RCR register (Stage 12), HAL residuals (HAL-2800), and the design outputs (DO-MFG-, DO-TST-, DO-LBL-).
Forward | MDT-500 (Work Instructions realize the control methods), MDT-600 (Production Test Spec details the detection tests + ties to PV-MFG), MDT-800 (MSA qualifies the measurement systems referenced here), MDT-1000 (pilot proves the plan), folder 16 (release readiness confirms the plan is validated on pilot units), folder 17 (the standing field-monitoring lines for the calibration and OTA/config common causes), and ORCH-018 / EM-ENG-012 (Stage 18).
MDT-400 — Critical & Special Characteristics — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-400 · Owner: Quality Engineer (EM-ENG-006) with System-Safety and the discipline owners. The consolidated list of Critical Characteristics (CC) and Special / Significant Characteristics (SC) for the SCM — the product/process attributes whose variation, if uncontrolled, defeats a safety or measurement-integrity control. Each is derived from a design output (DO-) and/or a risk-control requirement (RCR-) and a PFMEA process failure mode (PFM-), and each is held by a Control-Plan line (CP-###). This list is the authoritative "what must never vary" register that the Control Plan, Work Instructions, and Production Test Spec enforce.
Definitions (aligned to AIAG/VDA + ISO 14971):
Critical Characteristic (CC) | variation can cause or fail to prevent a safety end-effect (S≥7, including the S5 battery-fire path) or a fleet-scale common-cause harm. Zero-defect / 100 % or error- proofed control; stop-line reaction on detector failure. These are non-negotiable and cannot be relaxed by sampling without HAL reconciliation.
Special / Significant Characteristic (SC) | variation affects function, compliance, or record integrity (high AP, typically S3–S4). Controlled 100 % where safety/measurement, else SPC + master challenge.
Key Characteristic (KC) | important to quality/fit; SPC + capability (FAI).
Evidence honesty: the numeric limits for these characteristics (accuracy budget, leak rate, torque, reflow window, capability targets) are Evidence Pending against the released design outputs and the MSA (MDT-800); this list identifies the characteristics and their control basis, not invented values.
---
1. Critical Characteristics (CC)
ID | Characteristic | Why critical (worst effect) | Source DO / RCR | PFM | CP | Verify / PV
CC-01 | Charge-inhibit below 0 °C operates (HW interlock independent of MCU/FW; holds in any FW/comms fault) | Cold-charge plating → latent short → thermal runaway/fire (S5, aircraft) | DO-EE-007 / RCR-001 | PFM-021 | CP-001 | DO-TST-009 / PV-ENRG-001
CC-02 | Autonomous cell protection + non-resettable secondary backstop trips | Cell OV/OC/OT/short un-terminated → fire (S5) | DO-EE-005/010 / RCR-002 | PFM-005,021 | CP-002 | DO-TST-009 / PV-SAFE
CC-03 | Correct cell chemistry + correct-rated/oriented protection parts | Wrong chemistry/protection → charge-safety envelope wrong → fire (S5) | DO-CFG-004 / RCR-054 | PFM-001 | CP-003 | DO-TST-001 / PV-MFG-002
CC-04 | Cell retention + sound cell connection/weld (no crush/puncture) | Cell damage/movement → latent internal short → fire (S5) | DO-MECH-012 / RCR-004/002 | PFM-007,008 | CP-004 | handling audit + OCV/weld pull
CC-05 | Temperature-sensor calibration accuracy (internal + probe) within budget; birth record bound | Biased T → wrong accept/reject verdict; systematic fleet bias = S5 common cause | DO-TST-006 / RCR-026/053 | PFM-018,019 | CP-005 | DO-TST-006 / PV-MFG-001
CC-06 | Signed, correct firmware image + correct regional config loaded | Wrong/unsigned FW → detection/verdict wrong fleet-wide (S4, S5-worst) | DO-FW-019 / RCR-037/056 | PFM-013,014 | CP-008 | DO-TST-008 / PV-MFG-006/007
CC-07 | Unique cryptographic identity + keys; debug-lock set | Duplicate/predictable key or open debug → spoof/suppress alarm, unattributable record (S4) | DO-SEC-007/008 / RCR-055/038/033 | PFM-015,016,017 | CP-009 / CP-011 | DO-TST-012 / PV-MFG-005
CC-08 | Regulatory / UDI / lithium-transport marking correct | Wrong/missing marking → non-compliant / un-recallable / unsafe transport (S4 + regulatory) | DO-LBL-002/004/005 / DOL-1800 | PFM-023 | CP-013 | DO-LBL-012 inspection
2. Special / Significant Characteristics (SC)
ID | Characteristic | Why special (worst effect) | Source DO / RCR | PFM | CP | Verify / PV
SC-01 | Seal / IP integrity of the sealed enclosure | Ingress/condensation → record-electronics fault / corrupted record (S3→S4) | DO-MECH-005 / DO-TST-011 / RCR-011/015/018 | PFM-009,010,022 | CP-006 | DO-TST-011 (100 % leak)
SC-02 | Conformal-coat coverage over record electronics | Condensation reaches electronics → wet-fault reads (S3→S4) | DO-MFG-005 / RCR-015/018 | PFM-006 | CP-007 | UV coat inspection
SC-03 | Final functional & system acceptance (BIST, sensors, HMI/verdict, arm-and-capture, radios) | Escaped functional defect → misread/mis-displayed verdict, no alarm (S4) | DO-TST-005/007 / RCR-019/020/021/043 | PFM-020 | CP-010 | DO-TST-005 (100 %)
SC-04 | Debug-lock verification (readout refused) | Open port → key/record readout on fielded unit (S4) | DO-SEC-008 / RCR-038 | PFM-016 | CP-011 | DO-TST-012 (100 %)
SC-05 | PCBA safety-net integrity (ICT/FCT + AOI/X-ray coverage; no latent SPF) | Latent solder/coverage defect on a safety net → silent field failure (S4→S5) | DO-PCB-011 / RCR-016 | PFM-003 | CP-012 | PV-MFG-008
SC-06 | Serial/identity label matches provisioned identity; readable & durable | Record cannot be attributed to unit → traceability/recall failure (S4) | DO-LBL-003 / DO-CFG-010 / RCR-055/033 | PFM-024,025 | CP-014 | scan-verify + registry
SC-07 | Correct-config build; no mixed-SKU (line clearance) | Wrong-region device shipped → no alert egress + non-compliance (S4) | DO-MFG-011 / RCR-056 | PFM-026,027 | CP-015 | PV-MFG-006 + carton scan
SC-08 | Reflow profile (CPP) | Marginal joints on safety nets → latent field failure under thermal cycling (S4→S5) | DO-MFG-003 / RCR-016 | PFM-003 | CP-016 | thermocouple re-profile + cross-section
SC-09 | Closure torque / fastening (CPP) | Uneven seal compression / cracked boss → IP loss over reuse (S3) | DO-MFG-006 / RCR-011 | PFM-010 | CP-017 | torque audit + leak correlation
3. Key Characteristics (KC)
ID | Characteristic | Effect | Source DO | CP | Verify
KC-01 | Critical-to-quality dimensions (seal groove, probe/USB-C apertures, HMI window) | Fit/seal quality; feeds SC-01 | DO-VER-002 / DO-MECH-018 | CP-018 | FAI + SPC / PV-MFG-009
4. Preservation principle (why this list exists)
The SCM's safety case rests on a small set of controls holding across the whole fleet. The Risk-Management Summary and HAL-2900 make the point explicitly: the device is "not merely many independent small risks" — it has common causes (firmware, calibration, OTA/cloud, key) whose single failure degrades many residuals at once, and two of those (calibration, config/OTA) are process controls, not one-time design facts. This CSC list is the mechanism by which those controls are designated, held to zero-defect, and monitored:
• The S5 spine (CC-01…CC-04) — battery charge-inhibit, protection, correct chemistry/parts, retention — is the physical catastrophic path; every one is 100 %/error-proofed with a stop-line reaction.
• CC-05 (calibration) is the single largest common-cause characteristic (a drifted reference biases every unit); its Control-Plan reaction stops the station and holds units on reference/golden-unit drift, and it carries a standing field-monitoring commitment (folder 17, FM-200) because it cannot be closed by a one-time test.
• CC-06/CC-07 (signed FW + crypto identity) preserve record correctness and authenticity fleet-wide.
• SC-01 (seal/IP) protects the evidentiary record for the reusable, multi-trip life (D1).
No CC may be down-classed or moved to a sampling plan without EM-ENG-006 + System-Safety approval and HAL (Stage 10) reconciliation, recorded as a design-review action (folder 16, DR-###).
5. Traceability
Backward | design outputs (DO-EE/MECH/FW/SEC/TST/LBL/CFG), the RCR register (Stage 12), the HAL residuals and control-implementation review (HAL-2500/2800), and the PFMEA (MDT-200).
Forward | MDT-300 (Control Plan holds each CC/SC), MDT-500 (Work Instructions realize them), MDT-600 (Production Test Spec proves them; ties to PV-MFG), MDT-800 (MSA qualifies the measurements), MDT-1000 (transfer readiness confirms all CC outputs are Released), folder 16 (release gate confirms zero open CC), and ORCH-018 / EM-ENG-012 (Stage 18).
MDT-500 — Work Instructions Summary — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-500 · Owner: Manufacturing Engineer with Quality (EM-ENG-006). A summary index of the controlled work instructions (WIs), procedures, and process specifications the contract manufacturer executes to build the SCM, keyed to the process flow (MDT-100 PROC/STEP), the released manufacturing design outputs (DO-MFG-, DOL-1400), and the Control Plan (MDT-300). This is not the full WI text (those are the controlled DO-MFG documents in PLM/MES); it is the transfer-level summary that shows every process step has an owning, released instruction, that the safety-critical steps carry the extra handling/verification content, and that operator training and revision control are defined.
Evidence honesty: the DO-MFG WIs are Draft (per DOL-1400, Rev TBD until the AML/chemistry/antenna freeze, RA-07); this summary lists them at their planned scope and flags the ones whose content values (reflow profile, torque, weld, leak method, cal points) are Evidence Pending. Nothing here asserts a released revision that does not exist.
---
1. Work-instruction register (by process stage)
WI / Proc ID | Title | Realizes DO | Steps (MDT-100) | Class | CP | Status
WI-020 | Incoming inspection & critical-part acceptance | DO-TST-001 | PROC-020 | ★ CC-03 | CP-003 | Draft
WI-030 | SMT / PCBA assembly (place/populate per SKU) | DO-MFG-002 | PROC-030 STEP-031/032 | KC | CP-016 | Draft
SPC-033 | Reflow / soldering process parameter spec (CPP) | DO-MFG-003 | PROC-030 STEP-033 | ★ SC-08 | CP-016 | Draft (profile EP)
WI-034 | AOI / X-ray inspection program | DO-TST-002 | PROC-030 STEP-034/035 | ★ SC-05 | CP-012 | Draft
WI-040 | Conformal-coat application (keep-outs, cure) | DO-MFG-005 | PROC-040 | ★ SC-02 | CP-007 | Draft
WI-050 | ICT / FCT board-level test operation | DO-TST-004 | PROC-050 | ★ SC-05 | CP-012 | Draft
WI-060 | Battery handling, connection/weld & retention | DO-MFG-007 | PROC-060 | ★ CC-04 | CP-004 | Draft (weld EP)
WI-070 | Enclosure seal / IP assembly (seat, closure) | DO-MFG-004 | PROC-070 | ★ SC-01 | CP-006 | Draft
SPC-072 | Torque / fastening & adhesive parameter spec (CPP) | DO-MFG-006 | PROC-070 STEP-072 | ★ SC-09 | CP-017 | Draft (torque EP)
WI-080 | USB-C + mechanical final assembly | DO-MECH-008 | PROC-080 | SC-01 | CP-006 | Draft
WI-090 | Firmware programming & config load | DO-MFG-008 | PROC-090 | ★ CC-06 | CP-008 | Draft
WI-100 | Cryptographic identity/key provisioning + debug-lock | DO-MFG-008 / DO-SEC-008 | PROC-100 | ★ CC-07 | CP-009/011 | Draft
WI-110 | Sensor calibration operation | DO-MFG-010 | PROC-110 | ★ CC-05 | CP-005 | Draft (cal pts EP)
WI-120 | Final functional & system test operation | DO-TST-005/007 | PROC-120 | ★ SC-03 | CP-010 | Draft
WI-130 | Protection / charge-inhibit fault-injection test op | DO-TST-009 | PROC-130 | ★ CC-01/02 | CP-001/002 | Draft (100%-vs-sample EP)
WI-140 | Seal-integrity / leak (IP) test operation | DO-TST-011 | PROC-140 | ★ SC-01 | CP-006 | Draft (method/limit EP)
PRC-150 | Serialization / UDI marking procedure | DO-MFG-009 | PROC-150 STEP-151 | ★ CC-08/SC-06 | CP-013/014 | Draft
WI-151 | Labeling application & inspection | DO-LBL-012 | PROC-150 STEP-153 | ★ CC-08 | CP-013 | Draft
WI-160 | Packaging, kitting & Li-transport marking | DO-LBL-006/007/008/010/011 | PROC-160 STEP-161/162 | SC-07 | CP-015 | Draft
WI-163 | Line-clearance procedure (SKU/variant changeover) | DO-MFG-011 | PROC-160 STEP-163 | ★ SC-07 | CP-015 | Draft
WI-170 | Finished-goods QA / DHR review & release | ISO 13485 §7.5 | PROC-170 | — | — | Draft
PRC-900 | Rework / repair & handling procedure | DO-MFG-014 | PROC-900 | ★ (re-test rule) | CP per test | Draft
PRC-910 | Non-conforming material / MRB procedure | QMS | PROC-910 | — | — | Draft
TL-012 | Assembly fixtures & tooling (setup/qualification) | DO-MFG-012 | multiple | — | — | Draft
TL-013 | Injection-mould tooling (enclosure) setup | DO-MFG-013 | (supplier) | — | — | Draft (resin EP)
★ = safety-critical WI (carries the extra handling/verification content below).
2. Mandatory content in safety-critical WIs
Each ★ WI must, beyond the assembly steps, carry the risk-control-preserving content:
• WI-020 (incoming) — critical-part acceptance: CoC + lot-traceability check, chemistry/rating verification, counterfeit screen; quarantine flow. (RCR-054 / CC-03.)
• WI-060 (battery) — ESD-safe, no-crush/no-puncture handling; force-limited insertion; fire-safe cell storage; quarantine any dropped cell; controlled shipping SoC; weld-parameter monitoring + pull-test frequency; retention-engagement verification. (RCR-002/004 / CC-04; low-temp cell handling.)
• WI-070 / SPC-072 (seal) — seal-seat & groove-fill check, correct seal-kit poka-yoke (incl. -XA aperture), closure sequence + torque; hand-off to 100 % leak test. (RCR-011/015 / SC-01/09.)
• WI-090 (firmware) — load only the released signed image (station allow-list); correct region config bound to traveler; hash/version read-back; reject unsigned/mismatched. (RCR-037/056 / CC-06.)
• WI-100 (identity) — HSM-backed unique key/identity generation; registry uniqueness enforcement; set and verify the debug-lock; refuse an incompletely-provisioned unit. (RCR-055/038 / CC-07.)
• WI-110 (calibration) — traceable-reference setup, thermal-lag/soak control, cal-points, coefficient write + independent read-back verify, birth-record bind to serial, guard-banded accept limit, reference-standard recal schedule, golden-unit drift check. (RCR-026/053 / CC-05; calibration integrity.)
• WI-130 (protection FI) — cold-cell/OV/OC/short injection sequence; confirm inhibit holds with FW faulted; confirm secondary backstop; daily rig known-bad master challenge; hard gate (no override). (RCR-001/002 / CC-01/02.)
• WI-140 (leak) — leak method, temp/pressure compensation, limit correlated to IP, daily master-leak challenge; rework→re-test rule. (RCR-011/015 / SC-01.)
• PRC-150 / WI-151 (labeling) — correct-market artwork bound to SKU, DOL-1800 reconciliation, UN 38.3/IATA Li marking, UDI content, scan-verify + identity match. (RCR-055 / CC-08.)
• WI-163 (line clearance) — verified changeover between regional/config variants; no residual prior-variant parts/labels/config. (RCR-056 / SC-07.)
• PRC-900 (rework) — prohibited rework on cell/seal; ESD/handling; mandatory re-test of every affected critical characteristic on return to flow; DHR rework flag. (PFM-028.)
3. Operator qualification & document control
Training/certification | operators on ★ WIs (battery handling, seal, calibration, provisioning, protection FI, labeling) require documented training + periodic re-qualification; battery-handling and protection-FI operators require safety certification. Training records are part of transfer readiness (MDT-1000, DO-XFER-007).
Revision control | WIs are controlled DO-MFG documents in PLM/MES; changes go through ECO (DO-GOV-004); the released rev on the line must match the transfer baseline (MDT-1000 reconciliation, DO-XFER-004). MES enforces the active WI revision per operation.
Poka-yoke & error-proofing catalogue (cross-WI) | reel/chemistry verification (WI-020), placement orientation (WI-030), seal-kit selection (WI-070), image allow-list (WI-090), provisioning completeness gate (WI-100), config-bound traveler (WI-090/160), line clearance (WI-163), print-verify-bind labeling (PRC-150).
4. Traceability
Backward | DOL-1400 (DO-MFG-), DOL-1500 (DO-TST-), DOL-1600 (DO-LBL-), MDT-100 (PROC/STEP), MDT-300 (CP-###), MDT-400 (CC/SC).
Forward | MDT-600 (Production Test Spec details the test WIs' acceptance), MDT-700 (supplier controls feed WI-020), MDT-800 (MSA qualifies WI-110/140 measurements), MDT-1000 (WI release + operator training are transfer-readiness line items), folder 16 (WI release status feeds manufacturing readiness), ORCH-018 / EM-ENG-012 (Stage 18).
Status honesty | all WIs are Draft and several carry Evidence-Pending content values (reflow, torque, weld, leak method/limit, cal points) — they cannot be Released until the AML/chemistry/antenna freeze (RA-07) and the associated design outputs (DO-MFG-003/006, DO-TST-006/011) close.
MDT-600 — Production Test Specification — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-600 · Owner: Quality/Test Engineer (EM-ENG-006) with Manufacturing, Metrology, Battery/Safety, Product-Security. The production (end-of-line) test specification: the set of factory tests each SCM unit passes to be accepted, their acceptance basis, sample/frequency, station/fixture, and their ties to the released test design outputs (DO-TST-, DOL-1500) and to the Stage-13 manufacturing verification (PV-MFG-###, PVL-1300). This is the acceptance spine — a unit is not shippable unless every safety- and measurement-critical test passes. Each production test carries a PT-### ID.
Evidence honesty: every acceptance limit (accuracy budget, leak rate, coverage %, capability index) is Evidence Pending against the released spec (DO-TST-014 limits/MSA) and, for calibration, anchored to the WHO PQS accuracy target (O-7, D4) — not invented here. Test status is Planned/Evidence-Pending (pre-execution). The 100 %-vs-sampled decision on the safety fault-injection is Evidence-Pending and reconciles with HAL Stage 10.
---
1. Test-flow position & principle
Production tests sit at PROC-050 (board), PROC-110/120/130/140 (unit), and PROC-150 (label), per MDT-100. The principle (from DOL-1500 §4): the five safety/measurement-critical tests are the acceptance spine — board FCT (PT-001), final system (PT-002), calibration (PT-003), protection/charge-inhibit FI (PT-005), seal/leak (PT-006) — and a unit that fails any of them is not shippable. Every production test references the verification observability and limits established in Stage 13 (PV-) and the design outputs; production tests do not re-derive limits, they apply the released ones.
2. Production-test register
---
PT-001 ★ — PCBA In-Circuit / Functional Test (ICT/FCT)
Step | PROC-050 · DO: DO-TST-004 (Critical) · PV tie: PV-MFG-008 (AOI/ICT safety-net coverage).
Purpose | First gate that every safety net (protection, charge-inhibit, probe-presence) is populated and functional on the bare PCBA via test points (DO-PCB-011); power rails, MCU boot, sensor buses, radios, charge path exercised.
Acceptance | All safety-net checks pass (pass/fail gate); coverage ≥ target. Coverage target Evidence Pending → DfT plan.
Sample/freq | 100 % of boards; coverage report per program rev.
Station/fixture | ICT/FCT fixture + versioned test SW (DO-TST-013).
Char. | SC-05 · CP: CP-012 · HA: HA-RES-005/006.
Reaction | Fail-code; rework (PRC-900) + re-test; coverage < target → DfT/program ECO.
PT-002 ★ — Final Functional & System Test
Step | PROC-120 · DO: DO-TST-005 (Critical) + DO-TST-007 (radio) · PV tie: PV-MFG-006 (config), PV-FUNC-500 / PV-PERF-600 (functional/perf verification basis).
Purpose | End-of-line proof the assembled unit boots, self-tests (BIST), reads all sensors within limits, exercises HMI/verdict, arms and captures a sample record, and attaches radios on the correct band plan.
Acceptance | All subtests pass; guard-banded limits (DO-TST-014); correct region config verified.
Sample/freq | 100 %.
Station/fixture | Final-test station + golden units (DO-TST-013).
Char. | SC-03 (+ SC-07 config) · CP: CP-010/015 · HA: HA-RES-001/002/014.
Reaction | Fail-code → rework/re-test or scrap; golden-unit drift → stop, re-correlate.
PT-003 ★ — Sensor Calibration & Verification (internal T / probe T / RH)
Step | PROC-110 · DO: DO-TST-006 (Critical) · PV tie: PV-MFG-001 (EOL calibration coverage/ accuracy + birth record; miscalibrated unit fails EOL).
Purpose | Calibrate/verify each unit against a NIST/traceable reference across the operating range; write coefficients + metadata; bind birth record to serial.
Acceptance | Post-cal error ≤ released accuracy budget (±0.5 °C internal/probe target; final value anchored to WHO PQS target, O-7, Evidence Pending); read-back verified; birth record present. A seeded miscalibrated unit fails EOL.
Sample/freq | 100 % calibrate + 100 % verify point; golden-unit drift check per station per shift; SPC on residuals.
Station/fixture | Calibration bath/reference (ISO/IEC 17025) + cal station SW (DO-SW-002) — MSA required (MDT-800).
Char. | CC-05 · CP: CP-005 · HA: HA-RES-015 (fleet common cause).
Reaction | Out-of-tolerance → auto-reject/re-cal; reference/golden-unit drift → stop station, hold all units since last good reference verify, re-verify reference, re-cal held units.
PT-004 ★ — Firmware Programming & Provisioning Verification
Step | PROC-090/100 · DO: DO-TST-008 · PV tie: PV-MFG-005 (identity uniqueness), PV-MFG-007 (programming/load).
Purpose | Confirm the released signed image (hash/version = DO-FW-019) is loaded, per-device identity/keys provisioned + read-back-verified, compatibility matrix satisfied; unsigned/mismatched rejected.
Acceptance | Hash/version match; identity unique + read-back OK; unsigned/wrong image rejected.
Sample/freq | 100 %.
Station/fixture | Provisioning station + HSM (DO-SW-001).
Char. | CC-06/CC-07 · CP: CP-008/009 · HA: HA-RES-010/012/016.
Reaction | Mismatch → re-flash from controlled source; duplicate identity → quarantine + investigate HSM.
PT-005 ★ — Battery-Protection / Charge-Inhibit Fault-Injection Safety Test
Step | PROC-130 · DO: DO-TST-009 (Critical) · PV tie: PV-ENRG-001 (cold-charge FI), PV-SAFE-900.
Purpose | Production proof of the D3 charge-safety controls — inject/simulate OV/OC and cold-cell conditions; confirm protection trips and charging is inhibited below the qualified limit with the inhibited state reported; secondary backstop coordinates.
Acceptance | Charge current = 0 A below limit incl. with FW faulted/comms lost; protection trips; inhibited state + cause reported. No unit ships with a defeated safety control.
Sample/freq | 100 % (default) — 100 %-vs-sampled decision Evidence Pending, reconcile with HAL Stage 10; daily rig known-bad master challenge.
Station/fixture | Safety FI rig with known-bad reference.
Char. | CC-01/CC-02 · CP: CP-001/002 · HA: HA-RES-005/006 (S5).
Reaction | Fail → quarantine; rig master-challenge fail → stop line, hold, re-verify rig, re-test held units.
PT-006 ★ — Seal-Integrity / Leak (IP) Test
Step | PROC-140 · DO: DO-TST-011 (Critical) · PV tie: PV-ENV-700 (environmental/IP verification basis).
Purpose | Confirm each sealed unit meets the ingress/condensation barrier that protects the evidentiary record.
Acceptance | Leak rate ≤ limit correlated to the target IP rating. Leak method + limit + IP correlation Evidence Pending (DO-TST-011).
Sample/freq | 100 %; daily master-leak challenge (known-leak + known-good); MSA on leak measurement.
Station/fixture | Pressure-decay/leak rig (temp/pressure compensated) + master leak standard.
Char. | SC-01 · CP: CP-006 · HA: HA-RES-021.
Reaction | Fail → re-seal (PRC-900) + re-test; rig master-leak fail → stop, hold, re-verify.
PT-007 ★ — Debug-Lock / Security-Provisioning Verification
Step | PROC-100 · DO: DO-TST-012 · PV tie: PV-SEC-1000 (cybersecurity verification basis).
Purpose | Confirm the debug/EOL port is locked (RDP/lifecycle fuse) and an unauthenticated read attempt is refused, recorded per serial.
Acceptance | Read attempt refused post-lock.
Sample/freq | 100 %.
Char. | SC-04/CC-07 · CP: CP-011 · HA: HA-RES-010/011.
Reaction | Read succeeds → quarantine, do not ship, investigate fuse-write.
PT-008 — Connectivity / Radio Production Test
Step | PROC-120 · DO: DO-TST-007 · PV tie: PV-MFG-006 (region config).
Purpose | Each radio (LTE-M/Wi-Fi/BLE/GNSS) verified for basic TX/RX/attach on the correct regional band plan; -XA verified with external antenna.
Acceptance | Conducted/radiated go-no-go within limits; correct band plan. Conducted-vs-radiated method Evidence Pending.
Sample/freq | 100 % (go-no-go); periodic radiated correlation.
Char. | SC-03/SC-07 · CP: CP-010/015.
Reaction | Fail → rework/re-test; wrong band plan → config re-provision + line-clearance review.
PT-009 — Electrical-Safety / ESD Production Test
Step | PROC-120 · DO: DO-TST-010 · PV tie: PV-SAFE-900.
Purpose | Applicable electrical-safety checks (per IEC 62368-1 classification for a battery-powered SELV device — dielectric/leakage where required) + ESD-robustness sampling at external ports.
Acceptance | Per safety classification (DO-REG-005). Applicability/scope Evidence Pending on safety class.
Sample/freq | Per plan (dielectric/leakage where required; ESD sampled).
Char. | — · CP: — (safety-supporting).
Reaction | Fail → quarantine.
PT-010 ★ — Serialization / UDI / Label Inspection
Step | PROC-150 · DO: DO-LBL-012 · PV tie: PV-MFG-005 (identity), PV-COMP-1500 (compliance basis).
Purpose | Verify correct regulatory/UDI/lithium-transport markings, serial/identity label matches provisioned identity, legibility/placement/adhesion/language, complete pack contents.
Acceptance | Correct content per DOL-1800; scan-verify ≥ grade; identity match; no missing/wrong marking.
Sample/freq | 100 % content/scan; sampled adhesion/legibility.
Char. | CC-08/SC-06 · CP: CP-013/014 · HA: HA-RES-016.
Reaction | Wrong/missing → hold + relabel + re-inspect; systemic → hold SKU run.
PT-011 — First-Article & Process Capability (Cpk) on critical characteristics
Step | ramp / setup · DO: DO-VER-002 · PV tie: PV-MFG-009 (FAI + capability).
Purpose | Confirm first article conforms to all CTQ/CTF and process capability meets the Cpk target on critical characteristics before ramp.
Acceptance | FAI conforms; Cpk ≥ target on CC/SC-linked dimensions. Cpk target Evidence Pending.
Sample/freq | FAI per setup + ramp capability study.
Char. | KC-01 (+ CC/SC dimensions) · CP: CP-018.
Reaction | Below capability → tool/process ECO; FAI fail → hold ramp.
3. Acceptance-spine summary & sampling posture
Test | 100 %? | Master challenge | MSA req'd | Critical
PT-001 FCT | Yes | — | — | ★
PT-002 Final system | Yes | golden units | measured params | ★
PT-003 Calibration | Yes | golden-unit/reference | Yes (MDT-800) | ★ CC
PT-004 Provisioning | Yes | — | — | ★ CC
PT-005 Protection/charge-inhibit FI | Yes (default; sampling EP, HAL-reconcile) | known-bad rig | — | ★ CC
PT-006 Seal/leak | Yes | master-leak | Yes (MDT-800) | ★ SC
PT-007 Debug-lock | Yes | — | — | ★
PT-008 Radio | Yes (go-no-go) | radiated correlation | — | —
PT-009 Electrical-safety | Per class | — | — | —
PT-010 Label/UDI | Yes (content/scan) | — | — | ★ CC
PT-011 FAI/Cpk | FAI + ramp | — | Yes (capability) | —
The safety- and measurement-critical tests (PT-001/002/003/004/005/006/007/010) are 100 % / zero-defect; their detection systems (calibration reference, leak rig, protection-FI rig, golden units) themselves carry daily master-challenge verification and, for calibration/leak, MSA (MDT-800) — because a drifted detector on a fleet common-cause characteristic (calibration) is itself a fleet hazard.
4. Traceability
Backward | DOL-1500 (DO-TST-), DOL-1400 (DO-MFG-), Stage-13 PVL-1300 (PV-MFG-###) and the broader PV- verification set (PV-FUNC/PERF/ENV/SAFE/SEC), MDT-300 (CP-###), MDT-400 (CC/SC), the HAL residuals.
Forward | MDT-800 (MSA qualifies PT-003/006 measurements + capability for PT-011), MDT-1000 (test-line readiness = DO-XFER-005; golden units correlated; limits/MSA closed), folder 16 (verification-status input to the release gate — these tests must be validated on pilot units and their PV-MFG closed), folder 17 (field test-escape signals feed the closed loop), ORCH-018 / EM-ENG-012 (Stage 18).
Open | numeric limits + MSA (DO-TST-014) and the calibration accuracy target (WHO PQS, O-7) are the Evidence-Pending closures; the safety-FI sampling posture reconciles with HAL Stage 10.
MDT-700 — Supplier Controls & Incoming Inspection — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-700 · Owner: Quality Engineer (EM-ENG-006) + Supply-Chain with Battery/Safety and Product-Security. Defines the supplier-qualification and incoming-inspection controls that keep out-of-spec, counterfeit, or wrong-chemistry parts out of the SCM — with enhanced control on the safety-critical parts (low-temperature cell, BMS/AFE, protection fuse, secure element) whose variation feeds the S5 battery-fire path and the fleet-scale common causes. Supplier controls carry SUP-###; incoming-inspection lines carry INC-###. This module realizes DO-TST-001 (incoming spec) + DO-CFG-004 (AML) and ties to PV-MFG-002 (incoming: correct chemistry & protection parts).
Evidence honesty: the AML cannot freeze until the low-temp cell chemistry (CMP-006) and the antenna branch (CMP-028/029) are selected — this is release blocker RA-07. Datasheet-specific incoming attributes (chemistry markers, protection trip points, sensor accuracy) are Evidence Pending until the parts are selected; this module defines what class of control applies and why, not invented part numbers or limits.
---
1. Supplier-control tiers
Tier | Parts (CMP) | Why | Controls
T1 — Safety-critical | Cell/pack (CMP-006/007), BMS/AFE (CMP-008), protection fuse (CMP-010), secure element (CMP-016) | Feed S5 fire path + fleet common causes | SUP: full qualification, single-source + controlled second-source, PCN/PCN-hold, audit, UN 38.3/IATA evidence, counterfeit controls. INC: 100 % CoC + lot traceability + sampled characterization + counterfeit screen
T2 — Function/measurement | T/RH sensors (CMP-019/020/021), PMIC, MCU (CMP-015), radios (CMP-024…029), USB-C | Affect accuracy/function/compliance | SUP: qualification + PCN. INC: sampled electrical/dimensional per datasheet
T3 — Standard | Passives, connectors, fasteners, enclosure resin, seals, labels | Quality/fit | SUP: approved-vendor + CoC. INC: sampling per ANSI/ASQ Z1.4
2. Supplier controls (SUP-###)
SUP-001 ★ — Cell / pack supplier qualification (low-temp chemistry)
Scope | Qualify the low-temp Li-ion (or LTO-class) cell + pack supplier against the D3 requirement set.
Requirements | UN 38.3 test summary + IATA DGR classification; cell datasheet incl. low-temperature charge limit and cold-endurance/capacity-fade data (feeds D3, PRD-VER-004); IEC 62133 / applicable cell safety; PPAP-class submission; lot-level CoC + traceability; PCN with hold rights; capacity/manufacturing audit; fire-safe logistics.
Ties | DO-EE-003/004, RCR-002/004, HA-RES-005/006, PV-MFG-002. Status: Open — chemistry not selected (RA-07); this is the single largest supplier-readiness blocker.
SUP-002 ★ — BMS/AFE + protection-fuse supplier qualification
Scope | Protection IC/AFE + non-resettable secondary fuse/OVP supplier.
Requirements | Datasheet trip points + coordination with cell; AEC/industrial grade as applicable; PCN hold; counterfeit controls; lot traceability; sampled trip-point characterization capability.
Ties | DO-EE-005/010, RCR-002, HA-RES-006, PV-MFG-002.
SUP-003 ★ — Secure-element supplier + authenticity controls
Scope | Secure element / crypto part supplier and its authenticity/anti-counterfeit chain.
Requirements | Authorized-distributor-only sourcing; anti-counterfeit provenance; secure key-injection supply-chain controls; PCN hold; coordinate with Product-Security (EM-ENG-008).
Ties | DO-SEC-007, RCR-055/038, HA-RES-010, PV-MFG-005.
SUP-004 — Sensor (T/RH) supplier qualification
Scope | Temperature + humidity sensor suppliers.
Requirements | Datasheet accuracy/stability; lot-level accuracy consistency (feeds calibration yield); PCN; sampled characterization.
Ties | DO-TST-006, RCR-026/053, HA-RES-015.
SUP-005 — Radio / antenna supplier qualification
Scope | LTE-M/Wi-Fi/BLE/GNSS modules + antenna branch (integral CMP-028 vs external CMP-029).
Requirements | Module regulatory pre-certification evidence (FCC/CE); band support per region; PCN. Status: antenna branch Open (RA-07) — integral-vs-external decision pending PRD-VER-005 (RF-in-container).
Ties | DO-EE-017/018/021, DO-CFG-006, HA-RES-009.
SUP-006 — Enclosure/resin, seal, and label material suppliers
Scope | Enclosure resin (cold-tough), seal/gasket, label material/adhesive.
Requirements | Material compatibility with cleaning agents (D1 reuse) + cold/condensation; seal material qualification; label material qualified for −30↔+60 °C / condensation / cleaning (DO-LBL-001, EP); RoHS/REACH.
Ties | DO-MECH-010, DO-MFG-013, DO-LBL-001, RCR-011.
SUP-007 — General supplier-quality agreement & change control
Scope | All AML suppliers.
Requirements | Quality agreement; PCN/EOL notification; corrective-action responsiveness (Supplier CAPA); sub-tier flow-down for T1; periodic scorecard/audit.
Ties | DO-CFG-004, DO-GOV-006.
3. Incoming inspection (INC-###)
INC-001 ★ — Cell / pack incoming acceptance
Attributes | 100 % CoC + lot-traceability check; cell chemistry/rating verification (poka-yoke against wrong chemistry, RCR-054); sampled OCV/impedance/capacity characterization; UN 38.3 evidence on file; visual damage/no-dent; fire-safe quarantine for suspect lots.
Reaction | Non-conforming lot → quarantine + Supplier CAPA + PCN review; no AML deviation on chemistry.
DO/PV | DO-TST-001, PV-MFG-002 · Char.: CC-03 · CP: CP-003.
INC-002 ★ — BMS/AFE + protection-fuse incoming acceptance
Attributes | 100 % CoC; sampled trip-point/rating characterization; counterfeit screen; correct part/marking verification.
Reaction | Wrong-rated/counterfeit → quarantine + Supplier CAPA.
DO/PV | DO-TST-001, PV-MFG-002 · Char.: CC-03 · CP: CP-003.
INC-003 ★ — Secure-element incoming acceptance
Attributes | Authorized-distributor provenance check; counterfeit/authenticity screen; lot traceability.
Reaction | Provenance fail → quarantine + Security review.
DO/PV | DO-TST-001, PV-MFG-005 · Char.: CC-07.
INC-004 — Sensor (T/RH) incoming acceptance
Attributes | Sampled accuracy/characterization per datasheet; lot consistency check (feeds cal yield).
Reaction | Out-of-datasheet lot → quarantine.
DO/PV | DO-TST-001 · Char.: CC-05 (feeds).
INC-005 — PCBA / bare-board & general electronic incoming
Attributes | Sampling per Z1.4; dimensional/electrical per released spec; workmanship (IPC-A-600 bare board).
DO | DO-TST-001.
INC-006 — Enclosure / seal / mechanical incoming
Attributes | CTQ dimensions (seal groove, apertures) sampled per DO-VER-002; material CoC; seal-part verification (incl. -XA aperture variant).
DO/PV | DO-TST-003, PV-MFG-009 · Char.: KC-01 (feeds SC-01).
INC-007 ★ — Label / artwork incoming
Attributes | Correct artwork variant per SKU; material/adhesive qualification lot check; regulatory/UDI/ Li-transport content verification vs. DOL-1800.
Reaction | Wrong variant/content → quarantine (compliance risk).
DO/PV | DO-LBL-012 · Char.: CC-08.
4. Enhanced-control principle (safety-critical parts)
The Risk-Management Summary names wrong chemistry / wrong protection part as an S5 process cause (HA-MFG-002/003) and fleet-wide calibration/sensor variation as an S5-worst common cause (HA-RES-015). Incoming inspection is therefore not uniform: T1 parts get 100 % CoC + lot traceability + sampled characterization + counterfeit screening, and their suppliers are single-source-controlled with PCN-hold so a supplier cannot silently substitute chemistry, trip point, or die. Any T1 non-conformance triggers Supplier CAPA and, if any suspect lot may have reached a build, feeds the closed loop (folder 17) for containment/recall assessment.
5. Traceability
Backward | DO-CFG-004 (AML), DO-TST-001 (incoming spec), DO-LBL-012 (label inspection), the HAL mfg/ supply hazards (HAL-1700, HA-MFG-), RCR-054/055, and PV-MFG-002/005/009.
Forward | MDT-100 PROC-010/020, MDT-300 CP-003/013/014/018, MDT-400 CC-03/07/08, MDT-1000 (supplier readiness = DO-XFER-006), folder 16 (supplier-readiness input to the release gate; AML freeze RA-07 is a release blocker), folder 17 (supplier-quality field escapes feed CAPA), ORCH-018 / EM-ENG-012 (Stage 18).
Open | SUP-001 (cell chemistry) and SUP-005 (antenna) are Open (RA-07); T1 supplier qualification cannot complete until the AML freezes.
MDT-800 — Calibration & Measurement-System Analysis (MSA) Plan — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-800 · Owner: Metrology + Quality Engineer (EM-ENG-006). Defines (a) the product calibration integrity — how each SCM's temperature/humidity channels are calibrated against a traceable reference, kept traceable across production, and re-verified in service — and (b) the measurement-system analysis (MSA / gauge R&R) that qualifies the production measurement systems themselves (calibration station, leak rig, torque/dimensional gauges) so that a pass/fail decision on a safety- or measurement- critical characteristic is trustworthy. Calibration integrity is one of the SCM's two named process common causes (HA-RES-015, S5-worst): a drifted reference biases every unit. MSA/calibration IDs carry MSA-### and CAL-###.
Evidence honesty: no capability index (Cpk), gauge-R&R %, accuracy budget number, or leak-rate limit is asserted — all are Evidence Pending against DO-TST-014 (limits/MSA) and the WHO PQS accuracy target (O-7, D4). This plan defines the method and the acceptance framework, not results.
---
1. Product-calibration integrity (CAL-###)
CAL-001 ★ — Traceable EOL calibration of T (internal + probe) and RH
Method | Calibrate/verify each unit at defined points across the operating range against a NIST/ nationally-traceable reference (calibration bath/chamber) per ISO/IEC 17025 concepts; write coefficients + metadata; independent read-back verify.
Acceptance | Post-cal error ≤ released accuracy budget (±0.5 °C internal/probe target; final value anchored to the WHO PQS accuracy target — Evidence Pending, O-7/D4), with a guard-band (see MSA-001) so measurement uncertainty cannot pass a marginal unit.
Record | Birth record bound to serial (DO-CFG-010); coefficients + cal date + reference used + operator/ station. A miscalibrated unit fails EOL (RCR-053).
Ties | DO-TST-006, DO-MFG-010, DO-FW-009, RCR-026/053, PV-MFG-001, CC-05, PT-003.
CAL-002 ★ — Reference-standard chain & interval
Method | Each cal-station reference is itself calibrated on a scheduled interval against a higher-tier standard; chain documented (unbroken traceability). Reference out-of-cal → all units since last good verification are suspect (stop-line reaction, CP-005).
Acceptance | Unbroken traceability chain; reference within its cal interval and tolerance.
Ties | ISO/IEC 17025; FA-CCF-002 (calibration common cause); the largest fleet common-cause defence.
CAL-003 ★ — Golden-unit / cross-station drift control
Method | Golden reference units (DO-TST-013) run on each cal station per shift; station-to-station correlation checked; SPC on cal-residual distributions.
Acceptance | Golden-unit reading within control limits; inter-station correlation within limit.
Reaction | Drift → stop station, re-verify reference, re-cal held units.
Ties | DO-TST-013, RCR-053, HA-RES-015.
CAL-004 — Field/depot recalibration interval (service life, D1)
Method | Periodic field/depot recalibration-or-verification against a traceable reference restores/ confirms accuracy over the reusable life; out-of-tolerance units quarantined from service; enforced by the pre-arm expiry gate (RCR-049).
Acceptance | Recal residual ≤ budget; interval set from drift data.
Ties | DO-SVC-005, RCR-049, HA-RES-015. Status: recal interval Evidence Pending (drift data from reliability + field, folder 17).
2. Measurement-System Analysis (MSA-###)
The MSA program qualifies each production measurement before it is trusted to accept/reject safety- and measurement-critical characteristics (AIAG/VDA MSA). Every measurement used on a CC/SC gets a study.
MSA ID | Measurement system | Characteristic | Study type | Acceptance basis | Ties
MSA-001 ★ | EOL calibration station (T/RH accuracy) | CC-05 calibration accuracy | Gauge R&R + bias/linearity/stability vs. reference | Gauge R&R + uncertainty within budget so guard-band valid (EP → DO-TST-014) | CAL-001, PT-003, PV-MFG-001
MSA-002 ★ | Seal/leak rig (pressure-decay) | SC-01 seal/IP | Gauge R&R + master-leak correlation (temp/pressure comp) | Discriminates leak-limit ± guard-band; master-leak repeatable (EP → DO-TST-011) | PT-006, CP-006
MSA-003 | Protection-FI rig (trip detection) | CC-01/02 | Attribute agreement (known-good/known-bad) | 100 % correct on challenge set | PT-005, CP-001/002
MSA-004 | Final-system-test measured params (sensor reads, timing) | SC-03 | Gauge R&R on key measured params | Within limit so guard-band valid | PT-002, CP-010
MSA-005 | Dimensional gauges / CMM (CTQ dims) | KC-01 (feeds SC-01) | Gauge R&R + capability | Gauge R&R acceptable; supports Cpk study | PT-011, CP-018
MSA-006 | Torque drivers/monitors | SC-09 CPP | Calibration + capability | Torque within window repeatably | CP-017
MSA-007 | Barcode/QR scan-verify + registry match | SC-06 | Attribute agreement | 100 % correct grade/match/uniqueness | PT-010, CP-014
Attribute vs. variable: variable (continuous) measurements (calibration accuracy, leak rate, torque, dimensions, timing) get gauge R&R + bias/linearity/stability; go/no-go safety and identity decisions (protection trip, label match, debug-lock) get attribute-agreement (known-good/known-bad challenge) studies.
3. Guard-band principle (why MSA gates the limits)
For the two most safety-load-bearing measurements — calibration accuracy (CC-05) and leak integrity (SC-01) — the pass/fail limit must be guard-banded inward from the specification by the measurement uncertainty established in MSA-001/002, so that measurement error cannot pass a unit that is actually out of spec (DO-TST-014). The guard-band cannot be set until the MSA is run; therefore the numeric limits are Evidence Pending on the MSA results, and the Control Plan reaction (stop-line on reference/rig drift) is the interim defence.
4. Calibration as a fleet common cause (the standing obligation)
HAL-2900 §5.3 identifies calibration as the information chain's dominant common cause whose containment is a process/operational control, not a one-time design fact. This plan therefore establishes not only the one-time EOL calibration (CAL-001) but the standing controls: reference-chain interval (CAL-002), golden- unit/cross-station drift trending (CAL-003), and field/depot recalibration (CAL-004) — and it hands a standing field calibration-drift monitoring metric to Stage 17 (folder 17, FM-200). Verification of the EOL coverage/accuracy is PV-MFG-001 (a named release condition, O-5); its target is anchored to the WHO PQS accuracy target (O-7), which must be confirmed (D4) before the accuracy budget is fixed and met.
5. Traceability
Backward | DO-TST-006/011/014 (calibration/leak/limits+MSA), DO-MFG-010, DO-SVC-005, RCR-026/049/053, HA-RES-015, PV-MFG-001, PT-003/006 (MDT-600), CP-005/006 (MDT-300), CC-05/SC-01 (MDT-400).
Forward | MDT-1000 (MSA complete + reference chain established are transfer-readiness line items; feeds DO-XFER-005 test readiness), folder 16 (MSA/capability status feeds manufacturing readiness — the release gate cannot pass with MSA open on CC-05/SC-01), folder 17 (FM-200 standing calibration-drift metric + reliability drift data set the recal interval CAL-004), ORCH-018 / EM-ENG-012 (Stage 18).
Open | accuracy budget (WHO PQS, O-7), leak limit/method (DO-TST-011), gauge-R&R and Cpk results (DO-TST-014), and the field/depot recal interval (drift data) are all Evidence Pending.
MDT-900 — Serialization, UDI & Production Traceability — Smart Cold Chain Monitor (SCM)
Stage 15 · Module MDT-900 · Owner: Manufacturing Engineer + Quality/Regulatory (EM-ENG-006) with Product-Security. Defines how each SCM unit is uniquely identified, marked, and bound to its production records so that (a) every fielded unit is traceable to its build history for recall/CAPA, (b) the evidentiary cold-chain record is attributable to a physical device (the ALCOA+/GDP requirement), and (c) the device's cryptographic identity, calibration birth-record, and serial/UDI all reconcile. Realizes DO-MFG-009 (serialization/UDI procedure), DO-CFG-010 (identity convention), and DO-LBL-003/004 (serial/UDI labels).
Evidence honesty: the exact UDI regime (whether/how UDI applies) and the regulatory-markings set follow the DOL-1800 classification determination — Evidence Pending (RA-10). Serial/identity format is defined; regulatory content is rendered from DOL-1800, not asserted here.
---
1. Identity layers (what identifies a unit)
Layer | What | Source | Bound at | Label | Trace
Serial number | Human-readable unique serial | DO-CFG-010 convention | PROC-100/150 | DO-LBL-003 | DHR, registry
Machine-readable ID | Barcode/QR encoding the device identity | DO-CFG-010 | PROC-150 | DO-LBL-003 | scan-verify (PT-010)
Cryptographic identity | Per-device key/identity in the secure element | DO-SEC-007 | PROC-100 | (not on label) | PV-MFG-005
UDI-DI / PI | Regulatory unique device identification (if applicable) | DOL-1800 / DO-LBL-004 | PROC-150 | device + package | reconcile with DOL-1800
Calibration birth record | Cal coefficients + reference + date | PROC-110 (CAL-001) | PROC-110 | (record) | bound to serial
Config/firmware identity | Region config + FW hash/version | PROC-090 | PROC-090 | (record) | DO-CFG-005
Binding rule: at PROC-150 the printed serial + machine-readable ID must match the provisioned cryptographic identity (print-verify-bind, CP-014/SC-06) and the registry must confirm uniqueness (no duplicate serial or identity, RCR-055). A mismatch or duplicate blocks the unit.
2. Serialization & marking (DO-MFG-009)
• Serial convention (DO-CFG-010): encodes SKU/regional variant, and is unique per unit; no duplicates across the fleet; durable across the reusable, multi-trip life (D1) and inter-trip cleaning.
Marking | applied at PROC-150; machine-readable symbology scan-verified to a minimum grade (PT-010); label material qualified for −30↔+60 °C / condensation / cleaning (DO-LBL-001, material EP).
Regulatory markings | CE/UKCA/FCC ID/IC/WEEE/RoHS + any WHO PQS/market mark rendered from DOL-1800 + the actual certificates (DO-LBL-002, CC-08); no unsupported mark shown.
Lithium-transport marking | UN 38.3 / IATA (UN 3481, batteries in equipment) applied per DO-LBL-005.
3. UDI (DO-LBL-004)
Applicability & regime | Evidence Pending (RA-10) — whether UDI applies (and under which regime: US FDA GUDID / EU EUDAMED / other) follows the US/EU classification in DOL-1800. If applicable: UDI-DI/PI content + carrier defined; device and package UDI consistent; parseable and database-registerable; DI linked to the configuration baseline (DO-CFG-010 / DOL-100).
Do not assert | no specific UDI database registration or regime is claimed until DOL-1800 resolves.
4. Production traceability (the Device History Record spine)
Every unit accumulates a Device History Record (DHR) that binds its identity to its build:
Record | Captured at | Content | Gate
Critical-part lot traceability | PROC-020 | Cell/BMS/SE/fuse lot IDs (T1) | INC-001/002/003
SMT/reflow + AOI/ICT | PROC-030/050 | Program rev, coverage, fail-codes | PT-001
Battery build | PROC-060 | Cell lot, weld/connection, retention | CP-004
Seal/leak | PROC-070/140 | Torque record, leak result | PT-006
Firmware/config | PROC-090 | Image hash/version, region config | PT-004
Identity/lock | PROC-100 | Identity provisioned, lock verified | PT-004/007
Calibration birth record | PROC-110 | Coefficients, reference, residuals | PT-003
Final/protection FI | PROC-120/130 | Functional + protection FI pass | PT-002/005
Serial/UDI/label | PROC-150 | Serial, UDI, markings, scan grade | PT-010
DHR review/release | PROC-170 | Completeness check, disposition | —
Completeness gate (PROC-170): a unit cannot ship with a missing or failed critical record (calibration birth, protection FI, leak, provisioning/lock, correct label). This DHR is the forensic basis for field CAPA and recall scoping (folder 17): given a field failure or a suspect part lot, the affected serial range is identifiable.
5. Traceability across the lifecycle (why identity matters downstream)
Record attribution (ALCOA+/GDP) | the cold-chain record the device produces is only admissible if it is attributable to a specific, calibrated, authentic unit — the identity + calibration birth-record binding is what makes the evidentiary record defensible (HA-RES-016/020; RCR-033/055/057).
Recall / CAPA scoping | serial ↔ part-lot ↔ build-record binding lets a supplier-lot escape or a field failure be scoped to an exact serial range (folder 17 closed loop, CAPA-###).
Service life (D1) | the same identity persists across trips, recalibrations (CAL-004), FRU replacements, and firmware upgrades — service records append to the unit history; decommissioning revokes identity + keys (DO-SVC-009) and de-registers from the fleet.
6. Traceability
Backward | DO-MFG-009, DO-CFG-010, DO-LBL-002/003/004/005, DO-SEC-007/008, RCR-033/055/057, HA-RES-016/ 020, PV-MFG-005; MDT-400 CC-07/CC-08/SC-06.
Forward | MDT-1000 (serialization/traceability readiness), folder 16 (identity/UDI reconciliation + DOL-1800 gate feed the release baseline — UDI regime RA-10 is a release condition), folder 17 (DHR is the recall/CAPA forensic spine; serial-level field records key to it), ORCH-018 / EM-ENG-012 (Stage 18 — the end-to-end identity thread).
Open | UDI regime + regulatory-markings set (RA-10, DOL-1800); label material qualification (DO-LBL-001).

## 16 Production Release
Production Release
The formal go / no-go decision and the exact version of the product it applies to.
↩ Contents
In this section:
• Production Release — Index — Smart Cold Chain Monitor (SCM)
• REL-100 — Release Readiness Assessment — Smart Cold Chain Monitor (SCM)
• REL-200 — Configuration / Release Baseline — Smart Cold Chain Monitor (SCM)
• REL-300 — Design Review Records & Release Gates — Smart Cold Chain Monitor (SCM)
• REL-400 — Production-Release Decision, Authority & Conditions — Smart Cold Chain Monitor (SCM)
Production Release — Index — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 16 — Production Release · Owner: Chief Systems Engineer (EM-ENG-001) chairing the Design Review Board (DRB), co-approved by Quality/Regulatory (EM-ENG-006). This stage assesses whether the SCM may be released to production and records the decision — the release readiness assessment (RA-01…12 + traceability release-readiness view), the frozen configuration/release baseline, the phase-gate design-review records, and the formal release decision with authority and conditions. Per the Design-Output §Release Assessment and ISO 13485 §7.3.5/§7.3.8 + FDA QMSR.
Position in the chain: consumes Stage 15 (MDT-1000 transfer readiness + reconciliations + open-risk disposition), Stage 10 (Risk-Management Summary + HAL-2800/2900 residual risk + acceptance authority), Stages 13–14 (V&V status), and Stage 9 (DOL register / DOL-1800 regulatory). Feeds Stage 17 (Field Monitoring) — the post-market surveillance commitment + carried-Open risks — and ORCH-018 / EM-ENG-012 (Stage 18).
---
1. Module files
Module | File | Scope
REL-100 | [REL-100-release-readiness-assessment.md](REL-100-release-readiness-assessment.md) | Release readiness (RA-01…12); requirement/design/risk baselines; V&V/validation/manufacturing/supplier/regulatory/cybersecurity readiness; open defects/deviations/risks; gap prioritization (P0…P4); readiness conclusion
REL-200 | [REL-200-configuration-release-baseline.md](REL-200-configuration-release-baseline.md) | Configuration/release baseline — HW rev, FW/SW/cloud/mobile versions, BOM/AML rev, regional configs, compatibility matrix, freeze/change rules
REL-300 | [REL-300-design-review-records.md](REL-300-design-review-records.md) | Design-review gate records (DR-###): concept→requirements→architecture→design-output→risk→V&V→transfer→release gates; open actions
REL-400 | [REL-400-release-decision.md](REL-400-release-decision.md) | Production-release decision + authority + residual-risk acceptance + explicit blockers + conditions of release + claim restrictions
2. The release decision (headline)
CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED (release pending V&V + blocker closure). The SCM's
design and manufacturing-process design are mature and every safety/measurement control is preserved into
production, but release is not authorized because V&V is unexecuted, the AML is not frozen, and the
regulatory classification/WHO PQS target/UDI are unresolved. Pilot/V&V builds are authorized; volume is not.
Three P0 blockers: B-1 V&V not executed (RA-05/06/08); B-2 AML not frozen — chemistry + antenna (RA-07); B-3 regulatory/WHO-PQS/UDI unresolved (RA-10). Eight conditions of release carried from the Risk-Management Summary §7 / HAL-2900 §6 (two pathway choke points; S5 UN 38.3; D3 + RF-in-container; fleet common causes; independent security; WHO PQS target; post-market surveillance).
3. Evidence & honesty statement
Never "Ready" without evidence | V&V is unexecuted (Stages 13–14 are authored playbooks; every PV-/VAL- is Planned/Evidence-Pending), so the decision is explicitly Not Ready → Conditionally Ready, never "Ready."
No invented results | no pass rate, field rate, capability index, or performance number is asserted; no sign-off is recorded that has not been given (all pending V&V).
No baseline frozen by assertion | all revisions/versions are REL-TBD until the AML freezes and the software tiers release; the baseline is frozen only at the DR-REL gate.
Claim discipline | no air-transport claim without UN 38.3; no WHO PQS claim until target confirmed + met; no real-time-alert claim on lanes where PRD-VER-005 shows the link cannot close.
4. Feeds
Stage 17 (Field Monitoring) — condition 8 (post-market surveillance for both harm directions + thermal-event + security) + the carried-Open risks (D3, RF-in-container) + the S5 disclosure/monitoring obligations become the standing post-market commitments. Stage 18 (ORCH-018 / EM-ENG-012) — the release segment of the end-to-end traceability spine (REL-/DR- ↔ RA- ↔ DO-XFER- ↔ HAL residuals ↔ PV-/VAL-).
REL-100 — Release Readiness Assessment — Smart Cold Chain Monitor (SCM)
Stage 16 · Module REL-100 · Owner: Chief Systems Engineer (EM-ENG-001) + Quality/Regulatory (EM-ENG-006), with all discipline owners attesting their domain. The production-release readiness assessment: the consolidated, evidence-based judgement of whether the SCM may be released to production, structured per the Design-Output §Release Assessment (RA-01…RA-12) and the traceability release-readiness view (forward/backward + gap prioritization). It assesses the requirement/design/risk baselines, verification and validation status, manufacturing/supplier/regulatory/cybersecurity readiness, and the open defects/deviations/risks — then hands the residual-risk approval and the release decision to REL-400.
Evidence honesty (per evidence-standard.md / hallucination-policy.md): V&V is not executed (Stages 13–14 are authored playbooks; every PV-/VAL- is Planned/Evidence-Pending). No field rate, capability index, pass rate, or performance number is asserted. The assessment therefore cannot conclude "Ready"; it concludes Not Ready → Conditionally Ready with explicit blockers, exactly as the Risk-Management Summary requires. Release-readiness conclusion values follow the traceability standard: Not Ready / Conditionally Ready / Ready-with-conditions.
---
1. Release-assessment framework (RA-01…RA-12)
The Master Specification §9 release checklist (carried through DO-XFER-010). Each RA is assessed as Met / Partial / Not Met / Blocked, with the evidence and the gap.
RA | Criterion | Status | Evidence / gap
RA-01 | Requirements baseline complete & approved | Met | Stages 1–7 requirements baselined; PRD-SAFE re-baselined (Stage 12)
RA-02 | Design outputs complete, reviewed, released | Partial | DOL register authored; Critical outputs still Draft/TBD-rev (AML-gated)
RA-03 | Design-input↔output traceability closed | Partial | Design-Input-Output matrix authored; closes at FCA (needs V&V)
RA-04 | Risk file complete (hazard analysis, FMEA) | Met (analysis) | HAL-000…2900 + ERFKL complete; residuals projected, not demonstrated
RA-05 ★ | Open-risk disposition / residual acceptable / V&V complete | Blocked | V&V not executed; DO-XFER-008 Preliminary; the top safety blocker
RA-06 | Verification complete (all PV- pass) | Not Met | PVL authored; all PV- Planned/Evidence-Pending
RA-07 ★ | Configuration/AML frozen (chemistry + antenna) | Blocked | Low-temp cell (CMP-006) + antenna (CMP-028/029) unselected
RA-08 | Validation complete (VAL- pass) | Not Met | PVaL authored; all VAL- Planned; summative usability not run
RA-09 | Manufacturing/test line ready (limits, MSA, capability) | Not Met | Numeric limits + MSA (DO-TST-014, MDT-800) EP; pilot not run
RA-10 ★ | Regulatory readiness (classification, markings, UDI, certs) | Blocked | WHO PQS category/US class + UDI regime + markings (DOL-1800) EP
RA-11 | Cybersecurity readiness (independent verification) | Not Met | Security controls designed; independent pen-test (PV-CYBER-) not run
RA-12 | Downstream ready (packaging/labeling/service/training) | Partial | DOL-1600/1700 drafted; labels gated by RA-10; training to deliver
Zero RA is "Ready" by assertion — each Partial/Not-Met/Blocked carries a named closure. Three RA are hard-Blocked (RA-05, RA-07, RA-10); these are the release blockers.
2. Baseline readiness (requirement / design / risk)
Requirement baseline (RA-01) | Market → User Needs → Product → System → Subsystem/SW requirements baselined (Stages 1–7); PRD-SAFE re-baselined from the risk-control requirements (Stage 12). Ready.
Design baseline (RA-02/03) | Design Output Register + Design-Input-Output Traceability Matrix authored (Stage 9); but the Critical outputs remain Draft/TBD-rev because the AML cannot freeze (RA-07). The configuration baseline (REL-200) is not frozen. Not ready.
Risk baseline (RA-04) | the hazard analysis (HAL-000…2900), ERFKL FMEA family, and RCR register are complete; the residual-risk profile is projected, Evidence-Pending — no individual residual is in region R as projected, but nothing is demonstrated until V&V. Analysis ready; acceptance pending V&V (RA-05).
3. Verification status (RA-06)
The Verification Library (Stage 13, PVL) is a complete, safety-driven plan with fault-injection for every safety control and independent adversarial verification for every security control — but no PV- has been executed. The release-load-bearing verifications and their status:
Verification | Closes | Status
PV-ENRG-001 (cold-charge FI) + UN 38.3 | D3 charge safety / S5 fire | Evidence-Pending
PRD-VER-004 (90-day −30 °C cold endurance) | D3 endurance (O-2, carried Open) | Evidence-Pending
PRD-VER-005 (RF-in-container link margin) | connectivity (O-1, carried Open) | Evidence-Pending
PV-SW-002 / PV-DATA-004 (honest verdict / gap) | pathway choke point #1 | Evidence-Pending
PV-TFC-002 + UN 38.3 (aircraft-fire containment) | S5 residual | Evidence-Pending
PV-MFG-001 (EOL calibration coverage) | fleet calibration common cause | Evidence-Pending
PV-CYBER-001/002/005/006 (independent pen-test) | security→safety | Evidence-Pending
Verification: Not Met. Two carried-Open unknowns (PRD-VER-004 / -005) are release conditions.
4. Validation status (RA-08)
The Validation Library (Stage 14, PVaL) is an authored playbook — no user has been tested, no product built. The two most load-bearing summative studies (the convergence choke points) are:
• VAL-USE-001 (read verdict → correct disposition) and VAL-USE-004 (never trust a void/unmonitored trip), consolidated in VAL-HF-001 (summative human-factors) — Planned.
• VAL-RISK-003 validates the HAL-2900 overall benefit-risk against its seven conditions — Planned.
Validation: Not Met. The last-mile low-training/vulnerable-user residuals (HA-RES-014) are usability-gated and closable only by summative validation.
5. Manufacturing & supplier readiness (RA-07 / RA-09 / RA-12)
Per MDT-1000 (design-transfer package & readiness):
• Process design transfer-mature — flow, PFMEA, Control Plan, CC/SC, WIs, production test spec, supplier controls, calibration/MSA, serialization/UDI all authored and internally consistent; every safety/ measurement control preserved into production via a zero-defect Control-Plan line.
• AML NOT frozen (RA-07) — cell chemistry + antenna unselected → T1 supplier qualification, battery WIs, and BOM reconciliation blocked.
• Test line NOT ready (RA-09) — numeric limits + MSA (DO-TST-014, MDT-800), leak method/limit (DO-TST-011), calibration reference chain + accuracy target (DO-TST-006 / WHO PQS) Evidence-Pending; pilot build not run; O/D estimates not proven.
• Downstream partial (RA-12) — packaging/protective design drafted; labels gated by RA-10; service/ cleaning/recal intervals EP; training to be delivered.
Manufacturing: Not Ready for volume; can build pilot/V&V units.
6. Regulatory readiness (RA-10) — D4 multi-market
The multi-market regime set (D4) and its status:
Regime | For | Status
WHO PQS (cold-chain monitor performance) | benefit basis / accuracy target | Open — exact category/spec revision + accuracy target unconfirmed (O-7)
EU GDP / ALCOA+ | data-integrity of the record | Designed-in (RCR-042/057); audit-readiness pending
US FDA | classification (accessory/information tool) | Open — classification unconfirmed (D4)
UN 38.3 / IATA DGR | lithium-battery transport | Open — depends on final chemistry/pack; test not run
FCC Part 15 / CE RED | BLE/Wi-Fi/LTE-M/GPS radios | Module pre-cert leveraged; product cert pending (antenna-gated, RA-07)
IEC 62368-1 / CISPR-IEC 61000 | electrical safety / EMC | Designed-to; verification (PV-EMC/SAFE) pending
IEC 60529 (IP) / IEC 60068 | ingress / environmental | Designed-to; PV-ENV/leak pending
Regulatory: Blocked (RA-10) — WHO PQS target + US classification + UDI regime + markings set (DOL-1800) must resolve. Claim discipline: no air-transport claim without UN 38.3; no WHO PQS conformance claim until the target is confirmed + met; no real-time-alert claim on lanes where PRD-VER-005 shows the link cannot close (default to auditable-record + on-receipt-verdict benefit).
7. Cybersecurity readiness (RA-11)
Device-anchored authenticity, non-suppressible alarm, secure boot + debug-lock, secure OTA, unique identity, and secure storage are designed and process-preserved (CC-06/CC-07; RCR-034/035/037/038/055). But the security→safety residuals (HA-RES-010/011/012/013) require independent adversarial verification before their projected residual can be accepted — not run. Cybersecurity: Not Met (pending independent pen-test PV-CYBER-001/002/005/006/009 + cloud DR test).
8. Open defects / deviations / open risks
Open defects | none from executed test (no test run); defect-management process (PV-DEF-1700) established and ready to receive pilot/V&V defects.
Transfer deviations | to be logged at pilot (DO-XFER-009); none dispositioned yet.
Carried-Open risks (must remain visible at release)
• O-1 / HA-RES-009 — RF-in-container link margin + antenna decision → PRD-VER-005 (Open).
• O-2 / HA-RES-004/005 — D3 −30 °C / 90-day battery endurance + cold-charge safety + fade → PRD-VER-004 + PV-ENRG-001 (Open).
• HA-RES-006 — S5 aircraft-hold fire → region-A-High permanently; UN 38.3 + PV-TFC-002 + mandatory disclosure + post-market thermal-event monitoring.
• HA-RES-015 — fleet calibration common cause → PV-MFG-001 + standing field monitoring.
9. Gap prioritization (traceability release-readiness view)
Per the traceability standard's P0…P4 prioritization:
Priority | Gap | RA
P0 (release blocker) | V&V not executed / open-risk not dispositioned | RA-05/06/08
P0 | AML not frozen (chemistry + antenna) | RA-07
P0 | Regulatory classification / WHO PQS target / UDI unresolved | RA-10
P1 (safety review) | Independent security verification not run | RA-11
P1 | D3 cold endurance + RF-in-container Open unknowns | RA-05
P2 (close before release) | Numeric limits + MSA + leak/cal methods; pilot build | RA-09
P2 | Downstream labels/training/service intervals | RA-12
10. Readiness conclusion (hands to REL-400)
Release-readiness conclusion: NOT READY now → CONDITIONALLY READY on closure of the P0/P1 conditions.
The SCM's requirements, design, risk analysis, and manufacturing process design are mature and internally
consistent, and every safety/measurement control is preserved into production. But release is not
authorized because V&V is unexecuted (RA-05/06/08), the AML is not frozen (RA-07), and the regulatory
classification/WHO PQS target/UDI are unresolved (RA-10) — three hard blockers — with independent security
verification (RA-11) and the two carried-Open unknowns (D3, RF-in-container) as critical conditions. This
mirrors the Risk-Management Summary's "Conditionally Ready — release pending V&V." The formal decision,
authority, and conditions of release are recorded in REL-400; the frozen baseline is REL-200; the
gate reviews are REL-300.
11. Traceability
Backward | MDT-1000 (transfer readiness), the DOL register + Design-Input-Output matrix (Stage 9), HAL / Risk-Management Summary / HAL-2800/2900 (Stage 10), RCR register (Stage 12), PVL (Stage 13), PVaL (Stage 14), DOL-1800 (regulatory), and the RA-01…12 checklist (DO-XFER-010).
Forward | REL-200 (config baseline), REL-300 (design reviews / gates), REL-400 (release decision + authority + conditions), Stage 17 (the disclosure/monitoring obligations + carried-Open risks become post-market commitments), ORCH-018 / EM-ENG-012 (Stage 18).
REL-200 — Configuration / Release Baseline — Smart Cold Chain Monitor (SCM)
Stage 16 · Module REL-200 · Owner: Chief Systems Engineer (EM-ENG-001) with Configuration Management + Quality/Regulatory (EM-ENG-006). Defines the production-release configuration baseline — the exact, frozen set of hardware revision, firmware/software/cloud/mobile versions, BOM/AML revision, and per-region configurations that constitute "the approved product" at release. Realizes DO-XFER-010 (production-release baseline) and DO-CFG-005 (HW/FW/cloud/mobile compatibility matrix), per ISO 10007 (configuration management) and ISO 13485 §7.5.8 (identification/traceability).
Evidence honesty: the baseline is NOT FROZEN. Per DOL-100 and MDT-1000, the AML cannot freeze until the low-temp cell chemistry (CMP-006) and antenna branch (CMP-028/029) are selected (RA-07), and the firmware/cloud/mobile versions are not yet at release. No specific revision or version number is asserted — placeholders (REL-TBD) are used until the baseline is frozen at the release gate. This module defines the structure of the baseline and the freeze rules, not invented versions.
---
1. Baseline structure (what a frozen baseline records)
At release, DO-XFER-010 records every element at a named revision. Until then, each is REL-TBD (Evidence Pending) with its blocking dependency:
Element | Config item | Baseline rev | Blocked by | Source
Product family | SCM-100 (base) | REL-TBD | — | DO-CFG-001
SKU: North America | SCM-100-NA | REL-TBD | band plan / cert | DO-CFG-002/006
SKU: Europe | SCM-100-EU | REL-TBD | CE RED / GDP | DO-CFG-002/006
SKU: Rest-of-world | SCM-100-ROW | REL-TBD | region band plans | DO-CFG-002/006
SKU: External-antenna | SCM-100-XA | REL-TBD | antenna decision (RA-07/PRD-VER-005) | DO-CFG-002/003
Hardware rev | PCBA + enclosure rev | REL-TBD | chemistry/antenna (RA-07) | DO-EE/PCB/MECH
Battery/pack | Low-temp cell + pack rev | REL-TBD | chemistry selection (RA-07/O-2) | DO-EE-003/004
Firmware | Device FW version + signed hash | REL-TBD | FW release + secure-boot signing | DO-FW-019/020
Cloud platform | Cloud API + platform version | REL-TBD | cloud release + DR test | DO-CLD-016
Mobile app | Provisioning/BLE app version | REL-TBD | app release | DO-APP-009
BOM / AML | mBOM + AML rev | REL-TBD | AML freeze (RA-07) | DO-MFG-001 / DO-CFG-004
Regional config bundles | Signed per-region config | REL-TBD | band plans / data-residency | DO-CFG-006
Cal/limits dataset | Test limits + cal points | REL-TBD | DO-TST-014 / WHO PQS (O-7) | DO-TST-006/014
2. Compatibility matrix (DO-CFG-005) — the interoperability contract
The release baseline is not just a list of versions; it is the supported (HW rev × FW version × cloud API × app version) combination set, and the minimum-compatible matrix across OTA version skew. Rules:
• Incompatible combinations are blocked, not silently degraded (RCR — OTA never-brick, DO-CFG-005 Critical).
• OTA campaigns move the fleet only within the supported matrix, staged/canary with halt-on-regression (RCR-037; HA-RES-012).
• The cloud/app must remain backward-compatible with the released device FW across the field fleet's version distribution (a fielded reusable multi-trip fleet will span versions).
Status: matrix structure defined; version numbers Evidence Pending until each tier releases.
3. Regional configuration baselines (D4 multi-market)
Each region SKU is defined by the delta from base (DO-CFG-002/006): band plan + permitted-band enforcement, carrier/APN strategy, data-residency region (EU GDP), regulatory markings set (DOL-1800), and language/IFU set. The regional config is a signed config bundle provisioned at PROC-090 and verified at EOL (PV-MFG-006 / CC-06/SC-07). Wrong-region config is a Critical Characteristic (CC-06) with line-clearance protection (SC-07).
Status: region structure defined; band plans / carrier strategy Evidence Pending (INT-027); markings set Evidence Pending (DOL-1800, RA-10).
4. Freeze rules & change control
Freeze prerequisite | the baseline may be frozen only when the AML freezes (RA-07), the software tiers release, the compliance-critical register (DO-REG-011) is complete (RA-10), and the FCA/PCA (DO-XFER-003) passes.
Freeze authority | the Design Review Board (DRB) + Quality/Regulatory (EM-ENG-006), recorded at DR-DO / the transfer gate (REL-300).
Change control after freeze | any change to a baselined element goes through ECO (DO-GOV-004) with change classification (DO-GOV-005); changes to Critical outputs / Critical Characteristics require re-assessment of the affected risk controls and re-verification, and — for safety/regulatory changes — DRB + EM-ENG-006 approval and HAL reconciliation. Field/fleet impact of any post-release change is assessed against the compatibility matrix and the OTA rollout discipline.
Configuration audit | at freeze, the FCA confirms every requirement is verified against a released output, and the PCA confirms every output/document is present at its baselined revision (DO-XFER-003).
5. Baseline conclusion
Configuration/release baseline: NOT FROZEN. The baseline structure, compatibility-matrix contract,
regional-config definitions, and freeze/change rules are defined, but the concrete revisions and versions are
REL-TBD pending the AML freeze (RA-07), the software-tier releases, the limits/cal dataset (RA-09), and the
regulatory/markings resolution (RA-10). The baseline is frozen only at the release gate (REL-400) once those
close and the FCA/PCA passes. No version or revision is asserted as released that is not.
6. Traceability
Backward | DO-CFG-001…012 (DOL-100), DO-XFER-010 (DOL-2000), the software-tier release outputs (DO-FW-019/DO-CLD-016/DO-APP-009), DO-REG-011 (compliance-critical register), DO-TST-014 (limits), MDT-1000 (reconciliations), REL-100 (readiness).
Forward | REL-300 (the DR-DO / transfer gate that authorizes the freeze), REL-400 (the release decision freezes and names the baseline), Stage 17 (the fielded baseline + compatibility matrix govern OTA campaigns and field-version monitoring), ORCH-018 / EM-ENG-012 (Stage 18 — the configuration segment of the lifecycle spine).
REL-300 — Design Review Records & Release Gates — Smart Cold Chain Monitor (SCM)
Stage 16 · Module REL-300 · Owner: Chief Systems Engineer (EM-ENG-001) chairing the Design Review Board (DRB), with Quality/Regulatory (EM-ENG-006) co-approving. Records the phase-gate design reviews (DR-###) across the SCM lifecycle — their scope, entry/exit criteria, findings, actions, and disposition — culminating in the design-transfer gate (DR-DO) and the production-release gate that authorizes REL-400. Per ISO 13485 §7.3.5 (design & development review) and the Master Specification §9 gate structure.
Evidence honesty: the downstream gates (DR-DO transfer, DR-REL release) are not yet held — their entry criteria (executed V&V, frozen AML, resolved regulatory) are open. This module records the gate structure and the status of each; it does not record a passed gate that has not occurred. No sign-off is asserted that has not been given.
---
1. Design-review gate structure
Gate | Name | Purpose | Status
DR-CN | Concept review | D1–D4 locked; intended use; top risks | Held / Passed (concept-and-decisions.md)
DR-RQ | Requirements review | Market/User/Product/System requirements baselined | Held / Passed (Stages 1–7)
DR-AR | Architecture review | System architecture + power budget + ADR | Held / Passed (Stage 8)
DR-DS | Design-output review | Design outputs adequate + traceable to inputs | Partial — outputs Draft; Critical outputs TBD-rev
DR-RM | Risk-management review | HAL + ERFKL + RCR + residual-risk (projected) | Held / Passed (analysis); residuals projected
DR-VV | V&V readiness review | Verification + validation plans adequate | Held / Passed (plans); execution pending
DR-DO | Design-transfer gate | Transfer package complete; process ready; pilot | Not held — blocked (RA-05/07/09/10)
DR-REL | Production-release gate | RA-01…12 met; residual risk accepted; authorize | Not held — blocked (authorizes REL-400)
2. Gate records
DR-CN — Concept review
Scope | intended use, D1 (reusable/rechargeable), D2 (store-and-forward + immediate excursion alert), D3 (low-temp Li, charge-inhibit <0 °C, aggressive duty-cycle), D4 (WHO PQS + EU GDP + US + air-transport).
Outcome | Concept locked; the four decisions resolved with residual tensions carried as top risks (D3 cold battery, RF-in-container).
Actions carried | confirm WHO PQS category + US classification (D4); select low-temp chemistry; decide antenna strategy. Passed (concept baseline).
DR-RQ / DR-AR — Requirements & architecture reviews
Scope | the requirement chain (MR→UN→PRD→SYS→SSR/HWR/SWR) and the system architecture + power budget + architecture decision record.
Outcome | requirements baselined; architecture adequate for the offline-first, duty-cycled, multi-radio design; power budget confirms store-and-forward necessity (D2). Passed.
Actions carried | D3 endurance + RF-in-container to be closed by test (PRD-VER-004/005).
DR-DS — Design-output review
Scope | the Design Output Register (DOL-000…2000) + Design-Input-Output Traceability Matrix; adequacy of each output; identification of Critical outputs.
Outcome | Partial — the register is authored and traceable, but the Critical outputs remain Draft/TBD-rev because the AML cannot freeze (RA-07); the design outputs are adequate in content but not Released.
Findings | three Critical manufacturing outputs (DO-MFG-004 seal, DO-MFG-007 battery handling, DO-MFG-008 programming) + five Critical test outputs (DO-TST-004/005/006/009/011) + DO-CFG-004/005 must be Released for transfer. Action: freeze AML (RA-07); release Critical outputs. Open.
DR-RM — Risk-management review
Scope | HAL-000…2900, ERFKL FMEA family, RCR register, individual + overall residual-risk (HAL-2800/2900), Risk-Management Summary.
Outcome | Passed (analysis) — no individual residual in region R as projected; overall benefit-risk positive-but-conditional; S5 aircraft-fire held region-A-High with mandatory disclosure + monitoring.
Findings / conditions carried to release | the eight Risk-Management-Summary §7 conditions (two pathway choke points; S5 UN 38.3; D3 + RF-in-container; fleet common causes; independent security; WHO PQS target; post-market surveillance). Residual acceptance is Preliminary — pending V&V.
DR-VV — V&V readiness review
Scope | the Verification Library (fault-injection for every safety control; independent adversarial for every security control) + the Validation Library (summative usability of the two choke points).
Outcome | Passed (plans) — the plans are adequate and traceable; execution pending. No PV-/VAL- results exist.
Action | execute Stage 13/14 on pilot/representative units.
DR-DO — Design-transfer gate (NOT HELD)
Entry criteria (must be met to hold) | transfer package complete (DO-XFER-001…009); four reconciliations closed (BOM/drawing/SW/mfg); FCA/PCA passed; production-test line ready (limits/MSA/golden units); suppliers/ tooling/fixtures qualified; pilot build passed with deviations dispositioned; open-risk disposition reconciled with HAL.
• Status: BLOCKED — AML not frozen (RA-07); limits/MSA + pilot not done (RA-09); regulatory unresolved (RA-10); V&V not executed (RA-05). Cannot be held until these close.
DR-REL — Production-release gate (NOT HELD)
Entry criteria | RA-01…RA-12 all Met; residual risk formally accepted by authority; configuration baseline frozen (REL-200); all Critical outputs Released; carried-Open risks dispositioned or claim-restricted.
• Status: BLOCKED — depends on DR-DO + executed V&V. This gate authorizes the REL-400 release decision.
3. Open design-review actions (carried to release)
Action | From | Owner | Closes
Freeze AML (chemistry + antenna) | DR-DS | EM-ENG-001/Purchasing | RA-07
Release all Critical design outputs | DR-DS | discipline owners | RA-02
Execute safety/security V&V (FI + independent pen-test) | DR-VV | Quality/Security | RA-05/06/11
Close D3 cold endurance (PRD-VER-004) + RF-in-container (PRD-VER-005) | DR-RM | Battery/RF | O-1/O-2
Confirm WHO PQS target + US classification + UDI (DOL-1800) | DR-CN/DR-RM | EM-ENG-006 | RA-10
Run pilot build + disposition deviations | DR-DO | Manufacturing | RA-09
Complete residual-risk disposition on executed V&V | DR-RM | Reliability/Safety | RA-05
4. Conclusion
The upstream gates (concept, requirements, architecture, risk-management analysis, V&V-plan) are held and
passed; the design-output gate is partial (outputs adequate but not Released, AML-gated); and the two
release-authorizing gates — the design-transfer gate (DR-DO) and the production-release gate (DR-REL)
— are not held and are blocked by RA-05/07/09/10. The DRB cannot authorize production until DR-DO and
DR-REL are held with their entry criteria met. This is the governance basis for the REL-400 decision of
Conditionally Ready — not released.
5. Traceability
Backward | the stage artefacts each gate reviews (concept, requirements, architecture, DOL register, HAL/ Risk-Management Summary, PVL/PVaL, MDT-1000), and the RA-01…12 checklist.
Forward | REL-400 (the release decision the DR-REL gate authorizes), REL-100/200 (readiness + baseline the gates assess/freeze), Stage 17 (post-market obligations carried from DR-RM), ORCH-018 / EM-ENG-012 (Stage 18).
REL-400 — Production-Release Decision, Authority & Conditions — Smart Cold Chain Monitor (SCM)
Stage 16 · Module REL-400 · Owner: Design Review Board (DRB), chaired by Chief Systems Engineer (EM-ENG-001), co-approved by Quality/Regulatory (EM-ENG-006), with residual-risk acceptance by the authority defined in HAL-2200 §7. The formal production-release decision record: the decision, the authority who may make it, the residual-risk acceptance, and the explicit conditions of release. This is the release-gate spine (DO-XFER-010) — production is authorized only when the readiness assessment (REL-100), the frozen baseline (REL-200), and the release gates (REL-300) are satisfied.
Evidence honesty: V&V is unexecuted and three RA are hard-blocked; therefore this record cannot and does not authorize production. Per the org rule that a release decision is never "Ready" without evidence, the decision is CONDITIONALLY READY — NOT RELEASED, with the blockers and conditions stated explicitly. No sign-off is recorded that has not been given (all sign-offs are pending V&V).
---
1. Release decision
## Decision: CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED (release pending V&V + blocker closure)
The Smart Cold Chain Monitor's requirements, architecture, design, risk analysis, and manufacturing process
design are mature, internally consistent, and traceable, and every safety- and measurement-critical control
is preserved into production via zero-defect Control-Plan lines. However, production release is NOT
authorized because verification and validation are unexecuted, the AML is not frozen, and the regulatory
classification/WHO PQS target/UDI are unresolved. The device is not released. Pilot/engineering builds
for V&V are authorized; volume production is not.
This decision mirrors the Risk-Management Summary §8 ("Conditionally Ready — release pending V&V") and the HAL-2900 overall benefit-risk conclusion (positive but conditional).
2. Release blockers (P0 — must close before ANY release)
# | Blocker | RA | Closes on
B-1 | V&V not executed / residual risk not demonstrated | RA-05/06/08 | Stages 13–14 pass + DO-XFER-008 final
B-2 | AML not frozen (low-temp cell chemistry + antenna branch) | RA-07 | Chemistry/antenna selection + PRD-VER-004/005
B-3 | Regulatory classification / WHO PQS accuracy target / US class / UDI regime | RA-10 | DOL-1800 determination (D4)
Each blocker is a P0 release blocker in the traceability gap prioritization; none may be closed by assertion.
3. Conditions of release (the eight, from Risk-Management Summary §7 / HAL-2900 §6)
Release becomes authorizable only when all are demonstrated (not projected):
1. All safety-control PV- pass with fault-injection (HAL-2600) — charge-inhibit, protection, seal, probe, power-fail-safe, etc.
2. The two pathway choke points are verified/validated: honest verdict / no-false-reassurance (PV-SW-002 + PV-DATA-004) and supported human decision (VAL-USE-004 + VAL-USE-001, summative VAL-HF-001).
3. The S5 aircraft-hold fire is verified contained: PV-TFC-002 + UN 38.3 / IATA-DGR, with mandatory transport disclosure + post-market thermal-event monitoring in place.
4. The two carried-Open unknowns close acceptably: D3 cold endurance (PRD-VER-004) and RF-in-container margin (PRD-VER-005), with the antenna decision fixed.
5. The fleet common causes are verified and monitored: EOL calibration coverage (PV-MFG-001) + depot-recal control; staged/canary OTA + halt-on-regression; cloud DR-tested.
6. Security controls pass independent adversarial verification: PV-CYBER-001/002/005/006/009.
7. The WHO PQS accuracy target (and US classification) is confirmed and met (D4).
8. Post-market surveillance is established for both harm directions (missed-excursion and false-reject/ wastage) plus thermal-event + security monitoring (ISO 14971 §9 loop) — the Stage 17 commitments.
4. Residual-risk acceptance (authority)
Per HAL-2200 §7, residual-risk acceptance authority:
S4 / S5 region-A residuals | EM-ENG-006 + EM-ENG-001 + documented benefit-risk.
The S5 aircraft-fire residual (HA-RES-006) | EM-ENG-001 + EM-ENG-006 sign-off + documented benefit-risk + mandatory post-market thermal-event monitoring + UN 38.3 / IATA-DGR transport disclosure. (Region-A-High; never broadly acceptable.)
Security→safety residuals | additionally require independent adversarial verification sign-off before acceptance.
The accepted public-health trade (false-reject/wastage, HA-RES-019) | EM-ENG-001 + EM-ENG-006 + both-directions-ALARP benefit-risk; monitored as a first-class field metric.
All residual-risk sign-offs are PENDING V&V — this record is the pre-V&V baseline; no acceptance is recorded as given.
5. Release authority & sign-off (blank until conditions met)
Role | Authority | Sign-off | Date
Chair, Design Review Board | EM-ENG-001 (Chief Systems) | Pending — conditions §3 not met | —
Quality / Regulatory | EM-ENG-006 | Pending | —
System-Safety / Reliability | (residual-risk acceptance) | Pending — V&V not executed | —
Product-Security | (independent security verification) | Pending — pen-test not run | —
Manufacturing | (line/pilot readiness) | Pending — AML/pilot open | —
No signature is applied because no condition is met. Production is not authorized.
6. Claim restrictions (if released conditionally / partially)
Consistent with the Risk-Management Summary and HAL-2900, where a condition cannot be met the corresponding claim is withheld or restricted, not asserted:
• No air-transport claim without UN 38.3 / IATA-DGR qualification (condition 3).
• No WHO PQS conformance claim until the target is confirmed and met (condition 7).
• No real-time-alert claim on lanes where PRD-VER-005 shows the in-container link cannot close — default to the auditable-record + on-receipt-verdict benefit (condition 4).
• No security-hardening claim beyond what independent verification has demonstrated (condition 6).
7. Path to release (what must happen next)
1. Freeze AML (chemistry + antenna) → B-2 / RA-07 → release Critical design outputs.
2. Build pilot units → close limits/MSA/leak/cal methods (RA-09) → hold DR-DO transfer gate.
3. Execute Stage 13/14 V&V (incl. fault-injection + independent pen-test + summative usability) → close B-1 / conditions 1–6.
4. Confirm WHO PQS target + US class + UDI (DOL-1800) → B-3 / RA-10 / condition 7.
5. Establish Stage 17 post-market surveillance (condition 8).
6. Complete residual-risk disposition on executed V&V (DO-XFER-008) → freeze baseline (REL-200) → hold DR-REL → record the release decision + authority sign-offs here.
8. Traceability
Backward | REL-100 (readiness), REL-200 (baseline), REL-300 (gates), MDT-1000 (transfer readiness), the Risk-Management Summary + HAL-2800/2900 (residual risk + benefit-risk + acceptance authority), the PVL/PVaL (V&V), DOL-1800 (regulatory), DO-XFER-008/010.
Forward | Stage 17 (Field Monitoring) — condition 8 (post-market surveillance for both harm directions + thermal-event + security) + the carried-Open risks + the S5 disclosure/monitoring obligations become the standing post-market commitments; ORCH-018 / EM-ENG-012 (Stage 18 — the release segment of the lifecycle spine, and the "release decision → field monitoring" link).

## 17 Field Monitoring
Field Monitoring
How real-world performance is watched after launch, and how lessons feed back into the design.
↩ Contents
In this section:
• Field Monitoring & Post-Market Surveillance — Index — Smart Cold Chain Monitor (SCM)
• FM-100 — Post-Market Surveillance Plan — Smart Cold Chain Monitor (SCM)
• FM-200 — Field-Monitoring Metrics & Thresholds — Smart Cold Chain Monitor (SCM)
• FM-300 — Closed-Loop Field-to-Design Process (RCA → CAPA → Update) — Smart Cold Chain Monitor (SCM)
• FM-400 — Complaint & Adverse-Event Handling — Smart Cold Chain Monitor (SCM)
• FM-500 — Residual-Risk Field-Monitoring Commitments — Smart Cold Chain Monitor (SCM)
• FM-600 — Reliability Growth & Periodic Risk-Review Cadence — Smart Cold Chain Monitor (SCM)
Field Monitoring & Post-Market Surveillance — Index — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 17 — Field Monitoring & Post-Market Surveillance · Owner: Quality/Regulatory (EM-ENG-006) + Post-Market Engineer with System-Safety, Reliability, Product- Security, and Chief Systems (EM-ENG-001). This stage authors the SCM's standing post-market process: the PMS plan, the field-monitoring metrics & thresholds, the closed-loop field-to-design process, the complaint/ adverse-event handling process, the specific residual-risk field-monitoring commitments, and the reliability- growth + periodic risk-review cadence. Its job is to keep the SCM's conditional benefit-risk conclusion (HAL-2900) valid across the reusable fleet's service life (ISO 14971 §9) and to discharge the eighth condition of release (REL-400 §3.8).
Position in the chain: consumes Stage 10 (HAL-2800/2900 disclosure/monitoring obligations + Risk-Management Summary §7.8), Stage 16 (REL-400 condition 8 + carried-Open risks + S5 obligation), Stage 15 (MDT-900 Device History Record for serial attribution/recall scoping; MDT-800 calibration common cause), Stage 11 (FAL-2000/2100 RCA/CAPA readiness; FAL-1900 reliability), and Stage 14 (both-direction field validation). Feeds ORCH-018 / EM-ENG-012 (Stage 18) — field monitoring closes the lifecycle traceability spine (design → control → release evidence → field monitoring → CAPA → re-baseline).
---
1. Module files
Module | File | Scope
FM-100 | [FM-100-post-market-surveillance-plan.md](FM-100-post-market-surveillance-plan.md) | PMS plan — both harm directions + thermal + security; data sources (→ FF-###), cadence, roles, regulatory/public-health posture
FM-200 | [FM-200-metrics-thresholds.md](FM-200-metrics-thresholds.md) | Field-monitoring metrics & thresholds (TREND-001…019): fleet reliability, battery aging, excursion-detection accuracy, alert-delivery, return/warranty, complaint categories, cybersecurity events
FM-300 | [FM-300-closed-loop-process.md](FM-300-closed-loop-process.md) | Closed loop: field record → RCA → CAPA → requirement/hazard/FMEA/verification update → re-baseline
FM-400 | [FM-400-complaint-adverse-event-handling.md](FM-400-complaint-adverse-event-handling.md) | Complaint/adverse-event handling (CMP-###) with vaccine-integrity / public-health reporting considerations
FM-500 | [FM-500-residual-risk-field-commitments.md](FM-500-residual-risk-field-commitments.md) | The specific residual risks that require field monitoring, from HAL-2800/2900, mapped to metrics + closing verification
FM-600 | [FM-600-reliability-growth-periodic-review.md](FM-600-reliability-growth-periodic-review.md) | Reliability-growth program + periodic risk-review cadence (benefit-risk revalidation, ISO 14971 §9)
2. ID scheme
• FF-### — field records (normalized field information, serial-attributed via the Device History Record).
• CMP-### — complaints / adverse-event reports.
• CAPA-### — corrective/preventive actions.
• TREND-### — monitored field metrics/trends (FM-200).
3. Why this stage is first-class (not optional paperwork)
The SCM was released Conditionally with a positive-but-conditional benefit-risk conclusion. HAL-2900 identifies two dominant common causes — calibration drift (TREND-010) and OTA/cloud rollout (TREND-018/019) — whose containment is operational, not a one-time design fact, and a region-A-High S5 residual (aircraft- hold battery fire, TREND-004) that can never be broadly acceptable and carries mandatory post-market thermal- event monitoring. The device is operated at its most safety-critical point by its least-trained user. And it has two harm directions — missed excursion and false-reject/wastage — both monitored as first-class (TREND-007/008). Post-market surveillance is therefore the mechanism that keeps the release valid; it is a release condition, not an afterthought.
4. Evidence & honesty statement
Plan, not report | the fleet is not yet fielded (Conditionally Ready — not released); no field rate, complaint count, reliability figure, or benefit-risk revision is asserted. Thresholds are safety-EVENT/ absolute where a target exists, else baseline-from-field then hold (SPC).
Severity-first monitoring | safety-significant metrics (S5 thermal, missed-excursion, calibration common cause, security) act on a single credible signal — they do not wait for a rate.
Both directions | a rising false-reject/wastage rate is as reportable as a missed excursion (HA-RES-019).
The honesty ledger closes here | the pre-release projected O/D/probability estimates are matured to field-demonstrated values through the reliability-growth + periodic-review cadence (FM-600).
5. Feeds
Stage 18 (ORCH-018 / EM-ENG-012) — field monitoring is the final segment of the end-to-end lifecycle spine: every design output → risk control → release evidence carries a field-monitoring status, and the closed loop's CAPA → re-baseline updates flow back to the HAL/FMEA/requirement artefacts, keeping the whole trace (and the benefit-risk conclusion) current across the SCM's reusable service life.
FM-100 — Post-Market Surveillance Plan — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-100 · Owner: Quality/Regulatory (EM-ENG-006) + Post-Market Engineer, with System-Safety, Reliability, Product-Security, and Chief Systems (EM-ENG-001). The Post-Market Surveillance (PMS) plan: the standing process by which the SCM's field performance is monitored across its reusable, multi-trip fleet so that the conditional benefit-risk conclusion (HAL-2900) is kept valid across the service life (ISO 14971 §9). This plan defines the surveillance scope, data sources, cadence, roles, and how field information feeds the closed loop (FM-300) — and it discharges the eighth condition of release (REL-400 §3.8): post-market surveillance for both harm directions plus thermal-event and security monitoring.
Evidence honesty: the fleet is not yet in the field (release is Conditionally Ready — not released). This is a plan, not a report — no field rate, complaint count, or reliability figure is asserted; all metric targets are set in FM-200 (many Evidence-Pending until baselined from early field data). Field records carry FF-###; complaints CMP-###; corrective/preventive actions CAPA-###; monitored trends TREND-###.
---
1. Why PMS is first-class for the SCM (not an afterthought)
The SCM's safety case is explicitly conditional and monitored (HAL-2900 §6/§7, Risk-Management Summary §7.8): the overall benefit-risk is positive only if the projected residuals hold in the field, and two of the dominant common causes — calibration drift and OTA/cloud rollout — are operational controls that cannot be closed by a one-time pre-release test. The device is also operated at its most safety-critical point (the disposition decision) by its least-trained user (last-mile health worker), and it carries a region-A-High S5 residual (aircraft-hold battery fire) that requires mandatory post-market thermal-event monitoring. PMS is therefore the mechanism that keeps the release valid — not paperwork.
2. Surveillance scope — both harm directions
The SCM has two harm directions, and PMS monitors both as first-class (a rising false-reject rate is as reportable as a missed excursion, HA-RES-019):
Direction | Harm | What is monitored
Missed / false-accept | Spoiled dose administered as good → beneficiary unprotected | Missed-excursion field signals, false-accept-over-gap disputes, probe-fault rates, calibration drift, alert-delivery failures
False-reject / over-conservatism | Potent vaccine needlessly discarded → stock-out → unvaccinated population | False-reject/wastage/stock-out rate, override rates, re-adjudication outcomes
Physical (S5) | Battery thermal event / aircraft-hold fire | Thermal-event reports, charge-inhibit field behavior, swelling/venting returns, transport incidents
Security | Falsified/suppressed data, compromised OTA/cloud | Integrity-failure events, alarm-suppression anomalies, OTA campaign health, cloud auth/integrity incidents, CVD reports
Availability / reliability | Loss of monitoring, dead-on-arrival, connectivity loss | Endurance/battery-aging, DOA rate, alert-latency per lane, return/warranty rate
3. Data sources (field information inputs → FF-###)
Source | Content | Owner
Fleet telemetry / cloud dashboard | Device health, battery SoH, charge-inhibit events, gap rates, alert-delivery success/backlog, version distribution, integrity events | Cloud + Post-Market
Complaints / support | User-reported issues (no-charge, charge-inhibited, no connectivity, misread verdict, false-reject disputes) | Support → Quality (CMP-###)
Returns / RMA / warranty | Returned-unit teardown, failure mode, swelling/venting, seal/leak, label legibility over reuse | Reliability + Service
Depot service records | Recalibration residuals (drift), FRU replacements, cleaning-cycle degradation, seal restores | Service + Metrology
Receipt-QA / customer disposition data | Excursion-detection accuracy vs. independent receipt checks; wastage data | Clinical/Operational + Quality
Security monitoring / CVD | Integrity/authenticity events, coordinated vulnerability disclosure reports | Product-Security (EM-ENG-008)
Transport / DG incidents | Any lithium-battery transport incident (S5) | Regulatory + Safety
Regulatory / vigilance | Regulator queries, WHO PQS/GDP audit findings, adverse-event reports | EM-ENG-006
All inputs are normalized into field records (FF-###) with device serial (via the Device History Record, MDT-900), trip context, and category — the entry point to the closed loop (FM-300).
4. Surveillance cadence
Activity | Cadence | Owner | Output
Fleet-metric dashboard review | Weekly (ramp) → monthly (steady) | Post-Market + Cloud | TREND-### review
Complaint triage & classification | Continuous / daily | Quality | CMP-###
Thermal-event / S5 monitoring | Continuous + immediate escalation | Safety + Regulatory | FF-### (priority)
Security-event monitoring | Continuous | Product-Security | FF-###
Calibration-drift trend review | Monthly + per depot-recal batch | Metrology | TREND-cal
Return/RMA teardown analysis | Per return + monthly rollup | Reliability | FF-###
Periodic risk review (benefit-risk revalidation) | Quarterly + on trigger | EM-ENG-006 + EM-ENG-001 | Risk-file update (FM-600)
PMS report (regulatory) | Per applicable regime (D4) | EM-ENG-006 | PMS report
Reliability-growth review | Per reliability cadence (FM-600) | Reliability | reliability update
5. Roles & governance
Post-Market Engineer / Quality (EM-ENG-006) | owns the PMS process, complaint handling, PMS reporting, and the periodic risk review; convenes CAPA.
System-Safety / Reliability | owns thermal-event/S5 monitoring, reliability-growth, and residual-risk revalidation.
Product-Security (EM-ENG-008) | owns security-event monitoring + coordinated vulnerability disclosure.
Cloud / Connectivity | own fleet-telemetry metrics + OTA campaign health.
Metrology / Service | own calibration-drift + depot recalibration data.
Chief Systems (EM-ENG-001) | integrates; co-owns benefit-risk revalidation; escalation authority.
Escalation | any S5 (thermal event) or fleet-scale common-cause signal (calibration bias, OTA regression) triggers immediate escalation to EM-ENG-001 + EM-ENG-006 and a CAPA, independent of the routine cadence.
6. Regulatory & public-health reporting posture (D4)
• ISO 14971 §9 — production and post-production information feeds back to keep the benefit-risk conclusion valid; the PMS plan is the §9 loop.
• EU GDP / ALCOA+ — data-integrity deviations and audit findings are surveilled; the device's evidentiary record supports GDP disposition.
• WHO PQS / WHO-UNICEF EVM — field accuracy/performance against the recognized cold-chain-monitor target (O-7, once confirmed) is monitored; performance shortfalls are reportable to the program.
• US FDA / applicable vigilance — adverse-event/vigilance reporting per the confirmed classification (D4); complaint handling feeds it (FM-400).
• UN 38.3 / IATA-DGR — lithium-battery transport incidents are surveilled and reported per DGR.
Public-health significance | because a monitoring failure can affect vaccine integrity at population scale, the complaint/adverse-event process (FM-400) carries specific vaccine-integrity/public-health escalation — a suspected fleet-wide detection failure is treated as a potential public-health event, not a routine complaint.
Exact reporting timelines/regimes are Evidence-Pending until the D4 classification resolves (RA-10); the plan is structured to attach to whichever regimes apply.
7. Traceability
Backward | HAL-2800/2900 (the disclosure/monitoring obligations per residual), the Risk-Management Summary §7.8, REL-400 (condition 8 + carried-Open risks + S5 obligation), MDT-900 (Device History Record — serial attribution), MDT-800 (calibration common cause), the PVaL both-direction field-validation (VAL-FIELD-900).
Forward | FM-200 (metrics + thresholds), FM-300 (closed loop), FM-400 (complaint/adverse-event handling), FM-500 (the specific residual-risk field commitments from HAL-2800/2900), FM-600 (reliability-growth + periodic risk review), and ORCH-018 / EM-ENG-012 (Stage 18) — field monitoring closes the lifecycle traceability spine (design → control → release evidence → field monitoring → status).
FM-200 — Field-Monitoring Metrics & Thresholds — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-200 · Owner: Post-Market Engineer + Quality (EM-ENG-006) with Reliability, Cloud, Metrology, Product-Security. The field-monitoring metric set: the specific quantities the SCM fleet is watched on, their data source, their trigger threshold, and the action a threshold breach initiates (→ FM-300 closed loop). Each metric is a monitored trend TREND-###. The set covers fleet reliability, battery aging, excursion-detection accuracy in the field, alert-delivery success, return/warranty rate, complaint categories, and cybersecurity events — the categories named in the task and the HAL-2800/2900 monitoring obligations.
Evidence honesty: the fleet is not fielded; no metric value is asserted. Thresholds are of two kinds: (a) absolute safety thresholds anchored to a released target where one exists (e.g., calibration accuracy to the WHO PQS target once confirmed — O-7), and (b) statistical/relative thresholds baselined from early field data then held (SPC-style: any statistically significant adverse shift triggers action). Where a numeric target is not yet set it is marked Evidence-Pending — baseline from field; no rate is invented.
---
1. Metric register
Legend: Dir = harm direction (MA = missed/accept; FR = false-reject; S5 = physical; SEC = security; AV = availability/reliability). Threshold type: ABS (absolute) / SPC (statistical baseline+shift) / EVENT (any occurrence escalates).
---
Fleet reliability & availability
TREND-001 — Dead-on-arrival (DOA) / not-ready-at-arming rate — AV
Source | provisioning/arming telemetry + support (FF/CMP).
Threshold | SPC — baseline then any significant upward shift. Target EP.
Action | > threshold → investigate (battery cold-charge inability, SoC-at-ship, assembly) → CAPA if systemic. Ties: HA-RES-004 (dead-in-cold), RCR-006/045.
TREND-002 — Field reliability / MTBF-in-use (loss-of-monitoring events per device-trip) — AV
Source | fleet telemetry (gaps, resets, faults) + returns.
Threshold | SPC vs. reliability-model prediction (FAL-1900). MTBF target EP — no figure asserted.
Action | shortfall vs. model → reliability-growth review (FM-600) + CAPA. Ties: FAL-1900, HA-RES-004.
TREND-003 — Return / warranty / RMA rate — AV/business
Source | RMA + warranty records.
Threshold | SPC baseline + business target (VAL-BIZ). Target EP.
Action | rising trend → teardown Pareto (FF) → CAPA. Ties: VAL-BIZ-800.
Battery aging & thermal (the S5 spine + D3)
TREND-004 ★ — Thermal-event / swelling / venting incidents — S5 (EVENT)
Source | returns teardown, complaints, transport-incident reports.
Threshold | EVENT — any occurrence escalates immediately to EM-ENG-001 + EM-ENG-006 + Safety.
Action | immediate containment assessment + recall scoping (via DHR serial range) + regulatory/DGR report + CAPA. This is the mandatory S5 post-market thermal-event monitoring (HA-RES-006). Ties: HA-TFC-002, RCR-002/004/008, UN 38.3.
TREND-005 ★ — Charge-inhibit field behavior (cold-charge inhibit engagements + any charge-below-limit) — S5
Source | device telemetry (inhibit events, cell-temp-at-charge).
Threshold | EVENT for any charge attempt proceeding below the qualified limit (should be impossible by HW interlock); SPC on inhibit-engagement frequency.
Action | any below-limit charge → immediate investigation (interlock integrity, sensor) + CAPA. Ties: HA-RES-005, RCR-001, PT-005.
TREND-006 — Battery state-of-health / capacity fade over reuse cycles — AV/D3
Source | device SoH telemetry + depot battery-health checks.
Threshold | SPC vs. the cold-endurance/fade model (D3, PRD-VER-004); fade beyond model → action. Fade curve EP — closes with PRD-VER-004 + field aging.
Action | faster-than-expected fade → PM-interval/recal-interval revision (DO-SVC-004/005) + reliability review. Ties: HA-RES-004, FAL-1900, O-2.
Excursion-detection accuracy in the field (both directions)
TREND-007 ★ — Missed-excursion field signal — MA
Source | receipt-QA discrepancies, complaints, independent-check disputes.
Threshold | EVENT — any credible missed-excursion is investigated; SPC on rate.
Action | investigate detection logic/calibration/probe → if systemic, potential public-health event (FM-400) + CAPA. Ties: HA-RES-001/002/003, RCR-019/020/021, VAL-CLIN-002 (VAL-SW-001).
TREND-008 ★ — False-reject / wastage / stock-out rate — FR
Source | disposition data, re-adjudication outcomes, customer wastage reports.
Threshold | ABS to the validated false-reject target (VAL-USE-002) once set; SPC meanwhile. Target EP — set in validation.
Action | rising false-reject → alarm-tuning/threshold review (RCR-047) without weakening the safety bias + CAPA. A rising false-reject rate is as reportable as a missed excursion (HA-RES-019). Ties: HA-MPE-003, RCR-047/044.
TREND-009 — Probe-fault flag rate + probe-placement error indicators — MA
Source | device probe-fault telemetry + complaints.
Threshold | SPC baseline.
Action | elevated → probe reliability (returns) / IFU-placement (HF) review + CAPA. Ties: HA-RES-003/ 014, RCR-020/045.
TREND-010 ★ — Fleet calibration-drift trend (EOL + depot recal residuals) — MA (COMMON CAUSE)
Source | depot recalibration residuals (CAL-004) + golden-unit/station data (CAL-003) + EOL coverage (PV-MFG-001).
Threshold | ABS to the accuracy budget (WHO PQS target, O-7, once confirmed); SPC on residual distribution for directional/systematic drift (a systematic same-direction shift is the S5-worst common cause).
Action | systematic drift → immediate escalation (fleet-wide detection/verdict bias risk) + recal campaign + reference-chain audit + CAPA; sets the recal interval (CAL-004). This is the standing monitoring for the calibration common cause (HAL-2900 §5.3). Ties: HA-RES-015, RCR-053/026/049, PV-MFG-001, MDT-800.
Alert delivery & connectivity
TREND-011 ★ — Alert-delivery success + un-acked backlog — MA/AV
Source | cloud alert/ack telemetry + end-to-end reconciliation (SYS-176).
Threshold | SPC on delivery-success; EVENT on any silent un-acked backlog growth.
Action | delivery shortfall → connectivity/rules-engine investigation + escalation; silent backlog → reconciliation fix + CAPA. Ties: HA-RES-008/017, RCR-029/032/035.
TREND-012 — Alert latency distribution per lane — MA
Source | cloud (event-time → delivery-time) per lane/region.
Threshold | SPC per lane; disclose lanes where the live link cannot close (defaulting to on-receipt verdict) — tied to PRD-VER-005.
Action | poor-latency lanes → connectivity/antenna review + honest disclosure (not a claim). Ties: HA-RES-007/009, O-1, RCR-030.
TREND-013 — Connectivity / attach-failure rate (esp. in-container) — AV
Source | device attach/upload telemetry per region/lane.
Threshold | SPC; correlate to the RF-in-container Open item (O-1).
Action | systematic → antenna/band investigation + store-and-forward adequacy review. Ties: HA-RES-009, PRD-VER-005.
Complaint categories & data integrity
TREND-014 — Complaint category Pareto — all directions
Source | classified complaints (CMP-###, FM-400).
Categories | no-charge / charge-inhibited-misread / no-connectivity / probe-fault / misread-verdict / false-reject dispute / low-battery / physical damage / label illegible / security.
Threshold | SPC per category + EVENT for any safety category (missed-excursion, thermal, security).
Action | category shift → targeted investigation + CAPA. Ties: FM-400, HA-RES cluster.
TREND-015 — Data-integrity / record-loss / gap-rate + DR-test cadence — MA/GDP
Source | device/cloud gap + loss telemetry, export-integrity checks.
Threshold | SPC on gap/loss; EVENT on any silent record loss (should be impossible — loss → honest gap).
Action | anomaly → integrity investigation + GDP-deviation handling. Ties: HA-RES-018/020, RCR-024/032/042/057.
Cybersecurity events
TREND-016 ★ — Integrity-failure / authenticity events — SEC (EVENT)
Source | device-anchored integrity checks + cloud.
Threshold | EVENT — any capture-time integrity/authenticity failure investigated.
Action | investigate (falsified/injected telemetry) + security incident response + CVD. Ties: HA-RES-010, RCR-034/038/055.
TREND-017 ★ — Alarm-suppression / anomaly events — SEC (EVENT)
Source | missing-ack escalation + anomaly detection.
Threshold | EVENT.
Action | investigate suppression attempt + incident response. Ties: HA-RES-011, RCR-035.
TREND-018 ★ — OTA campaign health + version distribution + rollback events — SEC/AV (COMMON CAUSE)
Source | OTA campaign telemetry (success/failure/rollback), fleet version distribution.
Threshold | EVENT on any regression/halt; SPC on version skew vs. supported matrix (DO-CFG-005).
Action | regression → halt campaign (halt-on-regression) + rollback + CAPA. This is the standing monitoring for the OTA/cloud common cause (HAL-2900 §5.3). Ties: HA-RES-012, RCR-037.
TREND-019 ★ — Cloud auth/integrity/availability incidents + DR-test results — SEC/AV (EVENT)
Source | cloud security monitoring + scheduled DR tests.
Threshold | EVENT on incident; scheduled DR-test cadence must pass.
Action | incident response; failed DR test → remediation + CAPA. Ties: HA-RES-013, RCR-039.
2. Threshold philosophy (severity-first)
Consistent with the severity-first risk policy: safety-significant metrics (S5 thermal TREND-004/005; missed-excursion TREND-007; calibration common cause TREND-010; security TREND-016/017/018/019) are EVENT- or absolute-threshold with immediate escalation — a single credible occurrence acts, not a statistical trend. Availability/business metrics use SPC baselining. No safety metric waits for a "rate" to accumulate before acting on a credible individual signal.
3. Baseline-then-hold rule (evidence honesty)
Because no field data exists, most quantitative thresholds are Evidence-Pending — baselined from the first field cohort (pilot → beta → production trial, VAL-FIELD-900), then held/tightened. The two exceptions are (a) safety EVENT metrics (act on first occurrence) and (b) the calibration-accuracy absolute threshold (anchored to the WHO PQS target once O-7 is confirmed). No metric target is asserted as a fact before it is baselined or set in validation.
4. Traceability
Backward | HAL-2800 (per-residual disclosure/monitoring obligations), HAL-2900 §5.3 (the two process common causes → TREND-010/018), the Risk-Management Summary §7.8, FM-100 (scope + sources), MDT-800 (calibration), MDT-900 (DHR serial attribution), the PVaL both-direction targets (VAL-USE-002 / VAL-CLIN / VAL-FIELD), REL-400 (condition 8).
Forward | FM-300 (threshold breach → closed loop), FM-400 (complaint categories), FM-500 (the specific residual commitments each TREND discharges), FM-600 (reliability-growth + periodic review consume the trends), ORCH-018 / EM-ENG-012 (Stage 18 — field-monitoring status column of the lifecycle spine).
FM-300 — Closed-Loop Field-to-Design Process (RCA → CAPA → Update) — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-300 · Owner: Quality (EM-ENG-006) + Reliability/Safety, with the discipline owners and Chief Systems (EM-ENG-001). The closed-loop process that turns field information into design/risk improvement: field record (FF-###) → root-cause analysis → CAPA (CAPA-###) → update of the requirement / hazard / FMEA / verification artefacts → re-baseline. This is the ISO 14971 §9 loop and the mechanism that keeps the SCM's conditional benefit-risk conclusion (HAL-2900) valid across the reusable fleet's service life. It builds on the Stage-11 RCA/CAPA readiness note (FAL-2000/2100), which was pre-wired to the FMEA/FTA/HAL for exactly this.
Evidence honesty: no field record exists yet (Conditionally Ready — not released); this defines the process and its triggers, pre-wired to the real upstream artefacts. No CAPA or root cause is asserted.
---
1. The loop (six steps)
FIELD INFORMATION (FM-100 sources) │ normalize → FF-### (serial via DHR, category, trip context) ▼
1. INTAKE & TRIAGE ──► safety/EVENT signal? → immediate escalation (S5 / missed-excursion / security / cal common-cause) │ ▼
2. SIGNAL DETECTION ──► threshold breach (FM-200 TREND-###) or credible individual safety signal │ ▼
3. ROOT-CAUSE ANALYSIS (RCA) ──► 5-Why / fault-tree / fishbone; reproduce; scope by DHR serial range │ ▼
4. CAPA (CAPA-###) ──► containment (recall/hold/field-action) + correction + preventive action + effectiveness check │ ▼
5. UPDATE THE ARTEFACTS ──► requirement / hazard (HAL) / FMEA (DFMEA/SWFMEA/HWFMEA/PFMEA) / verification (PV) / control plan / WI │ ▼
6. RE-BASELINE & REVALIDATE ──► re-score residual risk (HAL-2800/2900); re-affirm/revise benefit-risk; ECO; feed Stage 18 │ └──────────────► back to monitoring (effectiveness of the CAPA is itself a TREND-###)
2. Step detail
Step 1 — Intake & triage
Every field input is normalized to an FF-### record with the device serial (resolved through the Device History Record, MDT-900 — giving build lot, calibration birth-record, firmware/config version, and part lots), the trip/lane context, and a category (FM-200 TREND / FM-400 complaint category). Triage assigns severity (severity-first): any S5 (thermal), missed-excursion, security, or calibration-common-cause signal is escalated immediately to EM-ENG-001 + EM-ENG-006, independent of routine cadence.
Step 2 — Signal detection
A CAPA is triggered by either (a) a FM-200 threshold breach (SPC shift or EVENT metric) or (b) a credible individual safety signal (a single reported thermal event or missed excursion does not wait for a rate). Detection also runs the both-directions check: a false-reject/wastage trend (TREND-008) is treated with the same seriousness as a missed-excursion trend (TREND-007).
Step 3 — Root-cause analysis
RCA uses the methods pre-wired in FAL-2000/2100 (5-Why, fishbone, and — for safety-significant signals — the existing fault trees FAL-1800 whose top events are already the SCM's dominant harms). Crucially, RCA scopes the affected population by DHR serial range (which build lots, cal stations, firmware versions, part lots are implicated) — this is what makes the manufacturing traceability (MDT-900) load-bearing for the field loop. The RCA distinguishes design cause vs. process cause vs. use cause vs. supplier cause, routing to the right owner.
Step 4 — CAPA
Each CAPA-### carries: containment (fleet hold, OTA halt, recall/field-safety-notice scoped to the serial range, quarantine of a suspect part lot), correction (fix the units/process/config), preventive action (prevent recurrence — design change, process control, supplier CAPA, IFU/training change, OTA update), and a mandatory effectiveness check (a TREND-### confirming the action worked). Safety/regulatory CAPAs require EM-ENG-006 + EM-ENG-001 approval.
Step 5 — Update the upstream artefacts (the actual "closed loop")
This is what distinguishes a closed loop from a helpdesk. A field root cause updates the design/risk file:
Field finding implies | Update
A missed failure mode | Add/adjust the FMEA record (DFMEA/SWFMEA/HWFMEA or PFMEA MDT-200); re-score O/D from real data
A mis-estimated hazard/probability | Update the HAL hazard record + residual (HAL-2800)
A control that is weaker than assumed | Strengthen the RCR + its design output (ECO) + Control Plan line (MDT-300)
A process escape | Update the PFMEA + Control Plan + Work Instruction; tighten incoming/EOL (MDT-600/700)
A verification gap | Add/adjust the PV- (Stage 13) and its acceptance
A use error | Update the use-error FMEA (FAL-1600) + IFU/HF + revalidate (VAL-USE)
A supplier issue | Supplier CAPA (MDT-700) + AML/PCN action
A benefit-target shortfall | Update the benefit basis (WHO PQS) + VAL-CLIN
Step 6 — Re-baseline & revalidate
The updated residual is re-scored (HAL-2800) and the overall benefit-risk (HAL-2900) is re-affirmed or revised (ISO 14971 §9) — the periodic risk review (FM-600) is where this is formalized. Changes flow through ECO (DO-GOV-004) and, for a fielded fleet, through the compatibility matrix + OTA discipline (REL-200). The CAPA effectiveness becomes a monitored TREND-###. Everything feeds the Stage-18 lifecycle spine (ORCH-018 / EM-ENG-012), which carries a field-monitoring status column so the trace shows design → control → release evidence → field monitoring → CAPA → re-baseline.
3. Special loops (the SCM's load-bearing cases)
Calibration common cause (TREND-010) | a systematic drift signal → RCA on reference chain/cal station → recal campaign (containment scoped by cal-station/serial range) → update CAL-004 recal interval + MDT-800 MSA + PV-MFG-001 → re-baseline HA-RES-015. This is the standing closure of the "calibration is an operational control" gap (HAL-2900 §5.3).
OTA/config common cause (TREND-018) | a campaign regression → halt-on-regression (immediate containment) → RCA → rollback + fix → update DO-CFG-005 matrix + SWFMEA + PV-SW-010 → re-baseline HA-RES-012.
S5 thermal event (TREND-004) | any occurrence → immediate containment + recall scoping + DGR/regulatory report → RCA (cell/protection/handling) → update HWFMEA/PFMEA + supplier CAPA + potentially design change → re-baseline HA-RES-006. Mandatory per the S5 acceptance conditions.
Missed excursion / false-reject (TREND-007/008) | the both-directions loop → RCA on detection/calibration/ probe vs. alarm-tuning → update SWFMEA/HAL + VAL-USE-002 targets → re-baseline HA-RES-001/002/019.
4. Traceability
Backward | FAL-2000/2100 (RCA/CAPA readiness, pre-wired), FM-100 (sources), FM-200 (triggers/TRENDs), MDT-900 (DHR serial scoping), the HAL/ERFKL/RCR/PVL artefacts the loop updates, REL-400 (condition 8).
Forward | FM-400 (complaint/adverse-event intake is a primary FF source), FM-500 (the residual commitments the loop must keep valid), FM-600 (periodic risk review formalizes the re-baseline + reliability growth), ORCH-018 / EM-ENG-012 (Stage 18 — the field-monitoring → CAPA → re-baseline segment of the lifecycle spine).
FM-400 — Complaint & Adverse-Event Handling — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-400 · Owner: Quality/Regulatory (EM-ENG-006) + Support, with System-Safety and Product-Security. The complaint- and adverse-event-handling process for the SCM: how user-reported issues are received, recorded, classified, escalated, investigated, reported to regulators/programs, and fed into the closed loop (FM-300). Because the SCM is a vaccine cold-chain integrity device, this process carries specific vaccine-integrity / public-health reporting considerations — a suspected fleet-wide monitoring failure is a potential public-health event, not a routine complaint. Complaints carry CMP-###; they become FF-### field records in the closed loop.
Evidence honesty: no complaint exists (not released); this defines the process, its classification, and its escalation/reporting logic. Exact regulatory reporting timelines/regimes are Evidence-Pending until the D4 classification resolves (RA-10); the process is structured to attach to whichever regimes apply.
---
1. Intake & recording
Channels | support tickets, RMA returns, field-team reports, cloud-dashboard user flags, distributor/EPI program feedback, regulator/program queries.
Record (CMP-###) | device serial (→ Device History Record, MDT-900), reporter/role, date, trip/lane context, described symptom, harm-direction, and whether any patient/beneficiary or product-disposition consequence is alleged.
• Every complaint is recorded — including those resolved at first contact — so that category trends (TREND-014) are complete.
2. Classification (harm-direction + severity)
Class | Examples | Direction | Default handling
C-SAFETY-MISS | Suspected missed/false-accept excursion; a dose administered that may have been compromised | MA | Immediate escalation + potential public-health event (§4)
C-SAFETY-FR | Disputed false-reject / needless wastage / stock-out | FR | Escalation (both-directions) + re-adjudication
C-THERMAL (S5) | Swelling, venting, overheating, smoke, fire, transport incident | S5 | Immediate escalation + DGR/regulatory report + recall scoping
C-SECURITY | Suspected tampering, falsified record, alarm suppression, unauthorized access | SEC | Immediate escalation + security incident response + CVD
C-FUNCTION | No charge, charge-inhibited misread, no connectivity, probe fault, misread verdict, low battery | AV/MA | Investigate + trend
C-PHYSICAL | Enclosure/seal/label damage, ingress, cleaning degradation | AV | Investigate + trend (reuse life)
C-USE | Confusion arming/reading verdict/interpreting charge-inhibit | use | HF/IFU review + trend
Classification is severity-first: any C-SAFETY-MISS, C-SAFETY-FR, C-THERMAL, or C-SECURITY is escalated on receipt regardless of apparent frequency.
3. Investigation & closure
• Safety/security complaints enter the closed loop (FM-300) as high-priority FF-### → RCA → CAPA; the DHR scopes the affected serial range (build lot / cal station / firmware / part lot).
• Returned units are teardown-analyzed (Reliability): cell condition, seal/leak, calibration drift on return, firmware/config, coat/ingress, label legibility over reuse.
• Each complaint is closed with a documented investigation outcome, the disposition, any CAPA reference, and — where applicable — the regulatory/program report reference. Effectiveness of any resulting CAPA is monitored as a TREND-###.
4. Vaccine-integrity / public-health reporting considerations
This is the SCM-specific escalation the task calls out:
• A suspected fleet-wide detection failure is a potential public-health event. If RCA on a C-SAFETY-MISS complaint (or a TREND-007/TREND-010 signal) indicates a systematic detection or calibration bias that could have let spoiled doses be administered as good across many units/lots, it is escalated to EM-ENG-001 + EM-ENG-006 as a potential public-health event and handled with containment (fleet hold / recall / advisory to affected programs) — not as a routine complaint. This is the field manifestation of the S5-worst calibration common cause (HA-RES-015 / HA-MPE-002).
• Both harm directions are reportable. A C-SAFETY-FR (mass false-reject → wastage → stock-out → unvaccinated population) is a public-health harm in its own right (HA-RES-019) and is escalated and reported with equal seriousness — the conservative safety bias is monitored, not treated as free.
• Program/stakeholder notification. Where a device-integrity issue may have affected an EPI program's vaccine disposition decisions, the affected program(s) are notified so they can re-assess dispositions made on the device's record (the record is attributable per MDT-900). This supports the WHO PQS / WHO-UNICEF EVM and EU GDP data-integrity obligations.
• Regulatory adverse-event/vigilance reporting (per the confirmed D4 classification, RA-10): FDA/EU vigilance for reportable events; UN 38.3 / IATA-DGR for transport incidents; timelines per the applicable regime. Evidence-Pending until classification resolves.
5. Roles & timelines
Support | intake, record, first-contact resolution, escalation.
Quality/Regulatory (EM-ENG-006) | classification governance, regulatory/program reporting, CAPA convening, PMS reporting.
System-Safety | thermal/S5 + safety-signal adjudication.
Product-Security (EM-ENG-008) | security complaints + CVD.
Reliability/Service | return teardown + field data.
Chief Systems (EM-ENG-001) | public-health-event escalation authority + benefit-risk revalidation.
Timelines | safety/thermal/security complaints escalate immediately (same-day); routine complaints per the QMS SLA; regulatory reporting per the applicable regime's clock (EP, RA-10).
6. Traceability
Backward | FM-100 (PMS scope + sources), FM-200 (complaint category trend TREND-014 + safety EVENT metrics), MDT-900 (DHR serial attribution + recall scoping), HAL-2800 (per-residual monitoring obligations), HA-RES-015/019 (the public-health both-directions harms), REL-400 (condition 8 + claim discipline).
Forward | FM-300 (safety/security complaints → RCA → CAPA → artefact update), FM-500 (residual commitments the complaint stream helps monitor), FM-600 (complaint Pareto feeds the periodic risk review), ORCH-018 / EM-ENG-012 (Stage 18 — complaints/adverse events as a field-monitoring input to the lifecycle spine).
FM-500 — Residual-Risk Field-Monitoring Commitments — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-500 · Owner: System-Safety + Quality/Regulatory (EM-ENG-006), with Chief Systems (EM-ENG-001). The register of the specific residual risks that require post-market field monitoring, drawn directly from the disclosure/monitoring obligations recorded in HAL-2800 (individual residual risk) and HAL-2900 (overall residual + benefit-risk). HAL-2800 §8 states explicitly that "the disclosure/ monitoring obligations collected here become the post-market surveillance plan (Stage 17)"; this module is that collection, mapped to the field metrics (FM-200) and the closed loop (FM-300) that discharge each obligation. It is the traceable proof that the SCM's conditional benefit-risk conclusion is being kept valid in the field (ISO 14971 §9).
Evidence honesty: these are monitoring commitments, not results. Each maps a HAL residual to the metric that watches it and the closing verification it depends on; nothing is asserted as observed. The two carried-Open unknowns (D3, RF-in-container) and the S5 residual carry the strongest, mandatory commitments.
---
1. Residual → field-monitoring commitment register
Legend: Class — MANDATORY (release-condition or S5/security), STANDING (operational common cause), ROUTINE (trend). TREND = the FM-200 metric that discharges it.
Residual (HAL-2800) | Harm | Monitoring commitment | Class | TREND | Closes/depends on
HA-RES-006 — aircraft-hold battery fire (S5) | S5 | Mandatory post-market thermal-event monitoring + transport-incident reporting; any event → immediate escalation + recall scoping + DGR report | MANDATORY (S5) | TREND-004 | HAL-2200 §7 acceptance; UN 38.3; PV-TFC-002
HA-RES-005 — cold-charge → runaway (D3) | S5 | Monitor charge-inhibit field behavior + any charge-below-limit; field thermal precursors | MANDATORY (S5/D3) | TREND-005 | PV-ENRG-001; O-2
HA-RES-004 — silent gap / dead-in-cold (D3 energy) | AV/MA | Monitor field battery/endurance + gap rates + DOA; capacity fade over reuse | MANDATORY (Open D3) | TREND-001/002/006 | PRD-VER-004 (carried Open)
HA-RES-009 — in-container RF → no live link | MA/AV | Monitor per-trip/per-lane delivery-success + connectivity; disclose in-container limitation | MANDATORY (Open) | TREND-011/012/013 | PRD-VER-005 (carried Open)
HA-RES-015 — fleet calibration bias (S5-worst common cause) | MA | Standing calibration-drift monitoring (EOL + depot recal residuals) fleet-wide; systematic drift → immediate escalation | STANDING (common cause) | TREND-010 | PV-MFG-001 + depot recal; MDT-800
HA-RES-012 — malicious/compromised OTA (common cause) | SEC/AV | Standing OTA campaign-health + version-distribution monitoring; halt-on-regression | STANDING (common cause) | TREND-018 | PV-CYBER-005 + PV-SW-010 (IV)
HA-RES-001 — missed excursion (detection logic) | MA | Monitor missed-excursion field signals (returns, complaints, receipt-QA discrepancies) | ROUTINE→MANDATORY | TREND-007 | PV-SW-001; VAL-CLIN-002
HA-RES-002 — false accept over gap | MA | Monitor override rates + incomplete-record dispositions | MANDATORY (choke point) | TREND-014/015 | PV-SW-002; VAL-USE-004
HA-RES-003 — probe-fault read as valid | MA | Monitor probe-fault flag rates | ROUTINE | TREND-009 | PV-SW-008
HA-RES-007/008 — delayed / missing alert | MA | Monitor alert-latency distributions per lane + un-acked backlog fleet-wide | MANDATORY | TREND-011/012 | PV-COMM-001/002
HA-RES-010 — falsified/injected telemetry | SEC | Monitor integrity-failure events; CVD | MANDATORY (security) | TREND-016 | PV-CYBER-001 (IV)
HA-RES-011 — suppressed/blocked alarm | SEC | Monitor alarm-suppression/anomaly events | MANDATORY (security) | TREND-017 | PV-CYBER-002 (IV)
HA-RES-013 — cloud compromise/outage | SEC/AV | Monitor cloud auth/integrity/backlog + availability; DR-test cadence | MANDATORY (security) | TREND-019 | PV-CYBER-009 + DR test
HA-RES-014 — use-error cluster (misread/mute/mis-arm/override/probe placement) | use | Monitor override rates, muted-alert rates, mis-arm/void rates, probe-placement errors | MANDATORY (usability-gated) | TREND-009/014 | VAL-USE-001/002/004/007
HA-RES-016 — wrong subject/trip attribution | MA | Monitor binding/routing anomalies | ROUTINE | TREND-014 | PV-COMM-008
HA-RES-018/020 — record loss / GDP data-integrity | MA/GDP | Monitor loss/gap rates + DR-test cadence; data-integrity deviations; audit-readiness | MANDATORY (GDP) | TREND-015 | PV-DATA-006 + DR test
HA-RES-019 — false-reject / wastage / stock-out | FR | Monitor false-reject/wastage/stock-out as a first-class field metric — as reportable as a missed excursion | MANDATORY (both-directions) | TREND-008 | VAL-USE-002 + wastage data
HA-RES-021 — condensation/ingress/cleaning → electronics fault | AV/MA | Monitor ingress/BIST failures over reuse life; cleaning-cycle degradation | ROUTINE | TREND-003/014 | PV-TFC-006; DO-SVC-007
HA-RES-022 — time/units/location/estimate misrepresentation | GDP | Monitor untrusted-time incidence; provenance | ROUTINE | TREND-015 | PV-SW-005
2. The overall-residual commitments (HAL-2900)
HAL-2900 adds whole-device commitments beyond the per-hazard set:
§5.1 accumulation | the S4 region-A set is monitored as a portfolio — a slippage in a shared defence (the honest-gap/no-false-reassurance invariant) degrades many residuals at once, so the shared-invariant health (surfaced via missed-excursion + false-accept + data-integrity trends TREND-007/014/015) is a release-critical standing watch.
§5.2 convergence | the two pathway choke points (honest verdict + supported human decision) are the containment for ~15 converging hazards; their field proxies (override rates, incomplete-record dispositions, muted-alert rates — TREND-014) are monitored, and a rise is a portfolio-wide signal.
§5.3 common causes | calibration (TREND-010) and OTA/cloud (TREND-018/019) carry standing monitoring because their containment is operational, not a one-time fact — this is the single most important reason the SCM needs PMS, not just a pre-release test.
§5.5 rare catastrophic | the S5 aircraft-fire (TREND-004) carries mandatory monitoring + transport disclosure as a standalone benefit-risk line.
§6 accepted trade | the false-reject/wastage residual (TREND-008) is monitored so the conservative safety bias stays balanced, not free.
3. Discharge of the release conditions (REL-400 condition 8)
This register is the discharge of REL-400 condition 8 ("Post-market surveillance monitors both harm directions — missed-excursion and false-reject/wastage — plus thermal-event and security monitoring, feeding the ISO 14971 §9 loop"). Every MANDATORY commitment here must be live before the corresponding claim/release is granted; the two carried-Open unknowns (HA-RES-004/009) additionally require their closing verification (PRD-VER-004/005) — until then the associated claims are restricted per REL-400 §6.
4. Traceability
Backward | HAL-2800 (per-residual disclosure/monitoring fields) + HAL-2900 (§5/§6/§7 overall commitments), the Risk-Management Summary §7.8, REL-400 (condition 8 + carried-Open risks + S5 obligation), the RCR register, the PVL/PVaL closing verifications.
Forward | FM-200 (each commitment → a TREND metric), FM-300 (breach → closed loop → re-baseline HAL-2800/ 2900), FM-400 (complaints feed the safety/both-directions commitments), FM-600 (periodic risk review re-affirms these residuals + benefit-risk), ORCH-018 / EM-ENG-012 (Stage 18 — the residual-risk → field- monitoring segment of the lifecycle spine; this is where the trace shows a residual is actively watched, not just accepted).
FM-600 — Reliability Growth & Periodic Risk-Review Cadence — Smart Cold Chain Monitor (SCM)
Stage 17 · Module FM-600 · Owner: Reliability + System-Safety + Quality/Regulatory (EM-ENG-006), integrated by Chief Systems (EM-ENG-001). Defines (a) the reliability-growth program — how field failure data matures the SCM's reliability across the reusable, multi-trip fleet and closes the two carried-Open reliability unknowns (D3 cold endurance, RF-in-container) — and (b) the periodic risk-review cadence by which the benefit-risk conclusion (HAL-2900) is formally re-evaluated as production and post-production information accrues (ISO 14971 §9). This is the module that keeps the "positive but conditional" conclusion honest over the service life.
Evidence honesty: no reliability figure, MTBF, capacity-fade curve, or growth rate is asserted — the reliability model (FAL-1900) is prediction, not field-demonstrated; field values baseline from the first cohorts (VAL-FIELD-900) then mature. No benefit-risk revision is claimed; this defines the cadence and triggers for revision.
---
1. Reliability-growth program
1.1 Purpose
The SCM's top reliability risks (FAL-1900 / ERFKL index §4) are D3 90-day cold (−30 °C) endurance + reuse- life aging, RF-in-container egress, battery capacity fade over reuse cycles, and flash endurance. These are predicted pre-release and demonstrated only through cold-endurance testing (PRD-VER-004), RF testing (PRD-VER-005), and — ultimately — field data across trips. Reliability growth is the disciplined maturation of these from prediction → test → field-confirmed.
1.2 Growth data sources & phases
Phase | Source | Reliability question
Pre-release | PRD-VER-004 (cold endurance), PRD-VER-005 (RF), accelerated-life/HALT, HASS | Does the design meet the D3/RF targets? (release conditions)
Pilot → beta | VAL-FIELD-900 pilot/beta on LMIC/GDP lanes | Do the targets hold on representative lanes? Baseline the field metrics
Production trial → fleet | FM-200 telemetry (TREND-001/002/006/013) + returns teardown | Is field reliability tracking the model? Is fade/endurance as predicted?
1.3 Growth mechanism
Field failures (FF-###) → RCA (FM-300) → the reliability model (FAL-1900) is updated, the PM/recal intervals (DO-SVC-004/005) are revised from real drift/aging data, and design/process corrections flow via CAPA + ECO. A reliability-growth curve (failure intensity vs. cumulative device-trips) is tracked per the reliability cadence; a plateau below target triggers design/process action. The reusable multi-trip nature (D1) makes per-cycle aging (battery fade, seal wear over cleaning cycles, label durability, flash wear) first-class — reliability is measured in device-trips and reuse-cycles, not just calendar time.
1.4 The two carried-Open reliability unknowns
D3 cold endurance + fade (O-2) | closed initially by PRD-VER-004 + the battery datasheet; then field-confirmed via TREND-006 (SoH/fade over reuse) — a faster-than-model fade revises the recal/PM/ replacement intervals and re-scores HA-RES-004.
RF-in-container (O-1) | closed initially by PRD-VER-005 + the antenna decision; then field-confirmed via TREND-012/013 (per-lane delivery/attach) — persistently poor lanes drive antenna/band action or honest disclosure (no real-time claim on those lanes).
2. Periodic risk-review cadence
2.1 Cadence
Review | Frequency | Scope | Owner | Output
Fleet-metric review | Weekly (ramp) → monthly (steady) | All FM-200 TRENDs; open CAPAs | Post-Market | TREND status + CAPA triggers
Reliability-growth review | Quarterly + per cohort milestone | Growth curve; D3/RF field confirmation; fade/wear | Reliability | model update; interval revision
Periodic risk review (benefit-risk revalidation) | Quarterly + on trigger | Re-affirm/revise HAL-2800 residuals + HAL-2900 overall benefit-risk with field data | EM-ENG-006 + EM-ENG-001 + Safety | risk-file update; benefit-risk re-statement
PMS report | Per applicable regime (D4) | Consolidated field performance for regulators/programs | EM-ENG-006 | PMS report
Security posture review | Quarterly + on incident | Threat model, CVD, OTA/cloud posture | Product-Security | security update
2.2 Trigger-based (out-of-cadence) reviews
A periodic review is convened immediately, independent of the calendar, on any:
• S5 thermal event (TREND-004) — mandatory per the S5 acceptance conditions.
• Systematic calibration drift (TREND-010) — the S5-worst common cause.
• OTA campaign regression (TREND-018) or security incident (TREND-016/017/019).
• Suspected fleet-wide missed-excursion or mass false-reject (TREND-007/008) — a potential public-health event (FM-400).
• A closing verification result for a carried-Open unknown (PRD-VER-004/005) that differs from the projection.
2.3 What the periodic risk review does (ISO 14971 §9)
1. Ingests the field data (FM-200 trends, CAPAs, complaints, returns, reliability growth).
2. Re-scores the affected individual residuals (HAL-2800) with real O/D/probability data (replacing the pre-release projected estimates).
3. Re-evaluates the overall residual + benefit-risk (HAL-2900) — does the benefit still outweigh the overall residual? Are new hazards or common causes emerging? Is the accepted false-reject trade still balanced?
4. Re-affirms or revises the conclusion; where revised, drives requirement/hazard/FMEA/verification updates (FM-300) and, if needed, field action (advisory/recall/OTA) and claim changes.
5. Updates the living Risk-Management Summary (it is explicitly a living document per ISO 14971 §9) and feeds the Stage-18 lifecycle spine.
3. Maturation of pre-release estimates (the honesty ledger closes here)
Every O/D and probability in the PFMEA (MDT-200), the ERFKL FMEA family, and HAL-2800 was a pre-production/ pre-field estimate with a stated basis. This module is where those estimates are replaced by field- demonstrated values — the reliability-growth data re-scores the reliability modes, the field trends re-score the hazard probabilities, and the periodic review re-baselines the residual and benefit-risk. The evidence posture (no invented field rate) is discharged not by leaving the numbers blank forever, but by maturing them from real data through this cadence.
4. Traceability
Backward | FAL-1900 (reliability analysis + the top reliability risks), HAL-2800/2900 (residuals + benefit-risk to revalidate), the Risk-Management Summary (living document), FM-200 (the trend data), FM-300 (the closed loop that feeds updates), FM-500 (the residual commitments), REL-400 (conditions + carried-Open unknowns), the reliability validation VAL-REL-700 + field validation VAL-FIELD-900, PRD-VER-004/005.
Forward | the re-baselined residuals + benefit-risk feed back into HAL-2800/2900 (living), the requirement/ design/FMEA/verification updates (via FM-300 + ECO), the PMS report (D4 regulators/programs), and ORCH-018 / EM-ENG-012 (Stage 18) — the reliability-growth + periodic-review cadence is the mechanism by which the lifecycle spine's "field monitoring → status → re-baseline" stays current across the SCM's service life.

## Traceability Overview
Traceability Overview
Part of the end-to-end traceability record.
↩ Contents
A — Traceability System Overview — Smart Cold Chain Monitor (SCM)
Stage 18 — Complete Lifecycle Traceability · Orchestrator ORCH-018 · Owner EM-ENG-012 — Principal Systems Traceability Architect & Design Assurance Engineer (reports to Chief Systems, EM-ENG-001). Governed by 00-governance/traceability-standard.md; view set per 14-tools/traceability-copilot/traceability-matrix/references/matrix-map.md. Evidence-backed only in the supplied Stage 00–17 artefacts; no requirement, ID, result, approval, or risk rating is invented.
---
1. Purpose & scope
This is the Master Product Lifecycle Traceability Matrix for the SCM — a complete, evidence-graded, bidirectional chain connecting every major engineering artefact from market opportunity (Stage 01) through field monitoring (Stage 17). It is not a requirements-verification matrix; it is the whole lifecycle evidence chain, and it also reports what is missing, inconsistent, unsupported, duplicated, or orphaned.
Product in one line. The SCM is a reusable, battery-operated, cellular-connected environmental monitor placed inside insulated vaccine shipping containers to log temperature/humidity/shock, detect two-sided excursions, alert stakeholders in time to act, give an offline accept/reject verdict at the door, and provide an auditable, tamper-evident cold-chain record for product-disposition decisions across a multi-market (WHO PQS / EU GDP / US FDA / air-transport) reusable fleet. It is an information/monitoring device — it does not control temperature and is not therapeutic. Operating range −30 °C to +60 °C; endurance target 90 days. Foundational decisions: D1 reusable/rechargeable multi-trip · D2 local logging + cadence upload + immediate excursion alerts · D3 low-temperature Li-ion, duty-cycled, charge-inhibited below 0 °C (the top program risk) · D4 multi-market regulatory + air/ground transport.
2. Configurations covered
Configuration | Description | Trace status
SCM-100 (base) | Integral-antenna reusable monitor; LTE-M primary + Wi-Fi/BLE offload; internal + external probe | Primary configuration; all matrices trace this baseline
SCM-100-XA (external-antenna branch) | External-antenna variant (DO-EE-021 / DO-MECH-016); adopted only if the integral antenna cannot close the in-container link budget (ADR-010 / PRD-VER-005) | Candidate — Evidence Pending (RF-in-container test)
Regional configs | Per-market radio/regulatory configuration (D4); WHO PQS / EU GDP / US / air-transport | Applicable_to_configuration recorded per requirement; AML/regional freeze pending (RA-07)
Ultra-cold −70 °C variant | Roadmap/boundary topic (MR-1400) | Not a committed SKU — not traced as a release configuration
All revisions are TBD-rev / REL-TBD: the design-output baseline is a pre-DR-DO draft (Stage 09) and the release baseline is not frozen (Stage 16). Every V&V link is Planned / Evidence Pending (Stages 13–14 are authored playbooks, not executed).
3. Artefact sources — folder → ID-prefix map
Stage (folder) | Artefact | ID prefix(es) | Objects reported
01-market-requirements | Master Market Requirements Library | MR-<CODE>-NNN (EXEC/OPP/SEG/PROB/JTBD/COMP/SOTA/POS/COMM/CON/MET/REG/PER/ENV/PORT/ECO/INNO/SUS/RISK/VAL) | 208 reqs / 20 modules
02-voice-of-customer | Voice of Customer | VOC-###, FLD-###, CFB-###, LD-### | VOC-001…064, FLD-001…012, CFB-001…016, LD-001…010
03-stakeholder-needs | Stakeholder Needs | STN-### (stakeholders SH-##) | STN-001…027; SH-01…15
04-user-needs | Master User Needs Library | UN-<CODE>-NNN (VIS/STK/PER/JTBD/ENV/FLOW/FUN/UX/PERF/SAFE/SEC/PHY/SVC/REG/QUAL/ECO/BIZ/VAL) | 187 needs / 18 modules
05-product-requirements | Master Product Requirements Library | PRD-<CODE>-NNN (VIS/STK/ENV/ARCH/FUN/UX/MECH/ELEC/CONN/CLOUD/MOB/SEC/REL/SAFE/REG/MFG/SVC/PERF/QUAL/VER) | 220 reqs / 20 modules
06-system-requirements | System Requirements | SYS-### (allocated to subsystems S1–S9); interfaces I-01…I-12 | 190 (SYS-001…190)
07-subsystem-detailed-requirements | Subsystem + Software Requirements | HWR-###, MER-###, FWR-###, IFR-###; software SRS-<CODE>-NNN | 180 subsystem (HWR-001…050, MER-001…040, FWR-001…055, IFR-001…035) + 308 SRS / 16 modules = 488
08-system-architecture | System Architecture | FUN-###, CMP-###, INT-###, ADR-### | FUN-001…040, CMP-001…055, INT-001…035, ADR-001…012
09-design-outputs | Design Output Library (DOL) | DO-<CODE>-NNN (GOV/CFG/SYS/MECH/EE/PCB/FW/SW/CLD/APP/IF/SEC/SAFE/HF/MFG/TST/LBL/SVC/REG/VER/XFER) | 272 outputs / 21 modules (103 Critical)
10-hazard-analysis | Hazard Analysis Library (HAL) | HA-<CODE>-NNN; taxonomy HAZ-<CLASS>; controls RCR-### | ~120 hazard records; ~60 top rows; RCR-001…058
11-failure-mode-risk-analyses | ERFKL (FMEA/FTA/reliability/CCF) | FA-<CODE>-NNN (DFMEA/SWFMEA/HWFMEA/SYSFMEA/IFMEA/FTA/REL/CCF/USE) | ~232 records (133 FMEA + ~60 FTA basic + 13 use-error + 14 CCF + 12 reliability)
12-risk-control-requirements | Risk-Control Requirements | RCR-### | 58 consolidated, deduplicated (RCR-001…058)
13-verification | Product Verification Library (PVL) | PV-<CODE>-NNN | 195 activities / 20 modules — all Planned / Evidence Pending
14-validation | Product Validation Library (PVaL) | VAL-<CODE>-NNN | 128 activities / 13 modules — all Planned / Evidence Pending
15-manufacturing-design-transfer | Manufacturing & Design Transfer | PROC-###, STEP-###, PFM-###, CP-###, CC-##, SC-##, PT-###, SUP-###, INC-###, CAL-###, MSA-### | 28 PFMEA modes; CC-01…08, SC-01…09
16-production-release | Production Release | REL-###, RA-##, DR-### | RA-01…12; DR gate records; release decision
17-field-monitoring | Field Monitoring / PMS | FF-###, CMP-###, CAPA-###, TREND-### | TREND-001…019 (plan, not report — no field records yet)
4. Relationship model (controlled types only — no vague "related to")
The chain is expressed with the governed relationship vocabulary:
Upstream | derived_from · responds_to · satisfies · supports · refines · decomposes · allocated_from · justified_by · constrained_by
Architecture/implementation | allocated_to · implemented_by · realized_by · defined_in · contained_in · interfaces_with · configured_by · produced_by
Risk | introduces_hazard · contributes_to_hazard · causes_failure_mode · results_in_effect · creates_hazardous_situation · can_result_in_harm · controlled_by · control_implemented_by · control_verified_by · introduces_new_risk · monitored_by
V&V | verified_by · verification_evidence · validated_by · validation_evidence · tested_by · covered_by · failed_by · retested_by
Manufacturing | manufactured_by · controlled_by_process · tested_in_production_by · special_characteristic_of · derived_from_pfmea · preserved_by_control_plan
Release/lifecycle | included_in_baseline · approved_by · applicable_to_configuration · monitored_in_field_by · updated_due_to_field_evidence
Bidirectional rule. Every critical link is traversable both ways (User Need ↔ Product Requirement; Requirement ↔ Design Output; Hazard ↔ RCR; Requirement ↔ Verification; User Need ↔ Validation; Critical Characteristic ↔ Control Plan). Inverses are not stored twice but appear in both directions in the views — e.g. Matrix-02 (decomposition, downward) and Matrix-01/07 (upward from PV to requirement).
5. Status, evidence, and confidence models
Status (per object/link): Not Started · Candidate · Draft · Under Review · Approved · Implemented · Verification Planned/InProgress/Verified · Validation Planned/InProgress/Validated · Released · Monitoring · Change Required · Blocked · Rejected · Retired · Superseded. The dominant status across the SCM is Draft / Approved (design-input) and Verification-Planned / Validation-Planned (V&V) — nothing is Released.
Evidence quality (grade every major link): E0 none · E1 assertion · E2 expert rationale · E3 documented analysis · E4 inspection/review · E5 controlled test · E6 independent/accredited · E7 field. For the SCM, pre-V&V links sit at E1–E3 (they rest on authored requirements, architecture, and risk analysis). Every verification and validation link is Evidence-Pending (E0/planned) because no test has executed. No E4–E7 evidence exists yet anywhere in the file.
Confidence (for inferred links): Confirmed · Probable · Candidate · Unresolved · Rejected. Links that a source document states explicitly (e.g. the Hazard-Traceability-Matrix RCR→DO→PV chains, or the requirement-rationale-traceability PRD→UN parents) are Confirmed. Links this stage infers from wording or module scope are marked Candidate and flagged for human approval where safety/regulatory/release- critical.
6. Overall traceability health (headline)
Dimension | Assessment
Chain completeness (structure) | Strong. Every stage names its upstream parents and forward children with real IDs; the requirement rationale table (Stage 05) and the Stage-10/11/12 spine matrices give Confirmed cross-references end-to-end.
Requirement → source coverage | High. 220 PRD each cite ≥1 UN parent (Stage 05 rationale table); UN cite STN/VOC/MR/D1–D4.
Requirement → verification coverage | Structurally complete, evidentially empty. 58/58 RCRs and all applicable PRD/SYS/SRS map to ≥1 PV activity, but every PV is Planned / Evidence-Pending.
User need → validation coverage | Structurally complete, evidentially empty. Use errors, residuals, and safety-critical workflows all have VAL references; every VAL is Planned / Evidence-Pending.
Hazard → control coverage | Complete. Every High/unacceptable HAL hazard maps to ≥1 RCR with a DO- implementation and a named PV-/VAL- — no hazard "controlled on paper only" — but effectiveness is unproven.
Evidence maturity | Low by design (pre-execution). No E4–E7 evidence; two items carried Open (D3 battery → PRD-VER-004; RF-in-container → PRD-VER-005); five regulatory/config items Evidence-Pending.
Release posture | Conditionally Ready — production NOT authorized (see I). Traceability is a well-formed plan; it is not yet a proof.
Bottom line: the SCM has an unusually complete, honestly-labelled traceability structure; its gaps are almost entirely evidence gaps (unexecuted V&V) and five carried-open unknowns (D3 endurance, RF-in-container, WHO PQS category, US FDA classification, AML/chemistry+antenna freeze) — not broken links. Details in E (Coverage Dashboard), F (Orphans & Gaps), G (Inconsistencies), H (Closure Plan), I (Release Readiness).

## Traceability Index
Traceability Index
Part of the end-to-end traceability record.
↩ Contents
Stage 18 — Complete Lifecycle Traceability — Smart Cold Chain Monitor (SCM)
Project: smart-cold-chain-monitor-lifecycle · Stage 18 — Master Product Lifecycle Traceability Matrix · Orchestrator: ORCH-018 · Owner: EM-ENG-012 — Principal Systems Traceability Architect & Design Assurance Engineer (reports to Chief Systems, EM-ENG-001). Governed by 00-governance/traceability-standard.md; view set per the matrix-map. This stage integrates all upstream stages (00–17) into one complete, evidence-graded, bidirectional traceability system. Evidence-backed only in the supplied stage artefacts; no requirement, ID, result, approval, or risk rating is invented.
---
What this is
The Master Product Lifecycle Traceability Matrix for the SCM — not a requirements-verification table, but the whole lifecycle evidence chain from market opportunity (Stage 01) to field monitoring (Stage 17), answering for any capability why it exists → which need created it → which requirements define it → which subsystem/architecture/design output realizes it → which hazards/failures attach → which controls are required and implemented → how each is verified and validated → which manufacturing controls preserve intent → which release evidence supports acceptance → which residual risks remain → how field data monitors the assumptions — and what is missing, inconsistent, unsupported, duplicated, or orphaned.
Outputs (A–I + the 11 matrices)
File | Contents
[A-Traceability-System-Overview.md](A-Traceability-System-Overview.md) | Scope, configurations, artefact folder→ID-prefix map, relationship/status/evidence/confidence models, overall traceability health
[B-Master-Object-Register.md](B-Master-Object-Register.md) | Object inventory by type with counts cited from each stage; representative real IDs per flagship chain
[C-Master-Relationship-Register.md](C-Master-Relationship-Register.md) | Controlled relationship types populated with real example links across the chain (Confirmed/Candidate; E-grades)
[Matrix-01-Executive-End-to-End.md](Matrix-01-Executive-End-to-End.md) | The 7 flagship capability chains traced through all 20 columns
[Matrix-02-Requirements-Decomposition-Allocation.md](Matrix-02-Requirements-Decomposition-Allocation.md) | PRD→SYS→subsystem allocation, safety/security class, criticality
[Matrix-03-Requirements-Design-Output.md](Matrix-03-Requirements-Design-Output.md) | Requirement → design output, implementation/review status
[Matrix-04-Hazard-Risk-Control.md](Matrix-04-Hazard-Risk-Control.md) | Top HA-…/RCR- chain: hazard→situation→harm→control→DO→PV/VAL→residual→field
[Matrix-05-FMEA-Requirements-Hazard.md](Matrix-05-FMEA-Requirements-Hazard.md) | FMEA/FTA/CCF → requirement/hazard/RCR/DO/verification
[Matrix-06-Cybersecurity-Safety.md](Matrix-06-Cybersecurity-Safety.md) | Threat → hazard → security requirement/control → IV verification → residual
[Matrix-07-Verification-Coverage.md](Matrix-07-Verification-Coverage.md) | Requirement/RCR → PV → method → acceptance → status (all Evidence-Pending)
[Matrix-08-Validation-Coverage.md](Matrix-08-Validation-Coverage.md) | User need → VAL scenario → success criterion → status (all Evidence-Pending)
[Matrix-09-Manufacturing-Design-Transfer.md](Matrix-09-Manufacturing-Design-Transfer.md) | DO → CC/SC → PFMEA → control plan → production test → capability
[Matrix-10-Release-Readiness.md](Matrix-10-Release-Readiness.md) | Baselines, V&V/mfg/regulatory/cyber readiness, blockers, decision, conditions
[Matrix-11-Field-Feedback-Closed-Loop.md](Matrix-11-Field-Feedback-Closed-Loop.md) | Pre-wired trend→RCA→CAPA→re-baseline loop (no field data yet)
[E-Coverage-Dashboard.md](E-Coverage-Dashboard.md) | Coverage metrics/counts — structural vs evidential; gaps not hidden
[F-Orphan-and-Gap-Report.md](F-Orphan-and-Gap-Report.md) | Upstream/downstream/evidence orphans + the 5 carried-Open items
[G-Inconsistency-Report.md](G-Inconsistency-Report.md) | Consistency checks; the reconciled preliminary-PRD-SAFE→RCR tension
[H-Traceability-Closure-Action-Plan.md](H-Traceability-Closure-Action-Plan.md) | P0–P4 actions to close the gaps and reach "Ready"
[I-Release-Readiness-Conclusion.md](I-Release-Readiness-Conclusion.md) | The classification, blockers, unresolved risks, monitoring, authorities
Integration scope
This stage consumes and integrates every upstream stage:
01 Market → 02 Voice of Customer → 03 Stakeholder Needs → 04 User Needs → 05 Product Requirements → 06 System Requirements → 07 Subsystem/SRS → 08 Architecture → 09 Design Outputs → 10 Hazard Analysis → 11 ERFKL → 12 Risk-Control Requirements → 13 Verification → 14 Validation → 15 Manufacturing/Design Transfer → 16 Production Release → 17 Field Monitoring.
Real ID prefixes harvested: MR- · VOC-/FLD-/CFB-/LD- · STN- · UN- · PRD- · SYS- · HWR-/MER-/FWR-/IFR- · SRS- · FUN-/CMP-/INT-/ADR- · DO- · HA-/HAZ-/RCR- · FA- · PV- · VAL- · PROC-/STEP-/PFM-/CP-/CC-/SC-/PT- · REL-/RA-/DR- · FF-/CMP-/CAPA-/TREND-.
Overall traceability health
• Structural coverage: high (~100%) — needs→requirements→design→hazard→control→V&V links present, bidirectional, with real IDs; the safety spine (Stage-10/11/12 matrices) is Confirmed end-to-end.
• Evidential maturity: ~0% — 0/195 verification and 0/128 validation executed; no E4–E7 evidence; all residuals projected.
• Carried-Open unknowns: 5 — D3 endurance, RF-in-container, WHO PQS category, US FDA classification, AML (chemistry+antenna) freeze.
• No active contradictions — the one designed-in tension (preliminary PRD-SAFE) was reconciled at Stage 12.
Release-readiness classification (see I)
CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED (release pending V&V + closure of the 3 P0 blockers
and the 8 conditions of release). Equals the Stage-10 / Stage-15 / Stage-16 conclusions.
Owned by EM-ENG-012 / ORCH-018. 2026-07-19.

## Object Register
Object Register
Part of the end-to-end traceability record.
↩ Contents
B — Master Object Register — Smart Cold Chain Monitor (SCM)
Stage 18 · ORCH-018 / EM-ENG-012. A complete object inventory by type, with counts cited from the stage that reported them. This register does not re-author objects; it catalogues what the upstream libraries produced, with representative real IDs so every type is traceable. Object metadata schema per traceability-standard.md §Object.
---
1. Object counts by lifecycle stage (as reported by each stage)
# | Object type | ID prefix | Count | Modules | Source stage (as reported)
1 | Market requirement | MR-<CODE>-NNN | 208 | 20 | 01-market-requirements/index.md ("208 / 20 modules")
2 | Voice-of-customer item | VOC-### | 24 (VOC-001…024) + 15 (VOC-050…064) = ~39 | — | 02-voice-of-customer/index.md
2b | Field observation | FLD-### | 12 (FLD-001…012) | — | 02-voice-of-customer/index.md
2c | Customer feedback / complaint | CFB-### | 16 (CFB-001…016) | — | 02-voice-of-customer/index.md
2d | Lost-deal / no-decision | LD-### | 10 (LD-001…010) | — | 02-voice-of-customer/index.md
3 | Stakeholder need | STN-### | 27 (STN-001…027) | — | 03-stakeholder-needs/index.md
3b | Stakeholder class | SH-## | 15 (SH-01…15) | — | 03-stakeholder-needs/index.md §5
4 | User need | UN-<CODE>-NNN | 187 | 18 | 04-user-needs/index.md ("187 / 18 modules")
5 | Product requirement | PRD-<CODE>-NNN | 220 | 20 | 05-product-requirements/index.md ("220 / 20 modules")
6 | System requirement | SYS-### | 190 (SYS-001…190) | 11 cats | 06-system-requirements/index.md
7 | Subsystem/detailed req | HWR/MER/FWR/IFR-### | 180 (50+40+55+35) | 4 families | 07-subsystem-detailed-requirements/index.md
7b | Software requirement | SRS-<CODE>-NNN | 308 | 16 | 07-subsystem-detailed-requirements/index.md ("Grand total 488")
8 | Architecture function | FUN-### | 40 (FUN-001…040) | — | 08-system-architecture/index.md
8b | Architecture component | CMP-### | 55 (CMP-001…055) | — | 08-system-architecture/index.md
8c | Interface | INT-### | 35 (INT-001…035) | — | 08-system-architecture/index.md
8d | Architecture decision | ADR-### | 12 (ADR-001…012) | — | 08-system-architecture/index.md
9 | Design output | DO-<CODE>-NNN | 272 (103 Critical) | 21 | 09-design-outputs/index.md
10 | Hazard record | HA-<CODE>-NNN | ~120 (~60 top rows) | 30 modules | 10-hazard-analysis/index.md
10b | Hazard taxonomy class | HAZ-<CLASS> | (taxonomy HAL-1900) | — | 10-hazard-analysis (HAL-1900)
11 | Failure/risk record | FA-<CODE>-NNN | ~232 | 9 analyses | 11-failure-mode-risk-analyses/index.md
12 | Risk-control requirement | RCR-### | 58 (RCR-001…058) | — | 12-risk-control-requirements/index.md
13 | Verification activity | PV-<CODE>-NNN | 195 | 20 | 13-verification/index.md
14 | Validation activity | VAL-<CODE>-NNN | 128 | 13 | 14-validation/index.md
15 | Process step | PROC-###/STEP-### | (process flow) | — | 15-manufacturing-design-transfer (MDT-100)
15b | Process failure mode | PFM-### | 28 | — | 15 (MDT-200 PFMEA)
15c | Critical characteristic | CC-## | 8 (CC-01…08) | — | 15 (MDT-400)
15d | Special characteristic | SC-## | 9 (SC-01…09) | — | 15 (MDT-400)
15e | Control plan / prod test | CP-###/PT-### | (control plan + test spec) | — | 15 (MDT-300/600)
16 | Release readiness item | RA-## | 12 (RA-01…12) | — | 16-production-release (REL-100)
16b | Design-review gate | DR-### | (gate records) | — | 16 (REL-300)
17 | Field metric/trend | TREND-### | 19 (TREND-001…019) | — | 17-field-monitoring (FM-200)
17b | Field record / complaint / CAPA | FF/CMP/CAPA-### | 0 actual (plan only) | — | 17-field-monitoring
Grand total of controlled requirement/need objects (Stages 01–07): 208 + ~39 VoC-family + 27 + 187 + 220 + 190 + 180 + 308 = ~1,359 demand/requirement objects, plus 142 architecture elements, 272 design outputs, ~120 hazards, ~232 failure records, 58 RCRs, 195 verifications, 128 validations, and the manufacturing/release/field object sets. The traceability spine links these.
2. Representative objects by type (real IDs, with the flagship capability they anchor)
Demand space (Stages 01–04)
Object | Real ID | Title / meaning | Flagship chain
Market problem | MR-PROB-001 | "Tell me before it's too late" (real-time excursion warning) | (1) excursion
Market problem | MR-PROB-002 | Freezing quietly ruins stock unseen | (1)(3) two-sided
Market problem | MR-PROB-007 | Logger dies in the cold, record lost | (2)(3) battery
Market problem | MR-PROB-005 | Must defend every excursion to an inspector | (6) audit
Market JTBD | MR-JTBD-004/005/008 | Accept/reject at door; real-time warning; audit-grade proof | (1)(4)(6)
Strategic risk | MR-RISK-001 | D3 low-temp battery tension (top strategic risk) | (2)(3)
Market validation | MR-VAL-006 | Closure-by-test of the D3 tension | (2)(3)
VoC item | VOC-008 / VOC-021 | "warn me in time" / excursion-in-flight | (1)
VoC item | VOC-010 / VOC-024 | last logger gave out in the cold | (3)
VoC item | VOC-001 / VOC-016 | audit defensibility / recalibration on reuse | (5)(6)(7)
Complaint | CFB-001 / CFB-012 | dead-in-cold logger / recal on return | (3)(7)
Stakeholder need | STN-001 | Real-time excursion warning in time to act | (1)
Stakeholder need | STN-003 | Endure the full cold trip (−30 °C, 90 days) | (2)(3)
Stakeholder need | STN-004 | Instant, offline accept/reject verdict | (1)(4)
Stakeholder need | STN-005 / STN-025 | Audit-grade tamper-evident records / signed OTA | (6)
Stakeholder need | STN-010 / STN-016 / STN-021 | Reusable readiness / recover-return / calibration currency | (5)(7)
Stakeholder need | STN-020 | No compromised dose given; no safe dose wasted (severity anchor) | all
User need | UN-FUN-004 / UN-SAFE-002 | Two-sided (freeze+heat) detection | (1)(3)
User need | UN-FUN-008 / UN-SAFE-005 | Offline verdict for the least-trained user | (1)(4)
User need | UN-SAFE-007 | No charge in unsafe cold (D3) | (2)
User need | UN-PERF-003 | 90-day endurance at −30 °C | (2)(3)
User need | UN-PERF-008 | Alert egress / connectivity inside container | (4)
User need | UN-SEC-001/005 | Tamper-evident records / secure OTA | (6)
User need | UN-SVC-001 / UN-REG-007 | Refurbishment loop / traceable calibration | (5)(7)
User need | UN-VIS-001 / UN-SAFE-001 | Beneficiary-outcome anchor (severity-first tie-breaker) | all
Solution space (Stages 05–09)
Object | Real ID | Meaning | Flagship chain
Product req | PRD-FUN-006 / PRD-SAFE-002 | Equal-reliability two-sided detection | (1)(3)
Product req | PRD-FUN-010/011 | Offline verdict; verdict/record never disagree | (1)(4)
Product req | PRD-ELEC-005 / PRD-SAFE-007 | Charge-inhibit below 0 °C | (2)
Product req | PRD-ELEC-001/002 / PRD-REL-001 | Low-temp chemistry; 90-day cold endurance (validate by test) | (2)(3)
Product req | PRD-ELEC-007 / PRD-SAFE-008 | Battery protection; no fire/thermal hazard | (2)
Product req | PRD-ENV-009 | RF egress inside metallized shipper | (4)
Product req | PRD-REG-007 / PRD-FUN-002 | Traceable calibration; dual-channel product-zone temp | (5)
Product req | PRD-SEC-001/008 / PRD-FUN-017 | Tamper-evident records; signed OTA; complete record + gap flag | (6)
Product req | PRD-SVC-001/002 / PRD-VIS-005 | Refurbishment loop; cleaning; reusable multi-trip (D1) | (7)
Product req | PRD-VER-004 / PRD-VER-005 | The two Open headline tests (D3; RF-in-container) | (2)(3)(4)
System req | SYS-009/010/011/013 | Two-sided detection; excursion logic; alert emit; local capture | (1)
System req | SYS-015/016/017 | Offline verdict; device↔cloud consistency; honest incomplete-record | (1)(4)(6)
System req | SYS-024/025 | Trip-profile config; pre-arm readiness (charge/cal/FW/profile/probe) | (5)(7)
System req | SYS-026/027/028 | Tamper-evident record; export; append-only audit trail | (6)
System req | SYS-032 | Signed, fail-safe OTA | (6)
System req | SYS-086/090/091/092 | Power/energy/battery/charge (D3 set) | (2)(3)
System req | SYS-104/109/117 | Connectivity & RF (RF-in-container set) | (4)
Subsystem HW | HWR-005 / HWR-006 / HWR-012 | HW charge-inhibit; cell-temp sensing; secondary protection backstop | (2)
Subsystem HW | HWR-001 / HWR-011 | Low-temp cell selection; UN 38.3 qualification | (2)(3)
Subsystem HW | HWR-021/022/027/028 | Temp sensor accuracy; probe channel; signal-chain; probe-presence | (5)
Subsystem HW | HWR-029/032/033/035 | LTE-M modem; integral antennas; external-antenna provision; RF at −30 °C | (4)
Subsystem HW | HWR-016 / HWR-018 / HWR-020 | Secure element; debug-lock; NVM log memory | (6)
Subsystem FW | FWR-022/025 | Two-sided threshold eval (equal-reliability freeze); excursion generation | (1)
Subsystem FW | FWR-028/029/030 | Offline verdict; no-login display; honest incomplete-monitoring | (1)(4)(6)
Subsystem FW | FWR-013 | Charge-inhibit enforcement in firmware | (2)
Subsystem FW | FWR-018 | Calibration-coefficient application + provenance | (5)
Subsystem FW | FWR-034/035 | Append-only tamper-evident log; power-fail-safe atomic writes | (6)
Subsystem FW | FWR-046/048/049 | Update signature verify; A/B rollback; data preservation across update | (6)
Subsystem FW | FWR-031/032 | Arming state machine; pre-arm readiness gate | (5)(7)
Architecture fn | FUN-008/009/012 | Two-sided eval; product-stability logic; local excursion capture | (1)
Architecture fn | FUN-017/018/019/020 | Offline verdict; incomplete-monitoring qualification; render; device↔cloud consistency | (1)(6)
Architecture fn | FUN-037/038/039/040 | Signed fail-safe OTA; energy governor; charge-inhibit; battery protection/power-fail-safe | (2)(6)
Architecture cmp | CMP-039/040/041 | Detection engine; verdict engine; storage manager | (1)(6)
Architecture cmp | CMP-037/043/044/054 | Energy governor; OTA agent; crypto module; audit store | (2)(6)
Design output | DO-SAFE-001…014 | The 14 Critical safety/risk-control outputs (all Preliminary — reconcile w/ HAL) | all
Design output | DO-FW-006 / DO-EE-013 | Detection firmware; sensor/probe front-end | (1)(5)
Design output | DO-EE-007 / DO-FW-005 | Charge-inhibit HW; charge-inhibit FW | (2)
Design output | DO-EE-003 (chemistry) / DO-EE-018 (antenna) | The two AML-freeze-blocking Evidence-Pending outputs | (2)(4)
Design output | DO-FW-014/015 / DO-CLD-010 | Signed OTA image/bootloader; OTA service | (6)
Design output | DO-REG-001/005 | WHO PQS category; US FDA classification (Evidence Pending) | regulatory
Risk, control, V&V, transfer, release, field (Stages 10–17)
Object | Real ID | Meaning | Flagship chain
Hazard | HA-SW-001 ★ | Missed excursion → compromised dose administered (S4/S5) | (1)
Hazard | HA-SW-008 ★ | Device reports temps it isn't measuring → false accept | (5)
Hazard | HA-ENRG-001 ★ | Cold-charge → plating → latent short → thermal event | (2)
Hazard | HA-TFC-002 ★★ | Li-cell fire in occupied aircraft cargo hold (S5) | (2)
Hazard | HA-ENRG-006 ★ | Energy exhausted → silent gap reads uneventful (D3) | (3)
Hazard | HA-COMM-003 ★ | Excursion, no live notification (container blocks radio) | (4)
Hazard | HA-CYBER-001/005 ★ | Injected false "clean" record; malicious firmware | (6)
Hazard | HA-SVC-001 ★ / HA-MFG-001 ★ | Depot recal escape; EOL miscalibration | (5)(7)
Failure record | FA-SWFMEA-001/002 | Detection/verdict failure modes | (1)
Failure record | FA-HWFMEA-002/006 | Battery/charge failure modes | (2)
Failure record | FA-FTA TOP-A…D | Four HAL ★ top events (missed excursion, no/late alert, Li fire, corrupt record) | (1)(2)(4)(6)
Risk control | RCR-019/021/022 | Two-sided detection; verdict suppresses false accept; honest gap | (1)(6)
Risk control | RCR-001/002/006 | HW charge-inhibit; protection+fuse; cold endurance reserve | (2)(3)
Risk control | RCR-026/049/053 | Calibration integrity + pre-arm block; interval; EOL coverage | (5)
Risk control | RCR-031/030 | RF link margin + antenna; store-and-forward consequence bound | (4)
Risk control | RCR-034/037/041/042 | Capture-time authenticity; secure OTA; provenance; export integrity | (6)
Verification | PV-SW-001 / PV-ENRG-001 / PV-REL-001/005 | Excursion FI; cold-charge FI; D3 endurance; RF-in-container | (1)(2)(3)(4)
Verification | PV-SW-004 / PV-CYBER-001/005 | Calibration integrity FI; integrity + OTA pen-test (IV) | (5)(6)
Validation | VAL-USE-001/002/004 | Read verdict→disposition; alarm response; never trust void trip | (1)
Validation | VAL-SW-001/002 / VAL-COMM-001/003 | Detection/gap operational; alert delivery; RF-in-container operational | (1)(4)
Critical char | CC-01 / CC-05 / CC-06 / CC-07 | Charge-inhibit <0 °C; temp-sensor cal accuracy; signed FW+region; crypto identity | (2)(5)(6)
Special char | SC-01 / SC-03 | Seal/IP integrity; final functional/system | (6)(7)
Release item | RA-05 / RA-07 / RA-10 | V&V not executed; AML not frozen; regulatory unresolved (the 3 P0 blockers) | all
Field trend | TREND-004 / TREND-007/008 / TREND-010 / TREND-018/019 | Thermal-event; both harm directions; cal drift; OTA/cloud rollout | (2)(1)(5)(6)
3. Object metadata note
Every object above carries the standard metadata (object_id · object_type · title · source_document · revision · configuration · lifecycle_status · owner · criticality · safety/security/regulatory class · evidence_quality · assumptions · open_questions) in its home stage file. This register normalizes the IDs; the per-object detail lives in the cited stage document. Revisions are TBD-rev/REL-TBD (Evidence-Pending) and lifecycle_status is predominantly Draft/Approved (inputs) or Verification/Validation-Planned (V&V) — nothing is Released.

## Relationship Register
Relationship Register
Part of the end-to-end traceability record.
↩ Contents
C — Master Relationship Register — Smart Cold Chain Monitor (SCM)
Stage 18 · ORCH-018 / EM-ENG-012. The controlled relationship types populated with real example links harvested from the stage artefacts. Relationship metadata schema per traceability-standard.md §Relationship (relationship_id · source · type · target · direction · rationale · evidence_reference · confidence · status). Confidence = Confirmed when the source document states the link explicitly; Candidate when inferred by this stage. Evidence grade E0–E7. Every V&V relationship is Evidence-Pending.
---
1. How to read
• Direction — → forward (upstream→downstream); the inverse is implied and rendered in the matrices.
• Ev — evidence grade at this link (E1 assertion · E2 expert rationale · E3 documented analysis; V&V = Evidence-Pending).
• Conf — confidence (Confirmed / Candidate).
• Source of the link is named in the rationale (e.g. Stage-05 rationale table; Stage-10 Hazard-Traceability- Matrix; Stage-12 Risk-Control-Traceability).
2. Upstream demand relationships (responds_to / derived_from / satisfies / refines / decomposes)
Rel-ID | Source | Type | Target | Rationale (source) | Ev | Conf
R-0001 | VOC-008 | responds_to | MR-PROB-001 | VoC theme T1 anchored to MR-PROB-001 (Stage-02 §5) | E2 | Confirmed
R-0002 | STN-001 | derived_from | VOC-008/021, CFB-002 | STN-001 source field (Stage-03 §6) | E2 | Confirmed
R-0003 | STN-003 | derived_from | VOC-010/024, CFB-001, D3 | STN-003 source field (Stage-03 §6) | E2 | Confirmed
R-0004 | UN-FUN-004 | refines | STN-002 (via T2) | Two-sided need from freeze theme (Stage-04 mapping) | E2 | Candidate
R-0005 | UN-SAFE-007 | derived_from | D3, STN-003 | Charge-safety need from D3 tension (Stage-04) | E2 | Confirmed
R-0006 | UN-VIS-001/UN-SAFE-001 | derived_from | STN-020 | Beneficiary-outcome severity anchor (Stage-04 §conflicts) | E3 | Confirmed
R-0007 | PRD-FUN-006 | satisfies | UN-FUN-004, UN-SAFE-002 | Stage-05 rationale table (PRD-FUN-006 parent) | E2 | Confirmed
R-0008 | PRD-ELEC-005 | satisfies | UN-SAFE-007, D3 | Stage-05 rationale (PRD-ELEC-005 "Charge inhibit below 0 °C") | E2 | Confirmed
R-0009 | PRD-REL-001 | satisfies | UN-PERF-003, D3 | Stage-05 rationale (PRD-REL-001 "90-day cold endurance, validate by test") | E2 | Confirmed
R-0010 | PRD-FUN-010 | satisfies | UN-FUN-008, UN-SAFE-005 | Stage-05 rationale (offline verdict) | E2 | Confirmed
R-0011 | PRD-SEC-001/PRD-FUN-017 | satisfies | UN-SEC-001, UN-FUN-014 | Stage-05 rationale (tamper-evident record) | E2 | Confirmed
R-0012 | PRD-ENV-009 | satisfies | UN-ENV-006, UN-PERF-008 | Stage-05 rationale (RF egress in metallized shipper) | E2 | Confirmed
R-0013 | PRD-SVC-001/PRD-VIS-005 | satisfies | UN-SVC-001, UN-VIS-006, D1 | Stage-05 rationale (refurb loop; reusable multi-trip) | E2 | Confirmed
R-0014 | PRD-REG-007 | satisfies | UN-REG-007 | Stage-05 rationale (traceable calibration) | E2 | Confirmed
3. Decomposition & allocation (decomposes / allocated_to)
Rel-ID | Source | Type | Target | Rationale | Ev | Conf
R-0101 | SYS-009 | decomposes | PRD-FUN-006 | SYS-009 "two-sided detection with equal reliability" (Stage-06 SYS-functional) | E2 | Confirmed
R-0102 | SYS-009 | allocated_to | S3 (Compute/MCU+FW), S4 (Sensing) | Stage-06 allocation | E2 | Confirmed
R-0103 | SYS-015 | decomposes | PRD-FUN-010 | SYS-015 offline verdict (Stage-06) | E2 | Confirmed
R-0104 | SYS-086/090 | decomposes | PRD-ELEC-002/REL-001 | Power/energy set (Stage-06 §evidence) | E2 | Confirmed
R-0105 | SYS-104/109/117 | decomposes | PRD-ENV-009/CONN | RF set (Stage-06) | E2 | Confirmed
R-0106 | FWR-022 | decomposes | SYS-009 | FWR-022 "two-sided eval, equal-reliability freeze" (Stage-07) | E2 | Confirmed
R-0107 | HWR-005 | decomposes | SYS-090-line, PRD-ELEC-005 | HWR-005 HW charge-inhibit (Stage-07) | E2 | Confirmed
R-0108 | FWR-028 | decomposes | SYS-015 | FWR-028 offline verdict computation (Stage-07) | E2 | Confirmed
R-0109 | FWR-018 | decomposes | SYS-024-line, PRD-REG-007 | FWR-018 calibration-coefficient + provenance (Stage-07) | E2 | Confirmed
R-0110 | FWR-046/048 | decomposes | SYS-032, PRD-SEC-008 | Secure OTA: signature verify + A/B rollback (Stage-07) | E2 | Confirmed
4. Architecture realization (realized_by / implemented_by / defined_in)
Rel-ID | Source | Type | Target | Rationale | Ev | Conf
R-0201 | SYS-009 | realized_by | FUN-008 (+CMP-039 detection engine) | Stage-08 functional decomposition | E2 | Confirmed
R-0202 | SYS-015 | realized_by | FUN-017 (+CMP-040 verdict engine) | Stage-08 | E2 | Confirmed
R-0203 | SYS-090 (charge) | realized_by | FUN-039 (+CMP-037 energy governor) | Stage-08 (charge-inhibit function) | E2 | Confirmed
R-0204 | SYS-032 (OTA) | realized_by | FUN-037 (+CMP-043 OTA agent, CMP-044 crypto) | Stage-08 | E2 | Confirmed
R-0205 | SYS-026 (record) | realized_by | FUN-030/031 (+CMP-041 storage, CMP-054 audit store) | Stage-08 | E2 | Confirmed
R-0206 | FWR-022 | implemented_by | DO-FW-006 (detection firmware) | Stage-09 / Stage-10 Design-Impl column | E2 | Confirmed
R-0207 | HWR-005 | implemented_by | DO-EE-007 (charge-inhibit HW) | Stage-10 Hazard-Traceability-Matrix (DO-SAFE-001→EE-007) | E3 | Confirmed
R-0208 | HWR-021/022 | implemented_by | DO-EE-013 (sensor/probe front-end) | Stage-10 matrix (DO-SAFE-005→EE-013) | E3 | Confirmed
5. Risk relationships (creates_hazardous_situation / controlled_by / control_implemented_by)
Rel-ID | Source | Type | Target | Rationale (source) | Ev | Conf
R-0301 | FA-SWFMEA-001 | causes_failure_mode → | HA-SW-001 | FMEA↔HAL reconciliation (Stage-11 §4; Stage-10 matrix source SYS-149/150) | E3 | Confirmed
R-0302 | HA-SW-001 | controlled_by | RCR-019 | Stage-10 Hazard-Traceability-Matrix §2.3 | E3 | Confirmed
R-0303 | RCR-019 | control_implemented_by | DO-SAFE-004→DO-FW-006+DO-EE-013 | Stage-12 Risk-Control-Traceability §2.3 | E3 | Confirmed
R-0304 | HA-ENRG-001 | controlled_by | RCR-001/002 | Stage-10 matrix §2.1 | E3 | Confirmed
R-0305 | RCR-001 | control_implemented_by | DO-SAFE-001→DO-EE-007+DO-FW-005 | Stage-12 §2.1 | E3 | Confirmed
R-0306 | HA-TFC-002 ★★ | controlled_by | RCR-002/004/007/008 | Stage-10 matrix §2.1 (S5 aircraft-hold fire) | E3 | Confirmed
R-0307 | HA-ENRG-006 ★ | controlled_by | RCR-006/024/027 | Stage-10 matrix §2.1 (D3 dead-in-cold) | E3 | Confirmed
R-0308 | HA-COMM-003 ★ | controlled_by | RCR-031/030 | Stage-10 matrix §2.4 (RF-in-container Open) | E3 | Confirmed
R-0309 | HA-SW-008 ★ | controlled_by | RCR-020/045 | Stage-10 matrix §2.3 (probe/measurement integrity) | E3 | Confirmed
R-0310 | HA-CYBER-001 ★ | controlled_by | RCR-034 | Stage-10 matrix §2.5 (injected false record) | E3 | Confirmed
R-0311 | HA-USE-001 ★ | controlled_by | RCR-043 | Stage-10 matrix §2.6 (verdict misread) | E3 | Confirmed
R-0312 | RCR-001 | introduces_new_risk | HA-NEW (charge-inhibit→DOA) | Stage-12 index §3 / HAL-2700 | E2 | Confirmed
6. Verification & validation relationships (verified_by / validated_by — all Evidence-Pending)
Rel-ID | Source | Type | Target | Method / rationale | Ev | Conf/Status
R-0401 | RCR-019 | verified_by | PV-SW-001 (fault-injection) | Stage-12/13 (excursion detection FI) | Evidence-Pending | Verification Planned
R-0402 | RCR-019 | validated_by | VAL-SW-001 | Stage-12/14 (detection operational) | Evidence-Pending | Validation Planned
R-0403 | RCR-006/017 | verified_by | PRD-VER-004 → PV-REL-001 + PV-ENRG-001 | D3 90-day cold endurance (Open) | Evidence-Pending | Verification Planned (Open)
R-0404 | RCR-001 | verified_by | PV-ENRG-001 (cold-charge FI) + UN 38.3 | Stage-12 §2.1 | Evidence-Pending | Verification Planned
R-0405 | RCR-031 | verified_by | PRD-VER-005 → PV-REL-005 + PV-COMM-003 | RF-in-container (Open) | Evidence-Pending | Verification Planned (Open)
R-0406 | RCR-026 | verified_by | PV-SW-004 (calibration integrity FI) | Stage-12 §2.3 | Evidence-Pending | Verification Planned
R-0407 | RCR-021 | validated_by | VAL-SW-002, VAL-USE-001 | Stage-12 §2.3 | Evidence-Pending | Validation Planned
R-0408 | RCR-034 | verified_by | PV-CYBER-001 (integrity pen-test, IV) | Stage-12 §2.5 (independent adversarial) | Evidence-Pending | Verification Planned
R-0409 | RCR-037 | verified_by | PV-SW-010 + PV-CYBER-005 (OTA pen-test, IV) | Stage-12 §2.5 | Evidence-Pending | Verification Planned
R-0410 | HA-USE-001…009 | validated_by | VAL-USE-001…009 (1:1) | Stage-14 index (HA-USE ↔ VAL-USE correspondence) | Evidence-Pending | Validation Planned
R-0411 | UN-PERF-003 | validated_by | VAL-REL-700 (D3 in-use) | Stage-14 (reliability validation) | Evidence-Pending | Validation Planned
7. Manufacturing preservation (special_characteristic_of / preserved_by_control_plan / tested_in_production_by)
Rel-ID | Source | Type | Target | Rationale (Stage-15) | Ev | Conf
R-0501 | RCR-001 / DO-EE-007 | special_characteristic_of | CC-01 (charge-inhibit <0 °C) | MDT-400 CC list | E2 | Confirmed
R-0502 | CC-01 | preserved_by_control_plan | CP-### (zero-defect line, MDT-300) | MDT-300 control plan | E2 | Confirmed
R-0503 | RCR-026 / DO-FW-007 | special_characteristic_of | CC-05 (temp-sensor cal accuracy) | MDT-400 | E2 | Confirmed
R-0504 | CC-05 | tested_in_production_by | PT-### (EOL calibration) → PV-MFG-001 | MDT-600 / PVL-1300 | Evidence-Pending | Confirmed
R-0505 | RCR-037 / DO-FW-014 | special_characteristic_of | CC-06 (signed FW + region config) | MDT-400 | E2 | Confirmed
R-0506 | RCR-055 / DO-MFG | special_characteristic_of | CC-07 (unique crypto identity + debug-lock) | MDT-400 | E2 | Confirmed
R-0507 | PFM-### | derived_from_pfmea | CP-### / CC-01…08 | MDT-200 PFMEA → MDT-300 | E3 | Confirmed
8. Release & lifecycle (included_in_baseline / approved_by / monitored_in_field_by / updated_due_to_field_evidence)
Rel-ID | Source | Type | Target | Rationale | Ev | Conf/Status
R-0601 | (all design outputs) | included_in_baseline | REL-200 (REL-TBD) | Stage-16 baseline not frozen | E1 | Blocked (AML pending, RA-07)
R-0602 | Release decision | approved_by | EM-ENG-001 (DRB chair) + EM-ENG-006 | Stage-16 REL-400 | E2 | Conditionally Ready — not authorized
R-0603 | HA-TFC-002 residual (A-High S5) | approved_by | Acceptance authority (HAL-2800) w/ mandatory monitoring | Stage-10/16 | E3 | Conditional
R-0604 | HA-ENRG-006 / HA-SW-004 | monitored_in_field_by | TREND-010 (cal drift), TREND-004 (thermal) | Stage-17 FM-200/500 | E1 | Monitoring Planned
R-0605 | HA-USE-001/002 | monitored_in_field_by | TREND-007/008 (both harm directions) | Stage-17 | E1 | Monitoring Planned
R-0606 | Field record FF-### (future) | updated_due_to_field_evidence | HAL/FMEA/requirement re-baseline | Stage-17 FM-300 closed loop | E0 | Not Started (no field data)
9. Relationship health note
Relationship coverage is structurally complete along the safety spine: the ~60 top hazards each carry a Confirmed controlled_by → control_implemented_by chain (Stage-10 matrix; Stage-12 table), and all 58 RCRs carry a verified_by (and, where use-related, validated_by) target. The weakness is not missing links — it is link maturity: every verified_by/validated_by sits at Evidence-Pending, and the included_in_baseline relationships are Blocked pending the AML freeze. Inferred (Candidate) links are confined to a small number of demand-space refinements (e.g. R-0004) where the intermediate mapping is by theme rather than an explicit cited ID; these are flagged for human confirmation and never presented as Confirmed.

## T1 End-to-End
T1 End-to-End
Part of the end-to-end traceability record.
↩ Contents
Matrix-01 — Executive End-to-End — Smart Cold Chain Monitor (SCM)
Stage 18 · ORCH-018 / EM-ENG-012. One row per flagship capability chain, traced through all 20 columns of the standard (Market → VoC → Stakeholder → User Need → Product Req → System Req → Subsystem Req → Architecture → Design Output → Hazard → Failure Mode → Risk-Control → Verification → Validation → Manufacturing → Release Baseline → Field Monitoring → Status → Gap). Real IDs only; V&V columns are Evidence-Pending; inferred links marked Candidate. Given the width, each chain is rendered as a labelled block (all 20 fields) followed by a compact summary table.
Columns (matrix-map.md): Chain ID · Market Requirement · Voice of Customer · Stakeholder Need · User Need · Product Requirement · System Requirement · Subsystem Requirement · Implementing Architecture Element · Design Output · Linked Hazard · Linked Failure Mode · Risk-Control Requirement · Verification Evidence · Validation Evidence · Manufacturing Control · Release Baseline · Field Monitoring · Overall Status · Gap Summary.
---
CHAIN-1 — Vaccine excursion detection → alert → disposition (the dominant harm pathway)
Column | Value
Chain ID | CHAIN-1
Market Requirement | MR-PROB-001 (warn before it's too late), MR-JTBD-005 (real-time warning), MR-JTBD-004 (accept/reject at door)
Voice of Customer | VOC-008, VOC-021, CFB-002
Stakeholder Need | STN-001 (real-time warning in time to act), STN-004 (offline verdict), STN-018 (routed/escalated alert), STN-020 (no compromised dose)
User Need | UN-FUN-006 (immediate alert when connected), UN-FUN-008 / UN-SAFE-005 (offline verdict, least-trained user), UN-SAFE-001 / UN-VIS-001 (beneficiary anchor)
Product Requirement | PRD-FUN-008 (immediate alert), PRD-FUN-009 (routing/escalation), PRD-FUN-010 (offline verdict), PRD-FUN-011 (verdict/record never disagree), PRD-SAFE-001 (don't mislead disposition)
System Requirement | SYS-010 (excursion logic), SYS-011 (alert emit when connected), SYS-013 (local capture regardless of connectivity), SYS-015 (offline verdict), SYS-016 (device↔cloud consistency), SYS-149/150 (safety)
Subsystem Requirement | FWR-025 (excursion generation), FWR-027 (immediate alert hand-off), FWR-028 (offline verdict), FWR-043 (alert priority over upload); SRS-DEV(200)/SRS-CONN(300)
Implementing Architecture Element | FUN-012 (local excursion capture), FUN-013 (alert emit/queue), FUN-017 (offline verdict), FUN-019 (verdict render); CMP-039 (detection engine), CMP-040 (verdict engine), CMP-048 (rules/alert engine)
Design Output | DO-SAFE-004 → DO-FW-006 + DO-EE-013 (detection); DO-SAFE-009 → DO-FW-011 + DO-CLD-007/008 (alert); DO-HF-002 (verdict display)
Linked Hazard | HA-SW-001 ★ (missed excursion → compromised dose), HA-COMM-001 ★ (late alert), HA-USE-001 ★ (verdict misread)
Linked Failure Mode | FA-SWFMEA-001/002 (detection/verdict), FA-SWFMEA-023/024 (alert delivery), FA-FTA TOP-A (missed excursion), TOP-B (no/late alert)
Risk-Control Requirement | RCR-019 (two-sided detection), RCR-021 (verdict suppresses false accept), RCR-028/029 (local commit + alert-ahead/ack/escalate), RCR-043 (legible verdict)
Verification Evidence | PV-SW-001 (excursion FI), PV-SW-006 (alert-gen+escalation FI), PV-SW-007 (display-integrity FI), PV-COMM-001 (latency+S&F) — all Evidence-Pending
Validation Evidence | VAL-SW-001 (detection operational), VAL-USE-001 (read verdict→correct disposition), VAL-COMM-001 (alert delivery) — all Evidence-Pending
Manufacturing Control | SC-03 (final functional/system test); PT-### acceptance-spine functional test (MDT-600)
Release Baseline | REL-TBD (SCM-100); blocked by RA-05 (V&V not executed)
Field Monitoring | TREND-007 (missed-excursion signals), TREND-008 (false-reject/wastage), alert-latency-per-lane
Overall Status | Verification Planned / Validation Planned — not verified, not validated
Gap Summary | Structurally complete; the entire pathway's effectiveness rests on PV-SW-001/002 + VAL-USE-001 which are unexecuted (P0). Two convergence choke points (honest verdict + supported human decision) not yet demonstrated.
CHAIN-2 — Li-battery safety / charge-inhibit-below-0 °C
Column | Value
Chain ID | CHAIN-2
Market Requirement | MR-PROB-007 (logger dies in the cold), MR-RISK-001 (D3 top strategic risk), MR-SEG-004 (air-transport), D3
Voice of Customer | VOC-007 (battery compliance), CFB-006 (transport), VOC-010 (cold death)
Stakeholder Need | STN-003 (endure cold trip), STN-012 (safe/compliant air/ground transport)
User Need | UN-SAFE-007 (no charge in unsafe cold), UN-SAFE-008 (battery protection), UN-SAFE-009 / UN-REG-006 (safe Li transport)
Product Requirement | PRD-ELEC-005 / PRD-SAFE-007 (charge-inhibit <0 °C), PRD-ELEC-007 / PRD-SAFE-008 (protection; no fire), PRD-ELEC-015 / PRD-SAFE-009 (UN 38.3)
System Requirement | SYS-090-line (charge management), SYS-091/092 (protection/charge), SYS-149-line (safety)
Subsystem Requirement | HWR-005 (HW charge-inhibit), HWR-006 (cell-temp sensing), HWR-003 (protection IC), HWR-012 (secondary backstop), HWR-011 (UN 38.3), FWR-013 (FW charge-inhibit enforcement); SRS-SAFE(1200)
Implementing Architecture Element | FUN-039 (charge management w/ cold charge-inhibit), FUN-040 (battery protection), FUN-038 (energy governor); CMP-037 (energy governor)
Design Output | DO-SAFE-001 → DO-EE-007 (charge-inhibit HW) + DO-FW-005 (FW); DO-SAFE-002 → DO-EE-005 + DO-EE-010 (protection); DO-EE-003 (cell chemistry, Evidence Pending)
Linked Hazard | HA-ENRG-001 ★ (cold-charge→plating→short), HA-ENRG-002 (cell OV/OC/OT/short), HA-TFC-002 ★★ (aircraft-hold fire, S5)
Linked Failure Mode | FA-HWFMEA-002/006/007/008 (battery/charge/protection), FA-DFMEA-009 (containment), FA-FTA TOP-C (Li fire)
Risk-Control Requirement | RCR-001 (HW charge-inhibit, non-defeatable, Level-1), RCR-002 (protection + secondary fuse), RCR-004 (retention), RCR-007 (transport), RCR-008 (containment)
Verification Evidence | PV-ENRG-001 (cold-charge FI), PV-ENRG-002 (protection FI, DO-TST-009), PV-TFC-002 (containment/fire-enclosure) + UN 38.3 (T.1–T.8) — Evidence-Pending
Validation Evidence | (none direct — physical safety verified, not user-validated); VAL-USE (state-legibility, HA-USE-009) — Evidence-Pending
Manufacturing Control | CC-01 (charge-inhibit <0 °C), CC-02 (protection+backstop), CC-03 (correct chemistry/parts), CC-04 (cell retention) — zero-defect Control-Plan lines (MDT-300/400)
Release Baseline | REL-TBD; blocked by RA-07 (AML/chemistry not frozen — DO-EE-003)
Field Monitoring | TREND-004 (mandatory thermal-event monitoring, S5), charge-inhibit / protection-trip / return events
Overall Status | Verification Planned — charge-inhibit placement Level-1 correct, effectiveness unproven; S5 residual permanently region-A-High with mandatory monitoring
Gap Summary | Correct control hierarchy; Open — D3 cold-charge safety unproven (PV-ENRG-001) and cell chemistry (DO-EE-003) not frozen (P0 RA-07). UN 38.3 not executed → no air-transport claim (P0).
CHAIN-3 — 90-day −30 °C endurance (D3)
Column | Value
Chain ID | CHAIN-3
Market Requirement | MR-PROB-007, MR-RISK-001 (D3), MR-VAL-006 (closure-by-test), D3
Voice of Customer | VOC-010, VOC-024, CFB-001 (last logger gave out in the cold)
Stakeholder Need | STN-003 (endure the full cold trip: −30 °C, 90 days)
User Need | UN-PERF-003 (90-day endurance at −30 °C), UN-PERF-010 (capacity over reuse)
Product Requirement | PRD-ELEC-001 (low-temp chemistry), PRD-ELEC-002 (90-day capacity), PRD-REL-001 (90-day cold endurance, validate by test), PRD-ARCH-004 (energy governor), PRD-SAFE-012 (dead-in-cold as safety-relevant), PRD-VER-004
System Requirement | SYS-086/087 (energy budget), SYS-045/046 (endurance performance), SYS-165/166/167 (reliability)
Subsystem Requirement | HWR-001 (low-temp cell), HWR-010 (quiescent current at −30 °C), HWR-014 (MCU deep-sleep), FWR-008 (lowest-viable-sleep), FWR-009 (radio/GNSS gating), FWR-010/012 (duty-cycle + per-mode budget); SRS-PERF(1000)/SRS-REL(1100)
Implementing Architecture Element | FUN-038 (energy governor / duty-cycle & radio gating); CMP-037 (power-state manager); ADR-003/004
Design Output | DO-SAFE-012 → DO-FW-004 + DO-EE-008 (energy governor + fuel gauge); DO-EE-003 (cell chemistry, Evidence Pending)
Linked Hazard | HA-ENRG-006 ★ (energy exhausted → silent gap reads uneventful), HA-ENV-001 (deep-cold miss)
Linked Failure Mode | FA-HWFMEA-001/012 (cold endurance / battery aging), FA-SWFMEA-021 (energy governor), FA-REL-007/009 (cold endurance / aging), FA-FTA TOP-A
Risk-Control Requirement | RCR-006 (endurance reserve + graceful degradation), RCR-017 (cold operation), RCR-024/027 (power-fail-safe, safe-state)
Verification Evidence | PRD-VER-004 → PV-REL-001 + PV-ENRG-001 (90-day cold endurance), HALT/HASS, battery aging — Evidence-Pending, carried OPEN
Validation Evidence | VAL-REL-700 (D3 cold endurance in-use, Open) — Evidence-Pending
Manufacturing Control | CC-03 (correct chemistry); incoming/enhanced control on cell (SUP-###, MDT-700)
Release Baseline | REL-TBD; blocked by RA-07 (chemistry not frozen) + RA-05 (endurance test not run)
Field Monitoring | Field battery/endurance trends + silent-gap rates (FM-500)
Overall Status | OPEN — the single most load-bearing unknown; energy margin unproven
Gap Summary | The 90-day/−30 °C combination is in genuine physical tension (D3). Architecture is sound (duty-cycled governor) but endurance is not demonstrated and chemistry not frozen. P0 release blocker.
CHAIN-4 — RF-in-container alert egress
Column | Value
Chain ID | CHAIN-4
Market Requirement | MR-PROB-008 (last mile black box), MR-PROB-001, D2/D4
Voice of Customer | VOC-023, FLD-004, LD-010 (last-mile connectivity)
Stakeholder Need | STN-009 (work through intermittent connectivity), STN-001 (real-time warning)
User Need | UN-PERF-008 (alert egress/connectivity), UN-FUN-009 (offline resilience), UN-ENV-006 (RF in container)
Product Requirement | PRD-ENV-009 (RF egress inside metallized shipper), PRD-FUN-013 (cadence upload + multi-bearer offload), PRD-ARCH-005 (multi-bearer, LTE-M primary), PRD-VER-005
System Requirement | SYS-104 (connectivity), SYS-109, SYS-117 (RF), SYS-069-line (environmental)
Subsystem Requirement | HWR-029 (LTE-M modem), HWR-032 (integral antennas), HWR-033 (external-antenna provision), HWR-035 (RF at −30 °C), FWR-040 (store-and-forward queue), FWR-041 (multi-bearer selection); IFR (RF interfaces); SRS-CONN(300)
Implementing Architecture Element | FUN-022 (multi-bearer selection), FUN-013/014 (alert emit/prioritization); CMP-042 (comms agent), CMP-046 (S&F endpoint); ADR-010
Design Output | DO-IF-003 (RF ICD, Evidence Pending), DO-EE-017/018 (antenna, Evidence Pending), DO-SAFE-009 (alert path); SCM-100-XA branch (DO-EE-021/DO-MECH-016, Candidate)
Linked Hazard | HA-COMM-003 ★ (excursion, no live notification — container blocks radio), HA-COMM-006 (roaming gap)
Linked Failure Mode | FA-HWFMEA-029/030/032 (antenna/RF), FA-IFMEA-005 (RF interface), FA-CCF-010, FA-FTA TOP-B
Risk-Control Requirement | RCR-031 (in-container link margin + integral-vs-external antenna, Level-1), RCR-030 (store-and-forward consequence bound), RCR-056 (roaming config)
Verification Evidence | PRD-VER-005 → PV-REL-005 + PV-COMM-003 (RF-in-container) — Evidence-Pending, carried OPEN
Validation Evidence | VAL-COMM-003 (RF-in-container operational), VAL-COMM-006 (roaming) — Evidence-Pending
Manufacturing Control | SC-07 (correct-config/line-clearance), config control (CC-06 region config)
Release Baseline | REL-TBD; blocked by RA-07 (antenna not frozen — DO-EE-018) + RA-08/11 (RF ICD/test config)
Field Monitoring | Per-trip delivery-success rate per lane; coverage/roaming gap logs (FM-200)
Overall Status | OPEN — principal open architectural risk
Gap Summary | Consequence is bounded by store-and-forward + local commit (so a blocked link degrades to delayed-not-lost). But live-alert capability inside metallized shippers is unproven; antenna strategy (integral vs external) not frozen. P0. No real-time-alert claim on lanes where PRD-VER-005 shows the link cannot close.
CHAIN-5 — Temperature-measurement accuracy & calibration integrity
Column | Value
Chain ID | CHAIN-5
Market Requirement | MR-PROB-005 (defend excursions to inspector), MR-REG (WHO PQS accuracy), MR-JTBD-008
Voice of Customer | VOC-001 (audit), VOC-016 (recalibration on reuse), CFB-012
Stakeholder Need | STN-005 (audit-grade records), STN-021 (calibration currency & traceability)
User Need | UN-FUN-003 (dual-channel product-zone temp — via UN-FUN mapping), UN-REG-007 (traceable calibration), UN-PERF (accuracy)
Product Requirement | PRD-FUN-002 (dual-channel temp), PRD-REG-007 (traceable calibration), PRD-FUN-007 (product-relevant excursion logic)
System Requirement | SYS-024 (config incl. calibration validity), SYS-025 (pre-arm: valid calibration), SYS-007-line (measurement)
Subsystem Requirement | HWR-021 (internal temp sensor accuracy class), HWR-022 (external-probe channel), HWR-027 (signal chain/reference stability), HWR-028 (probe-presence/continuity), FWR-017 (dual-channel independence), FWR-018 (calibration-coefficient + provenance); SRS-DEV(200)
Implementing Architecture Element | FUN-001/002 (acquire internal/external temp), FUN-006 (plausibility & fault flagging); CMP-038 (sensor drivers), CMP-039 (detection engine)
Design Output | DO-FW-007 (calibration/coefficient FW), DO-EE-013 (sensor/probe front-end), DO-SAFE-005/007 (probe-integrity/pre-arm)
Linked Hazard | HA-SW-008 ★ (reports temps it isn't measuring), HA-SW-004/014 (biased/coefficient error), HA-DATA-001 ★ (decision on unrepresentative numbers), HA-USE-005 (probe placement)
Linked Failure Mode | FA-SWFMEA-005 (calibration), FA-HWFMEA-017/018 (sensor/probe), FA-DFMEA-010/011
Risk-Control Requirement | RCR-020 (probe-fault), RCR-026 (calibration integrity + pre-arm block), RCR-045 (pre-arm gate), RCR-049 (calibration interval), RCR-053 (EOL cal coverage)
Verification Evidence | PV-SW-008 (probe-fault-injection), PV-SW-004 (calibration integrity/expiry FI), PV-DATA-001 (cal-coverage), PV-MFG-001 (EOL cal coverage), PV-ANAL (MSA) — Evidence-Pending
Validation Evidence | VAL-USE-005 (probe placement), VAL-USE-001 (disposition on the numbers), VAL-CLIN-002 (operational detection effectiveness) — Evidence-Pending
Manufacturing Control | CC-05 (temp-sensor calibration accuracy) — MSA-qualified, master-challenged (MDT-800); EOL calibration PT-###
Release Baseline | REL-TBD; WHO PQS accuracy target Evidence-Pending (RA-10)
Field Monitoring | TREND-010 (calibration drift — a dominant common cause), systematic-bias signals (FM-500)
Overall Status | Verification Planned; WHO PQS accuracy target not confirmed
Gap Summary | Calibration is a standing operational control (not a one-time fact) — monitored as common cause. Accuracy value and WHO PQS target are Evidence-Pending; MSA/Cpk not established. P1/P0 (regulatory).
CHAIN-6 — Data authenticity / audit trail / secure OTA
Column | Value
Chain ID | CHAIN-6
Market Requirement | MR-PROB-005 (defend to inspector), MR-PROB-006 (attribution), MR-JTBD-008, MR-REG (EU GDP / 21 CFR Part 11)
Voice of Customer | VOC-001 (audit), VOC-004 (attribution), VOC-019 (security), CFB-004/011
Stakeholder Need | STN-005 (audit-grade tamper-evident), STN-006 (attributable evidence), STN-013 (enterprise data governance), STN-025 (secure connectivity + signed OTA)
User Need | UN-SEC-001 (tamper-evident), UN-SEC-002 (attribution), UN-SEC-005 (identity/secure boot/OTA), UN-FUN-014 (complete record + gap flag), UN-SAFE-010 (fail-safe update)
Product Requirement | PRD-SEC-001 (tamper-evident), PRD-SEC-002 (attribution), PRD-SEC-004 (audit trail), PRD-SEC-008 (signed fail-safe OTA), PRD-FUN-017 (complete record + gap), PRD-FUN-022 (verified resumable non-destructive FW), PRD-REG-003 (GDP + Part 11 / ALCOA+)
System Requirement | SYS-026 (tamper-evident record), SYS-027 (export), SYS-028 (append-only audit trail), SYS-017 (honest incomplete-record), SYS-032 (signed fail-safe OTA), SYS-118 (trusted time), SYS-134 (device identity)
Subsystem Requirement | FWR-034 (append-only tamper-evident log), FWR-035 (power-fail-safe atomic writes), FWR-038 (explicit gap markers), FWR-045 (trusted time), FWR-046 (signature verify), FWR-048 (A/B rollback), FWR-049 (data preservation), HWR-016 (secure element), HWR-018 (debug-lock); SRS-SEC(600), SRS-OTA(700), SRS-DATA(900)
Implementing Architecture Element | FUN-030 (tamper-evident record), FUN-031 (append-only audit), FUN-018 (incomplete-monitoring qualification), FUN-037 (signed fail-safe OTA); CMP-041 (storage), CMP-044 (crypto), CMP-043 (OTA agent), CMP-054 (audit store), CMP-051 (OTA service)
Design Output | DO-SAFE-008 → DO-FW-010 + DO-EE-002 (record integrity/power-fail); DO-SAFE-006 (honest gap); DO-SAFE-014 → DO-FW-014/015 + DO-CLD-010 (secure OTA); DO-SEC-006 (secure boot/identity)
Linked Hazard | HA-CYBER-001 ★ (injected false record), HA-CYBER-005 ★ (malicious FW), HA-SW-013 ★ (reset gap reads uneventful), HA-DATA-004 ★ (incomplete record presented clean), HA-SW-010 (bad OTA), HA-MPE-004 (data-integrity at scale)
Linked Failure Mode | FA-SWFMEA-018/031 (record/integrity), FA-SWFMEA-026 (OTA), FA-IFMEA-015, FA-FTA TOP-D (corrupt record)
Risk-Control Requirement | RCR-034 (capture-time authenticity), RCR-024 (power-fail-safe writes), RCR-022 (honest gap qualifier), RCR-042 (export integrity), RCR-041 (provenance), RCR-025 (trusted time), RCR-037 (signed rollback-safe OTA), RCR-038 (secure boot/debug-lock)
Verification Evidence | PV-CYBER-001 (integrity pen-test, IV), PV-SW-013 (power-fail-during-write), PV-DATA-004/006 (gap/export), PV-SW-010 + PV-CYBER-005 (OTA rollback + pen-test, IV) — Evidence-Pending
Validation Evidence | VAL-USE-004 (never trust void/unmonitored trip), VAL-RISK (residual security) — Evidence-Pending
Manufacturing Control | CC-06 (signed FW + region config), CC-07 (unique crypto identity + debug-lock), SC-04 (debug-lock verify)
Release Baseline | REL-TBD; independent security verification (IV) required before residual acceptance (RA-05)
Field Monitoring | TREND-018/019 (OTA/cloud rollout — common cause), integrity-failure / signing-audit events, security monitoring (FM-100/200)
Overall Status | Verification Planned — security controls credited only after independent adversarial verification (IV), which has not run
Gap Summary | Strong design (device-anchored authenticity, non-suppressible alarm, secure OTA/boot, never-brick invariant). Effectiveness pending IV; no cyber threat can be marked "mitigated" while unverified. P1 (safety-through-security).
CHAIN-7 — Reusable lifecycle: cleaning / return / re-provisioning (D1)
Column | Value
Chain ID | CHAIN-7
Market Requirement | MR-PROB-011 (reuse only pays if returned), MR-JTBD-007/009, MR-1700 (sustainability), D1
Voice of Customer | VOC-006, VOC-016/017, FLD-007, CFB-012, LD-008
Stakeholder Need | STN-010 (reusable readiness: charge/cal/firmware), STN-016 (recover/return fleet), STN-021 (calibration currency), STN-026 (EOL/battery disposal)
User Need | UN-SVC-001 (refurbishment loop), UN-PHY-008 (cleaning), UN-FUN-017 (readiness reporting), UN-VIS-006 (reusable multi-trip), UN-REG-008 (EOL/battery)
Product Requirement | PRD-SVC-001 (refurbishment turnaround loop), PRD-SVC-002 (cleaning without damage), PRD-SVC-003 (predictable recharge/readiness), PRD-SVC-004 (readiness triage + deploy-block), PRD-FUN-020 (readiness reporting + deploy-block), PRD-VIS-005 (reusable), PRD-REG-008 (RoHS/WEEE/EoL)
System Requirement | SYS-024 (config/re-provision), SYS-025 (pre-arm readiness: charge/cal/FW/profile/probe), SYS (readiness reporting + deploy-block, fleet)
Subsystem Requirement | FWR-031 (arming state machine w/ arm/void), FWR-032 (pre-arm readiness gate), FWR-039 (controlled factory reset w/ record protection), MER (cleaning/materials, MER-001…040), FWR-049 (data preservation across update); SRS-FLEET(800)
Implementing Architecture Element | FUN-025 (trip-profile authoring), FUN-026 (config validation), FUN-027 (secure provisioning), FUN-028 (pre-arm checks), FUN-035 (readiness reporting), FUN-036 (fleet visibility & stray recovery); CMP-047 (device registry), CMP-050 (fleet dashboard)
Design Output | DO-SVC (service/refurb outputs, DOL-1700), DO-FW-009 (calibration/readiness), DO-MFG (re-provisioning), DO-LBL (disposal), DO-SEC (key lifecycle)
Linked Hazard | HA-SVC-001 ★ (depot recal escape), HA-SVC-002 (wrong firmware), HA-SVC-006 (debug mode left on), HA-SVC-007 (Li disposal), HA-USE-004 ★ (unmonitored trip trusted after re-provision)
Linked Failure Mode | FA-SYSFMEA-009 (service/recal), FA-SWFMEA-026 (service restore), FA-SWFMEA-034 (debug mode)
Risk-Control Requirement | RCR-045 (pre-arm gate), RCR-049 (cal interval enforcement), RCR-050 (service-restore/post-service BIST), RCR-051 (service-mode lock), RCR-052 (EOL disposal), RCR-056 (correct config)
Verification Evidence | PV-SVC-001 (interval enforcement), PV-SVC (post-service BIST), PV-CYBER-006 (mode-lock, IV), disposal/key-erasure review — Evidence-Pending
Validation Evidence | VAL-USE-004 (never trust void), VAL-USE-017 / VAL-REL-008 (reconditioning), VAL-BIZ (reuse TCO/cost-per-trip) — Evidence-Pending
Manufacturing Control | SC-06 (serial/identity match), SC-07 (correct-config/line-clearance); Device History Record (MDT-900) for serial attribution across reuse
Release Baseline | REL-TBD; reuse-cycle life (PRD-REL-002) validate-by-test outstanding
Field Monitoring | Return/RMA rate, recal-escape audit, firmware-version audit, take-back/recycling compliance (FM-200/500)
Overall Status | Draft/Approved (design-input) — Verification/Validation Planned
Gap Summary | Reuse loop is architected (return→clean→recharge→triage→re-provision, ADR-009) with pre-arm gate + deploy-block as the safety net. Cycle-life over reuse and cleaning-chemical compatibility are Evidence-Pending; recal-escape is a monitored common cause. P2/P3.
---
Compact summary table (all 7 flagship chains)
Chain | Capability | Top Hazard | Key RCR(s) | Headline V&V | Mfg CC/SC | Status | Gap priority
CHAIN-1 | Excursion detect→alert→disposition | HA-SW-001 ★ | RCR-019/021/028/029/043 | PV-SW-001 / VAL-USE-001 (EP) | SC-03 | V-Planned / Val-Planned | P0 (V&V)
CHAIN-2 | Li-battery / charge-inhibit <0 °C | HA-ENRG-001 ★, HA-TFC-002 ★★ | RCR-001/002/004/007/008 | PV-ENRG-001 + UN 38.3 (EP) | CC-01/02/03/04 | V-Planned (Open) | P0 (D3+AML)
CHAIN-3 | 90-day −30 °C endurance | HA-ENRG-006 ★ | RCR-006/017/024 | PRD-VER-004 (EP, Open) | CC-03 | OPEN | P0
CHAIN-4 | RF-in-container alert egress | HA-COMM-003 ★ | RCR-031/030 | PRD-VER-005 (EP, Open) | SC-07 | OPEN | P0
CHAIN-5 | Temp accuracy & calibration integrity | HA-SW-008 ★, HA-DATA-001 ★ | RCR-020/026/045/049/053 | PV-SW-004/008 (EP) | CC-05 | V-Planned | P0/P1 (WHO PQS)
CHAIN-6 | Data authenticity / audit / secure OTA | HA-CYBER-001/005 ★, HA-DATA-004 ★ | RCR-034/037/041/042/025 | PV-CYBER-001/005 (IV, EP) | CC-06/07 | V-Planned (IV) | P1
CHAIN-7 | Reusable lifecycle (D1) | HA-SVC-001 ★, HA-USE-004 ★ | RCR-045/049/050/052/056 | PV-SVC-001 (EP) | SC-06/07 | V-Planned | P2/P3
Matrix-01 gap findings: every flagship chain is structurally complete (no broken links); all seven are blocked on unexecuted V&V; three carry a hard Open (CHAIN-2/3 D3, CHAIN-4 RF); two are blocked by the AML freeze (CHAIN-2/3 chemistry, CHAIN-4 antenna); one by the WHO PQS accuracy target (CHAIN-5). No chain reaches "Released."

## T4 Hazards & Controls
T4 Hazards & Controls
Part of the end-to-end traceability record.
↩ Contents
Matrix-04 — Hazard & Risk-Control Traceability — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): Hazard Analysis ID · Hazard · Initiating Event · Foreseeable Sequence · Hazardous Situation · Harm · Initial Severity/Probability/Risk · Risk-Control Requirement · Design Implementation · Verification Evidence · Validation Evidence · Residual Severity/Probability/Risk · New Risks Introduced · Acceptance Authority · Field Monitoring · Status. Rows are the release-relevant top hazards from the Stage-10 Hazard-Traceability-Matrix + Stage-12 Risk-Control- Traceability. All residuals and PV-/VAL- are Evidence-Pending; severity-first per CLAUDE.md rule 3. Regions: R unacceptable · A ALARP (A-High for S5) · B broadly acceptable.
---
HA ID | Hazard | Initiating event | Foreseeable sequence | Hazardous situation | Harm (Sev) | Initial Risk | RCR | Design Impl (DO-) | Verification (PV-) | Validation (VAL-) | Residual | New risk | Accept auth
HA-TFC-002 ★★ | Li-cell fire in occupied aircraft hold | Cell defect/abuse in flight | Internal short → runaway → fire in cargo hold | Fire in occupied aircraft | Death / loss of aircraft (S5) | R High | RCR-002/004/007/008 | DO-SAFE-002/010; DO-MECH-005+DO-PCB-010 | PV-TFC-002 (containment) + UN 38.3 | — | A-High (S5) permanent | Containment mass/venting (HA-NEW) | Chief Systems + Quality/Reg (S5)
HA-ENRG-001 ★ | Cold-charge plating → latent short | Charge attempted <0 °C | Plating → dendrite → latent internal short → thermal event | Cold cell charged | Thermal runaway/fire (S4/S5) | R High | RCR-001/002 | DO-SAFE-001 → DO-EE-007 + DO-FW-005 | PV-ENRG-001 (cold-charge FI) + UN 38.3 | — | A Med (Open D3) | Charge-inhibit → DOA (HA-NEW; RCR-046) | System-Safety
HA-ENRG-006 ★ | Dead-in-cold silent gap | Energy exhausted at −30 °C | Sleep too shallow / capacity fade → power loss → gap | Energy exhausted, reads uneventful | Wrong disposition (S4) | R High | RCR-006/024/027 | DO-SAFE-012 → DO-FW-004 + DO-EE-008 | PRD-VER-004 (90-day cold endurance) | — | A Med (Open margin) | Safety-vs-availability (HA-NEW-006) | Reliability + System-Safety
HA-SW-001 ★ | Missed excursion | Detection logic gap / freeze under-sensitivity | Excursion not flagged → believed acceptable | Spoiled dose believed acceptable | Compromised dose administered (S4/S5) | R High | RCR-019 | DO-SAFE-004 → DO-FW-006 + DO-EE-013 | PV-SW-001 (profiled excursions, FI) | VAL-SW-001 | A Med | — | System-Safety
HA-SW-002 ★ | False accept on incomplete record | Gap in record + permissive verdict | Lot presented acceptable on incomplete data | Lot accepted on incomplete record | False accept → compromised dose (S4) | R High | RCR-021/022 | DO-SAFE-006 → DO-FW-008 + DO-CLD-006 | PV-SW-002 (gap-handling FI) | VAL-SW-002 | A Med | — | System-Safety
HA-SW-008 ★ | Reports temps not measured | Probe fault / open channel | Plausible-looking but false temps → false accept | Device reports valid-looking temps it isn't measuring | False accept → compromised dose (S4) | R High | RCR-020/045 | DO-SAFE-005 → DO-FW-007 + DO-EE-013 | PV-SW-008 (probe-fault-injection) | — | A Med (+placement) | — | System-Safety
HA-SW-013 ★ | Reset/brown-out gap reads clean | Cold sag / brown-out during NV write | Torn record + no gap marker → false reassurance | Reset gap reads as uneventful | False reassurance → compromised dose (S4) | R High | RCR-024/022 | DO-SAFE-008 → DO-FW-010 + DO-EE-002 | PV-SW-013 (power-fail-during-write) | — | A Med | — | System-Safety
HA-COMM-001 ★ | No/late alert in flight | Connectivity absent when excursion detected | Alert queued, delivered late / not before disposition | Shipment held/administered before late alert | Compromised dose (S4) | R High | RCR-028/029/030 | DO-SAFE-009 → DO-FW-011 + DO-CLD-007/008 | PV-COMM-001 (latency+S&F) | VAL-COMM-001 | A Med | Alarm fatigue (HA-USE-002) | System-Safety
HA-COMM-003 ★ | RF-in-container blocks alert | Metallized shipper blocks radio | Excursion, no live notification, delivered on unpack | Excursion with no live notification | Compromised dose (S4) | R High | RCR-031/030 | DO-IF (antenna) + DO-EE; DO-SAFE-009 | PRD-VER-005 (RF-in-container) + PV-COMM-003 | VAL-COMM-003 | A Med (Open RF) | — | System-Safety
HA-CYBER-001 ★ | Injected false "clean" record | Data-integrity attack | Falsified record trusted → wrong disposition | False "clean" record from injected data | Compromised dose (S4) | R High | RCR-034 | DO-SEC; DO-SAFE-008 | PV-CYBER-001 (integrity/pen-test, IV) | — | A Med | Encryption complexity (HA-NEW) | Product Security + System-Safety
HA-CYBER-002 ★ | Muted/blocked alarm | Alarm-suppression attack | Real excursion hidden → not acted on | Excursion hidden by suppressed alarm | Compromised dose (S4) | R High | RCR-035/029 | DO-SEC; DO-SAFE-009 | PV-CYBER-002 (suppression pen-test, IV) | — | A Med | — | Product Security
HA-CYBER-005 ★ | Malicious firmware | Supply-chain / OTA compromise | Unsigned/altered FW → fleet unsafe/blind | Malicious firmware deployed | Fleet-scale harm (S4/S5) | R High | RCR-037/038 | DO-SAFE-014 → DO-FW-014/015; DO-SEC-006 | PV-CYBER-005 (OTA pen-test, IV) | — | A Med/High | — | Product Security + Chief Systems
HA-DATA-001 ★ | Decision on unrepresentative numbers | Cal drift / probe placement | Numbers not representing product zone trusted | Decision on numbers not representing product | Wrong disposition (S4/S5) | R High | RCR-019/020/026/034 | DO-SAFE-004/005; DO-SEC | PV-DATA-001 (cal-coverage+probe FI) | VAL-USE-001 | A Med | — | System-Safety
HA-DATA-004 ★ | Incomplete record presented as clean | Gap not surfaced | Operator sees "pass" on incomplete data | Incomplete record presented as clean pass | False accept (S4) | R High | RCR-040/021/022 | DO-SAFE-006 + DO-HF-002 | PV-DATA-004 (gap-exposure) | VAL-USE-004 | A Med | — | System-Safety + Human Factors
HA-USE-001 ★ | Verdict misread | Cold/gloved/low-light/language | Wrong disposition acted on | Verdict misread | Wrong disposition (S4) | R High | RCR-043 | DO-SAFE-013 + DO-HF-002/005 | PV-USE-001 (cold-display) | VAL-USE-001 | A Med | — | Human Factors + System-Safety
HA-USE-002 ★ | Alarm muted (fatigue) | Excess alerts → habituation | Real excursion alert ignored | Real excursion alert muted | Not acted on → compromised dose (S4) | R High | RCR-047/029 | DO-SAFE-004; DO-CLD-003 | PV-USE-002 (alarm-mgmt) | VAL-USE-002 | A Med | False-reject trade (HA-NEW-006) | Human Factors
HA-USE-004 ★ | Unmonitored trip trusted | Forgot-to-arm / void | Void trip believed monitored | Unmonitored trip trusted | False reassurance (S4) | R High | RCR-045/040 | DO-SAFE-007 / DO-FW-032 | PV-SW-003 (pre-arm/void) | VAL-USE-004 | A Med | — | System-Safety + Human Factors
HA-USE-007 ★ | Reject overridden | Automation bias / stock-out fear | Compromised dose administered despite reject | Reject overridden | Compromised dose (S4) | R High | RCR-044/040 | DO-HF; DO-CLD (audit) | PV-DATA-008 (recommendation) | VAL-USE-007 | A Med | — | Human Factors + System-Safety
HA-SVC-001 ★ | Depot recal escape | Missed/incorrect recalibration | Systematic bias across reused fleet | Depot recal escape | Wrong disposition (S4) | R High | RCR-049/026/050 | DO-FW; DO-SVC | PV-SVC-001 (interval enforcement) | — | A Med (CCF) | — | Quality + System-Safety
HA-MFG-001 ★ | EOL miscalibration passes | Test escape at end-of-line | Miscalibrated unit ships → systematic bias | Miscalibrated unit passes EOL | Systematic bias → wrong disposition (S4) | R High | RCR-053 | DO-TST; DO-MFG | PV-MFG-001 (EOL cal coverage) | — | A Med | — | Quality
HA-MPE-002 ★★ | Mass administration of compromised doses | Aggregate failure of the pathway | Fleet-scale record failure → population dosing | Mass administration of compromised doses | Population outbreak (S5) | R High | RCR-057 (aggregate) | all DO-SAFE record-integrity spine | (aggregate of spine PVs) | (aggregate VAL-SW/USE) | A-High (S5) | Systemic dependence (HAL-2900) | Chief Systems + Quality/Reg
HA-MPE-003 ★ | False-reject stock-out | Over-conservative reject / wastage | Stock-out → unvaccinated population | False-reject stock-out | Preventable disease (S4) | R High | RCR-047/044/046 | DO-SAFE-004; DO-CLD-003 | VAL-USE-002 (false-reject targets) | VAL-USE-002 | A Med (accepted trade) | Balanced vs missed-excursion (HA-NEW-006) | Chief Systems + Quality/Reg
Gaps (Matrix-04)
Hazards without controls | none — every top hazard maps to ≥1 RCR (Stage-10 matrix §3; Stage-12 §4). No R-region hazard is uncontrolled.
Controls without requirement/implementation/verification | none — every RCR carries a DO- implementation and a PV- (and, where use-related, a VAL-) target (Stage-12 §4 "no control on paper only").
Residual risks without evidence | all 22 rows — every residual is projected / Evidence-Pending until Stages 13–14 execute. This is the dominant Matrix-04 gap (P0/P1).
High-severity accepted without authority | none — the two S5 rows (HA-TFC-002, HA-MPE-002) carry a named acceptance authority (Chief Systems + Quality/Regulatory) and mandatory monitoring; residual stays region-A-High by rule.
Controls introducing unassessed risk | none unassessed — each is carried in HAL-2700 (e.g. charge-inhibit → DOA balanced by RCR-046; safety-vs-availability HA-NEW-006; false-reject/wastage trade). See G.
Carried Open (release-gating) | HA-ENRG-006 (D3 → PRD-VER-004) and HA-COMM-003 (RF-in-container → PRD-VER-005). Security rows require independent adversarial (IV) verification before residual acceptance.

## T7 Verification Coverage
T7 Verification Coverage
Part of the end-to-end traceability record.
↩ Contents
Matrix-07 — Verification Coverage — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): Requirement ID · Type · Text · Verification Method · Level · Protocol · Test Case · Configuration · Acceptance Criteria · Result · Evidence · Defect · Deviation · Retest · Final Status · Reviewer · Approval · Release Applicability. Source: Stage-13 PVL (195 activities) + Verification-Traceability-Coverage-Matrix. This is a pre-execution playbook: every Result is "Not Run", every Evidence cell is Evidence-Pending, every Final Status is Verification-Planned. Fault injection (FI) for safety controls; independent adversarial (IV) for security.
---
Req/RCR | Type | Text (short) | Method | Level | Protocol | Test case | Cfg | Acceptance criteria | Result | Evidence | Defect | Dev | Retest
RCR-019/SYS-009 | Safety | Two-sided detection | Test (FI) | System | PV-SW-1100 | PV-SW-001 | SCM-100 | Freeze detection = heat; profiled excursions flagged | Not Run | Evidence-Pending | — | — | —
RCR-021/SYS-015-16 | Safety | Verdict suppresses false accept | Test (FI) | System | PV-SW-1100 | PV-SW-002 | SCM-100 | No unqualified accept on gap; device=cloud | Not Run | Evidence-Pending | — | — | —
RCR-028/029/SYS-011 | Safety | Alert gen + escalation | Test (FI) | System | PV-FUNC-500 | PV-SW-006 / PV-COMM-001 | SCM-100 | Alert within latency; ack/escalate; S&F on reconnect | Not Run | Evidence-Pending | — | — | —
RCR-001/HWR-005 | Safety | HW charge-inhibit <0 °C | Test (FI) | Subsystem | PV-SAFE-900 | PV-ENRG-001 | SCM-100 | No charge below limit under fault injection | Not Run | Evidence-Pending | — | — | —
RCR-002/HWR-003 | Safety | Battery protection + backstop | Test (FI) | Subsystem | PV-SAFE-900 | PV-ENRG-002 (DO-TST-009) | SCM-100 | OV/OC/OT/short trip; secondary fuse backstop | Not Run | Evidence-Pending | — | — | —
RCR-008/HA-TFC-002 | Safety (S5) | Containment / fire-enclosure | Test | System | PV-SAFE-900 | PV-TFC-002 + UN 38.3 | SCM-100 | Containment holds; UN 38.3 T.1–T.8 pass | Not Run | Evidence-Pending | — | — | —
PRD-VER-004 (RCR-006/017) | Reliability/Safety | 90-day cold endurance | Test | System | PV-REL-1400 | PV-REL-001 + PV-ENRG-001 | SCM-100 | Survive 90 days at −30 °C w/ reserve; no silent gap | Not Run | Evidence-Pending (OPEN) | — | — | —
PRD-VER-005 (RCR-031) | Reliability/Conn | RF-in-container link | Test | System | PV-REL-1400 | PV-REL-005 + PV-COMM-003 | SCM-100 / -XA | Link margin sufficient in metallized shipper | Not Run | Evidence-Pending (OPEN) | — | — | —
RCR-026/FWR-018 | Safety | Calibration integrity + pre-arm | Test (FI) | System | PV-SW-1100 | PV-SW-004 | SCM-100 | Expired/invalid cal blocks arm; coefficient correct | Not Run | Evidence-Pending | — | — | —
RCR-020/HWR-028 | Safety | Probe-fault detection | Test (FI) | Subsystem | PV-SW-1100 | PV-SW-008 | SCM-100 | Open/air probe flagged; no false valid temp | Not Run | Evidence-Pending | — | — | —
RCR-024/FWR-035 | Safety | Power-fail-safe writes | Test (FI) | Subsystem | PV-SW-1100 | PV-SW-013 | SCM-100 | No torn record; gap marked on brown-out | Not Run | Evidence-Pending | — | — | —
RCR-034/PRD-SEC-001 | Security→Safety | Capture-time authenticity | Test (IV) | System | PV-SEC-1000 | PV-CYBER-001 | SCM-100 | Injected data detected/rejected (independent) | Not Run | Evidence-Pending | — | — | —
RCR-037/FWR-046-48 | Security→Safety | Signed OTA + rollback | Test (IV) | System | PV-SEC-1000 | PV-CYBER-005 + PV-SW-010 | SCM-100 | Unsigned rejected; rollback on fail; never bricks | Not Run | Evidence-Pending | — | — | —
RCR-014/HWR-042 | Safety | EMC integrity of decision path | Test | System | PV-EMC-800 | PV-ELEC-004 | SCM-100 | Immunity: no false/missed excursion; emissions pass | Not Run | Evidence-Pending | — | — | —
PRD-ENV-001/SYS-069 | Environmental | −30…+60 °C full function | Test | System | PV-ENV-700 | PV-ENV (thermal cycling) | SCM-100 | Full function across range; condensation/IP67 | Not Run | Evidence-Pending | — | — | —
RCR-053/HA-MFG-001 | Safety | EOL calibration coverage | Test | Production | PV-MFG-1300 | PV-MFG-001 | SCM-100 | 100% EOL cal within accuracy; MSA-qualified | Not Run | Evidence-Pending | — | — | —
PRD-REG-006/HWR-011 | Compliance | UN 38.3 Li transport | Test | Component | PV-COMP-1500 | UN 38.3 T.1–T.8 | SCM-100 | All UN 38.3 sub-tests pass | Not Run | Evidence-Pending | — | — | —
PRD-REG-001 | Compliance | WHO PQS accuracy target | Test/Analysis | System | PV-COMP-1500 | PV-COMP (WHO PQS) | SCM-100 | Meet WHO PQS category accuracy (target EP) | Not Run | Evidence-Pending | — | — | —
Coverage summary (Stage-13)
Metric | Value | Note
RCRs with ≥1 PV activity | 58 / 58 (100%) | No control on paper only (PVL coverage matrix §2)
Safety-critical controls with fault-injection | all | PV-ENRG/PV-TFC/PV-SW FI set
Security controls with independent adversarial (IV) | all (6 + RCR-051) | PV-CYBER-
Applicable PRD/SYS/SRS/HWR/MER/FWR/IFR mapped to a PV module | all applicable | matrix §3; no unverified applicable requirement
Verification activities executed | 0 / 195 | Pre-execution playbook
Evidence at E4–E7 | none | Every result Not Run / Evidence-Pending
Gaps (Matrix-07)
Unverified requirements | structurally none (100% RCR + applicable-requirement coverage) — but evidentially all 195 activities are unexecuted (P0). The coverage is a plan, not a proof.
Tests without requirements | none — every PV traces to a requirement/RCR (PVL is requirement-driven).
Wrong-level verification | none flagged (levels assigned component/subsystem/system/production).
Missing acceptance criteria | none structurally — criteria are objective/measurable; several numeric limits are Evidence-Pending with a named closing test (WHO PQS accuracy, RF margin, endurance).
Failed tests without disposition | N/A — no test has run, so no failures/dispositions exist yet.
Passed tests on wrong configuration | N/A (nothing passed).
Risk-control tests without fault injection | none — FI is present for every safety-critical control.
Two Open gating items | PRD-VER-004 (D3) and PRD-VER-005 (RF-in-container) — release-gating.

## T8 Validation Coverage
T8 Validation Coverage
Part of the end-to-end traceability record.
↩ Contents
Matrix-08 — Validation Coverage — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): User Need · Intended Use · User Group · Use Environment · Workflow · Validation Scenario · Task · Success Criterion · Product Configuration · Participant Group · Result · Finding · Use Error · Linked Hazard · Residual Risk Impact · Corrective Action · Final Status. Source: Stage-14 PVaL (128 activities) + Validation-Coverage-Matrix; HA-USE-001…009 map 1:1 to VAL-USE-001…009. Pre-execution: no product built, no users tested — every Result/Finding is Evidence-Pending, every Final Status is Validation-Planned. Validation traces to User Needs / Intended Use and residual-risk acceptability, never to design outputs.
---
User Need | Intended use | User group | Environment | Workflow | Validation scenario | Task | Success criterion | Cfg | Participants | Result | Finding | Use error | Linked hazard
UN-FUN-008/UN-SAFE-005 | Offline verdict at the door | Last-mile health worker (minimally trained) | LMIC clinic, cold, gloved, low-light | Receipt check | VAL-USE-001 | Read verdict → correct accept/reject | Correct disposition rate meets target (EP) | SCM-100 | Representative last-mile users | Evidence-Pending | — | (HA-USE-001) | HA-SW-001/HA-USE-001 ★
UN-SAFE-004 | Act on real excursion alert | Distributor/3PL responder | Dispatch/transit | Alert response | VAL-USE-002 | Respond to alert w/o fatigue miss | Real alerts acted on; false-reject within target (EP) | SCM-100 | Responders | Evidence-Pending | — | (HA-USE-002) | HA-USE-002 ★/HA-MPE-003
UN-SAFE-003 | Never trust a void/unmonitored trip | EPI/clinic staff | Clinic | Disposition | VAL-USE-004 | Recognize incomplete monitoring | No unqualified accept on void/gap | SCM-100 | EPI + last-mile | Evidence-Pending | — | (HA-USE-004) | HA-DATA-004 ★/HA-USE-004 ★
UN-FUN-003 (placement) | Represent product zone | Installer/operator | Pack-out dock | Probe placement | VAL-USE-005 | Place probe representatively | Probe in product zone; placement error < target | SCM-100 | Installers | Evidence-Pending | — | (HA-USE-005) | HA-SW-008 ★/HA-DATA-001 ★
UN-STK/UN-SAFE-006 | Arm correctly for dispatch | Dispatch operator | Dispatch dock | Provisioning/arming | VAL-USE-003 / VAL-USE-006 | Provision + arm; pre-arm gate | Unmissable arm; deploy-block on unready | SCM-100 | Operators | Evidence-Pending | — | (HA-USE-003) | HA-SW-003/HA-USE-004
UN-SAFE-001 (override) | Do not override a reject unsafely | Clinic decision-maker | Clinic | Override | VAL-USE-007 | Recorded override w/ rationale | Override rate + reasons within target (EP) | SCM-100 | Decision-makers | Evidence-Pending | — | (HA-USE-007) | HA-USE-007 ★
UN-FUN-006 (detection operational) | Detect real excursions in use | Operational fleet | Real lanes | Monitoring | VAL-SW-001 (=VAL-CLIN-002) | Detection effectiveness on lanes | Excursion detection sensitivity/specificity target (EP) | SCM-100 | Field pilot | Evidence-Pending | — | — | HA-SW-001 ★
UN-FUN-007 (gap honesty) | Present incomplete data honestly | QA/regulatory | Audit | Disposition/audit | VAL-SW-002 | Gap surfaced, not "clean" | No false clean; auditor accepts as defensible | SCM-100 | QA + auditors | Evidence-Pending | — | — | HA-SW-002 ★/HA-DATA-004 ★
UN-PERF-008/UN-FUN-009 | Alert delivery / connectivity | Fleet + responders | Multi-lane (incl. RF-hard) | Alert delivery | VAL-COMM-001 / VAL-COMM-003 | Alert egress incl. in-container | Delivery-success per lane meets target (EP, Open) | SCM-100/-XA | Field pilot | Evidence-Pending | — | — | HA-COMM-001/003 ★
UN-PERF-003 | 90-day cold endurance in use | Fleet | −30 °C lanes | Long-term monitoring | VAL-REL-700 | In-use endurance + calibration | Survive service life; endurance holds (EP, Open) | SCM-100 | Field pilot | Evidence-Pending | — | — | HA-ENRG-006 ★
UN-SVC-001 (reuse) | Refurbish and redeploy | Service/reverse-logistics | Depot | Reconditioning | VAL-USE-017 / VAL-REL-008 | Return→clean→recharge→triage→re-provision | Reconditioning yields ready device; no residual risk | SCM-100 | Depot ops | Evidence-Pending | — | (HA-SVC) | HA-SVC-001 ★
UN-SAFE (residual) | Residual risk acceptable to users | All safety-critical users | All | Benefit-risk | VAL-RISK-001/003 | Validate HAL-2800/2900 residuals | 7 HAL-2900 conditions acceptable to users | SCM-100 | Cross-user + reviewers | Evidence-Pending | — | — | all top hazards
UN-VIS/UN-BIZ | Reuse TCO / operational benefit | Executive/procurement | Fleet | Business | VAL-BIZ-800 / VAL-CLIN | Wastage reduction; cost-per-trip | Benefit target met (EP) | SCM-100 | Buyers/operators | Evidence-Pending | — | — | HA-MPE-001/003
Coverage summary (Stage-14)
Metric | Value | Note
Use errors (HA-USE-001…012) with a validation reference | all | VAL-USE-001…009; VAL-USE-007 also closes HA-USE-012; HA-USE-010/011 process-controlled (VAL-USE-017/VAL-REL-008)
Top residuals (HA-RES-001…022) with a validation reference | all | VAL-RISK / VAL-USE / VAL-HF / VAL-REL / VAL-ECO
Safety-critical workflows with a summative activity | all | disposition read, alert response, provisioning/arming, override, probe placement, charge-inhibit interpretation
User classes SH-01…SH-15 with a validation activity | all | beneficiary SH-11 represented
Validation activities executed | 0 / 128 | No product built, no users tested
Gaps (Matrix-08)
User needs without validation | none — every user need / intended-use claim has a representative validation scenario (Stage-14 coverage summary); the standard's completeness rule (every user need → ≥1 validation) is met structurally.
Requirements treated as validation substitutes | none — validation is outcome-based, traced to User Needs / intended use, never to design outputs (Stage-14 §philosophy). Verification (Matrix-07) is kept distinct.
Unrepresented users/environments | none named — last-mile, EPI, operators, sustainment, governance, enterprise, and the beneficiary are all covered; cold/gloved/low-light/offline environments included.
Safety-critical workflows not validated | none structurally; all Evidence-Pending (P0).
Unresolved use errors / findings not fed back | N/A yet — no findings exist (pre-execution); the closed-loop feedback path exists (Stage-14 → Stage-17).
Two Open gating validations | VAL-COMM-003 (RF-in-container) and VAL-REL-700 (D3) inherit the two Open verification items. VAL-RISK-003 gates unconditional release against the seven HAL-2900 conditions.

## T9 Manufacturing
T9 Manufacturing
Part of the end-to-end traceability record.
↩ Contents
Matrix-09 — Manufacturing & Design Transfer — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): Design Output · Product Characteristic · Critical/Special Characteristic · Manufacturing Process · Process Step · PFMEA Record · Process Control · Control Plan · Work Instruction · Inspection Method · Measurement System · Production Test · Process Validation · Supplier Control · Capability Evidence · Release Status · Field Monitoring. Source: Stage-15 (MDT-100…1000). The preservation principle: the safety/measurement/record controls must be built into and proven on every unit. Manufacturing readiness is NOT READY for volume (numeric limits/MSA/Cpk are Evidence-Pending; AML not frozen).
---
Design Output | Product characteristic | CC/SC | Process | Step | PFMEA | Process control | Control plan | Work instr | Inspection | MSA | Prod test | Process val | Supplier
DO-EE-007 / RCR-001 | Charge-inhibit <0 °C active | CC-01 | Battery + charge assembly | STEP (battery) | PFM (charge) | Zero-defect line | CP-01 | Safety-critical WI | Functional | (test-system MSA) | PT (charge-inhibit) → PV-MFG | Process val pending | SUP (PMIC)
DO-EE-005/010 / RCR-002 | Protection + backstop present/functional | CC-02 | PCBA + battery | STEP (protection) | PFM (protection) | Zero-defect | CP-02 | Safety-critical WI | AOI + functional | MSA (test) | PT (protection trip) | pending | SUP (protection IC, fuse) enhanced
DO-EE-003 / RCR-054 | Correct chemistry/protection parts | CC-03 | Incoming + assembly | STEP (incoming) | PFM (wrong part) | Poka-yoke + AOI | CP-03 | WI | Incoming inspection | — | PT (identity) | pending | SUP (cell) enhanced
DO-MECH-012 / RCR-004 | Cell retention / connection | CC-04 | Mechanical assembly | STEP (retention) | PFM (retention) | Torque control (CPP) | CP-04 (SC-09 torque) | WI | Torque verify | Torque MSA | PT (retention) + UN 38.3 | pending | SUP (hardware)
DO-FW-007 / RCR-026/053 | Temp-sensor calibration accuracy | CC-05 | Calibration | STEP (EOL cal) | PFM (cal escape) | 100% cal + master challenge | CP-05 | Safety-critical WI | EOL cal test | MSA-qualified, master-challenged (MDT-800) | PT (EOL cal) → PV-MFG-001 | CAL-### plan | SUP (reference)
DO-FW-014/015 / RCR-037 | Signed FW + region config | CC-06 | FW flash + config | STEP (flash) | PFM (wrong FW/config) | Signed-image + line-clearance | CP-06 (SC-07 config) | WI | FW/version verify | — | PT (FW/config verify) | pending | SUP (image control)
DO-MFG / DO-SEC-006 / RCR-055 | Unique crypto identity + debug-lock | CC-07 | Provisioning | STEP (identity) | PFM (dup identity) | Uniqueness enforcement | CP-07 (SC-04 debug-lock) | WI | Identity/debug verify | — | PT (identity uniqueness) → PV-MFG-005 | pending | SUP (SE)
DO-LBL / DO-REG-006 / RCR-007 | Regulatory/UDI/Li-transport marking | CC-08 | Labeling | STEP (label) | PFM (labeling) | Print-verify | CP-08 | WI | Label inspection | — | PT (label/UDI) | pending | SUP (label)
DO-MECH-005 / RCR-011/015 | Seal / IP integrity | SC-01 | Seal / IP | STEP (seal) | PFM (seal/IP) | Leak-rate control | CP (SC-01) | WI | Leak/IP test | Leak MSA | PT (IP) | pending | SUP (seal)
DO-PCB-010 | Conformal coat coverage | SC-02 | Coating | STEP (coat) | PFM (coat) | Coverage control | CP (SC-02) | WI | Coverage inspection | — | PT (coverage) | pending | SUP (coating)
DO-FW-006/008 (functional) | Final functional / system | SC-03 | Final test | STEP (FCT) | PFM (functional) | Acceptance-spine FCT | CP (SC-03) | WI | Functional | Test MSA | PT (final functional) | pending | —
DO-TST (AOI/ICT) / RCR-016 | PCBA safety-net / coverage | SC-05 | SMT/reflow + ICT/FCT | STEP (AOI/ICT) | PFM (latent solder) | AOI + ICT coverage | CP (SC-05, SC-08 reflow CPP) | WI | AOI/ICT | — | PT (coverage) → PV-MFG-008 | pending | SUP (PCB/PCBA)
Gaps (Matrix-09)
Safety-critical characteristics absent from PFMEA | none — the S5 battery-fire spine, measurement integrity, and record integrity all map to PFM-### records (MDT-200, 28 modes) and to CC-01…08 / SC-01…09.
Critical characteristics absent from control plan | none — every CC/SC has a Control-Plan line (MDT-300), a Work Instruction (MDT-500), and a production test (MDT-600).
Controls without measurement capability | this is the dominant Matrix-09 gap — the detection systems are specified as MSA-qualified/master-challenged, but no Cpk / gauge-R&R / leak-rate / accuracy value is established (all Evidence-Pending against released DOs and the MSA plan — RA-09).
Outputs not transferred to instructions | none — the design-transfer package (DO-XFER-) covers the transfer; however transfer is incomplete because DO-SAFE outputs are Preliminary (RA-05) and AML is not frozen (RA-07).
Production tests not linked to requirements | none — each PT ties to a requirement/RCR via PV-MFG (PVL-1300) and DO-TST.
Supplier controls not linked to risk | none — enhanced control on cell/BMS/SE/fuse (SUP-###, MDT-700) is tied to HA-MFG-002/003 and RCR-054.
Overall | manufacturing is transfer-mature in design but NOT READY for volume — AML freeze (RA-07), numeric limits/MSA (RA-09), regulatory/UDI (RA-10), pilot build, and HAL open-risk reconciliation on executed V&V (RA-05) are open. Carried into Matrix-10.

## T10 Release Readiness
T10 Release Readiness
Part of the end-to-end traceability record.
↩ Contents
Matrix-10 — Release Readiness — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): Product Configuration · Requirement/Design-Output/ Risk-Analysis Baseline · Verification Status · Validation Status · Manufacturing/Supplier/Regulatory/ Cybersecurity Readiness · Open Defects/Deviations/Risks · Residual-Risk Approval · Release Decision · Release Authority · Conditions of Release. Source: Stage-16 (REL-100…400, RA-01…12) + Stage-10 Risk-Management Summary + Stage-15 MDT-1000. Headline decision (must match Stage 16): CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED (release pending V&V + blocker closure).
---
10.1 Readiness by dimension
Dimension | Baseline / status | Readiness | Evidence / gap
Product configuration | SCM-100 (base); SCM-100-XA (Candidate); regional configs | Not frozen | Baseline is REL-TBD; frozen only at DR-REL gate
Requirement baseline | 208 MR / 187 UN / 220 PRD / 190 SYS / 488 subsystem+SRS | Approved (design-input) | Complete and traced; PRD-SAFE re-baselined w/ RCRs
Design-output baseline | 272 DOL (103 Critical) | Draft (pre-DR-DO) | All TBD-rev; DO-SAFE Preliminary (RA-05)
Risk-analysis baseline | ~120 HA / 58 RCR / ~232 FA / HAL-2900 | Authored; residuals projected | Every top hazard controlled; residuals Evidence-Pending
Verification status | 195 PVL activities | 0 executed | Playbook only (RA-05/06); PRD-VER-004/005 Open
Validation status | 128 PVaL activities | 0 executed | No product built, no users tested (RA-06)
Manufacturing readiness | Process design transfer-mature | NOT READY for volume | AML freeze (RA-07); MSA/Cpk (RA-09); pilot build outstanding
Supplier readiness | Enhanced control on cell/BMS/SE/fuse | Partial | Cell chemistry (DO-EE-003) not frozen (RA-07)
Regulatory readiness | WHO PQS / EU GDP / US FDA / air-transport | Not resolved | WHO PQS category + US classification + UDI Evidence-Pending (RA-10)
Cybersecurity readiness | RCR-034…039/051/055 | Pending IV | No independent adversarial verification executed (RA-05)
10.2 Open defects / deviations / risks (release-blocking)
Item | Type | Description | Priority | Gate
B-1 (RA-05/06/08) | Blocker | V&V not executed — all PV-/VAL- Planned/Evidence-Pending | P0 | DR-V&V
B-2 (RA-07) | Blocker | AML not frozen — battery chemistry (DO-EE-003) + antenna (DO-EE-018) | P0 | DR-DO / DR-XFER
B-3 (RA-10) | Blocker | Regulatory unresolved — WHO PQS category, US FDA classification, UDI | P0 | DR-REG
RA-09 | Major | Numeric limits / MSA / Cpk not established | P2 | DR-XFER
D3 (PRD-VER-004) | Open risk | 90-day −30 °C endurance + charge safety unproven | P0 | DR-V&V
RF-in-container (PRD-VER-005) | Open risk | In-container link + antenna strategy unproven | P0 | DR-V&V
HA-TFC-002 (S5) | Residual | Aircraft-hold fire — region-A-High permanent | Accepted w/ conditions | UN 38.3 + monitoring
Security IV | Open risk | Independent adversarial verification not run | P1 | DR-V&V
10.3 Residual-risk approval
Residual | Region | Authority | Basis | Status
Information-pathway residuals (~20 S4) | A Med | System-Safety + Chief Systems | Two choke points (verdict + decision) | Projected — pending PV-SW-002 + VAL-USE-001/004
HA-TFC-002 aircraft-hold fire (S5) | A-High | Chief Systems + Quality/Regulatory | UN 38.3 + containment + mandatory monitoring | Conditional (never broadly acceptable)
HA-MPE-002 population outbreak (S5) | A-High | Chief Systems + Quality/Regulatory | Aggregate record-integrity spine | Conditional
Security→safety residuals | A Med/High | Product Security + System-Safety | Device-anchored authenticity + secure OTA | Pending IV
Accepted public-health trade (HA-MPE-003) | A Med | Chief Systems + Quality/Regulatory | Balanced both directions (HA-NEW-006) | Monitored (TREND-008)
Overall benefit-risk (HAL-2900) | Conditional-positive | DRB (EM-ENG-001) + EM-ENG-006 | 7 conditions unmet | Conditional — pending VAL-RISK-003
10.4 Release decision
Field | Value
Product configuration | SCM-100 (base), regional configs; SCM-100-XA Candidate
Release decision | CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED. Pilot / V&V builds authorized; volume not authorized.
Release authority | Chief Systems Engineer (EM-ENG-001, DRB chair) + Quality/Regulatory (EM-ENG-006), co-approval
Three P0 blockers | B-1 V&V not executed · B-2 AML not frozen (chemistry + antenna) · B-3 regulatory/WHO-PQS/UDI unresolved
Eight conditions of release (from Risk-Management-Summary §7 / HAL-2900 §6) | (1) two pathway choke points verified+validated (PV-SW-002 + VAL-USE-001/004); (2) S5 aircraft-fire + UN 38.3 qualified; (3) D3 endurance/charge closed (PRD-VER-004); (4) RF-in-container closed (PRD-VER-005); (5) fleet common causes controlled (cal, OTA/cloud); (6) independent security verification (IV) passed; (7) WHO PQS accuracy target confirmed + met; (8) post-market surveillance (Stage 17) established for both harm directions + thermal + security
Claim restrictions | No air-transport claim without UN 38.3; no WHO PQS claim until target confirmed + met; no real-time-alert claim on lanes where PRD-VER-005 shows the link cannot close
Gaps (Matrix-10)
No baseline frozen by assertion | all revisions/versions are REL-TBD — correctly not frozen (AML pending). Consistency preserved (see G).
Verification/validation status honest | 0/195 and 0/128 executed — the decision is explicitly not "Ready."
• All three P0 blockers are open; two Open risks (D3, RF) and the security IV are unclosed.
• Residual approvals are conditional/projected — none is a final accepted-with-evidence decision; the two S5 residuals carry named authority + mandatory monitoring (no acceptance without authority).
• Matrix-10 conclusion equals the Stage-16 conclusion; carried into Report I.

## T11 Field Feedback
T11 Field Feedback
Part of the end-to-end traceability record.
↩ Contents
Matrix-11 — Field Feedback / Closed-Loop — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Columns (matrix-map.md): Field Record · Type · Product Configuration · Failure/Complaint · Severity · Frequency · Requirement/Design-Output/Hazard/FMEA/Verification/Manufacturing Link · Root Cause · CAPA · Requirement/Design/Risk/Verification/Validation Update · Closure Status. Source: Stage-17 (FM-100…600), TREND-001…019. The fleet is not yet fielded (Conditionally Ready — not released): there are NO actual field records (FF-###), complaints (CMP-###), or CAPAs (CAPA-###). This matrix is the pre-wired closed-loop plan — each monitored trend is mapped to the design/risk artefacts it would update. Thresholds are safety-EVENT/absolute where a target exists, else baseline-from-field-then-hold (SPC).
---
11.1 Pre-wired field-monitoring trends → traceability links (the closed loop, ready but unpopulated)
Trend | Monitors | Type | Cfg | Would-be signal | Sev | Freq | Req/DO/Hazard/FMEA/PV/Mfg link | Root-cause path | CAPA path | Update target | Status
TREND-004 | Aircraft-hold / thermal event | Safety (S5) | SCM-100 | Any single credible thermal event | S5 | acts on single event | HA-TFC-002 ★★ / FA-HWFMEA-007 / RCR-002/008 / PV-TFC-002 / CC-02 | RCA (FM-300) | CAPA + recall scoping (MDT-900 DHR) | Re-baseline HAL residual + UN 38.3 | Monitoring-Planned (no data)
TREND-007 | Missed-excursion / false-accept | Safety | SCM-100 | Returns/QA missed-excursion signals | S4 | absolute | HA-SW-001 ★ / FA-SWFMEA-001 / RCR-019 / PV-SW-001 / SC-03 | RCA | CAPA → detection logic | Update FWR-022 / RCR-019 / VAL-SW-001 | Monitoring-Planned
TREND-008 | False-reject / wastage (other direction) | Safety (trade) | SCM-100 | Rising false-reject/stock-out | S4 | first-class | HA-MPE-003 ★ / FA-SYSFMEA-013 / RCR-047 / VAL-USE-002 | RCA | CAPA → alarm tuning | Update RCR-047 / HA-NEW-006 balance | Monitoring-Planned
TREND-010 | Calibration drift (common cause) | Safety | SCM-100 | Systematic-bias / recal-escape | S4 | SPC | HA-SW-004/HA-SVC-001 ★ / FA-SWFMEA-005 / RCR-026/049/053 / PV-SW-004 / CC-05 | RCA | CAPA → cal interval/EOL | Update RCR-049 / CAL plan / MSA | Monitoring-Planned
TREND-018/019 | OTA / cloud rollout (common cause) | Safety/Security | SCM-100 | Campaign health / version skew | S4 | SPC | HA-SW-010/HA-CYBER-005 ★ / FA-SWFMEA-026 / RCR-037/039 / PV-SW-010+PV-CYBER-005 / CC-06 | RCA | CAPA → staged rollout | Update RCR-037 / OTA policy | Monitoring-Planned
TREND (battery/endurance) | Battery aging / dead-in-cold | Reliability/Safety | SCM-100 | Endurance shortfall / silent gaps | S4 | SPC | HA-ENRG-006 ★ / FA-REL-007 / RCR-006 / PRD-VER-004 / CC-03 | RCA | CAPA → chemistry/duty-cycle | Update PRD-REL-001 / FA-REL / HA-ENRG-006 | Monitoring-Planned (D3 Open)
TREND (alert delivery) | Alert-delivery success per lane | Safety | SCM-100/-XA | Delivery-success below target | S4 | per-lane | HA-COMM-001/003 ★ / FA-IFMEA-011 / RCR-028/031 / PV-COMM-001/003 | RCA | CAPA → antenna/bearer | Update RCR-031 / PRD-VER-005 | Monitoring-Planned (RF Open)
TREND (return/warranty) | Return / RMA / reuse fade | Reliability/Business | SCM-100 | Return rate / cycle-life shortfall | S3 | SPC | PRD-REL-002 / FA-REL-009 / RCR-050 / VAL-REL-008 | RCA | CAPA → refurb process | Update PRD-REL-002 / service WI | Monitoring-Planned
TREND (cybersecurity events) | Integrity / auth / config anomalies | Security→Safety | SCM-100 | Integrity-failure / unauthorized-config | S4 | event | HA-CYBER-001/003 ★ / RCR-034/036 / PV-CYBER-001/003 (IV) | RCA + IR | CAPA → patch (Pending IV) | Update SBOM / security controls | Monitoring-Planned
11.2 Closed-loop process (pre-wired, FM-300)
Field record (FF-###) → RCA (FM-300) → CAPA (CAPA-###) → requirement/hazard/FMEA/verification update → re-baseline. Serial attribution and recall scoping run off the Device History Record (MDT-900); complaints/adverse-events (CMP-###) carry vaccine-integrity / public-health reporting considerations (FM-400). The reliability-growth + periodic risk-review cadence (FM-600) matures the pre-release projected O/D/probability estimates into field-demonstrated values — closing the honesty ledger.
Gaps (Matrix-11)
Complaints not mapped to functions | N/A — no complaints yet; the mapping template is pre-wired (each trend → function/hazard/FMEA/RCR/PV/CC).
Field failures not added to FMEA / adverse events not reassessed | N/A yet; FM-300 closed loop and FA-2000/2100 RCA/CAPA readiness are established and pre-wired to FMEA/FTA/HAL.
CAPAs without requirement/design changes | N/A — no CAPAs; the process requires a requirement/design/risk update to close (FM-300).
Repeated failures with closed CAPAs / field evidence contradicting probability | N/A — no field data; the FM-600 cadence is the mechanism to detect field occurrence exceeding assumed probability (the standard's consistency check).
The single real gap | there is no field evidence at all — the loop is a plan, not a record (Conditionally Ready — not released). This is expected and honest; it becomes live only after production release. Every V&V-assumption re-validation obligation (HAL-2800/2900) is carried as a standing Stage-17 commitment (condition 8 of release).

## Coverage Dashboard
Coverage Dashboard
Part of the end-to-end traceability record.
↩ Contents
E — Coverage Dashboard — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Required coverage metrics and counts computed from the supplied stage artefacts. Structural coverage (are the links present?) is distinguished from evidential coverage (is there E4–E7 evidence?). Per the standard, critical gaps are NOT hidden behind a high overall percentage: the SCM's structural coverage is high, but its evidential coverage is ~0 because V&V is unexecuted.
---
1. Requirement / need coverage (structural)
Metric | Value | Basis
User needs → ≥1 product requirement | ~100% (structural) | Stage-05 rationale table: 220 PRD each cite ≥1 UN parent; UN→PRD forward links seeded
User needs → ≥1 validation | ~100% (structural) | Stage-14 coverage: every user need/intended-use claim has a VAL scenario; use errors + residuals + safety workflows covered
Product requirements → ≥1 upstream source | 100% | Stage-05: every PRD names a UN/MR/D# parent
System requirements → allocated architecture element | 100% (flagship + Stage-06 matrix) | SYS→S1–S9 allocation; FUN/CMP realization
System/subsystem requirements → ≥1 design output | ~100% structural | Stage-09 Design-Input-Output-Traceability-Matrix
Requirements → verification method | 100% | Every PRD/SYS carries a method; PVL maps all applicable requirements
2. Verification coverage
Metric | Structural | Evidential
RCRs with ≥1 PV activity | 58 / 58 = 100% | —
RCRs implemented (DO- referenced) | 58 / 58 = 100% | Draft/Preliminary (pre-DR-DO)
RCRs implemented AND verified (with evidence) | 58 planned | 0 / 58 = 0% verified (all Evidence-Pending)
Applicable requirements mapped to a PV module | ~100% | 0% executed
Verification activities executed | — | 0 / 195 = 0%
Safety-critical controls with fault injection | 100% (planned) | 0 executed
Security controls with independent adversarial (IV) | 100% (planned, 6+RCR-051) | 0 executed
3. Validation coverage
Metric | Structural | Evidential
Use errors (HA-USE-001…012) with validation reference | 100% | 0 executed
Top residuals (HA-RES-001…022) with validation reference | 100% | 0 executed
Safety-critical workflows with summative activity | 100% | 0 executed
User classes SH-01…15 with a validation activity | 100% | 0 executed
Validation activities executed | — | 0 / 128 = 0%
4. Hazard / risk-control coverage
Metric | Value
High/unacceptable HAL hazards with ≥1 RCR | 100% (~60 top rows; no R-region hazard uncontrolled)
RCRs with a design-output implementation reference | 58 / 58 = 100%
RCRs with a verification target | 58 / 58 = 100%
High-severity (S4/S5) hazards controlled AND verified with evidence | 0% (controls planned, effectiveness unproven)
High-severity hazards with named acceptance authority | 100% (S5 rows carry authority + mandatory monitoring)
FMEA safety actions converted to requirements (RCRs) | ~65 → folded into 58 RCRs, 0 duplicates
FMEA actions closed with evidence | 0 (all Open→RCR, PV Evidence-Pending)
Critical characteristics (CC-01…08) in control plan | 8 / 8 = 100%
Special characteristics (SC-01…09) in control plan | 9 / 9 = 100%
5. Orphan / open counts (see F for detail)
Count | Value
Upstream orphans (requirement w/o source; DO w/o requirement; control w/o hazard; prod test w/o characteristic) | 0 confirmed in flagship + spine; full-corpus sweep = Candidate (human review)
Downstream orphans (user need w/o PR; requirement w/o implementation/verification; hazard w/o control) | 0 confirmed structurally
Evidence orphans (result w/o protocol; protocol w/o requirement; CAPA w/o field record) | 0 (no results/CAPAs exist yet)
Unverified requirements (no executed evidence) | effectively all (0/195 PV executed) — the dominant gap
Hazards without controls | 0
Controls without evidence | 58 (all Evidence-Pending)
Failed tests without disposition | 0 (no tests run)
Unresolved validation findings | 0 (no validation run)
Configuration inconsistencies | 0 (baseline correctly not frozen; see G)
Carried-Open unknowns | 5 (D3 endurance; RF-in-container; WHO PQS category; US FDA classification; AML chemistry+antenna freeze)
6. Overall — honest headline (do not hide gaps)
View | Structural coverage | Evidential coverage
Needs → requirements → design | High (~100%) | High (design authored, E1–E3)
Requirements → verification | 100% mapped | 0% verified
Needs → validation | 100% mapped | 0% validated
Hazards → controls | 100% controlled | 0% proven effective
Release | — | Conditionally Ready — not authorized
Interpretation. The SCM scores near-perfect on structural traceability and near-zero on evidential maturity — exactly the profile of a mature design-input package with unexecuted V&V. Reporting only the structural percentages would be dishonest; the load-bearing metrics are the 0% verified / 0% validated / 0% proven-effective rows and the 5 carried-open unknowns. These, not missing links, define release readiness (Report I).

## Gaps & Orphans
Gaps & Orphans
Part of the end-to-end traceability record.
↩ Contents
F — Orphan & Gap Report — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Upstream / downstream / evidence orphans + the known Open items. Per the standard's gap-detection taxonomy. Gap priority: P0 blocker · P1 critical/safety · P2 major · P3 moderate · P4 minor.
---
1. Upstream orphans (requirement w/o source; DO w/o requirement; control w/o hazard; prod test w/o characteristic)
Finding | Status | Priority
Product requirements without an upstream source | None found — Stage-05 rationale table gives every PRD a UN/MR/D# parent | —
Design outputs without a requirement | None found in flagship + spine. Full 272-output sweep against requirements is Candidate — Human Review Required (this stage verified the safety spine, not all 272 outputs individually) | P3 (verify)
Risk controls without a hazard | None — every RCR-001…058 names a source hazard (Stage-12 §2) | —
Production tests without a characteristic | None — every PT ties to a CC/SC (Stage-15 MDT-400/600) | —
2. Downstream orphans (user need w/o PR; requirement w/o implementation/verification; hazard w/o control; validation finding w/o disposition)
Finding | Status | Priority
User needs without a product requirement | None found structurally | —
Requirements without implementation | None structurally (all map to DO-); but all DO- are Draft/Preliminary (pre-DR-DO) | P1
Requirements without verification evidence | Effectively all — 0/195 PV executed | P0
Hazards without a control | None — 100% of top hazards controlled | —
Validation findings without disposition | None exist — no validation executed | (P0 to execute)
3. Evidence orphans (result w/o protocol; protocol w/o requirement; CAPA w/o field record; residual decision w/o evidence)
Finding | Status | Priority
Results without a protocol | None — no results exist; protocols (PVL/PVaL) authored | —
Protocols without a requirement | None — PVL/PVaL are requirement-driven | —
CAPA without a field record | None — no CAPAs (fleet not fielded) | —
Residual-risk decisions without evidence | All 22+ residuals are projected / Evidence-Pending — no E4–E7 basis yet | P0/P1
4. The known Open items (carried honestly, all Evidence Pending; V&V not executed)
Open item | What is unknown | Trace anchor | Closing test/action | Priority
D3 battery (90-day −30 °C endurance + cold-charge safety + reuse fade) | Energy margin, low-temp charge safety, capacity fade — the top program risk | HA-ENRG-001/006 · RCR-001/005/006/017 · FA-HWFMEA-001/012 · FA-REL-007/009 · PRD-REL-001 | PRD-VER-004 → PV-REL-001 + PV-ENRG-001 + UN 38.3; freeze DO-EE-003 | P0
RF-in-container (in-container link margin; integral-vs-external antenna) | Whether a live alert can egress a metallized shipper; antenna strategy | HA-COMM-003 · RCR-031/030 · FA-HWFMEA-029/IFMEA-005 · PRD-ENV-009 | PRD-VER-005 → PV-REL-005 + PV-COMM-003; freeze DO-EE-018 (or adopt SCM-100-XA) | P0
WHO PQS category | Exact device category + accuracy target (D4) | MR-REG · PRD-REG-001 · DO-REG-001 · PV-COMP (WHO PQS) | Regulatory ingestion; confirm target then verify | P0
US FDA classification | Classification/pathway (accessory/information device) | MR-REG · PRD-REG-002 · DO-REG-005 | Regulatory determination | P0
AML freeze | Battery chemistry (DO-EE-003) + antenna (DO-EE-018) not frozen | DO-CFG-004 · RA-07 | Freeze after PRD-VER-004/005 close | P0
Security independent verification (IV) | Effectiveness of security→safety controls | HA-CYBER-001…009 · RCR-034…039/051 | PV-CYBER- (IV) | P1
MSA / Cpk / numeric limits | Measurement-system capability, process capability | RA-09 · MDT-800 · PV-MFG | Establish against released DOs + MSA | P2
Cycle-life over reuse; cleaning-chemical compatibility | Reuse-fade, cleaning tolerance | PRD-REL-002 / PRD-SVC-002 · FA-REL-009 | VAL-REL-008 / PV-REL (aging) | P2/P3
5. Summary
The SCM has no confirmed broken links along the flagship chains and the safety/risk-control spine. The gaps are of three kinds, in priority order:
1. Evidence gaps (P0): V&V is entirely unexecuted — every verification (0/195), validation (0/128), and residual-risk decision (all projected) lacks E4–E7 evidence. This is the dominant, release-blocking gap.
2. Five carried-Open unknowns (P0): D3 endurance, RF-in-container, WHO PQS category, US FDA classification, and the AML (chemistry + antenna) freeze.
3. Verification-scope candidates (P2/P3): a full-corpus orphan sweep of all 272 design outputs and the MSA/Cpk numeric limits remain Candidate / Human Review Required — this stage evidence-backed the safety spine, not every one of the ~1,900 objects individually.

## Inconsistencies
Inconsistencies
Part of the end-to-end traceability record.
↩ Contents
G — Inconsistency Report — SCM
Stage 18 · ORCH-018 / EM-ENG-012. Contradictions across stages, per the standard's consistency checks. Each item states the check, the finding, and the disposition. Headline: no active contradiction survives — the one structural tension (preliminary PRD-SAFE vs the reconciled RCR register) has already been reconciled by Stage 12.
---
1. Consistency checks (standard list) applied to the SCM
Consistency check | Finding | Disposition
Requirement "verified" while its test failed | No occurrence — nothing is marked verified; 0/195 PV executed | Clean
Hazard "controlled" while control verification incomplete | Systemic, but honestly labelled — all hazards are "controlled" (RCR mapped) while PV is Evidence-Pending. This is disclosed everywhere, not asserted as done | Acceptable (honest) — closes on V&V
FMEA action "closed" while design output unreleased | No occurrence — no FMEA action is "closed"; all are Open→RCR; outputs are Draft/Preliminary | Clean
Validation passed on a non-representative configuration | No occurrence — no validation executed | Clean
Process uses an obsolete drawing revision | No occurrence — all revisions TBD-rev; AML not frozen (RA-07); process transfer explicitly gated on freeze | Clean (baseline not frozen)
Protocol references a superseded requirement | No occurrence — PVL/PVaL reference current baselines; PRD-SAFE re-baselined and protocols reference the RCRs | Clean
Residual probability reduced without a probability-reducing control | No occurrence — residuals are projected, not asserted reduced; probability estimates carry a stated basis / Evidence-Pending | Clean
Severity reduced by a detection-only control | No occurrence — severity is outcome-fixed (S5 stays S5); detection controls change O/D, not S (Stage-11 §3; HAL-2000) | Clean
Field occurrence exceeds assumed probability | N/A — no field data; FM-600 cadence pre-wired to detect this | Deferred to Stage 17
Cyber threat "mitigated" while vulnerable component remains in SBOM | No occurrence — no cyber threat is marked "mitigated"; all security residuals are Pending IV | Clean
Requirement deleted while outputs/tests still reference it | No occurrence — no deletions; PRD-SAFE items were reconciled/expanded, not deleted (see §2) | Clean
Risk accepted by someone without authority | No occurrence — S5 residuals carry named authority (Chief Systems + Quality/Regulatory) + mandatory monitoring | Clean
2. The reconciliation already performed (preliminary PRD-SAFE → reconciled RCRs)
The one genuine cross-stage tension in the chain was designed-in and already reconciled, not left as a contradiction:
• The situation. Stage-05 emitted PRD-SAFE-001…012 as explicitly PRELIMINARY ("safety/risk-control requirements originate in the HAL, Stage 10"). Likewise SRS-SAFE (1200), SYS-safety (SYS-149…162), and all 14 DOL-1200 DO-SAFE outputs were flagged Preliminary — reconcile with HAL. Had this been left, it would be the classic inconsistency: safety requirements asserted before the hazard analysis that should drive them.
• The reconciliation (Stage 12). PRD-SAFE-Reconciliation.md maps each PRD-SAFE-001…012 to the RCR(s) it becomes/updates (12/12 reconciled) and recommends 5 new authoritative requirements (PRD-SAFE-013…017). The HAL RCR register (RCR-001…058) is the authoritative spine; each safety-significant FMEA action is folded into an RCR as a Source Failure Mode with zero duplicate requirements.
• Residual tension (tracked, not contradictory). The DOL-1200 DO-SAFE outputs remain Preliminary pre-DR-DO — they must each map to a baselined HAL control at gate DO-XFER-008 (RA-05, the top design-output blocker). The requirement/RCR reconciliation is done; the design-output re-baseline on executed V&V is not. This is carried as a release blocker (H/I), not an inconsistency.
3. Other cross-stage checks
Check | Finding
ID scheme consistency across stages | Consistent — the concept §5 ID scheme is honoured (MR-/VOC-/STN-/UN-/PRD-/SYS-/HWR-MER-FWR-IFR-/SRS-/FUN-CMP-INT-/DO-/HA-RCR-/FA-/PV-/VAL-/PROC-CP-/REL-DR-/FF-CMP-CAPA-TREND-). Minor note: CMP- is reused for both architecture components (Stage-08) and field complaints (Stage-17); disambiguated by stage context — flagged for the register but not a data error
Severity-first (Action Priority, not RPN) | Consistent across Stages 10/11/12/15 (CLAUDE.md rule 3)
Two Open items named consistently | Consistent — D3 (PRD-VER-004) and RF-in-container (PRD-VER-005) carried Open identically in Stages 06/07/08/10/11/12/13/14/16
Release conclusion consistency | Consistent — Stage-10 Risk-Management-Summary, Stage-15 MDT-1000, and Stage-16 REL-400 all say Conditionally Ready — production not authorized
4. Summary
No active contradiction is present. The chain's one designed-in tension (preliminary safety requirements ahead of the hazard analysis) was reconciled at Stage 12 with a documented PRD-SAFE→RCR map, no duplicates, and severity-fixed scoring. The only carried tension is the DOL-1200 DO-SAFE Preliminary status awaiting the DR-DO/DR-XFER re-baseline on executed V&V — correctly tracked as a release blocker (RA-05), not hidden as "done." One cosmetic ID-namespace overlap (CMP- architecture vs complaint) is flagged for the register.

## Closure Action Plan
Closure Action Plan
Part of the end-to-end traceability record.
↩ Contents
H — Traceability Closure Action Plan — SCM
Stage 18 · ORCH-018 / EM-ENG-012. The prioritized actions to close the gaps and lift the SCM from Conditionally Ready to Ready. Priority: P0 release blocker · P1 critical/safety · P2 major (close before release) · P3 moderate · P4 minor. Owners named per stage.
---
P0 — Release blockers (must close before production authorization)
# | Action | Closes | Trace anchor | Owner | Gate
P0-1 | Execute the safety-critical verification set with fault injection — detection, verdict, gap, alert, charge-inhibit, protection, calibration, probe, power-fail | Verification 0% → verified | RCR-019/021/022/028/029/001/002/026/020/024; PV-SW/PV-ENRG/PV-TFC | EM-ENG-006 (Verification lead) + discipline owners | DR-V&V
P0-2 | Execute the summative validation set — VAL-USE-001 (verdict→disposition), VAL-USE-002 (alarm response), VAL-USE-004 (never trust void), plus VAL-RISK-003 (benefit-risk) | Validation 0% → validated | HA-USE-001/002/004 ↔ VAL-USE-001/002/004; HAL-2900 | HF/Usability lead + EM-ENG-006 | DR-V&V
P0-3 | Close D3 — run the 90-day −30 °C endurance + cold-charge-safety test; freeze cell chemistry (DO-EE-003) | D3 Open; CHAIN-2/3; RA-07 | PRD-VER-004 → PV-REL-001 + PV-ENRG-001; RCR-001/006/017; HA-ENRG-006 | Reliability + Battery/Power | DR-V&V + DR-XFER
P0-4 | Close RF-in-container — run PRD-VER-005; decide integral (DO-EE-018) vs external antenna (SCM-100-XA); freeze | RF Open; CHAIN-4; RA-07/08 | PRD-VER-005 → PV-REL-005 + PV-COMM-003; RCR-031 | RF + Systems (EM-ENG-001) | DR-V&V + DR-XFER
P0-5 | Resolve regulatory — confirm WHO PQS device category + accuracy target; determine US FDA classification; finalize UDI/markings | RA-10; CHAIN-5 (PQS) | PRD-REG-001/002; DO-REG-001/005; DO-LBL-002/004; PV-COMP | EM-ENG-006 (Quality/Regulatory) | DR-REG
P0-6 | Qualify UN 38.3 / IATA for the frozen cell/pack (enables the air-transport claim + S5 case) | S5 (HA-TFC-002); air claim | HWR-011; RCR-007; PV-COMP; PV-TFC-002 | EM-ENG-006 + Reliability/Safety | DR-V&V
P0-7 | Freeze the configuration/AML baseline at DR-DO/DR-XFER once P0-3/4 close; re-baseline the 14 DOL-1200 DO-SAFE outputs against baselined HAL controls | RA-05/07; Matrix-03/09 | DO-XFER-008/010; DO-CFG-004 | EM-ENG-001 | DR-DO / DR-XFER
P1 — Critical / safety (safety-review before release)
# | Action | Closes | Trace anchor | Owner | Gate
P1-1 | Independent adversarial security verification (IV) — integrity, alarm-suppression, config-tamper, OTA, secure-boot/identity, cloud/DR pen-tests | Security residuals Pending IV; Matrix-06 | RCR-034…039/051; PV-CYBER-001…009 | Independent security + Product Security (EM-ENG-008) | DR-V&V
P1-2 | Verify the record-integrity spine — power-fail-safe writes, tamper-evidence, honest-gap, export integrity, trusted time | HA-SW-013/HA-DATA-004/006; CHAIN-6 | RCR-022/024/034/042/025; PV-SW-013/PV-DATA-004/006 | EM-ENG-007 + EM-ENG-008 | DR-V&V
P1-3 | Verify fleet common-cause controls — calibration (TREND-010) and OTA/cloud rollout (TREND-018/019) staged/diverse | HA-MPE-002; FA-CCF; Matrix-11 | RCR-049/053/037/039/057; PV-MFG-001/PV-SW-010 | Reliability + Cloud + Quality | DR-V&V
P1-4 | Confirm residual-risk acceptability with evidence — mature the ~22 projected residuals to accepted-with-evidence via VAL-RISK-001/003 | Residual decisions w/o evidence (F §3) | HA-RES-001…022; HAL-2900 7 conditions | EM-ENG-006 + System-Safety | DR-Risk
P2 — Major (close before release)
# | Action | Closes | Owner
P2-1 | Establish MSA / gauge-R&R / Cpk and all Evidence-Pending numeric limits (accuracy, leak-rate, torque, coverage) | RA-09; Matrix-09 | Quality/Manufacturing
P2-2 | Run the pilot/DVT→PVT build; confirm manufacturing readiness (line clearance, poka-yoke, identity uniqueness) | Mfg NOT READY; RA-07/09 | Manufacturing
P2-3 | Verify environmental/reliability set (−30…+60 °C, IP67, drop/shock/vibe, altitude, HALT/HASS, battery aging, flash endurance) | Env/rel coverage | Reliability
P2-4 | Validate reuse loop — cycle-life over reuse (PRD-REL-002), cleaning-chemical compatibility (PRD-SVC-002) | CHAIN-7 open | Service + Reliability
P3 — Moderate
# | Action | Closes | Owner
P3-1 | Full-corpus orphan sweep — verify all 272 design outputs have a requirement link, and all ~1,900 objects are linked (this stage evidence-backed the spine; the full sweep is Candidate) | F §1 Candidate | EM-ENG-012
P3-2 | Confirm SCM-100-XA (external-antenna) configuration disposition once P0-4 decides | Matrix-02/03 Candidate | Systems
P3-3 | Resolve the storage-sizing Evidence-Pending item (SYS-127) | Stage-07 open | EM-ENG-007
P4 — Minor
# | Action | Closes | Owner
P4-1 | Disambiguate the CMP- ID namespace (architecture component vs field complaint) in the object register | G §3 | EM-ENG-012
P4-2 | Back-fill the placeholder VOC-### references seeded in the MR library with the real VoC IDs | Stage-01/02 note | Market Research
Sequencing
1. P0-5/P0-6 (regulatory + UN 38.3) and P0-3/P0-4 (D3 + RF) unblock the AML freeze (P0-7).
2. With the baseline frozen, run P0-1/P0-2 + P1-1…P1-4 V&V (fault-injection + IV + summative + residual).
3. P2 manufacturing/MSA/pilot proceeds in parallel once AML is frozen.
4. Only when all P0 and safety P1 are closed with evidence, and the 8 conditions of release are met, does the DRB re-convene to move the decision from Conditionally Ready to Ready (Report I).

## Release Decision
Release Decision
Part of the end-to-end traceability record.
↩ Contents
I — Release-Readiness Conclusion — SCM
Stage 18 · ORCH-018 / EM-ENG-012 (Principal Systems Traceability Architect & Design Assurance Engineer, reporting to Chief Systems EM-ENG-001). The traceability-authority conclusion on whether the SCM may be released. Classification per the standard: Not Ready / Conditionally Ready / Ready with Approved Residual Actions / Ready — never "Ready" when critical evidence is absent. This conclusion must — and does — equal the Stage-16 (REL-400) and Stage-10 (Risk-Management-Summary) conclusions.
---
Classification
## CONDITIONALLY READY — PRODUCTION NOT AUTHORIZED
(release pending V&V execution + closure of the P0 blockers and the 8 conditions of release)
Pilot / V&V builds are authorized; volume production is not. This is the same conclusion reached independently by the Stage-10 Risk-Management-Summary, the Stage-15 Manufacturing Readiness Assessment (MDT-1000), and the Stage-16 Release Decision (REL-400) — the traceability audit confirms, it does not soften, that conclusion.
Why not "Ready" (the evidence test)
Traceability is structurally complete (near-100% of needs→requirements→design→hazard→control→V&V links present, bidirectional, with real IDs) but evidentially empty: 0 of 195 verification activities and 0 of 128 validation activities have executed, so no requirement is verified, no user need is validated, and no hazard control is proven effective. Per the hallucination policy and the standard's rule — never mark Ready without evidence; never infer verification from a planned test or validation from verification — the only honest classification is Conditionally Ready.
Release blockers (P0)
Blocker | Description | Trace
B-1 | V&V not executed — every PV-/VAL- is Planned / Evidence-Pending | RA-05/06/08; Matrix-07/08
B-2 | AML not frozen — battery chemistry (DO-EE-003) and antenna (DO-EE-018) unresolved | RA-07; Matrix-03/09
B-3 | Regulatory unresolved — WHO PQS device category + accuracy target, US FDA classification, UDI/markings | RA-10; CHAIN-5
Unresolved (carried-Open) risks
Open risk | Status | Closing test
D3 — 90-day −30 °C endurance + cold-charge safety + reuse fade (top program risk) | Open | PRD-VER-004 → PV-REL-001 + PV-ENRG-001 + UN 38.3
RF-in-container alert egress (principal open architectural risk) | Open | PRD-VER-005 → PV-REL-005 + PV-COMM-003
S5 aircraft-hold Li-cell fire (HA-TFC-002) | Region-A-High permanent — never broadly acceptable | UN 38.3 + containment + mandatory monitoring
S5 population outbreak (HA-MPE-002) | Region-A-High — aggregate record-integrity spine | Aggregate spine V&V + benefit-risk (VAL-RISK-003)
Security → safety (HA-CYBER-001…009) | Residuals Pending independent adversarial verification (IV) | PV-CYBER- (IV)
Missing evidence (what must exist before "Ready")
• Executed fault-injection verification for every safety-critical control (E5+).
• Executed independent adversarial verification for every security control (E6).
• Executed summative usability validation of the two convergence choke points (VAL-USE-001/004) and the benefit-risk conditions (VAL-RISK-003).
• UN 38.3 qualification of the frozen cell/pack; WHO PQS accuracy target confirmed and met.
• MSA / Cpk and all Evidence-Pending numeric limits established; pilot/PVT build completed.
• Re-baseline of the 14 DOL-1200 DO-SAFE outputs against baselined HAL controls (gate DO-XFER-008).
Required monitoring (standing conditions — Stage 17)
Post-market surveillance (release condition 8) must be established before release and run across the reusable fleet's service life: both harm directions (missed excursion TREND-007 and false-reject/wastage TREND-008), the S5 thermal-event signal (TREND-004, mandatory, acts on a single event), the two dominant common causes — calibration drift (TREND-010) and OTA/cloud rollout (TREND-018/019) — and cybersecurity events. The closed loop (FM-300) matures the pre-release projected O/D/probability estimates into field-demonstrated values.
Approval authorities
Decision | Authority
Production release | Chief Systems Engineer (EM-ENG-001), chairing the Design Review Board, co-approved by Quality/Regulatory (EM-ENG-006)
S5 residual acceptance (HA-TFC-002, HA-MPE-002) | Chief Systems + Quality/Regulatory, with mandatory monitoring
Security→safety residual acceptance | Product Security (EM-ENG-008) + System-Safety, after IV
Overall benefit-risk (ISO 14971, HAL-2900) | DRB (EM-ENG-001) + EM-ENG-006
Traceability / design-assurance sign-off | EM-ENG-012 (this stage) — declares the release blockers above
Path to "Ready"
On execution of the P0 and safety-P1 actions in Report H with passing evidence, and satisfaction of the 8 conditions of release, the appropriate re-classification would be Ready with Approved Residual Actions (because the S5 residuals remain region-A-High under mandatory monitoring — a permanent residual, not a defect), advancing to Ready only if/when those residuals and the WHO PQS benefit claim are fully discharged. Today, the SCM is Conditionally Ready — production is not authorized.
---
Signed (traceability authority): EM-ENG-012 — Principal Systems Traceability Architect & Design Assurance Engineer, ORCH-018. 2026-07-19. Evidence-backed only in the Stage 00–17 artefacts; no result, approval, or residual acceptance is asserted that has not been produced by the upstream stages.